## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-710C | GDU-002 | 2008-01-01 | 2009-01-01
0 | 0 | 0 | 0
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 41304
NDUV Serial Number: | 217
FID Serial Number: | 177
Heated Filter Serial number: | 18185
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALMO23797 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.1 | Scott gas | ALMO24749 | ±1% | 2010-08-25 | 93.3
200 ppm CO | 200.2
6% CO2 | 6.07
300 ppm NO | 301.3
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6.01
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALMO31297 | ±1% | 2010-10-06
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | NO2 NDUV Serial Number: |  |  | 217 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-11-03 |  | 518.6 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 13:48
Test software: | 10.09 SP1
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0.8 | 0.8 |  |  | 0 | 0.1 | 0.1
90 | 105 | 1 | 51.86 | 51.3 | -0.6 | -1.16 |  | 2.09 | 2.1 | 0.01 | 0.48
150 | 165 | 2 | 103.72 | 103.5 | -0.2 | -0.19 |  | 4.18 | 4.1 | -0.08 | -1.91
210 | 225 | 3 | 155.58 | 156.1 | 0.5 | 0.32 |  | 6.27 | 6.2 | -0.07 | -1.12
270 | 285 | 4 | 207.44 | 208.5 | 1.1 | 0.53 |  | 8.36 | 8.3 | -0.06 | -0.72
330 | 345 | 5 | 259.3 | 261.4 | 2.1 | 0.81 |  | 10.45 | 10.4 | -0.05 | -0.48
390 | 405 | 6 | 311.16 | 314.3 | 3.1 | 1 |  | 12.54 | 12.4 | -0.14 | -1.12
450 | 465 | 7 | 363.02 | 365.2 | 2.2 | 0.61 |  | 14.63 | 14.5 | -0.13 | -0.89
510 | 525 | 8 | 414.88 | 416.8 | 1.9 | 0.46 |  | 16.72 | 16.6 | -0.12 | -0.72
570 | 585 | 9 | 466.74 | 468.5 | 1.8 | 0.39 |  | 18.81 | 18.7 | -0.11 | -0.58
630 | 645 | 10 | 518.6 | 520.7 | 2.1 | 0.4 |  | 20.9 | 20.9 | 0 | 0
690 | 705 | 9 | 466.74 | 470.7 | 4 | 0.86 |  | 18.81 | 18.8 | -0.01 | -0.05
750 | 765 | 8 | 414.88 | 419.1 | 4.2 | 1.01 |  | 16.72 | 16.6 | -0.12 | -0.72
810 | 825 | 7 | 363.02 | 367.6 | 4.6 | 1.27 |  | 14.63 | 14.5 | -0.13 | -0.89
870 | 885 | 6 | 311.16 | 316.2 | 5 | 1.61 |  | 12.54 | 12.5 | -0.04 | -0.32
930 | 945 | 5 | 259.3 | 264.1 | 4.8 | 1.85 |  | 10.45 | 10.4 | -0.05 | -0.48
990 | 1005 | 4 | 207.44 | 211.4 | 4 | 1.93 |  | 8.36 | 8.3 | -0.06 | -0.72
1050 | 1065 | 3 | 155.58 | 159.4 | 3.8 | 2.44 |  | 6.27 | 6.22 | -0.05 | -0.8
1110 | 1125 | 2 | 103.72 | 107 | 3.3 | 3.18 |  | 4.18 | 4.2 | 0.02 | 0.48
1170 | 1185 | 1 | 51.86 | 54.7 | 2.8 | 5.4 |  | 2.09 | 2.1 | 0.01 | 0.48
1230 | 1245 | 0 | 0 | 2.1 | 2.1 |  |  | 0 | 0.1 | 0.1
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | 1.6 | 5.2 |  | Intercept | 0.02 | 0.21
|  |  |  | Slope | 1 | 0.99 - 1.01 |  | Slope | 0.99 | 0.99 - 1.01
|  |  |  | SEE | 1.5407 | 5.186 |  | SEE | 0.0562 | 0.209
|  |  |  | r2 | 0.9999 | 0.998 |  | r2 | 0.9999 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.8 | 1.4 | 1.4 | 0 | 0 | 0.1 | 13 | 914 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 191 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8 | 27.7 | 1.4
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.8 | 1.3698 | 1.4 | 0 | 0 | 0.1 | 13 | 914 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 191.1 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.572 | 27.7 | 1.4
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.8 | 1.3 | 1.3 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.6852 | 27.7 | 1.3
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.8 | 1.3 | 1.3 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.6067 | 27.7 | 1.3
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.8 | 1.2698 | 1.3 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.5282 | 27.7 | 1.3
2008-11-03 | 13:48 | 0 | -0.0006 | -5.8232 | -0.8 | 1.2 | 1.2 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 29 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.6005 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.2 | 1.2 | 0 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 29.1 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.2 | 1.2 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 30 | 192 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.6786 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.2 | 1.2 | 0 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 30.054 | 192 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.00097 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.4431 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.2 | 1.2 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.4641 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.1698 | 1.2 | 0 | 0 | 0.1 | 13.1 | 912 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.5846 | 27.7 | 1.2
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.8 | 1.1 | 1.1 | 0 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191.9 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.3014 | 23.6588 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | 0 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.5009 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -0.8 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.5 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190.9 | 30 | 5 | 196 | 59.9 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.6986 | 23.5 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 913 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190 | 30 | 5 | 196 | 59.9302 | 0.0008 | 0.0009 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.5 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.5 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 913 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189.9 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.302 | 23.5 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.6405 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1.1 | 1.1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7 | 1.1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 905 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7051 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 12.9 | 914 | 905 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.784 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 189.1 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.6121 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 12.9 | 914 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.6233 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.7805 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.7023 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 190.1 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7 | 1
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 1 | 1 | -1 | 0 | 0.1 | 13 | 914 | 903 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.302 | 23.6451 | 27.7 | 1
2008-11-03 | 13:48 | 0 | -0.0004 | -4.0494 | -0.7 | 0.96976 | 1 | -1 | 0 | 0.1 | 13 | 914 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.6976 | 23.6334 | 27.7 | 1
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.7 | 0.9 | 0.9 | -1 | 0 | 0.1 | 13 | 914 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7 | 0.9
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.7 | 0.9 | 0.9 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191.1 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7 | 0.9
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.7 | 0.9 | 0.9 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.3032 | 23.7688 | 27.7 | 0.9
2008-11-03 | 13:48 | 0 | -0.0002 | -2.3514 | -0.7 | 0.9 | 0.9 | -1 | 0 | 0.1 | 13 | 913 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0007 | 0.0008 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8944 | 27.7 | 0.9
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.7 | 0.9 | 0.9 | -1 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.923 | 27.7 | 0.9
2008-11-03 | 13:48 | -0.001 | 0 | 0 | -0.6 | 0.8698 | 0.9 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7 | 27.7 | 0.9
2008-11-03 | 13:48 | -0.009 | 0 | 0 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | 0 | 0 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.004 | 0 | 0 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7395 | 27.7 | 0.8
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 192 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8 | 27.7 | 0.8
2008-11-03 | 13:48 | 0 | 0 | 0 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 913 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191.9 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8 | 27.7 | 0.8
2008-11-03 | 13:48 | 0 | -0.0004 | -4.0196 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 912 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8 | 27.7 | 0.8
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 913 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31 | 191 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.007 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 31.054 | 190.9 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8318 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 913 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.8694 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 196 | 59.9 | 0.0006 | 0.0007 | 0.0001 | 7e-05 |  | n/a |  | 33 | 60 | 1056 | 1017.6994 | 23.6341 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | -0.001 | -10 | -0.7 | 0.8 | 0.8 | -1 | 0 | 0.1 | 12.9 | 914 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 196 | 59.9 | 0.0005 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.6 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | -0.001 | -10 | -0.7 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189.9 | 30 | 5 | 196 | 59.8302 | 0.0005 | 0.00067 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.6455 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.01 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 12.9 | 915 | 903 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 196 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7239 | 27.7 | 0.8
2008-11-03 | 13:48 | -0.002 | -0.001 | -10 | -0.6 | 0.76956 | 0.8 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 196 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8 | 27.7 | 0.8
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 196 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8 | 27.7 | 0.7
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 915 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 195.8 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7407 | 27.7 | 0.7
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 914 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189.1 | 30 | 5 | 195.4 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.6623 | 27.7 | 0.7
2008-11-03 | 13:48 | 0 | -0.001 | -10 | -0.6 | 0.7302 | 0.7 | -1 | 0 | 0.1 | 13 | 915 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.6323 | 27.7 | 0.7
2008-11-03 | 13:49 | 0 | -0.001 | -10 | -0.6 | 0.8 | 0.8 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195.3 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.7892 | 27.7 | 0.8
2008-11-03 | 13:49 | 0 | -0.001 | -10 | -0.6 | 0.76911 | 0.8 | -1 | 0 | 0.1 | 12.9 | 915 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 32 | 190 | 30 | 5 | 195.7 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.727 | 27.7 | 0.8
2008-11-03 | 13:49 | 0 | -0.001 | -10 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 914 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190.1 | 30 | 5 | 196 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7 | 27.7515 | 0.7
2008-11-03 | 13:49 | 0 | -0.001 | -10 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 12.9 | 915 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191 | 30 | 5 | 196 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7598 | 27.7701 | 0.7
2008-11-03 | 13:49 | 0 | -0.0004 | -4.353 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 914 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191 | 30 | 5 | 196 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8917 | 27.7083 | 0.7
2008-11-03 | 13:49 | 0 | 0 | 0 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 12.9 | 915 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191 | 30 | 5 | 196 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8133 | 27.7867 | 0.7
2008-11-03 | 13:49 | 0 | 0 | 0 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 914 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191.1 | 30 | 5 | 195.6 | 59.9 | 0.0004 | 0.0005 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7349 | 27.8 | 0.7
2008-11-03 | 13:49 | 0 | 0 | 0 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195.3 | 59.9 | 0.000431 | 0.000531 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.7436 | 27.8 | 0.7
2008-11-03 | 13:49 | 0 | 0 | 0 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8 | 27.7779 | 0.7
2008-11-03 | 13:49 | 0 | 0 | 0 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017 | 23.8005 | 27.7 | 0.7
2008-11-03 | 13:49 | 0 | -0.0018 | -17.5184 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.879 | 27.7 | 0.7
2008-11-03 | 13:49 | 0 | -0.003 | -30 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195.2 | 59.9 | 0.0005 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8426 | 27.7574 | 0.7
2008-11-03 | 13:49 | 0 | -0.0001 | -1.4818 | -0.6 | 0.7 | 0.7 | -1 | 0 | 0.1 | 13 | 913 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195.5 | 59.9 | 0.0005 | 0.00057 | 0.0001 | 0 |  | n/a |  | 33 | 60 | 1056 | 1017.698 | 23.8 | 27.8 | 0.7
2008-11-03 | 13:49 | 0 | 0.0098 | 98.0396 | -0.6 | 9.6482 | 9.6 | -1 | 0 | 0.11 | 13 | 912 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 195.9 | 59.9 | 0.0005 | 0.00053 | 0.000551 | 0 |  | n/a |  | 33 | 60 | 1056 | 1018 | 23.8142 | 27.8 | 9.6
2008-11-03 | 13:49 | 0.006 | 0.0064 | 63.6457 | 0.5 | 42.4442 | 42.9 | -0.1 | 0 | 1.07 | 13 | 913 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 196 | 59.9 | 0.000721 | 0.000691 | 0.001802 | 0 |  | n/a |  | 33 | 60 | 1056.7389 | 1018 | 23.8926 | 27.8 | 42.4
2008-11-03 | 13:49 | 0.01 | 0.0068 | 67.5635 | 0.8 | 48.9076 | 49.7 | 3.8 | 3.7846 | 1.74 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 192 | 30 | 5 | 196 | 59.9 | 0.001604 | 0.001473 | 0.00263 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.8 | 48.9
2008-11-03 | 13:49 | 0.017 | 0.0064 | 64.1639 | 0.9 | 50.3302 | 51.3 | 3.2 | 3.1976 | 1.97 | 13 | 918 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191.9 | 30 | 5 | 196 | 59.9 | 0.0023 | 0.0022 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8011 | 27.8 | 50.3
2008-11-03 | 13:49 | 0.04 | 0.0044 | 44.0389 | 1 | 50.5 | 51.5 | 3 | 3 | 2 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 191 | 30 | 5 | 195.8 | 59.9 | 0.0024 | 0.00223 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7557 | 27.8 | 50.5
2008-11-03 | 13:49 | 0.04 | 0.002 | 20 | 1 | 50.5304 | 51.6 | 4 | 4 | 2 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190.9 | 30 | 5 | 195.5 | 59.9304 | 0.002613 | 0.002483 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8873 | 27.8 | 50.5
2008-11-03 | 13:49 | 0.033 | 0.002 | 20 | 1.2 | 50.6302 | 51.9 | 3.1 | 3.1042 | 2 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195.1 | 59.9 | 0.003221 | 0.003021 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7306 | 27.8 | 50.6
2008-11-03 | 13:49 | 0.03 | 0.002 | 20 | 1.3 | 50.8 | 52.1 | 2.1 | 2.1333 | 2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7632 | 27.8 | 50.8
2008-11-03 | 13:49 | 0.03 | 0.002 | 20 | 1.3 | 50.9 | 52.2 | 2 | 2 | 2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0036 | 0.0034 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1017.6986 | 23.8 | 27.8 | 50.9
2008-11-03 | 13:49 | 0.03 | 0.002 | 20 | 1.3 | 51.0302 | 52.3 | 2 | 2 | 2 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189.9 | 30 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1017 | 23.8401 | 27.8 | 51
2008-11-03 | 13:49 | 0.03 | 0.002 | 20 | 1.4 | 51.2 | 52.6 | 1 | 1 | 2 | 12.9 | 921 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9969 | 27.8 | 51.2
2008-11-03 | 13:49 | 0.03 | 0.002 | 20 | 1.4 | 51.2 | 52.6 | 1 | 1 | 2.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9231 | 27.8 | 51.2
2008-11-03 | 13:49 | 0.03 | 0.0019 | 19.0223 | 1.4 | 51.2 | 52.6 | 1 | 1 | 2.1 | 12.9 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.8 | 51.2
2008-11-03 | 13:49 | 0.019 | 0.0002 | 1.8696 | 1.4 | 51.2 | 52.6 | 0.2 | 0.1805 | 2.1 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189 | 30 | 5 | 195 | 59.9 | 0.00373 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8326 | 27.8 | 51.2
2008-11-03 | 13:49 | 0.006 | 0 | 0 | 1.4 | 51.2 | 52.6 | 0 | 0 | 2.1 | 12.9 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 189.1 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7242 | 27.8 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.4 | 51.2 | 52.6 | 0 | 0 | 2.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0035 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8812 | 27.8 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.4 | 51.2 | 52.6 | 0 | 0 | 2.1 | 12.9 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.00353 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.8 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.4 | 51.2 | 52.6 | 0 | 0 | 2.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.8 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.4 | 51.2 | 52.6 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9516 | 27.8258 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.2 | 52.7 | 0 | 0 | 2.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32 | 190.1 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 24.0917 | 27.9 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.2 | 52.7 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 32.054 | 191 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9348 | 27.9 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.2 | 52.7 | 0 | 0 | 2.1 | 13 | 920 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.2
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191.1 | 30 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 195 | 59.9 | 0.0039 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8823 | 27.8823 | 51.3
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 195.3 | 59.9 | 0.0039 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8038 | 27.8038 | 51.3
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 919 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 195.6 | 59.9 | 0.0039 | 0.0036 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9485 | 27.8742 | 51.3
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 196 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9472 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0.0006 | 5.6357 | 1.5 | 51.3 | 52.8 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 196 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0.001 | 10 | 1.6 | 51.3302 | 52.9 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 196 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0.001 | 10 | 1.6 | 51.3 | 52.9 | 0 | 0 | 2.1 | 13 | 918 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 196 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0.001 | 9.5056 | 1.6 | 51.3699 | 53 | 0 | 0 | 2.1 | 13 | 918 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 30 | 5 | 195.6 | 59.9 | 0.00393 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.3014 | 23.8338 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0.0001 | 0.83261 | 1.6 | 51.3302 | 52.9 | 0 | 0 | 2.1 | 13 | 918 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191.9 | 30 | 5 | 195.2 | 59.9 | 0.004 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.8446 | 27.9 | 51.3
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.6 | 51.4 | 53 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191 | 30 | 5 | 195.2 | 59.9 | 0.004 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.6 | 51.4 | 53 | 0 | 0 | 2.1 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191 | 30 | 5 | 195.5 | 59.9 | 0.004 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.6 | 51.4 | 53 | 0 | 0 | 2.1 | 13 | 919 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190.9 | 30 | 5 | 195.9 | 59.9 | 0.004 | 0.0037 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.6 | 51.4 | 53 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.004 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.6 | 51.4 | 53 | 0 | 0 | 2.1 | 12.9 | 920 | 905 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 12.9 | 920 | 905 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 189.9 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7558 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 189.1 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8009 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0041 | 0.0038 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 920 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.00413 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 12.9 | 920 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 30 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8577 | 27.9 | 51.4
2008-11-03 | 13:49 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190.1 | 30 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8207 | 27.9 | 51.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 12.9 | 919 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191 | 30 | 5 | 196 | 59.8698 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8992 | 27.9 | 51.4
2008-11-03 | 13:50 | 0 | 0.0001 | 0.9434 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191 | 30.1 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7448 | 27.9 | 51.4
2008-11-03 | 13:50 | -0.006 | 0.0018 | 18.0961 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 12.9 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191.1 | 31 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8677 | 27.9 | 51.4
2008-11-03 | 13:50 | -0.01 | 0.002 | 20 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 24.0342 | 27.9 | 51.4
2008-11-03 | 13:50 | -0.01 | 0.0014 | 14.3643 | 1.7 | 51.4 | 53.1 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 196 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 24.0875 | 27.9 | 51.4
2008-11-03 | 13:50 | -0.01 | 0.001 | 10 | 1.7 | 51.4302 | 53.1 | 0 | 0 | 2.1 | 13 | 918 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 195.6 | 59.9 | 0.0042 | 0.0039 | 0.0027 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 24.009 | 27.9 | 51.4
2008-11-03 | 13:50 | -0.01 | 0.001 | 10 | 1.7 | 63.3388 | 65 | 0 | 0 | 2.49 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 195.2 | 59.9 | 0.0042 | 0.0039 | 0.003351 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9306 | 27.9 | 63.3
2008-11-03 | 13:50 | -0.01 | 0.001 | 10 | 1.8 | 87.2642 | 89.1 | 0.8 | 0.8141 | 3.54 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 192 | 31 | 5 | 195.2 | 59.9698 | 0.004502 | 0.004242 | 0.004672 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 27.9 | 87.3
2008-11-03 | 13:50 | -0.01 | 0.001 | 10 | 2 | 101.1304 | 103.2 | 0.2 | 0.2108 | 3.97 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 195.5 | 59.9 | 0.00523 | 0.0049 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 101.1
2008-11-03 | 13:50 | -0.01 | 0.001 | 10 | 2.2 | 101.3302 | 103.5 | 0 | 0 | 4.01 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 195.9 | 59.9 | 0.00563 | 0.00523 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 101.3
2008-11-03 | 13:50 | -0.008 | 0.0004 | 4.3643 | 2.2 | 101.3 | 103.5 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 196 | 59.9 | 0.0057 | 0.00533 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 101.3
2008-11-03 | 13:50 | 0 | 0 | 0 | 2.3 | 101.4 | 103.7 | 0 | 0 | 4.1 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 192 | 31 | 5 | 196 | 59.8302 | 0.0058 | 0.0054 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9613 | 27.9 | 101.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 2.3 | 101.4 | 103.7 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191.9 | 31 | 5 | 196 | 59.9 | 0.0058 | 0.0055 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 24 | 27.9 | 101.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 2.4 | 101.4 | 103.8 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191 | 31 | 5 | 195.7 | 59.9 | 0.0062 | 0.0058 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 24.018 | 27.9 | 101.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 2.5 | 101.5 | 104 | 0 | 0 | 4.1 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191 | 31 | 5 | 195.3 | 59.9 | 0.006331 | 0.005931 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 24.0963 | 27.9 | 101.5
2008-11-03 | 13:50 | -0.008 | 0 | 0 | 2.6 | 101.5301 | 104.1 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191 | 31 | 5 | 195 | 59.9 | 0.006521 | 0.00609 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 24.0253 | 27.9 | 101.5
2008-11-03 | 13:50 | -0.02 | 0 | 0 | 2.6 | 101.8718 | 104.5 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190.9 | 31 | 5 | 195 | 59.9 | 0.0068 | 0.00633 | 0.0054 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.8408 | 27.9 | 101.9
2008-11-03 | 13:50 | -0.019 | 0 | 0 | 2.7 | 102.6301 | 105.3 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 31 | 5 | 195 | 59.9 | 0.00683 | 0.0064 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7943 | 27.9 | 102.6
2008-11-03 | 13:50 | -0.002 | 0.0002 | 2.0833 | 2.7 | 102.7 | 105.4 | 0 | 0 | 4.1 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 31 | 5 | 195.2 | 59.9 | 0.0069 | 0.0064 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9902 | 27.9 | 102.7
2008-11-03 | 13:50 | 0.007 | 0.001 | 9.5678 | 2.7 | 102.7302 | 105.5 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 189.9 | 31 | 5 | 195.5 | 59.9 | 0.007091 | 0.006591 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9118 | 27.9 | 102.7
2008-11-03 | 13:50 | 0.01 | 0.0001 | 0.92481 | 2.9 | 102.8 | 105.7 | 0 | 0 | 4.1 | 12.9 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 189 | 31 | 5 | 195.9 | 59.9 | 0.00763 | 0.00713 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8336 | 27.9 | 102.8
2008-11-03 | 13:50 | 0.01 | 0 | 0 | 3 | 102.9 | 105.9 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 189 | 31 | 5 | 195.8 | 59.9 | 0.0077 | 0.0072 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8448 | 27.9 | 102.9
2008-11-03 | 13:50 | 0.01 | 0 | 0 | 3.1 | 103 | 106.1 | 0 | 0 | 4.1 | 12.9 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 189 | 31 | 5 | 195.4 | 59.9 | 0.00773 | 0.0072 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 23.9 | 27.9 | 103
2008-11-03 | 13:50 | 0.01 | 0 | 0 | 3.2 | 103.1 | 106.3 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 189 | 31 | 5 | 195 | 59.9 | 0.0078 | 0.00723 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 103.1
2008-11-03 | 13:50 | 0.001 | 0 | 0 | 3.2 | 103.1 | 106.3 | 0 | 0 | 4.1 | 12.9 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 189.1 | 31 | 5 | 195.4 | 59.9 | 0.0078 | 0.0073 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 27.9 | 103.1
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.3 | 103.1 | 106.4 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 31 | 5 | 195.7 | 59.9 | 0.0078 | 0.0073 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 23.9583 | 27.9 | 103.1
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.3 | 103.1 | 106.4 | 0 | 0 | 4.1 | 12.9 | 921 | 903 | 895 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 31 | 5 | 196 | 59.9 | 0.0078 | 0.0073 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9634 | 27.9 | 103.1
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.3 | 103.1 | 106.4 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190 | 31 | 5 | 196 | 59.9 | 0.0079 | 0.0073 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 27.915 | 103.1
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.3 | 103.1302 | 106.4 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 190 | 31 | 5 | 196 | 59.9 | 0.0079 | 0.00733 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 27.9933 | 103.1
2008-11-03 | 13:50 | 0 | 0.0004 | 3.9518 | 3.4 | 103.2 | 106.6 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 190.1 | 31 | 5 | 196 | 59.9 | 0.0079 | 0.0074 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9718 | 27.9282 | 103.2
2008-11-03 | 13:50 | 0 | 0.0008 | 7.9399 | 3.4 | 103.2 | 106.6 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191 | 31 | 5 | 196 | 59.9 | 0.0079 | 0.0074 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.8998 | 27.9501 | 103.2
2008-11-03 | 13:50 | 0 | 0 | 0.39967 | 3.4 | 103.2301 | 106.6 | 0 | 0 | 4.1 | 13 | 921 | 904 | 895 | 0 | 0 | 0 | 995 | -40 | 0 | 33 | 191 | 31 | 5 | 196 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.6986 | 23.8284 | 27.9716 | 103.2
2008-11-03 | 13:50 | 0 | 0.0009 | 8.7261 | 3.4 | 103.3 | 106.7 | 0 | 0 | 4.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 33 | 191.1 | 31 | 5 | 196 | 59.9 | 0.00803 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9068 | 27.9068 | 103.3
2008-11-03 | 13:50 | 0 | 0.0003 | 2.8883 | 3.4 | 103.3 | 106.7 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 33.054 | 192 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9851 | 27.9851 | 103.3
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.4 | 103.3 | 106.7 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9365 | 28 | 103.3
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.4 | 103.3302 | 106.8 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9418 | 27.9582 | 103.3
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9799 | 27.9201 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9015 | 27.9985 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 195.6 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 28 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 195.2 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8448 | 28 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 5.1 | 195.2 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8336 | 28 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4304 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191.9 | 31 | 5.9 | 195.5 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 28 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4698 | 107 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191 | 31 | 5 | 195.9 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 23.9 | 28 | 103.5
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.4 | 106.9 | 0 | 0 | 4.1 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9 | 28 | 103.4
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.5 | 107 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190.9 | 31 | 5 | 196 | 59.9 | 0.0081 | 0.0076 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9 | 28 | 103.5
2008-11-03 | 13:50 | 0 | 0 | 0 | 3.5 | 103.5 | 107 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0082 | 0.0076 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.9 | 28 | 103.5
2008-11-03 | 13:50 | 0 | 0.0006 | 5.5925 | 3.6 | 103.5 | 107.1 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0082 | 0.0076 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.9 | 28 | 103.5
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.6 | 103.6 | 107.2 | 0 | 0 | 4.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189.9 | 31 | 5 | 196 | 59.9 | 0.0082 | 0.0076 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 28 | 103.6
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.6 | 103.6 | 107.2 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0082 | 0.0076 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9 | 28 | 103.6
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.6 | 103.7 | 107.3 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0083 | 0.0077 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8614 | 28 | 103.7
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.6 | 103.7 | 107.3 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.00833 | 0.00773 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8 | 28 | 103.7
2008-11-03 | 13:50 | 0 | 0.0017 | 17.2078 | 3.7 | 103.8 | 107.5 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.002 | 20 | 3.7 | 103.8 | 107.5 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.002 | 20 | 3.7 | 103.8 | 107.5 | 0 | 0 | 4.1 | 13 | 921 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7482 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.002 | 20 | 3.7 | 103.8 | 107.5 | 0 | 0 | 4.1 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189.1 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.791 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.0019 | 19.2874 | 3.7 | 103.8 | 107.5 | 0 | 0 | 4.1 | 12.9 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0 |  | n/a |  | 33 | 60 | 1057 | 1019 | 24 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.0011 | 11.281 | 3.7 | 103.8302 | 107.5 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1019 | 24 | 28 | 103.8
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.7 | 103.9 | 107.6 | 0 | 0 | 4.1 | 12.9 | 921 | 905 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.8694 | 28.0653 | 103.9
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.7 | 103.9 | 107.6 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190.1 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018.302 | 23.8 | 28.0563 | 103.9
2008-11-03 | 13:50 | 0 | 0.001 | 10 | 3.7 | 103.9 | 107.6 | 0 | 0 | 4.1 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018.698 | 23.844 | 28 | 103.9
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 3.8 | 103.9 | 107.7 | 0 | 0 | 4.1 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.9992 | 28 | 103.9
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 3.8 | 103.9 | 107.7 | 0 | 0 | 4.1 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0055 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.8426 | 28 | 103.9
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 3.8 | 104 | 107.8 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191.1 | 31 | 5 | 196 | 59.9 | 0.0084 | 0.0078 | 0.0056 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.743 | 28.057 | 104
2008-11-03 | 13:51 | 0 | 0.0003 | 2.9521 | 3.8 | 104 | 107.8 | 0 | 0 | 4.1 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 5 | 196 | 59.9 | 0.00843 | 0.0078 | 0.0056 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7 | 28.1 | 104
2008-11-03 | 13:51 | 0 | 0 | 0 | 3.8 | 104 | 107.8 | 0 | 0 | 4.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 5.1 | 196 | 59.9 | 0.0085 | 0.0078 | 0.0056 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018 | 23.7273 | 28.0864 | 104
2008-11-03 | 13:51 | 0 | 0 | 0 | 3.8 | 104 | 107.8 | 0 | 0 | 4.2 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5.9 | 195.8 | 59.9 | 0.0085 | 0.0078 | 0.0056 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1018.3006 | 23.884 | 28.008 | 104
2008-11-03 | 13:51 | 0 | 0.0002 | 2.1707 | 3.8 | 104 | 107.8 | 0 | 0 | 4.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5.1 | 195.4 | 59.9 | 0.0085 | 0.0078 | 0.0056 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.7594 | 28.0703 | 104
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 3.8 | 110.3186 | 114.1 | 0 | 0 | 4.66 | 13.1 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5.9 | 195 | 59.9 | 0.0085 | 0.00783 | 0.005911 | 0.0001 |  | n/a |  | 33 | 60 | 1057 | 1019 | 23.7973 | 28.1 | 110.3
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.1 | 148.418 | 152.5 | 1.5 | 1.5242 | 5.72 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 195.3 | 59.9 | 0.009491 | 0.00886 | 0.007921 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9268 | 28.0732 | 148.4
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.4 | 154.7 | 159.1 | 0.2 | 0.1897 | 6.1 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 5 | 195.7 | 59.9 | 0.010121 | 0.009391 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9954 | 28 | 154.7
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.4 | 154.6302 | 159.1 | 0 | 0 | 6.1 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 5.1 | 196 | 59.9 | 0.01043 | 0.00973 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9165 | 28 | 154.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.5 | 154.7 | 159.2 | 0 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191.9 | 31 | 6 | 196 | 59.9 | 0.01053 | 0.00983 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9 | 28.0619 | 154.7
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.6 | 154.8 | 159.4 | 0 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191 | 31 | 6 | 196 | 59.9 | 0.0107 | 0.0099 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8195 | 28.1 | 154.8
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 4.8 | 154.8 | 159.6 | 0 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191 | 31 | 6 | 196 | 59.9 | 0.011181 | 0.010381 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.7 | 28.1 | 154.8
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5 | 154.8 | 159.8 | 0 | 0 | 6.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190.9 | 31 | 6 | 196 | 59.9 | 0.012021 | 0.011221 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7 | 28.1 | 154.8
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.2 | 154.9302 | 160.1 | 0 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190 | 31 | 6 | 196 | 59.9 | 0.01233 | 0.0115 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8502 | 28.1 | 154.9
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.2 | 155 | 160.2 | 0 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.9 | 28.1 | 155
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.3 | 155 | 160.3 | -0.9 | 0 | 6.2 | 12.9 | 922 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189.9 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018.302 | 23.9 | 28.1 | 155
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.3 | 155 | 160.3 | -1 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8899 | 28.1 | 155
2008-11-03 | 13:51 | 0 | 0.001 | 9.7083 | 5.3 | 155.1 | 160.4 | -1 | 0 | 6.2 | 12.9 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.0082 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8116 | 28.1 | 155.1
2008-11-03 | 13:51 | 0 | 0.0001 | 1.375 | 5.4 | 155.1302 | 160.5 | -1 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.00823 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018.302 | 23.7333 | 28.1 | 155.1
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.4 | 155.2301 | 160.6 | -1 | 0 | 6.2 | 12.9 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0124 | 0.0115 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.745 | 28.1 | 155.2
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.5 | 155.3301 | 160.8 | -1 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 5.9 | 196 | 59.8301 | 0.0125 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8234 | 28.1 | 155.3
2008-11-03 | 13:51 | 0 | 0.0004 | 3.7945 | 5.6 | 155.4 | 161 | -1 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 5 | 196 | 59.9 | 0.0125 | 0.01157 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8931 | 28.1017 | 155.4
2008-11-03 | 13:51 | 0.009 | 0.001 | 10 | 5.6 | 155.4 | 161 | -1 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189.1 | 31 | 5 | 196 | 59.9 | 0.0125 | 0.0115 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.5799 | 28.18 | 155.4
2008-11-03 | 13:51 | 0.003 | 0.001 | 10 | 5.6 | 155.4606 | 161.1 | -1 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0125 | 0.01153 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7317 | 28.2 | 155.5
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.5698 | 161.3 | -1 | 0 | 6.2 | 13 | 922 | 905 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0125 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018.302 | 23.8638 | 28.2 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6 | 161.3 | -1 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 5 | 196 | 59.9 | 0.0125 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8 | 28.185 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6 | 161.3 | -1 | 0 | 6.2 | 12.9 | 922 | 905 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190.1 | 31 | 5 | 196 | 59.9 | 0.01253 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8 | 28.1067 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6 | 161.3 | -1 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 5 | 196 | 59.9 | 0.0126 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8 | 28.1 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6 | 161.3 | -1 | 0 | 6.2 | 12.9 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 5.1 | 196 | 59.9 | 0.0126 | 0.0116 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7503 | 28.1497 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6302 | 161.3 | -1 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 6 | 196 | 59.9 | 0.0126 | 0.01163 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1018.698 | 23.7559 | 28.2 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.7 | 155.6302 | 161.4 | -1 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 191.1 | 31 | 6 | 196 | 59.9 | 0.0126 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9 | 28.1934 | 155.6
2008-11-03 | 13:51 | 0 | 0.001 | 10 | 5.8 | 155.7 | 161.5 | -1 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 6 | 196 | 59.9 | 0.0126 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9 | 28.1151 | 155.7
2008-11-03 | 13:51 | 0 | 0.0005 | 4.5516 | 5.8 | 155.8 | 161.6 | -1 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 196 | 59.9 | 0.0126 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8369 | 28.1631 | 155.8
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.8 | 155.7301 | 161.5 | 0 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 196 | 59.9 | 0.01263 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7172 | 28.2 | 155.7
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.8 | 155.7 | 161.5 | 0 | 0 | 6.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 192 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.6197 | 28.2 | 155.7
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.8 | 155.8 | 161.6 | 0 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.698 | 28.2 | 155.8
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 155.9 | 161.8 | 0 | 0 | 6.2 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8527 | 28.2 | 155.9
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156 | 161.9 | 0 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 195.4 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9 | 28.2 | 156
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156 | 161.9 | 0 | 0 | 6.2 | 13.1 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 192 | 31 | 6 | 195 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9329 | 28.2 | 156
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156 | 161.9 | 0 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191.9 | 31 | 6 | 195 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9777 | 28.2 | 156
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.1 | 162 | 0 | 0 | 6.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 6 | 195 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.822 | 28.2 | 156.1
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.1 | 162 | 0 | 0 | 6.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 191 | 31 | 6 | 195 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7323 | 28.2 | 156.1
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.2 | 162.1 | 0 | 0 | 6.2 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190.9 | 31 | 6 | 195.4 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0.0001 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8386 | 28.2 | 156.2
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.2302 | 162.1 | 0 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 6 | 195.8 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.698 | 23.9755 | 28.2 | 156.2
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | -0.2 | 0 | 6.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 190 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.9 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189.9 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.9 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8407 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.7624 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.7159 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 196 | 59.8699 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.6986 | 23.7942 | 28.2 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3302 | 162.2 | 0 | 0 | 6.2 | 12.9 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 34 | 189 | 31 | 6 | 195.7 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.698 | 23.8 | 28.2726 | 156.3
2008-11-03 | 13:51 | 0 | 0 | 0 | 5.9 | 156.3 | 162.2 | 0 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 34 | 189 | 31 | 6 | 195.3 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.3034 | 23.8 | 28.2495 | 156.3
2008-11-03 | 13:51 | 0.009 | 0 | 0 | 6 | 156.4 | 162.4 | 0 | 0 | 6.2 | 12.9 | 922 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 34.0531 | 189.1 | 31 | 6 | 195 | 59.9 | 0.0127 | 0.0117 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8859 | 28.2 | 156.4
2008-11-03 | 13:51 | 0.01 | 0 | 0 | 6 | 156.4 | 162.4 | 0 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.4 | 59.9 | 0.0127 | 0.01173 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 24.0851 | 28.2074 | 156.4
2008-11-03 | 13:51 | 0.005 | 0 | 0 | 6 | 156.4 | 162.4 | 0 | 0 | 6.2 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190 | 31 | 6 | 195.8 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.9286 | 28.2857 | 156.4
2008-11-03 | 13:51 | 0 | 0 | 0 | 6 | 156.4302 | 162.4 | 0 | 0 | 6.2 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.9 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1018.302 | 23.9 | 28.3 | 156.4
2008-11-03 | 13:51 | 0 | 0 | 0 | 6 | 156.5 | 162.5 | 0 | 0 | 6.2 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190.1 | 31 | 6 | 195.5 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8578 | 28.3 | 156.5
2008-11-03 | 13:51 | 0 | 0 | 0 | 6 | 156.5302 | 162.5 | 0 | 0 | 6.2 | 13 | 922 | 904 | 895 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 195.1 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.698 | 23.7797 | 28.3 | 156.5
2008-11-03 | 13:52 | 0 | 0 | 0 | 6 | 156.6 | 162.6 | 0 | 0 | 6.2 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 195.2 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.7013 | 28.3 | 156.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 6 | 156.6302 | 162.7 | 0 | 0 | 6.2 | 13 | 921 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.1 | 31 | 6 | 195.6 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.8541 | 28.3 | 156.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.1 | 156.7 | 162.8 | 0 | 0 | 6.2 | 13 | 920 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 23.9 | 28.3 | 156.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.1 | 156.6 | 162.7 | 0 | 0 | 6.2 | 13 | 920 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.8302 | 0.0128 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8663 | 28.3 | 156.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.2 | 156.6 | 162.8 | 0 | 0 | 6.2 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.01283 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8238 | 28.3 | 156.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.2 | 156.7 | 162.9 | 0 | 0 | 6.2 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0129 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9795 | 28.3 | 156.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.2 | 156.7 | 162.9 | 0 | 0 | 6.2 | 13 | 920 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0129 | 0.0118 | 0.0083 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9317 | 28.3 | 156.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.4 | 168.646 | 175 | 0 | 0 | 6.83 | 13 | 920 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.013111 | 0.012011 | 0.008962 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8066 | 28.3 | 168.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 6.8 | 200.0878 | 206.9 | 0.8 | 0.8324 | 7.87 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.8302 | 0.014421 | 0.013291 | 0.010621 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.7 | 28.3 | 200.1
2008-11-03 | 13:52 | 0 | 0 | 0 | 7.1 | 206.4301 | 213.6 | 0 | 0 | 8.11 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.9 | 31 | 6 | 196 | 59.9 | 0.01473 | 0.01353 | 0.0109 | 0 |  | n/a |  | 33.5425 | 60 | 1058 | 1019 | 23.7099 | 28.3 | 206.4
2008-11-03 | 13:52 | 0 | 0 | 0 | 7.3 | 206.5301 | 213.9 | 0 | 0 | 8.2 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.014981 | 0.013781 | 0.0109 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1018.6986 | 23.9446 | 28.3 | 206.5
2008-11-03 | 13:52 | 0 | 0 | 0 | 7.6 | 206.9718 | 214.6 | 0 | 0 | 8.2 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.01613 | 0.0149 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 24 | 28.3 | 207
2008-11-03 | 13:52 | -0.009 | 0 | 0 | 7.7 | 207.7 | 215.4 | 0 | 0 | 8.26 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 24.0381 | 28.3 | 207.7
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 7.7 | 207.7 | 215.4 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.9 | 31 | 6 | 196 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018 | 24.0836 | 28.3 | 207.7
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 7.7 | 207.7 | 215.4 | 0 | 0 | 8.3 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1018.302 | 24.0054 | 28.3 | 207.7
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 7.7 | 207.7302 | 215.4 | 0 | 0 | 8.3 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8543 | 28.3 | 207.7
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 7.8 | 207.7302 | 215.5 | 0 | 0 | 8.3 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.9 | 31 | 6 | 196 | 59.9 | 0.01623 | 0.015 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8 | 28.3 | 207.7
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 7.8 | 207.8302 | 215.6 | 0 | 0 | 8.3 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8 | 28.3 | 207.8
2008-11-03 | 13:52 | -0.01 | 0.0002 | 1.8552 | 7.8 | 207.9 | 215.7 | 0 | 0 | 8.3 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.7923 | 28.3 | 207.9
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 7.8 | 207.9 | 215.7 | 0 | 0 | 8.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.7141 | 28.3 | 207.9
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 7.8 | 207.9304 | 215.8 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.7642 | 28.3 | 207.9
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 7.9 | 208 | 215.9 | 0 | 0 | 8.3 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.1 | 31 | 6 | 196 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.8422 | 28.3 | 208
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 7.9 | 208.0302 | 215.9 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.01633 | 0.015 | 0.011 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.9 | 28.3 | 208
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 7.9 | 208.1302 | 216 | 0 | 0 | 8.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.9 | 28.3 | 208.1
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 8 | 208.2 | 216.2 | 0 | 0 | 8.3 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.6687 | 28.3 | 208.2
2008-11-03 | 13:52 | -0.01 | 0.001 | 10 | 8 | 208.3 | 216.3 | 0 | 0 | 8.3 | 12.9 | 923 | 904 | 895 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.1 | 31 | 6 | 196 | 59.8699 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.7108 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.005 | 0.001 | 10 | 8 | 208.3 | 216.3 | 0 | 0 | 8.3 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018.3014 | 23.7663 | 28.3 | 208.3
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8 | 208.3302 | 216.3 | 0 | 0 | 8.3 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.01643 | 0.0151 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 23.6762 | 28.3 | 208.3
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8 | 208.3 | 216.3 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5208 | 28.3 | 208.3
2008-11-03 | 13:52 | 0 | 0.0006 | 6.3017 | 8.1 | 208.3 | 216.4 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.1 | 31 | 6 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.773 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.008 | 0 | 0 | 8.1 | 208.3 | 216.4 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.7131 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.1 | 208.3 | 216.4 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.01653 | 0.0152 | 0.011 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.5 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.1 | 208.3302 | 216.4 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0166 | 0.0152 | 0.011 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1018.698 | 23.5 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.1 | 208.3 | 216.4 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0166 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5 | 28.3 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.2 | 208.3 | 216.5 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.4 | 59.9 | 0.0166 | 0.0153 | 0.011 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.6789 | 28.3596 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.2 | 208.3302 | 216.5 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0166 | 0.0153 | 0.011 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.6482 | 28.4 | 208.3
2008-11-03 | 13:52 | -0.01 | 0 | 0 | 8.2 | 208.4 | 216.6 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.8698 | 0.01663 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4324 | 28.4 | 208.4
2008-11-03 | 13:52 | -0.003 | 0 | 0 | 8.2 | 208.4 | 216.6 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0167 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5889 | 28.4 | 208.4
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.4301 | 216.6 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.9 | 31 | 6 | 195 | 59.9 | 0.0167 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.6 | 28.4 | 208.4
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.4302 | 216.6 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195.4 | 59.9 | 0.0167 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.6 | 28.4 | 208.4
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.5 | 216.7 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 195.8 | 59.9 | 0.0167 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.6 | 28.4 | 208.5
2008-11-03 | 13:52 | 0 | 0.0004 | 3.7034 | 8.2 | 208.5 | 216.7 | 0 | 0 | 8.3 | 13 | 921 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.9 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.011 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5926 | 28.3926 | 208.5
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8.2 | 208.5 | 216.7 | 0 | 0 | 8.3 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5144 | 28.3144 | 208.5
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8.2 | 208.4698 | 216.7 | 0 | 0 | 8.3 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5 | 28.3638 | 208.5
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8.2 | 208.4301 | 216.6 | 0 | 0 | 8.3 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.458 | 28.4 | 208.4
2008-11-03 | 13:52 | 0 | 0.001 | 10 | 8.2 | 208.5 | 216.7 | 0 | 0 | 8.3 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4 | 28.4 | 208.5
2008-11-03 | 13:52 | 0 | 0.0005 | 4.5288 | 8.2 | 208.6 | 216.8 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.9 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0153 | 0.0111 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.4 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.6 | 216.8 | 0 | 0 | 8.3 | 12.9 | 923 | 905 | 892 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 189.1 | 31 | 6 | 195.7 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.5537 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.6 | 216.8 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.3 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.3797 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.6 | 216.8 | 0 | 0 | 8.3 | 12.9 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 33.457 | 60 | 1058 | 1019 | 23.3 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.6 | 216.8 | 0 | 0 | 8.3 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.4 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.4885 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.2 | 208.6302 | 216.9 | 0 | 0 | 8.3 | 12.9 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.8 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4107 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.3 | 208.6302 | 216.9 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.1 | 31 | 6 | 196 | 59.9 | 0.0167 | 0.0154 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4 | 28.4 | 208.6
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.3 | 208.7 | 217 | 0 | 0 | 8.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0168 | 0.0154 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.3075 | 28.4 | 208.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.3 | 208.7 | 217 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0168 | 0.0154 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.2735 | 28.4 | 208.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.4 | 208.7 | 217.1 | 0 | 0 | 8.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0168 | 0.01543 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5 | 28.4 | 208.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.4 | 208.7303 | 217.1 | 0 | 0 | 8.3 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191.1 | 31 | 6 | 196 | 59.9 | 0.01683 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.5 | 28.4 | 208.7
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.4 | 208.7698 | 217.2 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4408 | 28.4 | 208.8
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.4 | 208.8 | 217.2 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 195.7 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.3251 | 28.4 | 208.8
2008-11-03 | 13:52 | 0 | 0 | 0 | 8.4 | 208.8 | 217.2 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.3 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.2156 | 28.4 | 208.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 208.8 | 217.2 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 33.4575 | 60 | 1058 | 1019 | 23.2939 | 28.4 | 208.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 208.8 | 217.2 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 33 | 60 | 1058 | 1019 | 23.2279 | 28.4 | 208.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 208.8 | 217.2 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 33.543 | 60 | 1058 | 1019 | 23.2 | 28.4 | 208.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 208.9 | 217.3 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.2 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.2561 | 28.4 | 208.9
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 208.9 | 217.3 | 0 | 0 | 8.3 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 195.5 | 59.9 | 0.0169 | 0.0155 | 0.0111 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.4 | 28.4 | 208.9
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.4 | 209.957 | 218.4 | 0 | 0 | 8.3 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.9 | 31 | 6 | 195.8 | 59.9 | 0.0169 | 0.0155 | 0.01116 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.4 | 28.4 | 210
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.5 | 225.0824 | 233.6 | 0 | 0 | 8.57 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.01693 | 0.01553 | 0.011951 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.2735 | 28.4633 | 225.1
2008-11-03 | 13:53 | 0 | 0 | 0 | 8.6 | 247.6878 | 256.3 | 0.7 | 0.7469 | 9.74 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.017121 | 0.015691 | 0.012662 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.2829 | 28.4585 | 247.7
2008-11-03 | 13:53 | 0 | 0 | 0 | 9.1 | 258.9 | 268 | 0 | 0 | 10.15 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.9 | 31 | 6 | 196 | 59.9 | 0.018091 | 0.016591 | 0.0137 | 0 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.3413 | 28.4196 | 258.9
2008-11-03 | 13:53 | 0 | 0 | 0 | 9.5 | 259.7698 | 269.3 | 0 | 0 | 10.3 | 13 | 923 | 904 | 895 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0193 | 0.01773 | 0.0138 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 23.1082 | 28.4973 | 259.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 9.6 | 259.8 | 269.4 | 0 | 0 | 10.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0194 | 0.0179 | 0.0138 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018 | 23.0241 | 28.5 | 259.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 9.7 | 259.8302 | 269.5 | 0 | 0 | 10.3 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0195 | 0.0179 | 0.0138 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 23 | 28.5 | 259.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 9.7 | 260.8 | 270.5 | 0 | 0 | 10.3 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.9 | 31 | 6 | 195.6 | 59.9 | 0.0195 | 0.0179 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0976 | 28.4675 | 260.8
2008-11-03 | 13:53 | 0.007 | 0 | 0 | 9.7 | 260.8 | 270.5 | 0 | 0 | 10.3 | 12.9 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 195.2 | 59.9 | 0.0196 | 0.018 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.2679 | 28.4107 | 260.8
2008-11-03 | 13:53 | 0.01 | 0 | 0 | 9.7 | 260.8 | 270.5 | 0 | 0 | 10.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 195 | 59.9 | 0.019721 | 0.01809 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0337 | 28.4888 | 260.8
2008-11-03 | 13:53 | 0.01 | 0 | 0 | 9.9 | 260.8302 | 270.7 | 0 | 0 | 10.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.1 | 31 | 6 | 195 | 59.9 | 0.02 | 0.0184 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 260.8
2008-11-03 | 13:53 | 0.008 | 0 | 0 | 9.9 | 260.8 | 270.7 | 0 | 0 | 10.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195 | 59.9 | 0.0201 | 0.01843 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1817 | 28.5 | 260.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 10 | 260.8301 | 270.8 | 0 | 0 | 10.3 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.2 | 59.9 | 0.0201 | 0.0185 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.3528 | 28.5 | 260.8
2008-11-03 | 13:53 | 0 | 0 | 0 | 10 | 261 | 271 | 0 | 0 | 10.3 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.6 | 59.9 | 0.0202 | 0.0185 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.2036 | 28.4982 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10 | 261 | 271 | 0 | 0 | 10.3 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0202 | 0.0185 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.6994 | 23.3601 | 28.42 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10 | 261 | 271 | 0 | 0 | 10.35 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190.1 | 31 | 6 | 195.7 | 59.9 | 0.0202 | 0.0185 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4579 | 28.4579 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.1 | 261 | 271.1 | 0 | 0 | 10.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195.3 | 59.9 | 0.02023 | 0.0186 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.4278 | 28.4639 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.2 | 261 | 271.2 | 0 | 0 | 10.4 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195 | 59.9 | 0.0203 | 0.0186 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.2708 | 28.4146 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.3 | 261 | 271.3 | 0 | 0 | 10.4 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.1 | 31 | 6 | 195 | 59.9 | 0.0203 | 0.01863 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1145 | 28.4927 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.3 | 261 | 271.3 | 0 | 0 | 10.4 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.8698 | 0.0204 | 0.0187 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1 | 28.5 | 261
2008-11-03 | 13:53 | 0.007 | 0 | 0 | 10.3 | 261 | 271.3 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.8302 | 0.02043 | 0.01873 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1492 | 28.5 | 261
2008-11-03 | 13:53 | 0.005 | 0 | 0 | 10.3 | 261 | 271.3 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1453 | 28.5 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.3 | 261 | 271.3 | 0 | 0 | 10.4 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192.1 | 31 | 6 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0053 | 28.5 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.3 | 260.9698 | 271.3 | 0 | 0 | 10.4 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 193 | 31 | 6 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 23.0838 | 28.5 | 261
2008-11-03 | 13:53 | 0 | 0.0004 | 3.5679 | 10.4 | 261 | 271.4 | 0 | 0 | 10.4 | 13 | 921 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192.9 | 31 | 6 | 195 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 23.0379 | 28.5 | 261
2008-11-03 | 13:53 | 0 | 0.0008 | 8.1751 | 10.4 | 261 | 271.4 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 23 | 28.5 | 261
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.4 | 261.1302 | 271.5 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.3 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0184 | 28.5 | 261.1
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.4 | 261.2 | 271.6 | 0 | 0 | 10.4 | 13.1 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195.7 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0966 | 28.5 | 261.2
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.4 | 261.2301 | 271.6 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 196 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1 | 28.5 | 261.2
2008-11-03 | 13:53 | -0.002 | 0.0005 | 5.0042 | 10.4 | 261.3 | 271.7 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.9 | 31 | 6 | 196 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.047 | 28.5 | 261.3
2008-11-03 | 13:53 | -0.01 | 0.001 | 10 | 10.4 | 261.2699 | 271.7 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0206 | 0.0189 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.3
2008-11-03 | 13:53 | -0.01 | 0.001 | 10 | 10.4 | 261.3 | 271.7 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 191 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.01893 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9906 | 28.5 | 261.3
2008-11-03 | 13:53 | -0.01 | 0.001 | 10 | 10.4 | 261.3 | 271.7 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 190.9 | 31 | 6 | 196 | 59.8698 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9124 | 28.5 | 261.3
2008-11-03 | 13:53 | -0.005 | 0.001 | 10 | 10.4 | 261.3 | 271.7 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9 | 28.5 | 261.3
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.3 | 271.7 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9 | 28.5 | 261.3
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.3302 | 271.7 | 0 | 0 | 10.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 35 | 189.9 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9444 | 28.5 | 261.3
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.4302 | 271.8 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 23.1 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.5 | 271.9 | 0 | 0 | 10.4 | 12.9 | 923 | 904 | 895 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 23.1 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.5 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.1 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0.001 | 10 | 10.4 | 261.5 | 271.9 | 0 | 0 | 10.4 | 12.9 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189 | 31 | 6 | 196 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23.0653 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0.0003 | 3.0823 | 10.4 | 261.5 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 189.1 | 31 | 6 | 195.6 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.4 | 261.4698 | 271.9 | 0 | 0 | 10.4 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195.2 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.4 | 261.4302 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195 | 59.9 | 0.0207 | 0.019 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4 | 271.9 | 0 | 0 | 10.4 | 13 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 190.1 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.974 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.5 | 272 | 0 | 0 | 10.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9041 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4301 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195 | 59.8699 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.6994 | 22.9819 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.5 | 272 | 0 | 0 | 10.4 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 23 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4302 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 191.1 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 23 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4698 | 272 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 23 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.4 | 271.9 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35 | 192 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 23 | 28.5 | 261.4
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.5 | 272 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35.054 | 192 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8538 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.5 | 272 | 0 | 0 | 10.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35.946 | 192 | 31 | 6 | 195 | 59.9 | 0.0208 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.8 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.5 | 261.5302 | 272 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 35.054 | 192 | 31 | 6 | 195 | 59.9 | 0.02083 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8591 | 28.5 | 261.5
2008-11-03 | 13:53 | 0 | 0 | 0 | 10.6 | 261.6 | 272.2 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9923 | 28.5 | 261.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.6 | 272.2 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9141 | 28.5 | 261.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.6 | 272.2 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8359 | 28.5 | 261.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.6302 | 272.2 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8422 | 28.5 | 261.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.6698 | 272.3 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8797 | 28.5 | 261.7
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.6698 | 272.3 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 36 | 191 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8014 | 28.5 | 261.7
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.6 | 261.8762 | 272.5 | 0 | 0 | 10.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 31 | 6 | 195 | 59.9 | 0.0209 | 0.0191 | 0.013931 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8767 | 28.5 | 261.9
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.7 | 287.6146 | 298.3 | 0 | 0 | 10.86 | 13 | 922 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.021291 | 0.019491 | 0.015281 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9549 | 28.5 | 287.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 10.9 | 300.3784 | 311.3 | 0 | 0 | 11.89 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.02171 | 0.01991 | 0.01585 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9339 | 28.5 | 300.4
2008-11-03 | 13:54 | 0 | 0.0017 | 17.1034 | 11.2 | 310.3411 | 321.5 | 0 | 0 | 12.26 | 12.9 | 924 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.0226 | 0.0207 | 0.0165 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8112 | 28.5 | 310.3
2008-11-03 | 13:54 | 0 | 0.002 | 20 | 11.2 | 311.2718 | 322.5 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 31 | 6 | 195 | 59.9 | 0.02263 | 0.0208 | 0.01653 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.889 | 28.5 | 311.3
2008-11-03 | 13:54 | 0 | 0.0015 | 15.075 | 11.5 | 311.9609 | 323.5 | 0 | 0 | 12.4 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195 | 59.9 | 0.023691 | 0.021791 | 0.0166 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8325 | 28.5 | 312
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 11.7 | 312.1 | 323.8 | 0 | 0 | 12.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195 | 59.8698 | 0.024 | 0.0221 | 0.0166 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7542 | 28.5 | 312.1
2008-11-03 | 13:54 | 0 | 0.0008 | 8.1555 | 11.8 | 312.1302 | 323.9 | 0 | 0 | 12.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.2 | 59.9 | 0.0241 | 0.0221 | 0.0166 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.748 | 28.5 | 312.1
2008-11-03 | 13:54 | 0 | 0 | 0 | 11.8 | 312.2302 | 324 | 0 | 0 | 12.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.6 | 59.9 | 0.0241 | 0.02213 | 0.0166 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8958 | 28.5 | 312.2
2008-11-03 | 13:54 | 0 | 0 | 0 | 11.8 | 312.8738 | 324.7 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 31 | 6 | 196 | 59.9 | 0.02413 | 0.02223 | 0.01663 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7394 | 28.5 | 312.9
2008-11-03 | 13:54 | 0 | 0 | 0 | 11.9 | 314.2 | 326.1 | 0 | 0 | 12.4 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.7 | 59.9 | 0.02423 | 0.0223 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8168 | 28.5 | 314.2
2008-11-03 | 13:54 | 0 | 0.0005 | 5.0042 | 11.9 | 314.2 | 326.1 | 0 | 0 | 12.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.3 | 59.9 | 0.0243 | 0.022331 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.9 | 28.5365 | 314.2
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12 | 314.2 | 326.2 | 0 | 0 | 12.4 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.0244 | 0.0224 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8852 | 28.6 | 314.2
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12 | 314.1 | 326.1 | 0 | 0 | 12.4 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.4 | 59.9 | 0.02443 | 0.0225 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.8071 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.1 | 314.0698 | 326.2 | 0 | 0 | 12.4 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 31 | 6 | 195.8 | 59.9 | 0.0245 | 0.0225 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.7289 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.1 | 314.1 | 326.2 | 0 | 0 | 12.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.02453 | 0.0226 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7493 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.1 | 314.1 | 326.2 | 0 | 0 | 12.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0246 | 0.0226 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019.3026 | 22.7727 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.1 | 314.1 | 326.2 | 0 | 0 | 12.4 | 13 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 31 | 6 | 196 | 59.9 | 0.0246 | 0.0226 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.7 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.2 | 314.1 | 326.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0247 | 0.0227 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1020 | 22.7 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.2 | 314.1 | 326.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.4 | 59.9 | 0.0247 | 0.0227 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019.6986 | 22.8858 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.2 | 314.1 | 326.3 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0247 | 0.0227 | 0.0167 | 0 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7996 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.2 | 314.1 | 326.3 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.02473 | 0.02273 | 0.0167 | 3e-05 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5364 | 28.6 | 314.1
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.3 | 314.2 | 326.5 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192.1 | 31 | 6 | 195 | 59.9 | 0.0248 | 0.0228 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6928 | 28.6 | 314.2
2008-11-03 | 13:54 | 0 | 0.0005 | 4.8957 | 12.3 | 314.2302 | 326.5 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192.9 | 31 | 6 | 195 | 59.9 | 0.02483 | 0.02283 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7745 | 28.6 | 314.2
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.3 | 314.3 | 326.6 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.4 | 59.9 | 0.0249 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7473 | 28.6 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.3 | 314.3604 | 326.7 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0249 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 22.7 | 28.6 | 314.4
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.3 | 314.4698 | 326.8 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0249 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 22.7091 | 28.6 | 314.5
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.4 | 314.4 | 326.8 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.5 | 59.9 | 0.0249 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7872 | 28.6 | 314.4
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.4 | 314.4 | 326.8 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 31 | 6 | 195.1 | 59.9 | 0.0249 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6039 | 28.6 | 314.4
2008-11-03 | 13:54 | 0 | 0.001 | 10.1648 | 12.4 | 314.3 | 326.7 | 0 | 0 | 12.4 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 195 | 59.9 | 0.02493 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.6 | 314.3
2008-11-03 | 13:54 | 0 | 0.002 | 20 | 12.4 | 314.3 | 326.7 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 195 | 59.9 | 0.025 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.6 | 314.3
2008-11-03 | 13:54 | 0 | 0.002 | 20 | 12.4 | 314.3 | 326.7 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 31 | 6 | 195 | 59.9 | 0.025 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.6 | 314.3
2008-11-03 | 13:54 | 0 | 0.002 | 20 | 12.4 | 314.3301 | 326.7 | 0 | 0 | 12.4 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.3 | 59.9 | 0.025 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.678 | 314.3
2008-11-03 | 13:54 | 0 | 0.0012 | 11.7813 | 12.5 | 314.3 | 326.8 | 0 | 0 | 12.4 | 13 | 922 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.7 | 59.9 | 0.025 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7 | 314.3
2008-11-03 | 13:54 | 0.003 | 0.001 | 10 | 12.5 | 314.3 | 326.8 | 0 | 0 | 12.4 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.025 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7 | 314.3
2008-11-03 | 13:54 | 0.008 | 0.001 | 10 | 12.5 | 314.3301 | 326.8 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 996 | -40 | 0 | 36 | 190 | 31 | 6 | 195.6 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5122 | 28.7 | 314.3
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.5 | 314.3698 | 326.9 | 0 | 0 | 12.4 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 31 | 6 | 195.2 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5905 | 28.7 | 314.4
2008-11-03 | 13:54 | 0 | 0.0008 | 8.1951 | 12.5 | 314.3 | 326.8 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 31 | 6 | 195.2 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7375 | 28.6313 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0.065041 | 12.5 | 314.3 | 326.8 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.5 | 59.8697 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7063 | 28.6469 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.5 | 314.3698 | 326.9 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.9 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.625 | 28.7 | 314.4
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.3 | 326.9 | 0 | 0 | 12.4 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6969 | 28.7 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.3 | 326.9 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.023 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.619 | 28.7 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.3302 | 326.9 | 0 | 0 | 12.4 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.02303 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6592 | 28.7 | 314.3
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.4 | 327 | 0 | 0 | 12.4 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.0231 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.625 | 28.7 | 314.4
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.4302 | 327 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.0231 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7 | 314.4
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.5 | 327.1 | 0 | 0 | 12.4 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0251 | 0.0231 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7 | 314.5
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.6 | 314.6 | 327.2 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 31 | 6 | 196 | 59.9 | 0.0252 | 0.02313 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22.6441 | 28.7 | 314.6
2008-11-03 | 13:54 | 0 | 0 | 0 | 12.7 | 314.5 | 327.2 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.7 | 28.7 | 314.5
2008-11-03 | 13:54 | 0 | 0.0003 | 3.4761 | 12.7 | 314.5699 | 327.3 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.8 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.3014 | 22.7 | 28.7283 | 314.6
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.7 | 314.5 | 327.2 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.5 | 59.9302 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6936 | 28.7936 | 314.5
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.7 | 314.5 | 327.2 | 0 | 0 | 12.4 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.1 | 59.9699 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6155 | 28.7155 | 314.5
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.7 | 314.5 | 327.2 | 0 | 0 | 12.4 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.2 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6 | 28.7 | 314.5
2008-11-03 | 13:54 | 0 | 0.001 | 10 | 12.7 | 314.5 | 327.2 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.6 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6 | 28.7 | 314.5
2008-11-03 | 13:54 | 0 | 0.0005 | 5.1137 | 12.8 | 314.5699 | 327.4 | 0 | 0 | 12.4 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.3006 | 22.5811 | 28.7 | 314.6
2008-11-03 | 13:54 | 0 | -0.0003 | -3.359 | 12.8 | 314.5 | 327.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5031 | 28.7 | 314.5
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 12.8 | 314.5 | 327.3 | 0 | 0 | 12.4 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7 | 314.5
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 12.8 | 314.5302 | 327.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0232 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6066 | 28.7 | 314.5
2008-11-03 | 13:55 | 0 | -0.0002 | -1.8287 | 12.8 | 314.5698 | 327.4 | 0 | 0 | 12.4 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0253 | 0.0233 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.637 | 28.7 | 314.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 12.8 | 314.5 | 327.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 31 | 6 | 196 | 59.9 | 0.0254 | 0.0233 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5096 | 28.7 | 314.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 12.8 | 314.4698 | 327.3 | 0 | 0 | 12.4 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0254 | 0.0233 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5877 | 28.7 | 314.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 12.8 | 314.4 | 327.2 | 0 | 0 | 12.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0254 | 0.0233 | 0.0167 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5344 | 28.7 | 314.4
2008-11-03 | 13:55 | 0 | 0 | 0 | 12.8 | 319.6246 | 332.5 | 0 | 0 | 12.45 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0254 | 0.0233 | 0.016972 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7437 | 319.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 12.9 | 335.0522 | 348 | 0 | 0 | 13.47 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 31 | 6 | 196 | 59.9 | 0.02543 | 0.02333 | 0.017781 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7779 | 335.1
2008-11-03 | 13:55 | 0 | 0.0008 | 8.1417 | 13.3 | 360.9012 | 374.2 | 0 | 0 | 14.2 | 13 | 924 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.02659 | 0.02439 | 0.01916 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7002 | 360.9
2008-11-03 | 13:55 | 0 | 0.001 | 10 | 13.4 | 363.4 | 376.8 | 0 | 0 | 14.4 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.0269 | 0.0247 | 0.0193 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.7784 | 363.4
2008-11-03 | 13:55 | 0 | 0.001 | 10 | 13.5 | 363.4604 | 377 | 0 | 0 | 14.5 | 13 | 924 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 31 | 6 | 196 | 59.9 | 0.027021 | 0.024821 | 0.0193 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 363.5
2008-11-03 | 13:55 | 0 | 0.001 | 10 | 13.8 | 364.4302 | 378.2 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.027691 | 0.025491 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.4
2008-11-03 | 13:55 | 0 | 0.001 | 10 | 14 | 364.6699 | 378.7 | 0 | 0 | 14.5 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.02833 | 0.02603 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1018.3014 | 22.5 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0.001 | 10 | 14.1 | 364.6 | 378.7 | 0 | 0 | 14.5 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0284 | 0.0261 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.6
2008-11-03 | 13:55 | 0 | 0.0002 | 1.5503 | 14.1 | 364.7 | 378.8 | 0 | 0 | 14.5 | 12.9 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0285 | 0.0262 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.2 | 364.6 | 378.8 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 31 | 6 | 195.8 | 59.9 | 0.0289 | 0.0266 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4528 | 28.8 | 364.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.2 | 364.5698 | 378.8 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 195.5 | 59.9 | 0.02893 | 0.0266 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.8 | 364.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.3 | 364.4698 | 378.8 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 195.1 | 59.9 | 0.029 | 0.0267 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.4068 | 28.8 | 364.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.4 | 364.5302 | 378.9 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 31 | 6 | 195.2 | 59.9 | 0.02903 | 0.0267 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.5621 | 28.8 | 364.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.4 | 364.6 | 379 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.6 | 59.9 | 0.0291 | 0.0267 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.6 | 28.8 | 364.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.5 | 364.6 | 379.1 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0291 | 0.02673 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5623 | 28.7623 | 364.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.5 | 364.6 | 379.1 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0292 | 0.0268 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4841 | 28.7159 | 364.6
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.8 | 379.4 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0292 | 0.0268 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4059 | 28.7941 | 364.8
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.7698 | 379.4 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0292 | 0.0269 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4721 | 28.8 | 364.8
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.7 | 379.3 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.6 | 59.9 | 0.0292 | 0.0269 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.7 | 379.3 | 0 | 0 | 14.5 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.2 | 59.9 | 0.0292 | 0.0269 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4717 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.7302 | 379.3 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.2 | 59.9 | 0.0293 | 0.0269 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4065 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.6 | 364.7698 | 379.4 | 0 | 0 | 14.5 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.5 | 59.9 | 0.0293 | 0.0269 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.4846 | 28.8 | 364.8
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.7 | 364.7 | 379.4 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.9 | 59.9 | 0.02933 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 22.5 | 28.8 | 364.7
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.7 | 364.9 | 379.6 | 0 | 0 | 14.5 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.9
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.7 | 364.9 | 379.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.9
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.7 | 364.9 | 379.6 | 0 | 0 | 14.5 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.9
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.7 | 364.9 | 379.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 364.9
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.8 | 365.1 | 379.9 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.8 | 365.1
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 365.0698 | 380 | 0 | 0 | 14.5 | 12.9 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4686 | 28.8 | 365.1
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 365 | 379.9 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 31 | 6 | 195.6 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3905 | 28.8 | 365
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 364.9698 | 379.9 | 0 | 0 | 14.5 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.2 | 59.9 | 0.0294 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.3123 | 28.8 | 365
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 364.9604 | 379.9 | 0 | 0 | 14.5 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.2 | 59.9 | 0.02943 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.4311 | 28.8 | 365
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 365.0696 | 380 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.5 | 59.9 | 0.0295 | 0.027 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.3044 | 22.5 | 28.8 | 365.1
2008-11-03 | 13:55 | 0 | 0 | 0 | 14.9 | 365.1 | 380 | 0 | 0 | 14.5 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.9 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 22.5 | 28.8 | 365.1
2008-11-03 | 13:55 | 0 | -0.0005 | -4.6833 | 14.9 | 365.2 | 380.1 | 0 | 0 | 14.5 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.8 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4998 | 28.8001 | 365.2
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 14.9 | 365.3 | 380.2 | 0 | 0 | 14.5 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.4 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3437 | 28.8781 | 365.3
2008-11-03 | 13:55 | 0 | -0.0009 | -8.608 | 14.9 | 365.3 | 380.2 | 0 | 0 | 14.5 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 31 | 6 | 195 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4125 | 28.9 | 365.3
2008-11-03 | 13:55 | 0 | 0 | -0.068318 | 14.9 | 365.2698 | 380.2 | 0 | 0 | 14.5 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.4 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.5 | 28.9 | 365.3
2008-11-03 | 13:55 | 0 | -0.0008 | -8.0848 | 14.9 | 365.3 | 380.2 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195.7 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4878 | 28.8878 | 365.3
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 14.9 | 365.3 | 380.2 | 0 | 0 | 14.5 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4094 | 28.8094 | 365.3
2008-11-03 | 13:55 | 0 | -0.0005 | -5.3184 | 15 | 365.4 | 380.4 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 31 | 6 | 196 | 59.9 | 0.0295 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.8687 | 365.4
2008-11-03 | 13:55 | 0 | -0.0003 | -2.975 | 15 | 365.3699 | 380.4 | 0 | 0 | 14.5 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.02953 | 0.0271 | 0.0194 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4468 | 28.8532 | 365.4
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.2698 | 380.4 | 0 | 0 | 14.5 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.4751 | 28.8 | 365.3
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.2303 | 380.3 | 0 | 0 | 14.5 | 13 | 922 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.6974 | 22.403 | 28.8 | 365.2
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.3 | 380.4 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 31 | 6 | 196 | 59.9 | 0.0295 | 0.0271 | 0.0195 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4809 | 28.8 | 365.3
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.2 | 380.3 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.8 | 59.9 | 0.02953 | 0.0271 | 0.0195 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1020 | 22.5 | 28.8 | 365.2
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.2302 | 380.3 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.4 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0001 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.5 | 28.8 | 365.2
2008-11-03 | 13:55 | 0 | -0.001 | -10 | 15.1 | 365.3604 | 380.5 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.00013 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4846 | 28.8154 | 365.4
2008-11-03 | 13:55 | 0 | -0.0009 | -8.6489 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.3 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4064 | 28.8936 | 365.5
2008-11-03 | 13:55 | 0 | 0 | -0.30859 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.7 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.8283 | 365.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.8497 | 365.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4278 | 28.9 | 365.5
2008-11-03 | 13:55 | 0 | 0 | 0 | 15.1 | 365.5698 | 380.7 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4942 | 28.8942 | 365.6
2008-11-03 | 13:55 | 0 | -0.0003 | -3.2749 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.416 | 28.816 | 365.5
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0295 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.8 | 365.5
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.5 | 380.6 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.0295 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3598 | 28.8402 | 365.5
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.7301 | 380.8 | 0 | 0 | 14.5 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 31 | 6 | 196 | 59.9699 | 0.0295 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1018.6994 | 22.3 | 28.9 | 365.7
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.6397 | 380.7 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3 | 28.9 | 365.6
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.7 | 380.8 | 0 | 0 | 14.5 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3744 | 28.9 | 365.7
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 365.6 | 380.7 | 0 | 0 | 14.5 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 31 | 6 | 196 | 59.9 | 0.0296 | 0.0271 | 0.0195 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 365.6
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.1 | 379.0428 | 394.2 | 0 | 0 | 14.78 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.029691 | 0.02726 | 0.020181 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 379
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.2 | 390.775 | 406 | 0 | 0 | 15.99 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.02993 | 0.02743 | 0.020811 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3912 | 28.9 | 390.8
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.6 | 411.3818 | 427 | 0 | 0 | 16.44 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.031121 | 0.028621 | 0.021891 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.3131 | 28.9 | 411.4
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.8 | 415.6 | 431.4 | 0 | 0 | 16.5 | 12.9 | 924 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 196 | 59.9 | 0.0318 | 0.0292 | 0.0221 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.3649 | 28.9 | 415.6
2008-11-03 | 13:56 | 0 | -0.001 | -10 | 15.9 | 415.8 | 431.7 | 0 | 0 | 16.6 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.8 | 59.9 | 0.0319 | 0.02933 | 0.0221 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.357 | 28.9 | 415.8
2008-11-03 | 13:56 | 0 | -0.0004 | -3.6349 | 16 | 415.8 | 431.8 | 0 | 0 | 16.6 | 12.9 | 924 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 31 | 6 | 195.5 | 59.9 | 0.032391 | 0.029791 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.321 | 28.9 | 415.8
2008-11-03 | 13:56 | -0.016 | 0.001 | 10 | 16.1 | 415.8 | 431.9 | 0 | 0 | 16.6 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 31 | 6 | 195.1 | 59.9 | 0.0327 | 0.0301 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3992 | 28.9 | 415.8
2008-11-03 | 13:56 | -0.014 | 0.001 | 10 | 16.2 | 415.9 | 432.1 | 0 | 0 | 16.6 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.0327 | 0.0301 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3227 | 28.9 | 415.9
2008-11-03 | 13:56 | -0.01 | 0.001 | 10 | 16.2 | 416 | 432.2 | 0 | 0 | 16.6 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9698 | 0.0328 | 0.0302 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3554 | 28.9 | 416
2008-11-03 | 13:56 | -0.01 | 0.001 | 10 | 16.3 | 416.0698 | 432.4 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 31 | 6 | 195 | 59.9 | 0.0329 | 0.0302 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.1
2008-11-03 | 13:56 | -0.01 | 0.0005 | 5.1607 | 16.4 | 416.0302 | 432.4 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 31 | 6 | 195.3 | 59.9 | 0.0329 | 0.0303 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.5 | 416.1 | 432.6 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 195.7 | 59.9 | 0.032931 | 0.0303 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.1
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.5 | 416.1301 | 432.7 | 0 | 0 | 16.6 | 12.9 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 31 | 6 | 196 | 59.9 | 0.03303 | 0.03033 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4677 | 28.9 | 416.1
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.6 | 416.2 | 432.8 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 31.1 | 6 | 196 | 59.9 | 0.0331 | 0.0304 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4541 | 28.9 | 416.2
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.6 | 416.1 | 432.7 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.9 | 6 | 196 | 59.9 | 0.0331 | 0.0304 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.1
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.6 | 416.3 | 432.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 196 | 59.9 | 0.03313 | 0.03043 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.3
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.7 | 416.4 | 433.1 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.1 | 6 | 196 | 59.9 | 0.0332 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.4
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.7 | 416.5302 | 433.2 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 196 | 59.9 | 0.0332 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.5
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.7 | 416.6 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.9 | 6 | 196 | 59.9 | 0.0332 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.6
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.7 | 416.6 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.1 | 6 | 196 | 59.9 | 0.0332 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3857 | 28.9 | 416.6
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 16.7 | 416.6 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.9 | 6 | 196 | 59.9 | 0.03323 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3075 | 28.9 | 416.6
2008-11-03 | 13:56 | -0.007 | 0 | 0 | 16.7 | 416.6 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31 | 6 | 195.7 | 59.9 | 0.0333 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3706 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.5302 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 31.1 | 6 | 195.3 | 59.9 | 0.0333 | 0.0305 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3514 | 28.9 | 416.5
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.6 | 433.4 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.0333 | 0.0306 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3266 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.5697 | 433.4 | 0 | 0 | 16.6 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 32 | 6 | 195 | 59.9 | 0.0333 | 0.0306 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4044 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.5 | 433.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195 | 59.9 | 0.0333 | 0.0306 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4828 | 28.9 | 416.5
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.7 | 433.5 | 0 | 0 | 16.6 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.2 | 59.9 | 0.0333 | 0.0306 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4391 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.7302 | 433.5 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 32 | 6 | 195.5 | 59.9 | 0.0333 | 0.0306 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.7699 | 433.6 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195.9 | 59.9 | 0.0334 | 0.03063 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.8 | 416.7 | 433.5 | 0 | 0 | 16.6 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.7 | 433.6 | 0 | 0 | 16.6 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 32 | 6 | 195.4 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.7301 | 433.6 | 0 | 0 | 16.6 | 13 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3489 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.6 | 433.5 | 0 | 0 | 16.6 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3292 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.5604 | 433.5 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.3925 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.6302 | 433.5 | 0 | 0 | 16.6 | 12.9 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 32 | 6 | 195 | 59.9 | 0.0334 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.698 | 22.3144 | 28.9 | 416.6
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.7302 | 433.6 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195 | 59.9 | 0.0335 | 0.0307 | 0.0222 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3636 | 28.9 | 416.7
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.8302 | 433.7 | 0 | 0 | 16.6 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195 | 59.9 | 0.0335 | 0.0307 | 0.02227 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3583 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.9 | 433.8 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195 | 59.9 | 0.0335 | 0.0307 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3197 | 28.9 | 416.9
2008-11-03 | 13:56 | 0 | 0 | 0 | 16.9 | 416.9302 | 433.9 | 0 | 0 | 16.6 | 12.9 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195 | 59.9 | 0.0335 | 0.0307 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3977 | 28.9 | 416.9
2008-11-03 | 13:56 | 0 | 0 | 0 | 17 | 416.9302 | 433.9 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 32 | 6 | 195 | 59.9 | 0.0335 | 0.0307 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4 | 28.9 | 416.9
2008-11-03 | 13:56 | -0.001 | 0 | 0 | 17 | 416.9302 | 433.9 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.2 | 59.9 | 0.03353 | 0.0308 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3461 | 28.9 | 416.9
2008-11-03 | 13:56 | -0.01 | 0 | 0 | 17 | 416.9 | 433.9 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.6 | 59.9 | 0.0336 | 0.0308 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.364 | 28.9 | 416.9
2008-11-03 | 13:56 | -0.002 | 0 | 0 | 17 | 416.8 | 433.8 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 196 | 59.9 | 0.0336 | 0.0308 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.4799 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 32 | 6 | 196 | 59.9 | 0.0336 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.3237 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.7694 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 196 | 59.9 | 0.0336 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.3062 | 22.3657 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8301 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 196 | 59.9 | 0.0336 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.6986 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8698 | 434 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.4 | 28.9 | 416.9
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1058.9778 | 1020 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.9 | 434 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 416.9
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.8 | 433.9 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.4 | 28.9 | 416.8
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 416.9 | 434 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 32 | 6 | 195.6 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 416.9
2008-11-03 | 13:56 | 0 | 0 | 0 | 17.1 | 417.1 | 434.2 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 196 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 417.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 17.1 | 417.2 | 434.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 196 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3531 | 28.9 | 417.2
2008-11-03 | 13:57 | 0 | 0 | 0 | 17.1 | 417.2 | 434.3 | 0 | 0 | 16.6 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0337 | 0.0309 | 0.0223 | 0.0002 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3249 | 28.9 | 417.2
2008-11-03 | 13:57 | 0 | -0.0008 | -7.8934 | 17.2 | 417.2 | 434.4 | 0 | 0 | 16.6 | 13 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 196 | 59.9 | 0.03373 | 0.031 | 0.0223 | 0.00023 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.397 | 28.9 | 417.2
2008-11-03 | 13:57 | 0 | -0.0004 | -3.5436 | 17.2 | 417.2302 | 434.4 | 0 | 0 | 16.6 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0338 | 0.031 | 0.02233 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.319 | 28.9 | 417.2
2008-11-03 | 13:57 | 0 | 0 | 0 | 17.2 | 417.3302 | 434.5 | 0 | 0 | 16.6 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 32 | 6 | 195.2 | 59.9 | 0.0338 | 0.031 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3586 | 28.9 | 417.3
2008-11-03 | 13:57 | 0 | 0.0003 | 2.9135 | 17.2 | 417.4 | 434.6 | 0 | 0 | 16.6 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0338 | 0.031 | 0.0223 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 417.4
2008-11-03 | 13:57 | 0 | 0.001 | 10 | 17.2 | 417.6114 | 434.8 | 0 | 0 | 16.6 | 12.9 | 923 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0338 | 0.031 | 0.02233 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 417.6
2008-11-03 | 13:57 | 0 | 0.001 | 10 | 17.2 | 427.4 | 444.6 | 0 | 0 | 16.99 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195.9 | 59.9 | 0.0338 | 0.031 | 0.0229 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 427.4
2008-11-03 | 13:57 | 0 | 0.001 | 10 | 18.1 | 453.22 | 471.3 | 0 | 0 | 18.11 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 196 | 59.9 | 0.03403 | 0.0311 | 0.024251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3287 | 28.9 | 453.2
2008-11-03 | 13:57 | 0 | 0.001 | 10 | 18.6 | 466.2699 | 484.9 | 0 | 0 | 18.59 | 13 | 925 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 196 | 59.9 | 0.03519 | 0.03219 | 0.0249 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3988 | 28.9 | 466.3
2008-11-03 | 13:57 | 0 | 0.0005 | 5.4576 | 18.8 | 467.2 | 486 | 0 | 0 | 18.64 | 13 | 924 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 196 | 59.9 | 0.03573 | 0.0327 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4454 | 28.9 | 467.2
2008-11-03 | 13:57 | 0 | 0 | 0 | 18.9 | 467.1 | 486 | 0 | 0 | 18.7 | 13 | 925 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 32 | 6 | 195.7 | 59.9 | 0.03613 | 0.03303 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 18.9 | 467.1308 | 486 | 0 | 0 | 18.7 | 13 | 924 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195.3 | 59.9 | 0.036231 | 0.0332 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 18.9 | 467.2 | 486.1 | 0 | 0 | 18.7 | 12.9 | 924 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195 | 59.9 | 0.0363 | 0.0332 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.2
2008-11-03 | 13:57 | 0 | 0.0008 | 8.0259 | 18.9 | 467.3396 | 486.2 | 0 | 0 | 18.7 | 13 | 924 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 195.4 | 59.9 | 0.0364 | 0.0333 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3397 | 28.9 | 467.3
2008-11-03 | 13:57 | 0 | -0.0002 | -2.4547 | 19 | 467.2604 | 486.3 | 0 | 0 | 18.7 | 12.9 | 924 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.1 | 32 | 6 | 195.8 | 59.9 | 0.0364 | 0.0333 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 467.3
2008-11-03 | 13:57 | -0.001 | -0.0005 | -5.4796 | 19 | 467.4302 | 486.4 | 0 | 0 | 18.7 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.8 | 59.9 | 0.03643 | 0.0334 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 467.4
2008-11-03 | 13:57 | -0.01 | 0 | 0 | 19 | 467.5 | 486.5 | 0 | 0 | 18.7 | 12.9 | 924 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0365 | 0.0335 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 467.5
2008-11-03 | 13:57 | -0.01 | -0.0001 | -1.3288 | 19 | 467.4302 | 486.5 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.1 | 32 | 6 | 195.1 | 59.9 | 0.03663 | 0.03353 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3483 | 28.9 | 467.4
2008-11-03 | 13:57 | -0.004 | -0.001 | -9.8467 | 19.1 | 467.5 | 486.6 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.037 | 0.0339 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2701 | 28.9 | 467.5
2008-11-03 | 13:57 | 0 | -0.001 | -10 | 19.2 | 467.6 | 486.8 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.03703 | 0.03393 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2159 | 28.9 | 467.6
2008-11-03 | 13:57 | 0 | -0.001 | -10 | 19.2 | 467.6302 | 486.8 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195 | 59.9 | 0.0371 | 0.034 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3721 | 28.9 | 467.6
2008-11-03 | 13:57 | 0 | -0.0005 | -5.2992 | 19.2 | 467.7 | 486.9 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.4 | 59.9 | 0.0371 | 0.034 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3359 | 28.9 | 467.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.2 | 467.9 | 487.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0371 | 0.034 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.9
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.2 | 467.9 | 487.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 196 | 59.9 | 0.03713 | 0.034 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.9
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.2 | 467.9 | 487.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0372 | 0.034 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.9
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.2 | 468 | 487.2 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0372 | 0.0341 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 468
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.2 | 467.9302 | 487.2 | 0 | 0 | 18.7 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 192 | 32 | 6 | 195.2 | 59.9 | 0.03723 | 0.0341 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 467.9
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468.0699 | 487.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191.9 | 32 | 6 | 195.5 | 59.9 | 0.0373 | 0.0342 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 468.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468.1 | 487.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 195.9 | 59.9 | 0.0373 | 0.0342 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3103 | 28.9 | 468.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468.1 | 487.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 196 | 59.9 | 0.0374 | 0.0342 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3884 | 28.9 | 468.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468 | 487.3 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 191 | 32 | 6 | 196 | 59.9 | 0.0374 | 0.0343 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3336 | 28.9 | 468
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468.0604 | 487.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0374 | 0.0343 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3445 | 28.9 | 468.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.3 | 468.0302 | 487.4 | 0 | 0 | 18.7 | 13 | 924 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 196 | 59.9 | 0.03743 | 0.03433 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3775 | 28.9 | 468
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.2302 | 487.6 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 196 | 59.9 | 0.0376 | 0.0344 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3005 | 28.9 | 468.2
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.1 | 487.5 | 0 | 0 | 18.7 | 12.9 | 924 | 904 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.9 | 32 | 6 | 196 | 59.9 | 0.0376 | 0.03443 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3786 | 28.9 | 468.1
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.3604 | 487.8 | 0 | 0 | 18.7 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195.6 | 59.9 | 0.0376 | 0.0345 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 28.9 | 468.4
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.4 | 487.8 | 0 | 0 | 18.7 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0376 | 0.0345 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.331 | 28.9 | 468.4
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.4302 | 487.8 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36.054 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0376 | 0.0345 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2123 | 28.9 | 468.4
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.4 | 468.4695 | 487.9 | 0 | 0 | 18.7 | 12.9 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 36.946 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0376 | 0.0345 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2907 | 28.9 | 468.5
2008-11-03 | 13:57 | 0 | -0.0004 | -4.4787 | 19.5 | 468.5302 | 488 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 189.1 | 32 | 6 | 195.9 | 59.9 | 0.0377 | 0.0345 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 28.9 | 468.5
2008-11-03 | 13:57 | 0 | -0.001 | -10 | 19.5 | 468.5 | 488 | 0 | 0 | 18.7 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36 | 190 | 32 | 6 | 196 | 59.9 | 0.0377 | 0.0346 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2532 | 28.9468 | 468.5
2008-11-03 | 13:57 | 0 | -0.001 | -10 | 19.5 | 468.4603 | 488 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 36.0531 | 190 | 32 | 6 | 196 | 59.9 | 0.0378 | 0.0346 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2249 | 28.9751 | 468.5
2008-11-03 | 13:57 | 0 | -0.001 | -10 | 19.5 | 468.6 | 488.1 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0378 | 0.0346 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3029 | 28.9029 | 468.6
2008-11-03 | 13:57 | 0 | -0.0002 | -2.1171 | 19.5 | 468.7398 | 488.3 | 0 | 0 | 18.7 | 13 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.7 | 59.9 | 0.0378 | 0.0346 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.381 | 28.981 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.8 | 488.4 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.3 | 59.9 | 0.0378 | 0.03463 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2828 | 29 | 468.8
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0 | 0 | 18.7 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9302 | 0.0378 | 0.0347 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2366 | 29 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9 | 0.0378 | 0.0347 | 0.025 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.8 | 59.9 | 0.0378 | 0.0347 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0379 | 0.0347 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.5396 | 488.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0379 | 0.0347 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019.302 | 22.3 | 29 | 468.5
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.4302 | 488 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0379 | 0.0348 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.3 | 29 | 468.4
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.5 | 488.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.038 | 0.0348 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.3 | 29 | 468.5
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.5 | 488.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9 | 0.038 | 0.0348 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.3 | 29 | 468.5
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.5 | 488.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.038 | 0.03483 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.3 | 29 | 468.5
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.3 | 487.9 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.038 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.2607 | 29 | 468.3
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.5305 | 488.1 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019.6952 | 22.2173 | 29 | 468.5
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.6 | 488.2 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2953 | 29 | 468.6
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0 | 0 | 18.7 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3732 | 29 | 468.7
2008-11-03 | 13:57 | 0 | 0 | 0 | 19.6 | 468.8 | 488.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3487 | 29 | 468.8
2008-11-03 | 13:58 | 0 | 0 | 0 | 19.6 | 468.8 | 488.4 | 0 | 0 | 18.7 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.8 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3294 | 29 | 468.8
2008-11-03 | 13:58 | 0 | 0 | 0 | 19.6 | 468.8 | 488.4 | 0 | 0 | 18.7 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.5 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3926 | 29 | 468.8
2008-11-03 | 13:58 | 0 | -0.0016 | -15.5974 | 19.6 | 468.7698 | 488.4 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.1 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3145 | 29 | 468.8
2008-11-03 | 13:58 | 0 | -0.002 | -20 | 19.6 | 468.7698 | 488.4 | 0 | 0 | 18.7 | 13 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.2 | 59.9 | 0.0381 | 0.0349 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.427 | 29 | 468.8
2008-11-03 | 13:58 | 0 | -0.0016 | -15.5038 | 19.7 | 468.8396 | 488.5 | 0 | 0 | 18.7 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.6 | 59.9 | 0.0381 | 0.03493 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4169 | 29 | 468.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 19.7 | 469.7268 | 489.4 | 0 | 0 | 18.7 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0381 | 0.035 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019.302 | 22.3 | 29 | 469.7
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 19.7 | 483.6044 | 503.3 | 0 | 0 | 18.95 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.7 | 59.9 | 0.03813 | 0.035 | 0.025851 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019.6986 | 22.3 | 29 | 483.6
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 19.8 | 502.1 | 521.9 | 0 | 0 | 20.12 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.3 | 59.9 | 0.0385 | 0.0353 | 0.0268 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1491 | 29 | 502.1
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.2 | 513.761 | 534 | 0 | 0 | 20.64 | 12.9 | 926 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.039493 | 0.036262 | 0.027391 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2072 | 29 | 513.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.5 | 518.6302 | 539.1 | 0 | 0 | 20.8 | 13 | 926 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.4 | 59.9 | 0.0405 | 0.03713 | 0.0277 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29 | 518.6
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.5 | 518.6 | 539.1 | 0 | 0 | 20.8 | 12.9 | 926 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.8 | 59.9 | 0.0405 | 0.0372 | 0.0277 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3096 | 29.0096 | 518.6
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.6 | 519.6698 | 540.3 | 0 | 0 | 20.8 | 13 | 925 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0406 | 0.0373 | 0.0277 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3877 | 29.0877 | 519.7
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.7 | 519.8 | 540.5 | 0 | 0 | 20.9 | 12.9 | 925 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.5 | 59.9 | 0.04073 | 0.0374 | 0.0277 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.4 | 29.1 | 519.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.7 | 519.7 | 540.4 | 0 | 0 | 20.9 | 13 | 925 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.1 | 59.9 | 0.0409 | 0.0375 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3563 | 29.1 | 519.7
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.8 | 519.7 | 540.5 | 0 | 0 | 20.9 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.2 | 59.9 | 0.04093 | 0.0376 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29.1 | 519.7
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.8 | 519.7302 | 540.5 | 0 | 0 | 20.9 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.6 | 59.9 | 0.041 | 0.0376 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29.1 | 519.7
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.8 | 520 | 540.8 | 0 | 0 | 20.9 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.04103 | 0.0377 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29.1 | 520
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.9 | 519.8603 | 540.8 | 0 | 0 | 20.9 | 13 | 924 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 196 | 59.9 | 0.04113 | 0.0378 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1885 | 29.1 | 519.9
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.9 | 520 | 540.9 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0412 | 0.0378 | 0.02777 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1338 | 29.1 | 520
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.9 | 520 | 540.9 | 0 | 0 | 20.9 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0412 | 0.0379 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2 | 29.1 | 520
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 20.9 | 520 | 540.9 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9698 | 0.04123 | 0.0379 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2 | 29.1 | 520
2008-11-03 | 13:58 | 0 | -0.0004 | -3.8095 | 21 | 520.0302 | 541 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0413 | 0.03793 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1324 | 29.1 | 520
2008-11-03 | 13:58 | 0 | 0 | 0 | 21 | 520.1 | 541.1 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0414 | 0.038 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.1459 | 520.1
2008-11-03 | 13:58 | 0 | 0 | 0 | 21 | 520.1 | 541.1 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0414 | 0.038 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1478 | 29.2 | 520.1
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.1 | 520.1302 | 541.2 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0414 | 0.038 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29.198 | 520.1
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.1 | 520.2 | 541.3 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.0415 | 0.0381 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3 | 29.12 | 520.2
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.1 | 520.3 | 541.4 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9698 | 0.0415 | 0.0381 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3578 | 29.1 | 520.3
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.2 | 520.1604 | 541.4 | 0 | 0 | 20.9 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0416 | 0.03813 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.3642 | 29.1 | 520.2
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.2 | 520.3 | 541.5 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.3 | 59.9 | 0.0416 | 0.0382 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2721 | 29.1 | 520.3
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.3 | 520.4 | 541.7 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.7 | 59.9 | 0.0416 | 0.0382 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1161 | 29.1 | 520.4
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.3 | 520.4 | 541.7 | 0 | 0 | 20.9 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0416 | 0.0382 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2399 | 29.1 | 520.4
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.3 | 520.4 | 541.7 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0417 | 0.0382 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2041 | 29.1 | 520.4
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.3 | 520.4902 | 541.8 | 0 | 0 | 20.9 | 12.9 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0417 | 0.03823 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1519 | 29.1 | 520.5
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.3 | 520.5 | 541.8 | 0 | 0 | 20.9 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9 | 0.04183 | 0.03833 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2921 | 29.104 | 520.5
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.4 | 520.4697 | 541.9 | 0 | 0 | 20.9 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.5 | 59.9 | 0.04193 | 0.0385 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1363 | 29.1819 | 520.5
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.4 | 520.4 | 541.8 | 0 | 0 | 20.9 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.1 | 59.9698 | 0.042 | 0.0386 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.14 | 520.4
2008-11-03 | 13:58 | 0 | -0.0001 | -1.2652 | 21.4 | 520.4 | 541.8 | 0 | 0 | 20.9 | 13 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0421 | 0.03863 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.1381 | 520.4
2008-11-03 | 13:58 | 0 | -0.0009 | -9.3755 | 21.5 | 520.4604 | 542 | 0 | 0 | 20.9 | 13 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.04213 | 0.0387 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1161 | 29.2 | 520.5
2008-11-03 | 13:58 | 0 | -0.0002 | -2.3083 | 21.5 | 520.6 | 542.1 | 0 | 0 | 20.9 | 13 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.0387 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1941 | 29.2 | 520.6
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.5 | 520.6302 | 542.1 | 0 | 0 | 20.9 | 13 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.0387 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.2 | 29.2 | 520.6
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.5 | 520.8 | 542.3 | 0 | 0 | 20.9 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.0387 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1501 | 29.2 | 520.8
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.6 | 520.9396 | 542.5 | 0 | 0 | 20.9 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.03873 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1018.302 | 22.1 | 29.1721 | 520.9
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.6 | 520.8 | 542.4 | 0 | 0 | 20.9 | 13 | 922 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1061 | 29.1061 | 520.8
2008-11-03 | 13:58 | 0 | 0 | 0 | 21.6 | 520.8 | 542.4 | 0 | 0 | 20.9 | 12.9 | 923 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0422 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1841 | 29.1841 | 520.8
2008-11-03 | 13:58 | 0 | -0.0008 | -7.748 | 21.6 | 520.8302 | 542.4 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.2 | 59.9 | 0.0422 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1379 | 29.2 | 520.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.6 | 520.8394 | 542.4 | 0 | 0 | 20.9 | 12.9 | 923 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1401 | 29.2 | 520.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.6 | 520.8699 | 542.5 | 0 | 0 | 20.9 | 13 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.9 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1819 | 29.2 | 520.9
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.6 | 520.8 | 542.4 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 196 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1039 | 29.2 | 520.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.6 | 520.8 | 542.4 | 0 | 0 | 20.9 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 520.8
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 520.9 | 542.6 | 0 | 0 | 20.9 | 13 | 921 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 60 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 520.9
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 520.9302 | 542.6 | 0 | 0 | 20.9 | 13 | 921 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 520.9
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 521.0607 | 542.8 | 0 | 0 | 20.9 | 13 | 921 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0423 | 0.0388 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1018.6966 | 22.0919 | 29.2 | 521.1
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 521.2 | 542.9 | 0 | 0 | 20.9 | 13 | 921 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.04233 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1018 | 22.0139 | 29.2 | 521.2
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 521 | 542.7 | 0 | 0 | 20.9 | 13 | 921 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9698 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1924 | 29.2 | 521
2008-11-03 | 13:58 | 0 | -0.001 | -10 | 21.7 | 520.9 | 542.6 | 0 | 0 | 20.9 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9698 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.258 | 29.2 | 520.9
2008-11-03 | 13:58 | -0.012 | -0.001 | -10 | 21.7 | 520.8 | 542.5 | 0 | 0 | 20.9 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9302 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1801 | 29.2 | 520.8
2008-11-03 | 13:58 | -0.016 | -0.0002 | -2 | 21.7 | 520.8 | 542.5 | 0 | 0 | 20.9 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.102 | 29.2 | 520.8
2008-11-03 | 13:58 | -0.01 | -0.0006 | -6.1223 | 21.7 | 520.6302 | 542.3 | 0 | 0 | 20.9 | 13 | 921 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.0239 | 29.2 | 520.6
2008-11-03 | 13:58 | -0.01 | -0.001 | -10 | 21.7 | 520.5 | 542.2 | 0 | 0 | 20.9 | 13 | 921 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.1083 | 29.2 | 520.5
2008-11-03 | 13:58 | -0.01 | -0.001 | -10 | 21.7 | 520.7 | 542.4 | 0 | 0 | 20.9 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1020 | 22.1679 | 29.2 | 520.7
2008-11-03 | 13:59 | -0.01 | -0.001 | -10 | 21.7 | 520.8 | 542.5 | 0 | 0 | 20.9 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 520.8
2008-11-03 | 13:59 | -0.01 | -0.001 | -10 | 21.8 | 521 | 542.8 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 521
2008-11-03 | 13:59 | -0.006 | -0.001 | -10 | 21.8 | 521 | 542.8 | 0 | 0 | 20.9 | 13 | 922 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9698 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 521
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 21.8 | 520.8604 | 542.7 | 0 | 0 | 20.9 | 12.9 | 923 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0424 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1 | 29.2 | 520.9
2008-11-03 | 13:59 | 0 | -0.0006 | -5.6917 | 21.8 | 520.86 | 542.7 | 0 | 0 | 20.9 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.4 | 59.9 | 0.0425 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1059 | 1019 | 22.1221 | 29.2 | 520.9
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.8 | 521 | 542.8 | 0 | 0 | 20.9 | 12.9 | 923 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.8 | 59.9 | 0.0425 | 0.0389 | 0.0278 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1998 | 29.2 | 521
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.8 | 517.2552 | 539.1 | 0 | 0 | 20.83 | 13 | 922 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.8 | 59.9 | 0.0425 | 0.0389 | 0.027619 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0437 | 29.2 | 517.3
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.8 | 492.403 | 514.2 | 0 | 0 | 19.77 | 12.9 | 923 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.5 | 59.9 | 0.042339 | 0.038839 | 0.026318 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 492.4
2008-11-03 | 13:59 | 0 | -0.0008 | -7.7425 | 21.5 | 474.1698 | 495.7 | 0 | 0 | 19.06 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.1 | 59.9 | 0.041958 | 0.038458 | 0.025749 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0341 | 29.2 | 474.2
2008-11-03 | 13:59 | 0 | -0.0004 | -4.0149 | 21.5 | 473.1906 | 494.7 | 0 | 0 | 18.9 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.04137 | 0.0379 | 0.0253 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1 | 29.2 | 473.2
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.4 | 473.2698 | 494.7 | 0 | 0 | 18.86 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0412 | 0.03777 | 0.0253 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1 | 29.2 | 473.3
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.4 | 472.8282 | 494.2 | 0 | 0 | 18.8 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0411 | 0.03767 | 0.02527 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0319 | 29.2 | 472.8
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.3 | 472.1698 | 493.5 | 0 | 0 | 18.8 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.7 | 59.9 | 0.041 | 0.03757 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0921 | 29.2 | 472.2
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.3 | 472.1302 | 493.4 | 0 | 0 | 18.8 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.3 | 59.9302 | 0.0409 | 0.0375 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 22.1521 | 29.2 | 472.1
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.3 | 472.2302 | 493.5 | 0 | 0 | 18.8 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9698 | 0.04087 | 0.0374 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22.0041 | 29.2 | 472.2
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.3 | 472.2392 | 493.5 | 0 | 0 | 18.8 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9304 | 0.04077 | 0.03737 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22.1601 | 29.2 | 472.2
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.2 | 471.9792 | 493.2 | 0 | 0 | 18.8 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.8 | 59.9 | 0.04067 | 0.0372 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 22.1421 | 29.2 | 472
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.2 | 471.8698 | 493 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.04057 | 0.0371 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0641 | 29.2 | 471.9
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.1 | 471.8 | 492.9 | 0 | 0 | 18.8 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0404 | 0.037 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0141 | 29.2 | 471.8
2008-11-03 | 13:59 | 0 | 0 | 0 | 21.1 | 471.8 | 492.9 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0404 | 0.037 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0921 | 29.2 | 471.8
2008-11-03 | 13:59 | 0 | 0.0008 | 7.7498 | 21.1 | 471.7399 | 492.8 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0404 | 0.037 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.2401 | 29.2 | 471.7
2008-11-03 | 13:59 | 0 | 0.001 | 10 | 21.1 | 471.7 | 492.8 | 0 | 0 | 18.8 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.040369 | 0.036969 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.2039 | 29.2 | 471.7
2008-11-03 | 13:59 | 0 | 0.001 | 10 | 21 | 471.7302 | 492.7 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0402 | 0.0368 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 22.1 | 29.2 | 471.7
2008-11-03 | 13:59 | 0 | 0.0007 | 7.3087 | 21 | 471.6 | 492.6 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9301 | 0.0401 | 0.03677 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0961 | 29.2 | 471.6
2008-11-03 | 13:59 | 0 | 0 | 0 | 21 | 471.3397 | 492.3 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0401 | 0.0367 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0187 | 29.2 | 471.3
2008-11-03 | 13:59 | 0 | 0 | 0 | 21 | 471.1698 | 492.2 | 0 | 0 | 18.8 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0401 | 0.0367 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 471.2
2008-11-03 | 13:59 | 0 | 0 | 0 | 21 | 471.1 | 492.1 | 0 | 0 | 18.8 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0401 | 0.0367 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.962 | 29.2 | 471.1
2008-11-03 | 13:59 | 0 | 0 | 0 | 21 | 471.1302 | 492.1 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.2 | 59.9 | 0.0401 | 0.0367 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.916 | 29.2 | 471.1
2008-11-03 | 13:59 | 0 | 0 | 0 | 21 | 471 | 492 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.04007 | 0.0367 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.994 | 29.2 | 471
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.9 | 470.9396 | 491.8 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.04 | 0.0366 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019.302 | 22.072 | 29.2 | 470.9
2008-11-03 | 13:59 | 0 | -0.0001 | -1.3011 | 20.9 | 471.0302 | 491.9 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.03997 | 0.0366 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0501 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -9.6414 | 20.9 | 471 | 491.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0399 | 0.0365 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0277 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.9 | 471 | 491.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0399 | 0.0365 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0946 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.9 | 471 | 491.9 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.03987 | 0.0365 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0162 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.9 | 471 | 491.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.7 | 59.9 | 0.0398 | 0.0364 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.9 | 471 | 491.9 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.3 | 59.9 | 0.0398 | 0.0364 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.8 | 471.1 | 491.9 | 0 | 0 | 18.8 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 60 | 0.0397 | 0.03637 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.2 | 471.1
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.8 | 470.9697 | 491.8 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.4 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 471
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.8 | 470.8396 | 491.6 | 0 | 0 | 18.8 | 12.9 | 919 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.8 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.8
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.8 | 470.7 | 491.5 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.7
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.8 | 470.7 | 491.5 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0299 | 29.2 | 470.7
2008-11-03 | 13:59 | 0 | -0.0007 | -7.3466 | 20.8 | 470.6396 | 491.4 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1 | 29.2 | 470.6
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.8 | 470.6 | 491.4 | 0 | 0 | 18.8 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0397 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1 | 29.2 | 470.6
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.8 | 470.6 | 491.4 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.4 | 59.9 | 0.0396 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1638 | 29.2 | 470.6
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.8 | 470.6604 | 491.4 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0396 | 0.0363 | 0.0252 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.1165 | 29.2 | 470.7
2008-11-03 | 13:59 | 0.003 | 0 | 0 | 20.7 | 470.7 | 491.4 | 0 | 0 | 18.8 | 13 | 919 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.03957 | 0.0362 | 0.02517 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.7
2008-11-03 | 13:59 | 0.01 | 0 | 0 | 20.7 | 470.6698 | 491.4 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.0395 | 0.0362 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.7
2008-11-03 | 13:59 | 0.01 | 0 | 0 | 20.7 | 470.5 | 491.2 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.5
2008-11-03 | 13:59 | 0.01 | 0 | 0 | 20.7 | 470.4 | 491.1 | 0 | 0 | 18.8 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9696 | 0.0395 | 0.0361 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.4
2008-11-03 | 13:59 | 0.01 | 0 | 0 | 20.7 | 470.6 | 491.3 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9698 | 0.03947 | 0.0361 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.6
2008-11-03 | 13:59 | 0.006 | 0 | 0 | 20.7 | 470.6 | 491.3 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.0394 | 0.0361 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.6
2008-11-03 | 13:59 | 0 | -0.0006 | -5.9262 | 20.7 | 470.4 | 491.1 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0394 | 0.03607 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.4
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.7 | 470.4302 | 491.1 | 0 | 0 | 18.8 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.1 | 59.9 | 0.0394 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9345 | 29.2 | 470.4
2008-11-03 | 13:59 | 0 | -0.001 | -10 | 20.7 | 470.5 | 491.2 | 0 | 0 | 18.8 | 13 | 918 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0394 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9435 | 29.2 | 470.5
2008-11-03 | 13:59 | 0 | -0.0009 | -8.749 | 20.7 | 470.3 | 491 | 0 | 0 | 18.8 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.03937 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.3
2008-11-03 | 13:59 | 0 | 0 | -0.40867 | 20.6 | 470.3 | 490.9 | 0 | 0 | 18.8 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.3
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.6 | 470.3 | 490.9 | 0 | 0 | 18.8 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.7 | 59.9 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0775 | 29.2 | 470.3
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.6 | 470.5 | 491.1 | 0 | 0 | 18.8 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.3 | 59.9302 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22.0445 | 29.2 | 470.5
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.6 | 470.4698 | 491.1 | 0 | 0 | 18.8 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.1 | 59.9 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.5
2008-11-03 | 13:59 | 0 | 0 | 0 | 20.6 | 470.4 | 491 | 0 | 0 | 18.8 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.4 | 59.9302 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 470.4
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.6 | 491.1 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9698 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.2 | 470.6
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.5396 | 491 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.8 | 59.9 | 0.0393 | 0.036 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 21.9326 | 29.2674 | 470.5
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.4 | 490.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0393 | 0.03597 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9454 | 29.2546 | 470.4
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.4 | 490.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.1 | 59.9 | 0.0393 | 0.0359 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9766 | 29.2234 | 470.4
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.4 | 490.9 | 0 | 0 | 18.8 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9013 | 29.2987 | 470.4
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 470.4 | 490.9 | 0 | 0 | 18.8 | 13 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 60 | 0.0393 | 0.0359 | 0.0251 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9793 | 29.2207 | 470.4
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.5 | 457.2404 | 477.7 | 0 | 0 | 18.75 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9302 | 0.0392 | 0.0359 | 0.024449 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2 | 457.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.2 | 434.1586 | 454.4 | 0 | 0 | 17.68 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.038598 | 0.035298 | 0.023228 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9299 | 29.2 | 434.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 20.1 | 423.1692 | 443.2 | 0 | 0 | 16.96 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9 | 0.037808 | 0.034508 | 0.0226 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8128 | 29.2128 | 423.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 19.8 | 421.0699 | 440.8 | 0 | 0 | 16.8 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.03711 | 0.03391 | 0.0225 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8911 | 29.2911 | 421.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 19.7 | 420.9302 | 440.6 | 0 | 0 | 16.7 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.03687 | 0.03367 | 0.0225 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2308 | 420.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 19.6 | 420.9699 | 440.5 | 0 | 0 | 16.7 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.03661 | 0.03351 | 0.0225 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9472 | 29.2 | 421
2008-11-03 | 14:00 | 0 | -0.0002 | -1.7459 | 19.4 | 420.7287 | 440.2 | 0 | 0 | 16.7 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.2 | 59.9 | 0.036189 | 0.033019 | 0.02247 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.9498 | 29.2 | 420.7
2008-11-03 | 14:00 | 0 | -0.0019 | -19.032 | 19.3 | 420.1 | 439.4 | 0 | 0 | 16.7 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0357 | 0.03257 | 0.0224 | 0.0002 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8 | 29.2031 | 420.1
2008-11-03 | 14:00 | 0 | -0.002 | -20 | 19.3 | 420.0399 | 439.3 | 0 | 0 | 16.7 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.9 | 59.9 | 0.03567 | 0.0325 | 0.0224 | 0.00023 |  | n/a |  | 34 | 60 | 1058 | 1018.3006 | 21.8 | 29.281 | 420
2008-11-03 | 14:00 | 0 | -0.002 | -20 | 19.2 | 419.7 | 438.9 | 0 | 0 | 16.7 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0356 | 0.0325 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.859 | 29.3 | 419.7
2008-11-03 | 14:00 | 0 | -0.002 | -20 | 19.1 | 419.5302 | 438.6 | 0 | 0 | 16.7 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 196 | 59.9302 | 0.03557 | 0.03247 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.937 | 29.3 | 419.5
2008-11-03 | 14:00 | 0 | -0.0017 | -17.3295 | 19.1 | 419.5698 | 438.6 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9698 | 0.0355 | 0.0324 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.285 | 419.6
2008-11-03 | 14:00 | 0 | -0.0009 | -9.191 | 19 | 419.3698 | 438.4 | 0 | 0 | 16.7 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0355 | 0.0324 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.2071 | 419.4
2008-11-03 | 14:00 | 0 | -0.0001 | -0.85071 | 19 | 419.2698 | 438.3 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.03547 | 0.0324 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9291 | 29.2 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 19 | 419.2 | 438.2 | 0 | 0 | 16.7 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0354 | 0.0324 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 19 | 419.2302 | 438.2 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0354 | 0.0324 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 19 | 419.2698 | 438.2 | 0 | 0 | 16.7 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.03537 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2048 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.3 | 438.2 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2828 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.3 | 438.2 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9302 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9607 | 29.3 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.2 | 438.1 | 0 | 0 | 16.7 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.8 | 59.9302 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9226 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.2 | 438.1 | 0 | 0 | 16.7 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 60 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.2 | 438.1 | 0 | 0 | 16.7 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.3 | 438.2 | 0 | 0 | 16.7 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0353 | 0.0323 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9451 | 29.3 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.9 | 419.1302 | 438 | 0 | 0 | 16.7 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0353 | 0.0322 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8994 | 29.2497 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.8 | 419.2 | 438 | 0 | 0 | 16.7 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0353 | 0.0322 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2282 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.8 | 419.2 | 438 | 0 | 0 | 16.62 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9698 | 0.0352 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8128 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.8 | 419.2 | 438 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.03517 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9689 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.7 | 419.3 | 438 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8751 | 29.2376 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.7 | 419.3 | 438 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8404 | 29.2404 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419.2698 | 437.9 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9302 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 21.8817 | 29.3 | 419.3
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419.2302 | 437.8 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 21.8037 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419.2 | 437.8 | 0 | 0 | 16.6 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419.2 | 437.8 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9037 | 29.3 | 419.2
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419 | 437.6 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9095 | 29.3 | 419
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.6 | 419.1 | 437.7 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7243 | 29.2919 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419.1 | 437.6 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9308 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9581 | 29.214 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419.1 | 437.6 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 196 | 59.9698 | 0.0351 | 0.0321 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.872 | 29.264 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419.1 | 437.6 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.6 | 59.9 | 0.03507 | 0.03207 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419.1 | 437.6 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.2 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8199 | 29.2801 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419.0698 | 437.6 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8977 | 29.2023 | 419.1
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.9697 | 437.5 | 0 | 0 | 16.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9697 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8242 | 29.2758 | 419
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.9 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 418.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.9 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9302 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2682 | 418.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.8698 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2097 | 418.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.8 | 437.3 | 0 | 0 | 16.6 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2877 | 418.8
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.9 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 418.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419 | 437.5 | 0 | 0 | 16.6 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.035 | 0.032 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8866 | 29.3 | 419
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 418.9302 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.03497 | 0.03197 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9577 | 29.3 | 418.9
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.5 | 419 | 437.5 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8009 | 29.3 | 419
2008-11-03 | 14:00 | 0 | 0 | 0 | 18.4 | 419 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8775 | 29.3 | 419
2008-11-03 | 14:00 | 0 | -0.0008 | -7.7145 | 18.4 | 419 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.2 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2445 | 419
2008-11-03 | 14:00 | 0 | -0.0004 | -4.1505 | 18.4 | 419 | 437.4 | 0 | 0 | 16.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8331 | 29.2334 | 419
2008-11-03 | 14:00 | -0.001 | 0 | 0 | 18.3 | 418.9698 | 437.3 | 0 | 0 | 16.6 | 13 | 918 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.9 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7226 | 29.3 | 419
2008-11-03 | 14:00 | -0.009 | -0.0002 | -2.4806 | 18.3 | 418.8 | 437.1 | 0 | 0 | 16.6 | 13 | 918 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8777 | 29.3 | 418.8
2008-11-03 | 14:01 | -0.01 | -0.0009 | -8.9112 | 18.3 | 418.8301 | 437.1 | 0 | 0 | 16.6 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8329 | 29.2329 | 418.8
2008-11-03 | 14:01 | -0.004 | -0.0001 | -0.53601 | 18.3 | 418.7698 | 437.1 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8452 | 29.2 | 418.8
2008-11-03 | 14:01 | 0 | -0.0008 | -7.7164 | 18.3 | 418.6699 | 437 | 0 | 0 | 16.6 | 12.9 | 919 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.7 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2 | 418.7
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 18.3 | 418.6 | 436.9 | 0 | 0 | 16.6 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.3 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2011 | 418.6
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 18.3 | 418.6 | 436.9 | 0 | 0 | 16.6 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2791 | 418.6
2008-11-03 | 14:01 | 0 | -0.0008 | -7.6147 | 18.3 | 418.6302 | 436.9 | 0 | 0 | 16.6 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0349 | 0.0319 | 0.0224 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2429 | 418.6
2008-11-03 | 14:01 | 0 | -0.0001 | -0.73917 | 18.2 | 404.1345 | 422.3 | 0 | 0 | 16.5 | 12.9 | 920 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.03487 | 0.0319 | 0.021649 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9 | 29.2351 | 404.1
2008-11-03 | 14:01 | 0 | -0.0009 | -9.2353 | 17.5 | 375.1692 | 392.6 | 0 | 0 | 15.39 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.2 | 59.9 | 0.034309 | 0.031509 | 0.020109 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.861 | 29.287 | 375.2
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 17.3 | 371.4 | 388.7 | 0 | 0 | 14.77 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.5 | 59.9 | 0.034009 | 0.031209 | 0.0199 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.6271 | 29.209 | 371.4
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 17 | 369.3 | 386.3 | 0 | 0 | 14.62 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.9 | 59.9 | 0.032409 | 0.029609 | 0.0198 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7378 | 29.2689 | 369.3
2008-11-03 | 14:01 | 0 | -0.0006 | -5.7804 | 16.8 | 369.4302 | 386.2 | 0 | 0 | 14.6 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0318 | 0.0291 | 0.0198 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8468 | 29.2532 | 369.4
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.8 | 369.0584 | 385.8 | 0 | 0 | 14.6 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.4 | 59.9 | 0.03177 | 0.02907 | 0.01977 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.9248 | 29.2248 | 369.1
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.7 | 368.4697 | 385.2 | 0 | 0 | 14.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9697 | 0.03167 | 0.02897 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.3 | 368.5
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.7 | 368.2698 | 385 | 0 | 0 | 14.6 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.3 | 59.9302 | 0.0316 | 0.0289 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 22 | 29.3 | 368.3
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.7 | 368.2 | 384.9 | 0 | 0 | 14.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.7 | 59.9698 | 0.0316 | 0.0289 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7659 | 29.3 | 368.2
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.7 | 368.2302 | 384.9 | 0 | 0 | 14.6 | 12.9 | 918 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.03157 | 0.02887 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 21.6365 | 29.3 | 368.2
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.7 | 368.2397 | 384.9 | 0 | 0 | 14.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0315 | 0.0288 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.3014 | 21.7292 | 29.2854 | 368.2
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.6 | 368.1396 | 384.7 | 0 | 0 | 14.6 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 196 | 59.9 | 0.0315 | 0.0288 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.885 | 29.2075 | 368.1
2008-11-03 | 14:01 | 0 | -0.0001 | -0.69551 | 16.6 | 368.0301 | 384.6 | 0 | 0 | 14.6 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.03147 | 0.02877 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9705 | 29.2705 | 368
2008-11-03 | 14:01 | 0 | -0.0009 | -9.1773 | 16.5 | 368.0301 | 384.5 | 0 | 0 | 14.6 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6.1 | 195.5 | 59.9301 | 0.0314 | 0.0287 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9516 | 29.2516 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.5 | 368.0302 | 384.5 | 0 | 0 | 14.6 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195.1 | 59.9 | 0.0314 | 0.0287 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8737 | 29.2263 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.5 | 368.0698 | 384.6 | 0 | 0 | 14.6 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195 | 59.9 | 0.03137 | 0.0287 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2957 | 368.1
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.5 | 368.1 | 384.6 | 0 | 0 | 14.6 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195 | 59.9699 | 0.0313 | 0.0286 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2181 | 368.1
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.4 | 368.2 | 384.6 | 0 | 0 | 14.6 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195 | 59.9 | 0.03127 | 0.0286 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.86 | 29.2 | 368.2
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.4 | 368.0302 | 384.4 | 0 | 0 | 14.6 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195.3 | 59.9 | 0.0312 | 0.0286 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.698 | 21.8619 | 29.2381 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.4 | 368 | 384.4 | 0 | 0 | 14.6 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195.7 | 59.9 | 0.0312 | 0.02857 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 368 | 384.3 | 0 | 0 | 14.6 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 7 | 196 | 59.9 | 0.0312 | 0.0285 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 368 | 384.3 | 0 | 0 | 14.6 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 7 | 196 | 59.9 | 0.03117 | 0.0285 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.944 | 29.3 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 367.9396 | 384.2 | 0 | 0 | 14.6 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 7 | 196 | 59.9302 | 0.0311 | 0.0284 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9002 | 29.3 | 367.9
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 367.8302 | 384.1 | 0 | 0 | 14.6 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 7 | 195.8 | 59.9302 | 0.0311 | 0.0284 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 367.8
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 367.9 | 384.2 | 0 | 0 | 14.5 | 13 | 918 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 7 | 195.5 | 59.9 | 0.0311 | 0.0284 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8104 | 29.3 | 367.9
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 368 | 384.3 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 7 | 195.1 | 59.9 | 0.0311 | 0.0284 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.9673 | 29.3 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.3 | 367.9699 | 384.2 | 0 | 0 | 14.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 7 | 195 | 59.9 | 0.03107 | 0.0284 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.3 | 368
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.8698 | 384.1 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 7 | 195 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.3 | 367.9
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.8 | 384 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 7 | 195 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.2824 | 367.8
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.7699 | 384 | 0 | 0 | 14.5 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6.9 | 195.3 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 22 | 29.2044 | 367.8
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.6691 | 383.9 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.7 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.3089 | 21.8536 | 29.2732 | 367.7
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.5699 | 383.8 | 0 | 0 | 14.5 | 12.9 | 919 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 367.6
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.5302 | 383.7 | 0 | 0 | 14.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.6 | 59.9 | 0.031 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7706 | 29.2706 | 367.5
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.2 | 367.6 | 383.8 | 0 | 0 | 14.5 | 12.9 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.03097 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7073 | 29.2073 | 367.6
2008-11-03 | 14:01 | 0 | -0.0006 | -5.9506 | 16.2 | 367.7 | 383.9 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6.1 | 195.2 | 59.9 | 0.0309 | 0.0283 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.7853 | 29.2853 | 367.7
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.2 | 367.6696 | 383.9 | 0 | 0 | 14.5 | 13 | 918 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6.9 | 195.5 | 59.9 | 0.03087 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.3044 | 21.8 | 29.3 | 367.7
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.7 | 383.8 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6.1 | 195.8 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8 | 29.3 | 367.7
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.7302 | 383.8 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6.9 | 195.8 | 59.9698 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8 | 29.3 | 367.7
2008-11-03 | 14:01 | 0 | 0.0008 | 7.6684 | 16.1 | 367.7698 | 383.9 | 0 | 0 | 14.5 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6.1 | 195.4 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8 | 29.3 | 367.8
2008-11-03 | 14:01 | 0 | 0.001 | 10 | 16.1 | 367.7302 | 383.8 | 0 | 0 | 14.5 | 13 | 919 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 7 | 195 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018 | 21.8 | 29.2251 | 367.7
2008-11-03 | 14:01 | 0 | 0.001 | 10 | 16.1 | 367.5698 | 383.7 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 7 | 195 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1018.302 | 21.8529 | 29.2 | 367.6
2008-11-03 | 14:01 | 0 | 0.0007 | 7.4324 | 16.1 | 367.5 | 383.6 | 0 | 0 | 14.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9309 | 29.2309 | 367.5
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.4 | 383.5 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9912 | 29.2912 | 367.4
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.4 | 383.5 | 0 | 0 | 14.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195.4 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9133 | 29.2132 | 367.4
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.3 | 383.4 | 0 | 0 | 14.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 195.8 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7059 | 29.2647 | 367.3
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.3 | 383.4 | 0 | 0 | 14.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 196 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7278 | 29.2574 | 367.3
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.4 | 383.5 | 0 | 0 | 14.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 196 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.859 | 29.2205 | 367.4
2008-11-03 | 14:01 | 0 | 0 | 0 | 16.1 | 367.3301 | 383.4 | 0 | 0 | 14.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 196 | 59.9 | 0.0308 | 0.0282 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.703 | 29.2985 | 367.3
2008-11-03 | 14:01 | 0 | -0.0001 | -0.60189 | 16.1 | 367.3698 | 383.5 | 0 | 0 | 14.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 196 | 59.9 | 0.0308 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7765 | 29.3 | 367.4
2008-11-03 | 14:01 | 0 | -0.0009 | -9.2003 | 16.1 | 367.3308 | 383.4 | 0 | 0 | 14.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 7 | 196 | 59.9 | 0.0308 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.2456 | 367.3
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.1 | 367.4 | 383.5 | 0 | 0 | 14.5 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 7 | 196 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8646 | 29.2323 | 367.4
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.1 | 367.3698 | 383.5 | 0 | 0 | 14.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 7 | 196 | 60 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.3 | 367.4
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16.1 | 367.3302 | 383.4 | 0 | 0 | 14.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 7 | 196 | 59.9698 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 22 | 29.3 | 367.3
2008-11-03 | 14:01 | 0 | -0.001 | -10 | 16 | 367.3302 | 383.3 | 0 | 0 | 14.5 | 13 | 918 | 903 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6.9 | 196 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.934 | 29.234 | 367.3
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.4 | 383.4 | 0 | 0 | 14.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9439 | 29.2 | 367.4
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.3698 | 383.4 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9559 | 29.2221 | 367.4
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.4 | 383.4 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.8 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 367.4
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.4 | 383.4 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0307 | 0.0281 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.9559 | 29.2221 | 367.4
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.2302 | 383.2 | 0 | 0 | 14.5 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.1 | 59.9 | 0.03067 | 0.02807 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8883 | 29.2558 | 367.2
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 367.1698 | 383.2 | 0 | 0 | 14.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0306 | 0.028 | 0.0197 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.8 | 29.3 | 367.2
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 16 | 356.7443 | 372.7 | 0 | 0 | 14.43 | 13 | 919 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0306 | 0.028 | 0.01915 | 0.0003 |  | n/a |  | 34 | 60 | 1058 | 1019 | 21.7884 | 29.3 | 356.7
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 15.8 | 326.0806 | 341.9 | 0 | 0 | 13.36 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9699 | 0.030119 | 0.027519 | 0.017479 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7107 | 29.3 | 326.1
2008-11-03 | 14:02 | 0 | -0.0003 | -2.7951 | 15.6 | 321.1584 | 336.8 | 0 | 0 | 12.68 | 12.9 | 917 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.3 | 59.9 | 0.029458 | 0.026889 | 0.0172 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7674 | 29.3 | 321.2
2008-11-03 | 14:02 | 0 | -0.0006 | -5.647 | 15.4 | 318.5 | 333.9 | 0 | 0 | 12.6 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.7 | 59.9 | 0.02851 | 0.02601 | 0.0171 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.6994 | 21.7089 | 29.2544 | 318.5
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 15.2 | 318.3302 | 333.5 | 0 | 0 | 12.5 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.027809 | 0.025309 | 0.0171 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.647 | 29.2235 | 318.3
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 15.1 | 318.3698 | 333.4 | 0 | 0 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.02757 | 0.02507 | 0.0171 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8014 | 29.3 | 318.4
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 14.9 | 317.998 | 332.9 | 0 | 0 | 12.5 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.2 | 59.9 | 0.0274 | 0.02497 | 0.01707 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.8793 | 29.3 | 318
2008-11-03 | 14:02 | 0 | -0.001 | -10 | 14.9 | 317.1698 | 332 | 0 | 0 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0273 | 0.0249 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.9 | 29.3 | 317.2
2008-11-03 | 14:02 | -0.004 | -0.001 | -10 | 14.8 | 317.1302 | 331.9 | 0 | 0 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0273 | 0.0249 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.9 | 29.3 | 317.1
2008-11-03 | 14:02 | -0.01 | -0.0004 | -4.1708 | 14.8 | 316.9698 | 331.8 | 0 | 0 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.02727 | 0.02487 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8748 | 29.3 | 317
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.8 | 316.9698 | 331.8 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0272 | 0.02477 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.718 | 29.3 | 317
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.8 | 316.8698 | 331.7 | 0 | 0 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0271 | 0.0247 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.769 | 29.3 | 316.9
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.7 | 316.9 | 331.6 | 0 | 0 | 12.5 | 13 | 916 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0271 | 0.0247 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8469 | 29.3 | 316.9
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.7 | 316.8698 | 331.5 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6.1 | 196 | 59.9 | 0.027 | 0.0246 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.9249 | 29.3 | 316.9
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.6 | 316.7 | 331.3 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6.9 | 196 | 60 | 0.027 | 0.0246 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.9942 | 29.3 | 316.7
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.6 | 316.7302 | 331.3 | 0 | 0 | 12.5 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.027 | 0.0246 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8385 | 29.3 | 316.7
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.6 | 316.7 | 331.3 | 0 | 0 | 12.5 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.02697 | 0.0246 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.9166 | 29.3 | 316.7
2008-11-03 | 14:02 | -0.01 | 0 | 0 | 14.5 | 316.6397 | 331.1 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0269 | 0.0246 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.3014 | 21.9638 | 29.3 | 316.6
2008-11-03 | 14:02 | -0.003 | 0 | 0 | 14.5 | 316.5 | 331 | 0 | 0 | 12.5 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.8 | 59.9302 | 0.0269 | 0.02457 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.9 | 29.3 | 316.5
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.5 | 316.4 | 330.9 | 0 | 0 | 12.5 | 13 | 916 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0269 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.9 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.5 | 316.4 | 330.9 | 0 | 0 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.1 | 59.9 | 0.0269 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8296 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.4 | 330.8 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.0269 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8483 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.4 | 330.8 | 0 | 0 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0268 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8738 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.5 | 330.9 | 0 | 0 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9698 | 0.0268 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8 | 29.3 | 316.5
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.4 | 330.8 | 0 | 0 | 12.5 | 12.9 | 918 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0268 | 0.0245 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.4 | 330.8 | 0 | 0 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0268 | 0.02447 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8598 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.4 | 330.8 | 1.1 | 1.0872 | 12.5 | 12.9 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0268 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.7859 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.4 | 316.3698 | 330.7 | 2 | 2 | 12.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9302 | 0.0268 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.6319 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.2302 | 330.5 | 2 | 2 | 12.5 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9698 | 0.0268 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7877 | 29.3 | 316.2
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.3 | 330.6 | 3 | 3 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6564 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.2697 | 330.6 | 3 | 3 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.5 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6497 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.2302 | 330.5 | 3.9 | 3.9 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.9 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6726 | 29.3 | 316.2
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.3 | 330.6 | 4 | 4 | 12.5 | 13 | 917 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6048 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.3 | 330.6 | 4 | 4 | 12.5 | 12.9 | 918 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6833 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.3 | 330.6 | 4 | 4 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7614 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.4 | 330.7 | 4 | 4 | 12.5 | 12.9 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.7 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.8 | 29.3 | 316.4
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.3 | 316.3 | 330.6 | 4 | 4 | 12.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.3 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.7654 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.2 | 316.3 | 330.5 | 4 | 4 | 12.5 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6095 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.2 | 316.3 | 330.5 | 3.2 | 3.2045 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.4 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.7464 | 29.3 | 316.3
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.1 | 316.2 | 330.3 | 3 | 3 | 12.5 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.0267 | 0.0244 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8 | 29.3 | 316.2
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.1 | 316.1 | 330.2 | 3 | 3 | 12.5 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8289 | 29.3 | 316.1
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.1 | 316.1303 | 330.2 | 3 | 3 | 12.5 | 13 | 916 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.8653 | 29.3 | 316.1
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.1 | 316 | 330.1 | 2 | 2 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.4759 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14.1 | 316 | 330.1 | 2 | 2 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7137 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316.1 | 330.1 | 2 | 2 | 12.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7779 | 29.3 | 316.1
2008-11-03 | 14:02 | 0 | -0.0005 | -5.4825 | 14 | 316.0698 | 330.1 | 1 | 1 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6186 | 29.3 | 316.1
2008-11-03 | 14:02 | 0 | -0.0006 | -6.1917 | 14 | 316 | 330 | 1 | 1 | 12.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6966 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316 | 330 | 1 | 1 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6256 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316 | 330 | 0 | 0 | 12.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6523 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316 | 330 | 0 | 0 | 12.5 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.6 | 59.9699 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 315.9 | 329.9 | 0 | 0 | 12.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.2 | 59.9 | 0.0266 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6917 | 29.3 | 315.9
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316 | 330 | 0 | 0 | 12.5 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.02657 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6138 | 29.3 | 316
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316.0697 | 330.1 | 0 | 0 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.3026 | 21.6641 | 29.3 | 316.1
2008-11-03 | 14:02 | 0 | 0 | 0 | 14 | 316.1302 | 330.1 | 0 | 0 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.658 | 29.3 | 316.1
2008-11-03 | 14:03 | 0 | 0 | 0 | 14 | 316.2 | 330.2 | 0 | 0 | 12.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0243 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3 | 316.2
2008-11-03 | 14:03 | 0 | 0 | 0 | 14 | 316.1698 | 330.2 | 0 | 0 | 12.5 | 13 | 919 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3 | 316.2
2008-11-03 | 14:03 | 0 | 0 | 0 | 14 | 316.1 | 330.1 | 0 | 0 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4486 | 29.3757 | 316.1
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.9 | 316.0302 | 329.9 | 0 | 0 | 12.5 | 12.9 | 918 | 904 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.5074 | 29.3463 | 316
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.9 | 316.1 | 330 | 0 | 0 | 12.5 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.02647 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6633 | 29.3 | 316.1
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.9 | 315.9698 | 329.9 | 0 | 0 | 12.5 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9 | 0.0264 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.7808 | 29.3096 | 316
2008-11-03 | 14:03 | 0 | 0 | 0 | 14 | 315.9 | 329.9 | 0 | 0 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.02643 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6251 | 29.3875 | 315.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 14 | 314.2692 | 328.3 | 0 | 0 | 12.5 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0265 | 0.0242 | 0.017 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6 | 29.4 | 314.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.9 | 301.8 | 315.7 | 0 | 0 | 12.33 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0264 | 0.0242 | 0.0162 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6 | 29.4 | 301.8
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.8 | 287.1552 | 300.9 | 0 | 0 | 11.14 | 13 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.5 | 59.9 | 0.026009 | 0.023779 | 0.015419 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3789 | 287.2
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.6 | 266.6564 | 280.3 | 0 | 0 | 10.59 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.9 | 59.9 | 0.02537 | 0.02317 | 0.01437 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.6 | 29.3009 | 266.7
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.4 | 265.4 | 278.8 | 0 | 0 | 10.5 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.7 | 59.9 | 0.025119 | 0.022919 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3 | 265.4
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.2 | 265.3 | 278.5 | 0 | 0 | 10.4 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9 | 0.024279 | 0.022109 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3 | 265.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 13.1 | 265.3 | 278.4 | 0 | 0 | 10.4 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.024 | 0.02187 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3 | 265.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 13 | 265.3 | 278.3 | 0 | 0 | 10.4 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0239 | 0.0218 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5891 | 29.3 | 265.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 13 | 265.2 | 278.2 | 0 | 0 | 10.4 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0239 | 0.0218 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1017.698 | 21.5112 | 29.3 | 265.2
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.9 | 265.1699 | 278.1 | 0 | 0 | 10.4 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.02387 | 0.0217 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1017 | 21.5665 | 29.3 | 265.2
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.8 | 265.0302 | 277.8 | 0 | 0 | 10.4 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9 | 0.023609 | 0.021509 | 0.0143 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1017.698 | 21.6445 | 29.3445 | 265
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.6 | 264.9698 | 277.6 | 0 | 0 | 10.4 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.02337 | 0.02127 | 0.01427 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1017.302 | 21.6775 | 29.3775 | 265
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.5 | 264.8698 | 277.4 | 0 | 0 | 10.4 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.02317 | 0.02117 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5991 | 29.3005 | 264.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.4 | 264.8 | 277.2 | 0 | 0 | 10.4 | 13 | 915 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0231 | 0.0211 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4433 | 29.3783 | 264.8
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.4 | 264.7302 | 277.1 | 0 | 0 | 10.4 | 13 | 916 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0231 | 0.0211 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1017.698 | 21.4563 | 29.4 | 264.7
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.4 | 264.7 | 277.1 | 0 | 0 | 10.4 | 13 | 915 | 901 | 890 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.2 | 59.9 | 0.0231 | 0.0211 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.3658 | 264.7
2008-11-03 | 14:03 | 0 | 0 | 0 | 12.2 | 264.7 | 276.9 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.6 | 59.9 | 0.02307 | 0.02107 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5241 | 29.3121 | 264.7
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.9 | 264.6 | 276.5 | 0 | 0 | 10.4 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.02297 | 0.021 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6796 | 29.3898 | 264.6
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.9 | 264.6 | 276.5 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.7 | 59.9 | 0.0229 | 0.02097 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6321 | 29.4 | 264.6
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.9 | 264.5 | 276.4 | 0 | 0 | 10.4 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.3 | 59.9 | 0.0229 | 0.0209 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5082 | 29.4 | 264.5
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.9 | 264.4698 | 276.4 | 0 | 0 | 10.4 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0229 | 0.0209 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4477 | 29.4 | 264.5
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.4302 | 276.2 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0229 | 0.0209 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 264.4
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.4 | 276.2 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0228 | 0.0209 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 264.4
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.3 | 276.1 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0228 | 0.0209 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.3424 | 264.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.3 | 276.1 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0228 | 0.02087 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6354 | 29.3 | 264.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.3 | 276.1 | 0 | 0 | 10.4 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.9 | 59.9 | 0.0228 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6604 | 29.3132 | 264.3
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.8 | 264.4 | 276.2 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4262 | 29.3913 | 264.4
2008-11-03 | 14:03 | 0 | -0.0004 | -4.0864 | 11.7 | 264.4 | 276.1 | 0 | 0 | 10.4 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6075 | 29.3308 | 264.4
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.7 | 264.3698 | 276.1 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6529 | 29.3471 | 264.4
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.7 | 264.3 | 276 | 0 | 0 | 10.4 | 12.9 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9698 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5498 | 29.4 | 264.3
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.7 | 264.3 | 276 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4059 | 29.4 | 264.3
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.7 | 264.2 | 275.9 | 0 | 0 | 10.4 | 12.9 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5618 | 29.4 | 264.2
2008-11-03 | 14:03 | 0 | -0.0015 | -15.4996 | 11.7 | 264.2 | 275.9 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.6 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.6 | 29.4 | 264.2
2008-11-03 | 14:03 | 0 | -0.0016 | -16.1905 | 11.7 | 264.1698 | 275.9 | 0 | 0 | 10.4 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.2 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0003 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6 | 29.4 | 264.2
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.7 | 264.1302 | 275.8 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.00027 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.5854 | 29.3854 | 264.1
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 264.1 | 275.7 | 0 | 0 | 10.4 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.5075 | 29.3075 | 264.1
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 264.0302 | 275.6 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.5704 | 29.3704 | 264
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 264.1 | 275.7 | 0 | 0 | 10.4 | 13 | 916 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5034 | 29.4 | 264.1
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 264 | 275.6 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0227 | 0.0208 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.6994 | 21.4525 | 29.4 | 264
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 264.0699 | 275.7 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0226 | 0.0207 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5959 | 29.4 | 264.1
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.6 | 263.9698 | 275.6 | 0 | 0 | 10.4 | 13 | 917 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0226 | 0.0207 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5181 | 29.4 | 264
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.5 | 263.9 | 275.4 | 0 | 0 | 10.4 | 13 | 916 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0225 | 0.0206 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | -0.001 | -10 | 11.5 | 263.8301 | 275.3 | 0 | 0 | 10.4 | 13 | 916 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0225 | 0.0206 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.6986 | 21.5379 | 29.4 | 263.8
2008-11-03 | 14:03 | 0 | -0.0003 | -2.8956 | 11.5 | 264 | 275.5 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0225 | 0.0206 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 264
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.8698 | 275.4 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195 | 59.9 | 0.0225 | 0.0206 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.6 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.9 | 275.4 | 0 | 0 | 10.4 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0225 | 0.0206 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4567 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.9 | 275.4 | 0 | 0 | 10.4 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0224 | 0.02057 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4495 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.9 | 275.4 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.9 | 275.4 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4947 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.8 | 275.3 | 0 | 0 | 10.4 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0142 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.4168 | 29.4 | 263.8
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.8698 | 275.4 | 0 | 0 | 10.4 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0224 | 0.0205 | 0.01417 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5224 | 29.4 | 263.9
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.8 | 275.3 | 0 | 0 | 10.4 | 12.9 | 918 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.3 | 59.9698 | 0.0224 | 0.0205 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 263.8
2008-11-03 | 14:03 | 0 | 0 | 0 | 11.5 | 263.8 | 275.3 | 0 | 0 | 10.4 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.7 | 59.9 | 0.0223 | 0.0205 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5829 | 29.4 | 263.8
2008-11-03 | 14:04 | 0 | 0 | 0 | 11.4 | 263.8 | 275.2 | 0 | 0 | 10.4 | 12.9 | 917 | 904 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.9 | 59.9 | 0.0223 | 0.0204 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5051 | 29.4 | 263.8
2008-11-03 | 14:04 | 0 | 0 | 0 | 11.4 | 263.8 | 275.2 | 0 | 0 | 10.4 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.6 | 59.9 | 0.0223 | 0.0204 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4272 | 29.4 | 263.8
2008-11-03 | 14:04 | 0 | 0 | 0 | 11.4 | 263.7 | 275.1 | 0 | 0 | 10.4 | 12.9 | 918 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0223 | 0.0204 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5014 | 29.4 | 263.7
2008-11-03 | 14:04 | 0 | 0 | 0 | 11.4 | 263.7 | 275.1 | 0 | 0 | 10.4 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9 | 0.0223 | 0.0204 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 263.7
2008-11-03 | 14:04 | 0 | -0.0004 | -3.8889 | 11.4 | 263.6699 | 275.1 | 0 | 0 | 10.4 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0223 | 0.0204 | 0.0141 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.3006 | 21.6 | 29.4 | 263.7
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 11.4 | 258.4461 | 269.8 | 0 | 0 | 10.4 | 13 | 918 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0223 | 0.0204 | 0.01401 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6 | 29.4 | 258.4
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 11.2 | 230.9384 | 242.2 | 0 | 0 | 9.65 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.021868 | 0.019998 | 0.012419 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 230.9
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10.6 | 215.9295 | 226.5 | 0 | 0 | 8.71 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195 | 59.9 | 0.02038 | 0.01861 | 0.0116 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 215.9
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10.4 | 212.5 | 222.9 | 0 | 0 | 8.4 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.02 | 0.0183 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 212.5
2008-11-03 | 14:04 | 0 | -0.0004 | -4.4075 | 10.4 | 212.4 | 222.8 | 0 | 0 | 8.4 | 12.9 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.02 | 0.01827 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 212.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 10.3 | 212.4 | 222.7 | 0 | 0 | 8.4 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.019809 | 0.018109 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 212.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 10.2 | 212.2698 | 222.5 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.0195 | 0.0178 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.4969 | 29.4 | 212.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 10.2 | 212.2 | 222.4 | 0 | 0 | 8.3 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0194 | 0.01777 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.4597 | 29.4 | 212.2
2008-11-03 | 14:04 | 0 | 0 | 0 | 10.2 | 212.1698 | 222.4 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.01937 | 0.0177 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.6 | 29.4 | 212.2
2008-11-03 | 14:04 | 0 | -0.0007 | -7.0897 | 10.2 | 212.1 | 222.3 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0193 | 0.0176 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.6 | 29.4 | 212.1
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10.2 | 212.1 | 222.3 | 0 | 0 | 8.3 | 13 | 915 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0192 | 0.0176 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6 | 29.4 | 212.1
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10.1 | 212.0302 | 222.1 | 0 | 0 | 8.3 | 13 | 915 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.9 | 59.9 | 0.0192 | 0.0175 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.5172 | 29.4 | 212
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10.1 | 212.1698 | 222.3 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0191 | 0.0175 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4192 | 29.4 | 212.2
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10 | 212.1 | 222.1 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0191 | 0.0175 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4966 | 29.4 | 212.1
2008-11-03 | 14:04 | 0 | -0.001 | -10 | 10 | 212.1 | 222.1 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.01907 | 0.0174 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.575 | 29.4 | 212.1
2008-11-03 | 14:04 | 0 | -0.0003 | -2.9153 | 9.9 | 212.1 | 222 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.019 | 0.0174 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.547 | 29.4 | 212.1
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.9 | 212.0301 | 221.9 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.019 | 0.0174 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4691 | 29.4 | 212
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.8 | 212 | 221.8 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.018809 | 0.017209 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4089 | 29.4 | 212
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.7 | 211.9 | 221.6 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4868 | 29.4 | 211.9
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.7 | 211.8698 | 221.6 | 0 | 0 | 8.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.4 | 211.9
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.7 | 211.8 | 221.5 | 0 | 0 | 8.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5424 | 29.4 | 211.8
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.7 | 211.8 | 221.5 | 0 | 0 | 8.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5799 | 29.4 | 211.8
2008-11-03 | 14:04 | -0.002 | 0 | 0 | 9.7 | 211.8698 | 221.6 | 0 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.01847 | 0.01687 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5016 | 29.4 | 211.9
2008-11-03 | 14:04 | -0.01 | 0 | 0 | 9.6 | 211.8 | 221.4 | 0 | 0 | 8.3 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.8 | 59.9 | 0.01837 | 0.0168 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.4762 | 211.8
2008-11-03 | 14:04 | -0.01 | 0 | 0 | 9.6 | 211.8 | 221.4 | 0 | 0 | 8.3 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.4 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5542 | 29.4458 | 211.8
2008-11-03 | 14:04 | -0.01 | 0 | 0 | 9.6 | 211.7698 | 221.4 | 0 | 0 | 8.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5679 | 29.4 | 211.8
2008-11-03 | 14:04 | -0.01 | 0 | 0 | 9.6 | 211.6698 | 221.2 | 0 | 0 | 8.3 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.3 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5 | 29.4 | 211.7
2008-11-03 | 14:04 | -0.006 | 0 | 0 | 9.5 | 211.6 | 221.1 | 0 | 0 | 8.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.7 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.6966 | 21.5 | 29.4 | 211.6
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.6 | 221.1 | 0 | 0 | 8.3 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.6312 | 29.4656 | 211.6
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.6 | 221.1 | 0 | 0 | 8.3 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0183 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.613 | 29.5 | 211.6
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.5698 | 221.1 | 0 | 0 | 8.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0182 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.4784 | 29.4784 | 211.6
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.5 | 221 | 0 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.6994 | 21.4005 | 29.4005 | 211.5
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.4 | 220.9 | 0 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0167 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.4 | 29.4 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.5 | 211.4 | 220.9 | 0 | 0 | 8.3 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.3448 | 29.4552 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.3662 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.511 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.3302 | 220.7 | 0 | 0 | 8.3 | 13 | 916 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0182 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.5888 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 916 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9302 | 0.0181 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.5336 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.4556 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0114 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.4005 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.3302 | 220.7 | 0 | 0 | 8.3 | 13 | 915 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.01807 | 0.0165 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.4784 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.4 | 211.4 | 220.8 | 0 | 0 | 8.3 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.018 | 0.0165 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4437 | 29.4437 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.3 | 211.4 | 220.7 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.018 | 0.0165 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4684 | 29.4342 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.3 | 211.3698 | 220.7 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.9 | 59.9 | 0.018 | 0.0165 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.588 | 29.5 | 211.4
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.3 | 211.3 | 220.6 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.018 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5103 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.3 | 211.3 | 220.6 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.018 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4322 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.3 | 211.3 | 220.6 | 0 | 0 | 8.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.2 | 211.2699 | 220.5 | -0.8 | 0 | 8.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.7 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.2 | 211.2 | 220.4 | -1 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.3 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.2
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.2 | 211.3 | 220.5 | -1 | 0 | 8.3 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.3
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.2 | 211.2 | 220.4 | -1 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.4 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2852 | 29.5 | 211.2
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.1 | 211.1698 | 220.3 | -1 | 0 | 8.3 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.8 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3056 | 29.5 | 211.2
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | -1 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.487 | 29.5 | 211.2
2008-11-03 | 14:04 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | -1 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.409 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | -1 | 0 | 8.3 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | -1 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3064 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | -0.2 | 0 | 8.3 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.4 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2494 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | 0 | 0 | 8.3 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | 0 | 0 | 8.3 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 211.2
2008-11-03 | 14:05 | 0 | 0 | 0 | 9.1 | 211.2 | 220.3 | 0 | 0 | 8.3 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195 | 59.9 | 0.0179 | 0.0164 | 0.0113 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3416 | 29.5 | 211.2
2008-11-03 | 14:05 | -0.007 | 0 | 0 | 9 | 190.2062 | 199.2 | 0 | 0 | 8.17 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.01787 | 0.01637 | 0.010219 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3 | 29.5 | 190.2
2008-11-03 | 14:05 | -0.01 | 0 | 0 | 8.6 | 168.786 | 177.4 | 0 | 0 | 6.96 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9 | 0.016807 | 0.015338 | 0.009079 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3425 | 29.5 | 168.8
2008-11-03 | 14:05 | -0.006 | 0 | 0 | 8.3 | 163 | 171.3 | 0 | 0 | 6.49 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.8 | 59.9 | 0.0158 | 0.0144 | 0.0088 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5762 | 29.5 | 163
2008-11-03 | 14:05 | 0 | 0 | 0 | 8.2 | 162.3273 | 170.5 | 0 | 0 | 6.34 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.9 | 59.9 | 0.01541 | 0.01411 | 0.00877 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4601 | 29.5 | 162.3
2008-11-03 | 14:05 | 0 | 0 | 0 | 8.1 | 160.9 | 169 | 0 | 0 | 6.3 | 13 | 913 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0151 | 0.01377 | 0.0087 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 160.9
2008-11-03 | 14:05 | 0 | 0 | 0 | 8 | 160.8699 | 168.9 | 0 | 0 | 6.3 | 13 | 913 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.1 | 59.9 | 0.015 | 0.01367 | 0.0087 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 160.9
2008-11-03 | 14:05 | 0.007 | 0 | 0 | 7.9 | 160.7699 | 168.7 | 0 | 0 | 6.3 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0149 | 0.0136 | 0.0087 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 160.8
2008-11-03 | 14:05 | 0.01 | 0 | 0 | 7.9 | 160.6698 | 168.6 | 0 | 0 | 6.3 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.01487 | 0.01357 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 160.7
2008-11-03 | 14:05 | 0.01 | 0 | 0 | 7.8 | 160.6 | 168.4 | 0 | 0 | 6.3 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.01477 | 0.0135 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3408 | 29.5 | 160.6
2008-11-03 | 14:05 | 0.006 | 0 | 0 | 7.8 | 160.5698 | 168.4 | 0 | 0 | 6.3 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0147 | 0.01347 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2627 | 29.5 | 160.6
2008-11-03 | 14:05 | -0.01 | 0 | 0 | 7.8 | 160.5 | 168.3 | 0 | 0 | 6.3 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.01467 | 0.01337 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2304 | 29.5 | 160.5
2008-11-03 | 14:05 | -0.01 | -0.0002 | -1.8552 | 7.7 | 160.4 | 168.1 | -0.8 | 0 | 6.3 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.01457 | 0.0133 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3861 | 29.5 | 160.4
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.7 | 160.3 | 168 | 0 | 0 | 6.3 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0145 | 0.0132 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4709 | 29.5 | 160.3
2008-11-03 | 14:05 | -0.005 | -0.001 | -10 | 7.7 | 160.2698 | 168 | 0 | 0 | 6.3 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195 | 59.9 | 0.0145 | 0.0132 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.5488 | 29.5 | 160.3
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.7 | 160.1 | 167.8 | 0 | 0 | 6.3 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.2 | 59.9 | 0.01447 | 0.0132 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4933 | 29.5 | 160.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.6 | 160.1 | 167.7 | -0.4 | 0 | 6.3 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0144 | 0.01317 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2088 | 29.5044 | 160.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.6 | 160.1 | 167.7 | -1 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.9 | 59.9 | 0.0144 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3639 | 29.5819 | 160.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.6 | 160.0699 | 167.7 | -1 | 0 | 6.3 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0144 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5399 | 160.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.6 | 159.9 | 167.5 | -1 | 0 | 6.3 | 12.9 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0144 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3617 | 29.5 | 159.9
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.5 | 159.9 | 167.4 | -1 | 0 | 6.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.0144 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3324 | 29.5162 | 159.9
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.5 | 159.8698 | 167.4 | -1 | 0 | 6.3 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0143 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4882 | 29.5941 | 159.9
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.5 | 159.7699 | 167.3 | 0 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0143 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.428 | 29.528 | 159.8
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.5 | 159.7 | 167.2 | 0 | 0 | 6.3 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0143 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.6986 | 21.3502 | 29.5 | 159.7
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.5 | 159.6699 | 167.1 | -0.8 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.6 | 59.9 | 0.0143 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018.3014 | 21.3552 | 29.5276 | 159.7
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.4 | 159.6698 | 167 | -0.2 | 0 | 6.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195.2 | 59.9 | 0.0143 | 0.0131 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4948 | 29.6 | 159.7
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.3 | 159.6 | 166.9 | -0.8 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.01427 | 0.013 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4166 | 29.6 | 159.6
2008-11-03 | 14:05 | -0.004 | -0.001 | -10 | 7.3 | 159.6 | 166.9 | -1 | 0 | 6.3 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.5 | 59.9 | 0.0142 | 0.013 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.6 | 159.6
2008-11-03 | 14:05 | -0.01 | -0.0005 | -4.6274 | 7.3 | 159.6 | 166.9 | -1 | 0 | 6.3 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.9 | 59.9 | 0.0142 | 0.013 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5607 | 159.6
2008-11-03 | 14:05 | -0.01 | 0 | 0 | 7.3 | 159.6 | 166.9 | -1 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.8 | 59.9 | 0.0142 | 0.013 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 159.6
2008-11-03 | 14:05 | -0.002 | -0.0002 | -2.0662 | 7.3 | 159.6 | 166.9 | -0.2 | 0 | 6.3 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9 | 0.0142 | 0.013 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5 | 159.6
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.3 | 159.5698 | 166.9 | 0 | 0 | 6.3 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.01417 | 0.01297 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.5729 | 159.6
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.3 | 159.5 | 166.8 | -0.4 | 0 | 6.3 | 12.9 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.3 | 59.8302 | 0.0141 | 0.0129 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3493 | 29.6 | 159.5
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.5 | 166.7 | -1 | 0 | 6.2 | 13 | 915 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.7 | 59.9 | 0.0141 | 0.0129 | 0.0086 | 0.0002 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3286 | 29.6 | 159.5
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.5 | 166.7 | -1 | 0 | 6.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0141 | 0.0129 | 0.0086 | 0.00017 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.6 | 159.5
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.5 | 166.7 | 0 | 0 | 6.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0141 | 0.0129 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1017.698 | 21.4 | 29.6 | 159.5
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.5 | 166.7 | 0 | 0 | 6.2 | 13 | 915 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0141 | 0.0129 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4 | 29.6 | 159.5
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.4 | 166.6 | 0 | 0 | 6.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.01407 | 0.01287 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3196 | 29.5598 | 159.4
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.2 | 159.3699 | 166.6 | -1 | 0 | 6.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2361 | 29.5 | 159.4
2008-11-03 | 14:05 | -0.003 | -0.001 | -10 | 7.2 | 159.3 | 166.5 | -1 | 0 | 6.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.1 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.3014 | 21.3919 | 29.5 | 159.3
2008-11-03 | 14:05 | -0.01 | -0.0005 | -4.5333 | 7.2 | 159.2 | 166.4 | -1 | 0 | 6.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.3267 | 29.5733 | 159.2
2008-11-03 | 14:05 | -0.01 | 0 | 0 | 7.2 | 159.1 | 166.3 | -1 | 0 | 6.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2486 | 29.6 | 159.1
2008-11-03 | 14:05 | -0.01 | -0.0002 | -1.846 | 7.2 | 159.1302 | 166.3 | -1 | 0 | 6.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2593 | 29.6 | 159.1
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.00857 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.3925 | 29.6 | 159.2
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3146 | 29.6 | 159.2
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2368 | 29.6 | 159.2
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.014 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2411 | 29.6 | 159.2
2008-11-03 | 14:05 | -0.01 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 916 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9302 | 0.014 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.281 | 29.581 | 159.2
2008-11-03 | 14:05 | -0.002 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.014 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2033 | 29.5033 | 159.2
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.5 | 59.9 | 0.0139 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5742 | 159.2
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.1698 | 166.3 | -1 | 0 | 6.2 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.9 | 59.9 | 0.01393 | 0.0128 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.2 | 29.6 | 159.2
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.1 | 166.2 | -1 | 0 | 6.2 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.014 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2 | 29.6 | 159.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.1301 | 166.2 | -1 | 0 | 6.2 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.01397 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.6986 | 21.2085 | 29.6 | 159.1
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 13 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0139 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2863 | 29.6 | 159.2
2008-11-03 | 14:05 | 0 | -0.001 | -10 | 7.1 | 159.2 | 166.3 | -1 | 0 | 6.2 | 12.9 | 917 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.0139 | 0.0128 | 0.0086 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2358 | 29.5358 | 159.2
2008-11-03 | 14:06 | 0 | -0.001 | -9.9084 | 7.1 | 159.1 | 166.2 | -1 | 0 | 6.2 | 13 | 917 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0139 | 0.0127 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5 | 159.1
2008-11-03 | 14:06 | 0 | -0.0002 | -1.582 | 7.1 | 159.1 | 166.2 | -1 | 0 | 6.2 | 12.9 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0139 | 0.0127 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5199 | 159.1
2008-11-03 | 14:06 | 0 | -0.0014 | -13.9655 | 7.1 | 159.1 | 166.2 | -1 | 0 | 6.2 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0139 | 0.0127 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5974 | 159.1
2008-11-03 | 14:06 | 0 | -0.002 | -20 | 7.1 | 159.0698 | 166.2 | -1 | 0 | 6.2 | 13 | 917 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0139 | 0.0127 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.2 | 29.6 | 159.1
2008-11-03 | 14:06 | 0 | -0.0016 | -16.3274 | 7.1 | 159 | 166.1 | -1 | 0 | 6.2 | 13 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195 | 59.9 | 0.0139 | 0.0127 | 0.0085 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.6 | 159
2008-11-03 | 14:06 | 0 | -0.0004 | -4.4847 | 7.1 | 158.2633 | 165.3 | 0 | 0 | 6.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0139 | 0.0127 | 0.00847 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2314 | 29.6 | 158.3
2008-11-03 | 14:06 | 0 | 0.002 | 19.8335 | 7 | 141.3278 | 148.3 | 0 | 0 | 6.19 | 12.9 | 917 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.01387 | 0.0127 | 0.007749 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.698 | 21.2722 | 29.6 | 141.3
2008-11-03 | 14:06 | 0 | 0.0003 | 3.1807 | 6.5 | 114.6088 | 121.1 | 0 | 0 | 4.98 | 13 | 916 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.2 | 59.9 | 0.012119 | 0.011019 | 0.006209 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.0383 | 29.6 | 114.6
2008-11-03 | 14:06 | 0 | 0 | 0 | 6.3 | 110.0542 | 116.3 | 0 | 0 | 4.4 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.011369 | 0.010269 | 0.005969 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.3255 | 29.6 | 110.1
2008-11-03 | 14:06 | 0 | 0 | 0 | 6.2 | 108.6 | 114.8 | -0.8 | 0 | 4.3 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.01127 | 0.0102 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.4141 | 29.6 | 108.6
2008-11-03 | 14:06 | 0 | 0 | 0 | 6.1 | 108.4698 | 114.6 | -1 | 0 | 4.3 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0112 | 0.0102 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2792 | 29.6 | 108.5
2008-11-03 | 14:06 | 0 | 0 | 0 | 6.1 | 108.3692 | 114.5 | -1 | 0 | 4.2 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.011169 | 0.010169 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2013 | 29.6 | 108.4
2008-11-03 | 14:06 | 0 | 0 | 0 | 6 | 108.3 | 114.3 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.01107 | 0.01007 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1234 | 29.6 | 108.3
2008-11-03 | 14:06 | 0 | 0 | 0 | 6 | 108.2 | 114.2 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9 | 0.011 | 0.01 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1544 | 29.5456 | 108.2
2008-11-03 | 14:06 | 0 | 0 | 0 | 6 | 108.1 | 114.1 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.8 | 59.9 | 0.011 | 0.01 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.103 | 29.5323 | 108.1
2008-11-03 | 14:06 | 0 | 0 | 0 | 6 | 108.1 | 114.1 | -0.2 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.01097 | 0.00997 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 20.9102 | 29.5898 | 108.1
2008-11-03 | 14:06 | 0 | 0 | 0 | 6 | 108.1 | 114.1 | 0 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.0109 | 0.0099 | 0.0059 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 20.988 | 29.512 | 108.1
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.9 | 108 | 113.9 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0109 | 0.0099 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1311 | 29.5655 | 108
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.9 | 107.9699 | 113.8 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.8 | 59.9 | 0.01087 | 0.00987 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5565 | 108
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.8 | 107.8 | 113.6 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195.4 | 59.9 | 0.010769 | 0.0098 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5217 | 107.8
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.7 | 107.7 | 113.4 | -1 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0107 | 0.00977 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5996 | 107.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.7 | 107.6698 | 113.4 | -1 | 0 | 4.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.3 | 59.9 | 0.01067 | 0.0097 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 20.9677 | 29.6 | 107.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.7 | 107.6 | 113.3 | -1 | 0 | 4.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195.7 | 59.9 | 0.0106 | 0.0097 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.0659 | 29.5447 | 107.6
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.6 | 107.6 | 113.2 | -1 | 0 | 4.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0106 | 0.00967 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.5332 | 107.6
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.5 | 107.5 | 113 | -1 | 0 | 4.2 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0105 | 0.0096 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.189 | 29.6 | 107.5
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.4 | 107.5 | 112.9 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0105 | 0.0096 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1112 | 29.6 | 107.5
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.4 | 107.5 | 112.9 | -1 | 0 | 4.2 | 12.9 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1 | 29.6 | 107.5
2008-11-03 | 14:06 | 0.002 | -0.0005 | -5.3935 | 5.4 | 107.5 | 112.9 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1 | 29.6 | 107.5
2008-11-03 | 14:06 | 0.01 | -0.001 | -10 | 5.3 | 107.4 | 112.7 | -1 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1225 | 29.6 | 107.4
2008-11-03 | 14:06 | 0.002 | -0.0006 | -6.2647 | 5.3 | 107.4 | 112.7 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1995 | 29.6 | 107.4
2008-11-03 | 14:06 | 0 | 0.001 | 10 | 5.3 | 107.3698 | 112.7 | -1 | 0 | 4.2 | 12.9 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1217 | 29.6 | 107.4
2008-11-03 | 14:06 | 0 | 0.001 | 10 | 5.3 | 107.3 | 112.6 | -1 | 0 | 4.2 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0104 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.0438 | 29.6 | 107.3
2008-11-03 | 14:06 | 0 | 0.0003 | 3.3139 | 5.3 | 107.3 | 112.6 | -1 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0103 | 0.0095 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21 | 29.6 | 107.3
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.3 | 107.3 | 112.6 | -1 | 0 | 4.2 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 196 | 59.9 | 0.0103 | 0.00947 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 20.9882 | 29.6 | 107.3
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.3 | 107.3 | 112.6 | -0.9 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0103 | 0.0094 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 20.9103 | 29.6 | 107.3
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.3 | 107.2 | 112.5 | 0 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.6 | 59.9 | 0.0103 | 0.0094 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.0352 | 29.6 | 107.2
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.3 | 107.2 | 112.5 | 0 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.2 | 59.9 | 0.0103 | 0.0094 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.0546 | 29.6 | 107.2
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.2 | 107.2 | 112.4 | 0 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0102 | 0.0094 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.0466 | 29.6 | 107.2
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.2 | 107.1697 | 112.4 | -0.8 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0102 | 0.00937 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.6 | 107.2
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.2 | 107.1 | 112.3 | -1 | 0 | 4.2 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.9 | 59.9 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.2 | 29.6 | 107.1
2008-11-03 | 14:06 | 0 | -0.0004 | -3.75 | 5.2 | 107.1 | 112.3 | -1 | 0 | 4.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018 | 21.1431 | 29.6 | 107.1
2008-11-03 | 14:06 | 0 | -0.0008 | -7.9984 | 5.2 | 107.0698 | 112.3 | -1 | 0 | 4.2 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1018.302 | 21.1346 | 29.6 | 107.1
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.2 | 107 | 112.2 | -0.1 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2122 | 29.6 | 107
2008-11-03 | 14:06 | 0 | 0 | 0 | 5.1 | 107 | 112.1 | 0 | 0 | 4.2 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.8698 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.2905 | 29.6 | 107
2008-11-03 | 14:06 | 0 | 0.0007 | 6.6694 | 5.1 | 107 | 112.1 | 0 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0102 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.3 | 29.6 | 107
2008-11-03 | 14:06 | 0 | 0.001 | 10 | 5.1 | 107 | 112.1 | 0 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.01017 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1057 | 1019 | 21.1613 | 29.6 | 107
2008-11-03 | 14:06 | 0 | 0.0006 | 6.3636 | 5.1 | 106.9699 | 112.1 | 0 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0101 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056.9943 | 1018.3006 | 21.0241 | 29.6 | 107
2008-11-03 | 14:06 | 0 | 0 | 0 | 5 | 106.9 | 111.9 | -0.8 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.2 | 59.9 | 0.0101 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1019 | 21.1 | 29.6 | 106.9
2008-11-03 | 14:06 | 0 | 0 | 0 | 5 | 106.9 | 111.9 | -0.2 | 0 | 4.2 | 13 | 915 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0101 | 0.0093 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1 | 29.6 | 106.9
2008-11-03 | 14:06 | 0 | 0 | 0 | 5 | 106.8698 | 111.9 | -0.8 | 0 | 4.2 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0101 | 0.0092 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1 | 29.6 | 106.9
2008-11-03 | 14:06 | 0 | 0 | 0 | 5 | 106.8 | 111.8 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.01 | 0.0092 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0646 | 29.6 | 106.8
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.8 | 111.7 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.01 | 0.0092 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0132 | 29.6 | 106.8
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.8 | 111.7 | -1 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.01 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0913 | 29.6 | 106.8
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7698 | 111.7 | -1 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.302 | 21.1692 | 29.6 | 106.8
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.7 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1019 | 21.153 | 29.6 | 106.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195.3 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.698 | 21.1 | 29.6 | 106.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 195 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.698 | 21.1027 | 29.6 | 106.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 195 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.302 | 21.1805 | 29.6 | 106.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.084 | 29.6 | 106.7
2008-11-03 | 14:06 | 0 | 0 | 0 | 4.9 | 106.7 | 111.6 | -1 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 106.7
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.9 | 106.6698 | 111.5 | -1 | 0 | 4.2 | 13 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.5 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0141 | 29.6 | 106.7
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.8 | 106.6 | 111.4 | -0.3 | 0 | 4.2 | 12.9 | 916 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.9 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.092 | 29.6 | 106.6
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.8 | 106.6 | 111.4 | -0.8 | 0 | 4.2 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 196 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1699 | 29.6 | 106.6
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.8 | 106.6 | 111.4 | 0 | 0 | 4.2 | 12.9 | 916 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0099 | 0.0091 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0568 | 29.6 | 106.6
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.8 | 106.6 | 111.4 | 0 | 0 | 4.2 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.00987 | 0.00907 | 0.0058 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9768 | 29.6 | 106.6
2008-11-03 | 14:07 | 0 | 0 | 0 | 4.8 | 105.4826 | 110.3 | 0 | 0 | 4.2 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0098 | 0.009 | 0.00574 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1966 | 29.6 | 105.5
2008-11-03 | 14:07 | 0 | 0 | 0 | 4 | 82.0016 | 86 | 0 | 0 | 3.9 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 196 | 59.8306 | 0.009486 | 0.008816 | 0.004447 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1189 | 29.6 | 82
2008-11-03 | 14:07 | 0 | 0 | 0 | 3.1 | 63.4204 | 66.6 | 0 | 0 | 2.79 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.008089 | 0.007519 | 0.003479 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0412 | 29.6 | 63.4
2008-11-03 | 14:07 | -0.003 | 0 | 0 | 2.9 | 56.2584 | 59.2 | 0 | 0 | 2.3 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.6 | 59.9 | 0.0073 | 0.0068 | 0.0031 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 21.0367 | 29.6 | 56.3
2008-11-03 | 14:07 | -0.01 | 0 | 0 | 2.9 | 55.6698 | 58.6 | 0 | 0 | 2.2 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.2 | 59.9 | 0.0072 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 21.1 | 29.6 | 55.7
2008-11-03 | 14:07 | -0.01 | 0 | 0 | 2.9 | 55.5698 | 58.5 | 0 | 0 | 2.2 | 13 | 911 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.1 | 59.9 | 0.00717 | 0.00667 | 0.0031 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1 | 29.6 | 55.6
2008-11-03 | 14:07 | -0.01 | -0.0002 | -1.8115 | 2.9 | 55.5 | 58.4 | 0 | 0 | 2.2 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.5 | 59.9 | 0.00707 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.302 | 21.0294 | 29.6 | 55.5
2008-11-03 | 14:07 | -0.003 | -0.001 | -9.935 | 2.8 | 55.4698 | 58.2 | -0.8 | 0 | 2.2 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.9 | 59.9 | 0.006909 | 0.006409 | 0.0031 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.5
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.7 | 55.4 | 58.1 | -1 | 0 | 2.2 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0066 | 0.0062 | 0.0031 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.4
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.7 | 55.3698 | 58.1 | -1 | 0 | 2.2 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0065 | 0.00607 | 0.0031 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.302 | 21 | 29.6 | 55.4
2008-11-03 | 14:07 | 0 | -0.0005 | -4.8787 | 2.7 | 55.3 | 58 | -1 | 0 | 2.2 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 196 | 59.9 | 0.00647 | 0.006 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.302 | 21 | 29.6 | 55.3
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.3 | 57.9 | -1 | 0 | 2.2 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.0064 | 0.006 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.3
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.3 | 57.9 | -1 | 0 | 2.2 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 196 | 59.9 | 0.00637 | 0.0059 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.3
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.2302 | 57.8 | -1 | 0 | 2.2 | 13 | 912 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 196 | 59.9 | 0.00627 | 0.00587 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.0312 | 29.6 | 55.2
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.2 | 57.8 | -1 | 0 | 2.2 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0062 | 0.0058 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21.1869 | 29.6 | 55.2
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.2 | 57.8 | -1 | 0 | 2.2 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0062 | 0.0058 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9861 | 29.6 | 55.2
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.1 | 57.7 | -1 | 0 | 2.1 | 12.9 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0062 | 0.0058 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9491 | 29.6 | 55.1
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.6 | 55.1 | 57.7 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 196 | 59.9 | 0.0062 | 0.00577 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.1
2008-11-03 | 14:07 | 0 | -0.0002 | -1.8115 | 2.6 | 55.0699 | 57.7 | -1 | 0 | 2.1 | 12.9 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0062 | 0.0057 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55.1
2008-11-03 | 14:07 | 0 | -0.001 | -9.935 | 2.6 | 55 | 57.6 | -1 | 0 | 2.1 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0061 | 0.0057 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 55
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.6 | 55 | 57.6 | -1 | 0 | 2.1 | 12.9 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0061 | 0.00567 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9394 | 29.6 | 55
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.6 | 55 | 57.6 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189 | 32 | 6 | 196 | 59.9 | 0.0061 | 0.0056 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8615 | 29.6 | 55
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.6 | 55 | 57.6 | -1 | 0 | 2.1 | 12.9 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.1 | 32 | 6 | 196 | 59.9 | 0.006 | 0.0056 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8163 | 29.6 | 55
2008-11-03 | 14:07 | 0 | -0.001 | -10 | 2.5 | 54.9302 | 57.4 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.006 | 0.0056 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8942 | 29.6 | 54.9
2008-11-03 | 14:07 | 0 | -0.0008 | -8.1885 | 2.5 | 55 | 57.5 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.006 | 0.00557 | 0.003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017.698 | 20.972 | 29.6 | 55
2008-11-03 | 14:07 | -0.007 | 0 | -0.064988 | 2.5 | 55 | 57.5 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.6 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1017 | 21 | 29.6 | 55
2008-11-03 | 14:07 | -0.005 | 0 | 0 | 2.5 | 55 | 57.5 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195.2 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 21 | 29.6 | 55
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.5 | 55 | 57.5 | -1 | 0 | 2.1 | 12.9 | 914 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.1 | 32 | 6 | 195.1 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017 | 21 | 29.6 | 55
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.5 | 54.9 | 57.4 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.5 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017 | 21 | 29.6 | 54.9
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.5 | 54.9 | 57.4 | -1 | 0 | 2.1 | 12.9 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.8 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1017 | 20.9387 | 29.6 | 54.9
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.5 | 54.8699 | 57.4 | -0.2 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195.8 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.3006 | 20.9392 | 29.6 | 54.9
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.5 | 54.8 | 57.3 | 0 | 0 | 2.1 | 13 | 914 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.4 | 59.9 | 0.0059 | 0.0055 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.617 | 54.8
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.8 | 57.2 | -0.9 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0059 | 0.00547 | 0.003 | 0.0001 |  | n/a |  | 34.5425 | 60 | 1056 | 1018 | 21 | 29.6949 | 54.8
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.8 | 57.2 | -1 | 0 | 2.1 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34.4575 | 60 | 1056 | 1018 | 20.9273 | 29.6273 | 54.8
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.8 | 57.2 | -1 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.6504 | 54.8
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.8 | 57.2 | -1 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.7 | 54.8
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.7 | 57.1 | -1 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.7008 | 20.9063 | 29.6937 | 54.7
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.4 | 54.7 | 57.1 | -1 | 0 | 2.1 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1019 | 20.9841 | 29.6159 | 54.7
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.3 | 54.7 | 57 | -1 | 0 | 2.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1019 | 20.876 | 29.662 | 54.7
2008-11-03 | 14:07 | 0 | 0 | 0 | 2.3 | 54.6698 | 57 | -1 | 0 | 2.1 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.0058 | 0.0054 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018.302 | 20.8797 | 29.7 | 54.7
2008-11-03 | 14:07 | 0 | -0.0005 | -5.099 | 2.3 | 54.6 | 56.9 | -0.2 | 0 | 2.1 | 13 | 913 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9698 | 0.00577 | 0.00537 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6823 | 54.6
2008-11-03 | 14:07 | 0 | -0.0007 | -6.568 | 2.3 | 54.6 | 56.9 | 0 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.8 | 59.9 | 0.0057 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6044 | 54.6
2008-11-03 | 14:07 | 0 | -0.0002 | -1.8101 | 2.3 | 54.7 | 57 | -0.8 | 0 | 2.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.4 | 59.9 | 0.0057 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6734 | 54.7
2008-11-03 | 14:07 | 0 | -0.001 | -9.9269 | 2.3 | 54.7 | 57 | -1 | 0 | 2.1 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0057 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9488 | 29.7 | 54.7
2008-11-03 | 14:07 | 0 | -0.0002 | -1.7964 | 2.3 | 54.7 | 57 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.0057 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.698 | 20.9 | 29.671 | 54.7
2008-11-03 | 14:07 | -0.003 | 0 | 0 | 2.3 | 54.6698 | 57 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0056 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.302 | 20.9 | 29.6 | 54.7
2008-11-03 | 14:07 | -0.01 | -0.0005 | -4.9198 | 2.3 | 54.6 | 56.9 | -1 | 0 | 2.1 | 13 | 914 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 195 | 59.9 | 0.0056 | 0.0053 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.6 | 54.6
2008-11-03 | 14:07 | -0.01 | -0.001 | -10 | 2.3 | 54.6 | 56.9 | -1 | 0 | 2.1 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0056 | 0.00527 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8374 | 29.6626 | 54.6
2008-11-03 | 14:07 | -0.01 | -0.0006 | -6.3769 | 2.3 | 54.6 | 56.9 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37.054 | 189 | 32 | 6 | 195 | 59.9 | 0.0056 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8405 | 29.6595 | 54.6
2008-11-03 | 14:07 | -0.003 | 0.001 | 9.87 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 12.9 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195.2 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.6 | 54.6
2008-11-03 | 14:07 | 0 | 0.001 | 10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195.5 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.6 | 54.6
2008-11-03 | 14:07 | 0 | 0.001 | 10 | 2.2 | 54.6698 | 56.9 | -1 | 0 | 2.1 | 12.9 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37.946 | 189.1 | 32 | 6 | 195.9 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.302 | 20.9 | 29.6737 | 54.7
2008-11-03 | 14:07 | 0 | 0.001 | 10 | 2.2 | 54.7 | 56.9 | -1 | 0 | 2.1 | 13 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.9 | 29.6483 | 54.7
2008-11-03 | 14:07 | 0 | 0.0007 | 6.6015 | 2.2 | 54.6699 | 56.9 | -1 | 0 | 2.1 | 12.9 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.3014 | 20.9298 | 29.6 | 54.7
2008-11-03 | 14:07 | 0 | -0.0002 | -1.8115 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9924 | 29.6076 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -9.935 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190 | 32 | 6 | 196 | 59.9 | 0.0055 | 0.0052 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9145 | 29.6855 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37.054 | 190.1 | 32 | 6 | 196 | 59.9 | 0.0055 | 0.0051 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.7 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.8698 | 0.0055 | 0.0051 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9 | 29.6588 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 915 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 195.6 | 59.8302 | 0.0055 | 0.0051 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.881 | 29.619 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37.946 | 191 | 32 | 6 | 195.2 | 59.9 | 0.0055 | 0.0051 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8031 | 29.6969 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.2 | 54.6 | 56.8 | -1 | 0 | 2.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.1 | 32 | 6 | 195 | 59.9 | 0.0055 | 0.0051 | 0.003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.7254 | 29.6254 | 54.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2.1 | 52.5162 | 54.6 | -1 | 0 | 2.1 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0055 | 0.0051 | 0.002909 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8049 | 29.6525 | 52.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 2 | 22.778 | 24.8 | -1 | 0 | 1.35 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.005189 | 0.004889 | 0.001319 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8696 | 29.7 | 22.8
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.6 | 5.3 | 6.9 | -1 | 0 | 0.5 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.003679 | 0.003409 | 0.0004 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8083 | 29.7 | 5.3
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.4 | 3.4 | 4.8 | -1 | 0 | 0.2 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0034 | 0.0032 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8861 | 29.7 | 3.4
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.4 | 3.2699 | 4.6 | -1 | 0 | 0.1 | 13 | 910 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195 | 59.9 | 0.0033 | 0.0031 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.964 | 29.636 | 3.3
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.3 | 3.1699 | 4.4 | -1 | 0 | 0.1 | 13 | 910 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.3 | 59.9 | 0.003179 | 0.00291 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 21 | 29.6 | 3.2
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.2 | 3.0698 | 4.2 | -1 | 0 | 0.1 | 13 | 910 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 195.7 | 59.9 | 0.002779 | 0.002609 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9804 | 29.6196 | 3.1
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.1 | 3 | 4.1 | -1 | 0 | 0.1 | 13 | 910 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 192 | 32 | 6 | 196 | 59.9 | 0.00247 | 0.00227 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.903 | 29.697 | 3
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 1.1 | 3 | 4.1 | -1 | 0 | 0.1 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191.9 | 32 | 6 | 195.6 | 59.9 | 0.00237 | 0.0022 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8248 | 29.7 | 3
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.9 | 2.9 | 3.8 | -1 | 0 | 0.1 | 13 | 910 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195.2 | 59.9 | 0.0023 | 0.0022 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9064 | 29.7 | 2.9
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.9 | 2.9 | 3.8 | -1 | 0 | 0.1 | 13 | 911 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 191 | 32 | 6 | 195 | 59.9 | 0.0023 | 0.00217 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.9068 | 29.6689 | 2.9
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.9 | 2.8698 | 3.7 | -1 | 0 | 0.1 | 13 | 911 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 37 | 190.9 | 32 | 6 | 195 | 59.9 | 0.00227 | 0.0021 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.7089 | 29.6089 | 2.9
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.8 | 2.7698 | 3.6 | -1 | 0 | 0.1 | 13 | 912 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 37.054 | 190 | 32 | 6 | 195 | 59.9 | 0.0022 | 0.0021 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.7868 | 29.6868 | 2.8
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.8 | 2.7 | 3.5 | -1 | 0 | 0.1 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195 | 59.9 | 0.00217 | 0.002 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8 | 29.7 | 2.7
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.7 | 2.5698 | 3.3 | -1 | 0 | 0.1 | 12.9 | 912 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189.9 | 32 | 6 | 195 | 59.9 | 0.0021 | 0.002 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.8424 | 29.7 | 2.6
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.7 | 2.5 | 3.2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.0021 | 0.002 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8797 | 29.7 | 2.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.7 | 2.5 | 3.2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.0021 | 0.002 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.802 | 29.7 | 2.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.6 | 2.5 | 3.1 | -1 | 0 | 0.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.002 | 0.0019 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8759 | 29.7 | 2.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.6 | 2.5 | 3.1 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.002 | 0.0019 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017.698 | 20.9 | 29.7 | 2.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.6 | 2.5 | 3.1 | -1 | 0 | 0.1 | 12.9 | 914 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.0019 | 0.0018 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1017 | 20.9 | 29.6684 | 2.5
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.6 | 2.4 | 3 | -1 | 0 | 0.1 | 13 | 913 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195 | 59.9 | 0.0019 | 0.0018 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017 | 20.9095 | 29.6 | 2.4
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.6 | 2.4 | 3 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189.1 | 32 | 6 | 195.2 | 59.8 | 0.0018 | 0.0017 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.9873 | 29.6 | 2.4
2008-11-03 | 14:08 | 0 | -0.001 | -10 | 0.5 | 2.4 | 2.9 | -1 | 0 | 0.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195.6 | 59.9 | 0.00177 | 0.00167 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8697 | 29.6651 | 2.4
2008-11-03 | 14:08 | 0 | -0.0009 | -8.5424 | 0.5 | 2.3699 | 2.8 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 196 | 59.8699 | 0.0017 | 0.0016 | 0.0003 | 0.0001 |  | n/a |  | 34.5425 | 60 | 1056 | 1018 | 20.8 | 29.6572 | 2.4
2008-11-03 | 14:08 | 0 | 0 | -0.067797 | 0.4 | 2.3 | 2.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195.7 | 59.9 | 0.0016 | 0.0016 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1017.6986 | 20.8 | 29.6206 | 2.3
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.3 | 2.7 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190.1 | 32 | 6 | 195.3 | 59.9 | 0.0016 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8 | 29.6987 | 2.3
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.3 | 2.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 195 | 59.9 | 0.0016 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.8766 | 29.7 | 2.3
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.3 | 2.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 195 | 59.8 | 0.0016 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017 | 20.8456 | 29.7 | 2.3
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.3 | 2.7 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191.1 | 32 | 6 | 195 | 59.9 | 0.00157 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.8 | 29.6678 | 2.3
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.2 | 2.6 | -1 | 0 | 0.1 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.1 | 59.8302 | 0.0015 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.81 | 29.61 | 2.2
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.2 | 2.6 | -1 | 0 | 0.1 | 13 | 912 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.5 | 59.9 | 0.00153 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8879 | 29.6879 | 2.2
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.2 | 2.6 | -1 | 0 | 0.1 | 13 | 911 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.8 | 59.9 | 0.0015 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017.698 | 20.9656 | 29.7 | 2.2
2008-11-03 | 14:08 | 0 | 0 | 0 | 0.4 | 2.2 | 2.6 | -1 | 0 | 0.1 | 13 | 912 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.8 | 59.8302 | 0.0015 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8708 | 29.7 | 2.2
2008-11-03 | 14:08 | -0.001 | 0 | 0 | 0.4 | 2.2 | 2.6 | -1 | 0 | 0.1 | 13 | 911 | 901 | 890 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.5 | 59.9 | 0.0015 | 0.0015 | 0.0003 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.7 | 29.7 | 2.2
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.4 | 2.1 | 2.5 | -1 | 0 | 0.1 | 13 | 912 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.1 | 59.9 | 0.0015 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.7 | 29.7 | 2.1
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.4 | 2.1 | 2.5 | -1 | 0 | 0.1 | 13 | 911 | 902 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.7 | 29.7 | 2.1
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.4 | 2 | 2.4 | -1 | 0 | 0.1 | 13 | 912 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.755 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | -0.0006 | -6.347 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.3 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.8 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | -0.001 | -10 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191.9 | 32 | 6 | 195.7 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.8 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | -0.001 | -10 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.8 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | -0.0009 | -8.6007 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1017.698 | 20.844 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | 0 | -0.068259 | 0.3 | 2 | 2.3 | -1 | 0 | 0.1 | 13 | 912 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190.9 | 32 | 6 | 196 | 59.9 | 0.0015 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017 | 20.878 | 29.7 | 2
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 196 | 59.9 | 0.0014 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017 | 20.8002 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 196 | 59.8 | 0.0014 | 0.00137 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017 | 20.8777 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189.9 | 32 | 6 | 196 | 59.8 | 0.0013 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.7889 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 196 | 59.8698 | 0.0013 | 0.00127 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.7 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 12.9 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 196 | 59.9 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017.698 | 20.7112 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.3 | 1.9 | 2.2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 196 | 59.9 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.7888 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189 | 32 | 6 | 195.7 | 59.9 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.8 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 13 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 189.1 | 32 | 6 | 195.3 | 59.9 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.8 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 13 | 913 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195.1 | 59.9 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.8226 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 12.9 | 914 | 903 | 894 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195.4 | 59.8698 | 0.0013 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8996 | 29.7 | 1.9
2008-11-03 | 14:08 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190 | 32 | 6 | 195.8 | 59.9 | 0.0012 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.8218 | 29.7 | 1.9
2008-11-03 | 14:09 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 12.9 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 190.1 | 32 | 6 | 196 | 59.8698 | 0.0012 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.856 | 29.7 | 1.9
2008-11-03 | 14:09 | -0.01 | 0 | 0 | 0.2 | 1.9 | 2.1 | -1 | 0 | 0.1 | 13 | 913 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.8302 | 0.0012 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 34.5425 | 60 | 1056 | 1018 | 20.7984 | 29.7 | 1.9
2008-11-03 | 14:09 | -0.01 | 0 | 0 | 0.2 | 1.8 | 2 | -1 | 0 | 0.1 | 12.9 | 914 | 903 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.9 | 0.0012 | 0.0012 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.6117 | 29.7 | 1.8
2008-11-03 | 14:09 | -0.009 | 0 | 0 | 0.2 | 1.8 | 2 | -1 | 0 | 0.1 | 13 | 913 | 903 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191 | 32 | 6 | 196 | 59.9 | 0.0012 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.6895 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.2 | 1.8 | 2 | -1 | 0 | 0.1 | 13 | 913 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 191.1 | 32 | 6 | 196 | 59.8301 | 0.0012 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1018 | 20.7 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.1 | 1.8 | 1.9 | -1 | 0 | 0.1 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 196 | 59.9 | 0.0012 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 35 | 60 | 1056 | 1017.698 | 20.7453 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.1 | 1.8 | 1.9 | -1 | 0 | 0.1 | 13 | 912 | 902 | 893 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.7 | 59.9 | 0.0011 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.7538 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.1 | 1.8 | 1.9 | -1 | 0 | 0.1 | 13 | 912 | 902 | 892 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195.3 | 59.9 | 0.0011 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 34.543 | 60 | 1056 | 1018 | 20.6019 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.1 | 1.8 | 1.9 | -1 | 0 | 0.1 | 13 | 912 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195 | 59.9 | 0.0011 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 34.457 | 60 | 1056 | 1018 | 20.7575 | 29.7 | 1.8
2008-11-03 | 14:09 | 0 | 0 | 0 | 0.1 | 1.8 | 1.9 | -1 | 0 | 0.1 | 13 | 912 | 901 | 891 | 0 | 0 | 0 | 995 | -40 | 0 | 38 | 192 | 32 | 6 | 195 | 59.8 | 0.0011 | 0.0011 | 0.0002 | 0.0001 |  | n/a |  | 34 | 60 | 1056 | 1018 | 20.7434 | 29.7 | 1.8
Summary Information:
Post Processor DLL Version | 6.01
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-11-03
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 50
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 28261
Version | 6.907
CO Span(%) | 8.05
CO2 Span(%) | 11.9
C6H14 Span(ppm) | 2005
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 131
Version | 1.611
NO Span(ppm) | 3004
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 9642
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | Backhoe J Deere MN 700212 310G
License Plate
Engine Displacement | 4.5
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 13:48
Test End Time | 14:09
Test Duration (s) | 1264
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 995.144
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iAMBII_RPM
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
Faults:
0X0000 - 11/03/2008 13:47:36.498 - None Found
Warnings:
0X0000 - 11/03/2008 13:47:36.498 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1264
Brake Specific Dropout Limit Count | 1264
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-11-03 | 13:48 | [RELEASE_VER=2.018 BUILD=50 BDATE=07/18/2008 IP=10.10.1.55]
No Information Available

## NO2 Linearity
Test Date:
Test Time:
11/3/2008
13:47:37
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
217

## NDUV
Test Date:
Test Time:
11/3/2008
13:47:37
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
217

## O2 Linearity
Test Date:
Test Time:
11/3/2008
13:47:37
SEMTECH Serial Number:
G05-SDS02

## O2
Test Date:
Test Time:
11/3/2008
13:47:37
SEMTECH Serial Number:
G05-SDS02