## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 22
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 37249 |  |  | CO - AMBII Serial Number: |  |  | 37249 |  | HC - AMBII Serial Number: |  |  | 37249 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-10-20 |  | 11.99 | X | 0.1 | 3 |  |  | 1196 |  | 50 | 3 |  | 319 |  | 4 | 3 |  | 1504 |  | 15 | 3 |  | 957 | 1000 | 5 | 2
Test Time: | 16:18
Test Software: | 2.018 22
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0.01 | 0.01 |  |  | 0 | 7 | 7 |  |  | 0 | 1003.9 | 1003.9 |  |  | 0 | 1.2 | 1.2 |  |  | 0 | 0.4 | 0.4
90 | 105 | 1 | 1.199 | 1.183 | 1.161 | -0.022 | -1.86 |  | 119.6 | 130 | 10 | 8.36 |  | 31.9 | 1004 | 972.1 | 3047.34 | Fail | 150.4 | 151.8 | 1.4 | 0.93 |  | 95.7 | 97 | 1.3 | 1.36
150 | 165 | 2 | 2.398 | 2.37 | 2.37 | 0 | 0 |  | 239.2 | 250 | 11 | 4.6 |  | 63.8 | 1004.3 | 940.5 | 1474.14 | Fail | 300.8 | 301.5 | 0.7 | 0.23 |  | 191.4 | 193.4 | 2 | 1.04
210 | 225 | 3 | 3.597 | 3.56 | 3.602 | 0.042 | 1.18 |  | 358.8 | 360 | 1 | 0.28 |  | 95.7 | 1004.9 | 909.2 | 950.05 | Fail | 451.2 | 451.7 | 0.5 | 0.11 |  | 287.1 | 287.7 | 0.6 | 0.21
270 | 285 | 4 | 4.796 | 4.754 | 4.844 | 0.09 | 1.89 |  | 478.4 | 480 | 2 | 0.42 |  | 127.6 | 1004.3 | 876.7 | 687.07 | Fail | 601.6 | 600.3 | -1.3 | -0.22 |  | 382.8 | 382.3 | -0.5 | -0.13
330 | 345 | 5 | 5.995 | 5.951 | 6.069 | 0.118 | 1.98 |  | 598 | 590 | -8 | -1.34 |  | 159.5 | 1004.9 | 845.4 | 530.03 | Fail | 752 | 751.1 | -0.9 | -0.12 |  | 478.5 | 477.1 | -1.4 | -0.29
390 | 405 | 6 | 7.194 | 7.152 | 7.266 | 0.114 | 1.59 |  | 717.6 | 710 | -8 | -1.11 |  | 191.4 | 1005 | 813.6 | 425.08 | Fail | 902.4 | 902 | -0.4 | -0.04 |  | 574.2 | 573.2 | -1 | -0.17
450 | 465 | 7 | 8.393 | 8.356 | 8.466 | 0.11 | 1.32 |  | 837.2 | 830 | -7 | -0.84 |  | 223.3 | 1004.7 | 781.4 | 349.93 | Fail | 1052.8 | 1052.1 | -0.7 | -0.07 |  | 669.9 | 670 | 0.1 | 0.01
510 | 525 | 8 | 9.592 | 9.564 | 9.645 | 0.081 | 0.85 |  | 956.8 | 942 | -15 | -1.57 |  | 255.2 | 1004.9 | 749.7 | 293.77 | Fail | 1203.2 | 1201.6 | -1.6 | -0.13 |  | 765.6 | 766.6 | 1 | 0.13
570 | 585 | 9 | 10.791 | 10.775 | 10.77 | -0.005 | -0.05 |  | 1076.4 | 1070 | -6 | -0.56 |  | 287.1 | 1004.9 | 717.8 | 250.02 | Fail | 1353.6 | 1351.9 | -1.7 | -0.13 |  | 861.3 | 863.6 | 2.3 | 0.27
630 | 645 | 10 | 11.99 | 11.99 | 11.899 | -0.091 | -0.76 |  | 1196 | 1185 | -11 | -0.92 |  | 319 | 1004.4 | 685.4 | 214.86 | Fail | 1504 | 1507.1 | 3.1 | 0.21 |  | 957 | 960.7 | 3.7 | 0.39
690 | 705 | 9 | 10.791 | 10.775 | 10.787 | 0.012 | 0.11 |  | 1076.4 | 1060 | -16 | -1.49 |  | 287.1 | 1004.7 | 717.6 | 249.95 | Fail | 1353.6 | 1357 | 3.4 | 0.25 |  | 861.3 | 863.2 | 1.9 | 0.22
750 | 765 | 8 | 9.592 | 9.564 | 9.643 | 0.079 | 0.83 |  | 956.8 | 949 | -8 | -0.84 |  | 255.2 | 1004.7 | 749.5 | 293.69 | Fail | 1203.2 | 1205.6 | 2.4 | 0.2 |  | 765.6 | 766.5 | 0.9 | 0.12
810 | 825 | 7 | 8.393 | 8.356 | 8.47 | 0.114 | 1.36 |  | 837.2 | 830 | -7 | -0.84 |  | 223.3 | 1004.9 | 781.6 | 350.02 | Fail | 1052.8 | 1053.9 | 1.1 | 0.1 |  | 669.9 | 669.9 | 0 | 0
870 | 885 | 6 | 7.194 | 7.152 | 7.267 | 0.115 | 1.61 |  | 717.6 | 721 | 3 | 0.42 |  | 191.4 | 1004.7 | 813.3 | 424.92 | Fail | 902.4 | 904.4 | 2 | 0.22 |  | 574.2 | 573.1 | -1.1 | -0.19
930 | 945 | 5 | 5.995 | 5.951 | 6.072 | 0.121 | 2.03 |  | 598 | 593 | -5 | -0.84 |  | 159.5 | 1004.5 | 845 | 529.78 | Fail | 752 | 752.1 | 0.1 | 0.01 |  | 478.5 | 477 | -1.5 | -0.31
990 | 1005 | 4 | 4.796 | 4.754 | 4.84 | 0.086 | 1.81 |  | 478.4 | 470 | -8 | -1.67 |  | 127.6 | 1004.7 | 877.1 | 687.38 | Fail | 601.6 | 603 | 1.4 | 0.23 |  | 382.8 | 382 | -0.8 | -0.21
1050 | 1065 | 3 | 3.597 | 3.56 | 3.612 | 0.052 | 1.46 |  | 358.8 | 354 | -5 | -1.39 |  | 95.7 | 1004.6 | 908.9 | 949.74 | Fail | 451.2 | 454.1 | 2.9 | 0.64 |  | 287.1 | 288.1 | 1 | 0.35
1110 | 1125 | 2 | 2.398 | 2.37 | 2.369 | -0.001 | -0.04 |  | 239.2 | 239 | 0 | 0 |  | 63.8 | 1004.8 | 941 | 1474.92 | Fail | 300.8 | 303 | 2.2 | 0.73 |  | 191.4 | 192.7 | 1.3 | 0.68
1170 | 1185 | 1 | 1.199 | 1.183 | 1.179 | -0.004 | -0.34 |  | 119.6 | 120 | 0 | 0 |  | 31.9 | 1004.6 | 972.7 | 3049.22 | Fail | 150.4 | 154 | 3.6 | 2.39 |  | 95.7 | 96.8 | 1.1 | 1.15
1230 | 1245 | 0 | 0 | 0 | 0.003 | 0.003 |  |  | 0 | -7 | -7 |  |  | 0 | 1004 | 1004 |  |  | 0 | 2.4 | 2.4 |  |  | 0 | -0.2 | -0.2
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | 0.033 | 0.12 |  |  |  | Intercept | 4 | 12 |  |  | Intercept | 1004 | 3 | Fail |  | Intercept | 1.6 | 15 |  |  | Intercept | 0.1 | 9.6
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.99 | 0.99 - 1.01 |  |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0606 | 0.1199 |  |  |  | SEE | 5.4367 | 11.96 |  |  | SEE | 0.2727 | 3.19 |  |  | SEE | 1.6835 | 15.04 |  |  | SEE | 1.3391 | 9.607
|  |  | r2 | 0.9998 | 0.998 |  |  |  | r2 | 0.9998 | 0.998 |  |  | r2 | 0.339 | 0.998 | Fail |  | r2 | 1 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.9862 | 2.8 | 0.5 | 0.4901 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.001 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 |  |  | 1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.8 | 2.6 | 0.4 | 0.4 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.001 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.8 | 27.1 | 0.8
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.8 | 2.6 | 0.5 | 0.4865 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.001 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.743 | 27.1 | 0.8
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.7862 | 2.6 | 0.4 | 0.4175 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.001 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1003.862 | 10.7 | 27.1342 | 0.8
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.7 | 2.5 | 0.6 | 0.5632 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.000986 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7114 | 27.1886 | 0.7
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.7 | 2.5 | 0.4 | 0.4 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.286 | 0.0009 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7887 | 27.1113 | 0.7
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.7 | 2.5 | 0.4 | 0.4 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.602 | 27.1 | 0.7
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.8 | 0.6862 | 2.5 | 0.4 | 0.4 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.5865 | 27.1432 | 0.7
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.9 | 0.6 | 2.5 | 0.6 | 0.5577 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7 | 27.2 | 0.6
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.9 | 0.6 | 2.5 | 0.5 | 0.5 | 0.2 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7 | 27.2 | 0.6
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.9 | 0.6 | 2.5 | 0.6 | 0.55 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7 | 27.1252 | 0.6
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.9 | 0.6 | 2.5 | 0.5 | 0.5352 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7518 | 27.1 | 0.6
2008-10-20 | 16:18 | 0.01 | 0.0005 | 5.4762 | 1.9 | 0.6 | 2.5 | 0.5 | 0.5 | 0.2 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0008 | 0.000486 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.8291 | 27.1 | 0.6
2008-10-20 | 16:18 | 0.01 | 0.0002 | 2.3902 | 1.9 | 0.6 | 2.5 | 0.6 | 0.5938 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.8796 | 27.1 | 0.6
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.9 | 0.6 | 2.5 | 0.5 | 0.5 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.6478 | 27.1 | 0.6
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.9 | 0.58603 | 2.5 | 0.6 | 0.5639 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.6613 | 27.1613 | 0.6
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.9 | 0.5 | 2.4 | 0.5 | 0.5 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7 | 27.1614 | 0.5
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.5 | 2.3 | 0.5 | 0.5 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7317 | 27.1158 | 0.5
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.5 | 2.3 | 0.5 | 0.4533 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.8861 | 27.193 | 0.5
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.5 | 2.3 | 0.6 | 0.6413 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0008 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.1377 | 10.9 | 27.2 | 0.5
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.48617 | 2.3 | 0.5 | 0.4626 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.000786 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.8617 | 10.8049 | 27.1524 | 0.5
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.4 | 2.2 | 0.5 | 0.5058 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0007 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003 | 10.7248 | 27.1 | 0.4
2008-10-20 | 16:18 | 0.01 | -0.001 | -10 | 1.8 | 0.4 | 2.2 | 0.4 | 0.4311 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003 | 10.8021 | 27.1021 | 0.4
2008-10-20 | 16:18 | 0.01 | 0.0004 | 4.1878 | 1.8 | 0.4 | 2.2 | 0.4 | 0.4 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.138 | 10.8794 | 27.1794 | 0.4
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.8 | 0.4 | 2.2 | 0.5 | 0.4803 | 0.2 | 13.1 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.1434 | 0.4
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.8 | 0.38586 | 2.2 | 0.3 | 0.3 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.1339 | 0.4
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.8 | 0.3 | 2.1 | 0.4 | 0.3854 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.286 | 0.0007 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.8777 | 27.2 | 0.3
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.8 | 0.3 | 2.1 | 0.3 | 0.307 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7232 | 27.2 | 0.3
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.8 | 0.3 | 2.1 | 0.4 | 0.3623 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.000686 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039.6051 | 1004 | 10.8314 | 27.1343 | 0.3
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.4 | 0.2862 | 1.7 | 0.3 | 0.2558 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0006 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.1 | 0.3
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.4 | 0.2 | 1.6 | 0.4 | 0.4361 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0006 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.1202 | 0.2
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.3 | 0.1 | 1.4 | 0.2 | 0.2252 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0006 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.1974 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.2 | 0.1 | 1.3 | 0.4 | 0.3735 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.2 | 0.1 | 1.3 | 0.3 | 0.3347 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7965 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.1 | 0.1 | 1.2 | 0.3 | 0.2742 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7575 | 27.1713 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.1 | 0.1 | 1.2 | 0.4 | 0.3901 | 0.2 | 13.1 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.1058 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1 | 0.1 | 1.1 | 0.3 | 0.3 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.1837 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.1 | 0.1 | 1.2 | 0.4 | 0.3848 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0005 | 0.000286 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0.001 | 10 | 1.1 | 0.2 | 1.3 | 0.3 | 0.3179 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.2
2008-10-20 | 16:18 | 0.01 | 0.0009 | 9.375 | 1.1 | 0.2 | 1.3 | 0.4 | 0.3657 | 0.1 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.869 | 27.2 | 0.2
2008-10-20 | 16:18 | 0.01 | 0.0001 | 1.3622 | 1.1 | 0.18615 | 1.3 | 0.2 | 0.2182 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7145 | 27.2 | 0.2
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.1 | 0.1 | 1.2 | 0.5 | 0.5267 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.8617 | 10.8399 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.1 | 0.1 | 1.2 | 0.4 | 0.3623 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.138 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.4 | 0.4086 | 0.1 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.4 | 0.3522 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.862 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.5 | 0.4506 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.862 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.5 | 0.4648 | 0.1 | 13.1 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1003.1414 | 10.7875 | 27.1437 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.4 | 0.4 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.7671 | 27.1335 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.4 | 0.4 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.3 | 0.3187 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004.862 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:18 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.5 | 0.4646 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.3 | 0.3148 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9425 | 27.2 | 0.1
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.2 | 0.1 | 1.3 | 0.6 | 0.5511 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9802 | 27.2 | 0.1
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0.1 | 1.4 | 0.4 | 0.3623 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9029 | 27.2 | 0.1
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0.1 | 1.4 | 0.3 | 0.3105 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0 | 1.3 | 0.3 | 0.3274 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9515 | 27.2 | 0
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0 | 1.3 | 0.3 | 0.3 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9713 | 27.2 | 0
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0 | 1.3 | 0.3 | 0.3441 | 0.1 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.000186 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.2 | 0
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0 | 1.3 | 0.2 | 0.2 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1039 | 1004 | 10.9 | 27.2 | 0
2008-10-20 | 16:19 | 0.01 | 0 | 0 | 1.3 | 0 | 1.3 | 0.4 | 0.3937 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0
2008-10-20 | 16:19 | 0.174 | 0.0016 | 15.8101 | 1.3 | 0 | 1.3 | 0.2 | 0.2376 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0
2008-10-20 | 16:19 | 0.879 | 0.0065 | 64.6754 | 5.7 | 0 | 5.7 | 0.9 | 0.9061 | 0.1 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.001287 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0
2008-10-20 | 16:19 | 1.123 | 0.0128 | 128.2917 | 65.7 | 0.1 | 65.8 | 41.4 | 41.3605 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.01303 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.17 | 0.012 | 120 | 98.6 | 0.1 | 98.7 | 83.4 | 83.4145 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.019235 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.17 | 0.012 | 120 | 143.8 | 0.1 | 143.9 | 93 | 93.0179 | 0.1 | 13 | 959 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.027466 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.159 | 0.012 | 120 | 149.7 | 0.1 | 149.8 | 95.4 | 95.4254 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.028514 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.14 | 0.012 | 120 | 150.2 | 0.1 | 150.3 | 95.8 | 95.8 | 0.1 | 13 | 959 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2857 | 0.0286 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.14 | 0.0123 | 123.3694 | 150.8 | 0.11411 | 151 | 96 | 96.0253 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.028714 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.1
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.1 | 0.2 | 151.3 | 96.1 | 96.1228 | 0.1 | 13 | 958 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.2
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.1 | 0.2138 | 151.3 | 96.4 | 96.3868 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.138 | 10.9 | 27.2 | 0.2
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.1 | 0.3 | 151.4 | 96.6 | 96.6264 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 10.9 | 27.2 | 0.3
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.2 | 0.3 | 151.5 | 96.6 | 96.6 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.3
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.2 | 0.4 | 151.6 | 96.8 | 96.7639 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9649 | 27.2 | 0.4
2008-10-20 | 16:19 | 1.14 | 0.013 | 130 | 151.3 | 0.4 | 151.7 | 96.8 | 96.7756 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9579 | 27.2 | 0.4
2008-10-20 | 16:19 | 1.148 | 0.013 | 130 | 151.4 | 0.4138 | 151.8 | 97.2 | 97.1507 | 0.1 | 13 | 958 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.4
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.4 | 0.5138 | 151.9 | 96.9 | 96.9371 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 0.5
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 1.6138 | 152.9 | 97.1 | 97.1078 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 1.6
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 1.7141 | 153 | 97 | 97.0257 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.1414 | 10.9 | 27.2 | 1.7
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 1.814 | 153.1 | 97.1 | 97.0854 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.8603 | 10.9283 | 27.2 | 1.8
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 1.9138 | 153.2 | 97 | 97.0232 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 1.9
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 2.0138 | 153.3 | 97 | 97.013 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 2
2008-10-20 | 16:19 | 1.15 | 0.013 | 130 | 151.3 | 2.1138 | 153.4 | 97 | 97.0398 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0288 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 2.1
2008-10-20 | 16:19 | 1.155 | 0.013 | 130 | 151.4 | 2.2 | 153.6 | 96.9 | 96.9375 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.028814 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9633 | 27.2 | 2.2
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.4 | 2.215 | 153.6 | 96.9 | 96.9277 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.000115 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 2.2
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.4 | 2.3 | 153.7 | 96.8 | 96.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 2.3
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.4 | 2.3 | 153.7 | 97 | 96.975 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.8309 | 27.2 | 2.3
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.4 | 2.4 | 153.8 | 96.8 | 96.8267 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.8464 | 27.2 | 2.4
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.5 | 2.4 | 153.9 | 97 | 96.9554 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 2.4
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.6 | 2.4138 | 154 | 96.8 | 96.8 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 2.4
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 2.7257 | 154.4 | 96.9 | 96.8873 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003.8603 | 10.9 | 27.2 | 2.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 5.5 | 157.3 | 96.8 | 96.8401 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 5.5
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.7 | 159.5 | 96.8 | 96.8 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.7 | 159.5 | 96.9 | 96.9475 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.7 | 159.5 | 96.9 | 96.91 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.7 | 159.5 | 97 | 96.9887 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9644 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.7 | 159.5 | 96.8 | 96.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9584 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.6862 | 159.4 | 97 | 97 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.7
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.5862 | 159.3 | 96.8 | 96.8255 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.6
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.5 | 159.2 | 97 | 97.013 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2859 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9733 | 27.2 | 7.5
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.4861 | 159.2 | 97 | 97.0251 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9494 | 27.2 | 7.5
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.3862 | 159.2 | 97.3 | 97.2553 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.4
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.2864 | 159 | 97.1 | 97.1456 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2864 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.1363 | 10.9 | 27.2 | 7.3
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.7 | 7.1862 | 158.9 | 97.2 | 97.1742 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.2
2008-10-20 | 16:19 | 1.16 | 0.013 | 130 | 151.8 | 7.0861 | 158.9 | 97.3 | 97.2643 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7.1
2008-10-20 | 16:19 | 1.168 | 0.013 | 130 | 151.8 | 6.9862 | 158.8 | 97.1 | 97.1 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 7
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.9 | 158.7 | 97.2 | 97.2279 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 6.9
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.8862 | 158.7 | 97.1 | 97.1083 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.028914 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 10.9 | 27.2 | 6.9
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.8 | 158.6 | 97.2 | 97.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9684 | 27.2 | 6.8
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.8 | 158.6 | 97 | 97.0272 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9542 | 27.2 | 6.8
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.8 | 158.6 | 97.1 | 97.1279 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.8
2008-10-20 | 16:19 | 1.17 | 0.013 | 130 | 151.8 | 6.786 | 158.6 | 96.9 | 96.9474 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.000386 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.8
2008-10-20 | 16:20 | 1.17 | 0.013 | 130 | 151.8 | 6.7 | 158.5 | 97 | 96.9809 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003.862 | 10.9 | 27.2 | 6.7
2008-10-20 | 16:20 | 1.165 | 0.013 | 130 | 151.8 | 6.6 | 158.4 | 97 | 96.9588 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.029 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.6
2008-10-20 | 16:20 | 1.16 | 0.013 | 130 | 151.8 | 6.5861 | 158.4 | 97 | 96.9751 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.6
2008-10-20 | 16:20 | 1.16 | 0.013 | 130 | 151.8 | 6.4 | 158.2 | 96.9 | 96.9252 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003.8614 | 10.9093 | 27.2 | 6.4
2008-10-20 | 16:20 | 1.16 | 0.013 | 130 | 151.7 | 6.3862 | 158.1 | 96.9 | 96.875 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9865 | 27.2 | 6.4
2008-10-20 | 16:20 | 1.16 | 0.013 | 130 | 151.4 | 6.2862 | 157.7 | 97 | 97.0252 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9362 | 27.2 | 6.3
2008-10-20 | 16:20 | 1.167 | 0.013 | 130 | 151.3 | 6.2 | 157.5 | 96.8 | 96.8125 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.2
2008-10-20 | 16:20 | 1.17 | 0.013 | 130 | 151.3 | 6.2 | 157.5 | 97.1 | 97.0731 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.2
2008-10-20 | 16:20 | 1.17 | 0.013 | 130 | 151.3 | 6 | 157.3 | 96.9 | 96.9311 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2859 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6
2008-10-20 | 16:20 | 1.17 | 0.013 | 130 | 151.2 | 5.9 | 157.1 | 97 | 97.0321 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9727 | 27.2 | 5.9
2008-10-20 | 16:20 | 1.17 | 0.014 | 140.3641 | 151.2 | 5.9 | 157.1 | 96.9 | 96.9261 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0289 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 11 | 27.2 | 5.9
2008-10-20 | 16:20 | 1.972 | 0.0174 | 174.3131 | 154.6 | 5.9 | 160.5 | 98 | 97.9809 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.029493 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 10.9728 | 27.2 | 5.9
2008-10-20 | 16:20 | 2.272 | 0.0249 | 249.2506 | 179.7 | 5.9 | 185.6 | 133.9 | 133.9396 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.033821 | 0.0002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 5.9
2008-10-20 | 16:20 | 2.345 | 0.0242 | 241.7569 | 267.7 | 6.2187 | 273.9 | 177.6 | 177.6433 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.044367 | 0.000214 | 0.000314 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 6.2
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 294.3 | 8.3 | 302.6 | 190.2 | 190.2268 | 0.1 | 13 | 959 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0526 | 0.0003 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 8.3
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 294.8 | 8.4138 | 303.3 | 192.3 | 192.2768 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.052714 | 0.000314 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003.862 | 10.9 | 27.2 | 8.4
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 295.7 | 8.6 | 304.3 | 192.9 | 192.9263 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0529 | 0.0004 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003 | 10.8738 | 27.2 | 8.6
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 295.8 | 9.5138 | 305.3 | 193 | 193 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0529 | 0.0004 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7188 | 27.2 | 9.5
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 295.8 | 9.738 | 305.5 | 193.4 | 193.4471 | 0.1 | 13 | 958 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0529 | 0.0004 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9032 | 27.2 | 9.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 295.9 | 10.6 | 306.5 | 193.3 | 193.2751 | 0.1 | 13 | 958 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.052914 | 0.0004 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.2 | 10.6 | 306.8 | 193.7 | 193.6588 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.053 | 0.000414 | 0.000514 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9776 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.2 | 10.6862 | 306.9 | 193.4 | 193.4391 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0531 | 0.0005 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.138 | 10.9004 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.3 | 10.7 | 307 | 193.5 | 193.5331 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2864 | 0.0531 | 0.0005 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.3 | 10.7 | 307 | 193.5 | 193.539 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.0532 | 0.0006 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.3 | 10.7 | 307 | 193.5 | 193.5435 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0532 | 0.0006 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 296.3 | 10.7 | 307 | 193.6 | 193.5881 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.053214 | 0.0007 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.35 | 0.024 | 240 | 297 | 10.7 | 307.7 | 193.5 | 193.5 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0534 | 0.0007 | 0.0006 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.352 | 0.024 | 240 | 298.3 | 10.6 | 308.9 | 193.5 | 193.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0536 | 0.0007 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.774 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.0243 | 243.1693 | 298.7 | 10.6 | 309.3 | 193.4 | 193.4145 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0537 | 0.0007 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.7806 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.0248 | 248.4904 | 299 | 10.6141 | 309.6 | 193.6 | 193.579 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0537 | 0.0007 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.8586 | 10.9176 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.024 | 240.1501 | 299.2 | 10.7 | 309.9 | 193.4 | 193.4279 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.214 | 0.053714 | 0.0007 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9948 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.36 | 0.024 | 240 | 299.5 | 10.6862 | 310.2 | 193.7 | 193.6548 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.053814 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.138 | 11 | 27.2 | 10.7
2008-10-20 | 16:20 | 2.36 | 0.0247 | 246.5304 | 299.5 | 10.6 | 310.1 | 193.4 | 193.4352 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9507 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.025 | 250 | 299.6 | 10.6 | 310.2 | 193.7 | 193.6519 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.025 | 250 | 299.6 | 10.6 | 310.2 | 193.5 | 193.5477 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.025 | 250 | 299.6 | 10.6 | 310.2 | 193.5 | 193.5435 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.36 | 0.025 | 250 | 299.6 | 10.6 | 310.2 | 193.4 | 193.4294 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.365 | 0.025 | 250 | 299.8 | 10.5862 | 310.4 | 193.4 | 193.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0539 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9354 | 27.2 | 10.6
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.2 | 10.2637 | 310.5 | 193.4 | 193.4359 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2864 | 0.054 | 0.0008 | 0.0005 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9873 | 27.2 | 10.3
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.5 | 8.1588 | 308.7 | 193.4 | 193.4 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.054 | 0.0008 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9101 | 27.2 | 8.2
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.7 | 7.186 | 307.9 | 193.6 | 193.5706 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0541 | 0.0008 | 0.0004 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9668 | 27.2 | 7.2
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.7 | 6.862 | 307.6 | 193.5 | 193.5 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0541 | 0.0008 | 0.000386 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 6.9
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.7 | 6 | 306.7 | 193.5 | 193.4627 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0541 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 6
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.8 | 5.8861 | 306.7 | 193.4 | 193.4 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2861 | 0.0541 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2 | 5.9
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.8 | 5.7 | 306.5 | 193.6 | 193.5637 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0541 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9239 | 27.2 | 5.7
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 300.9 | 5.6 | 306.5 | 193.4 | 193.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0541 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 5.6
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 301.7 | 5.5 | 307.2 | 193.5 | 193.4848 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.054214 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 5.5
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.1 | 5.3862 | 307.5 | 193.4 | 193.4243 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0543 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 10.9 | 27.2 | 5.4
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.1 | 5.3 | 307.4 | 193.3 | 193.3124 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.967 | 190 | 27 | 5 | 195 | 59.2862 | 0.0543 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.138 | 10.9 | 27.2 | 5.3
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.2 | 5.3 | 307.5 | 193.4 | 193.3883 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.033 | 190 | 27 | 5 | 195 | 59.2 | 0.0543 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 5.3
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.7 | 5.2 | 307.9 | 193.4 | 193.3729 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9395 | 27.2 | 5.2
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.7 | 5.186 | 307.9 | 193.4 | 193.4279 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.214 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.1397 | 10.9499 | 27.2 | 5.2
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.8 | 5.1 | 307.9 | 193.3 | 193.3201 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.8623 | 10.7181 | 27.2 | 5.1
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.9 | 5 | 307.9 | 193.5 | 193.5 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9848 | 27.2 | 5
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.9 | 5 | 307.9 | 193.4 | 193.4476 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.0032 | 27.2 | 5
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 303 | 5 | 308 | 193.5 | 193.5471 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9769 | 27.2 | 5
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 303 | 4.9862 | 308 | 193.4 | 193.4301 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.1914 | 27.2 | 5
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 303 | 4.7864 | 307.8 | 193.5 | 193.5104 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2136 | 0.0544 | 0.0008 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9598 | 27.2 | 4.8
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 303 | 4.7 | 307.7 | 193.6 | 193.6 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.054414 | 0.000814 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.0147 | 27.2 | 4.7
2008-10-20 | 16:20 | 2.37 | 0.025 | 250 | 302.9 | 4.7 | 307.6 | 193.6 | 193.6 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0544 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.0308 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.37 | 0.025 | 250 | 302.9 | 4.7 | 307.6 | 193.5 | 193.538 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0544 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9118 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.37 | 0.025 | 250 | 302.9 | 4.7 | 307.6 | 193.4 | 193.4255 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2141 | 0.0544 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.989 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.37 | 0.025 | 250 | 302.8 | 4.7 | 307.5 | 193.6 | 193.5746 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2139 | 0.054414 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9337 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.37 | 0.025 | 250 | 302.9 | 4.7 | 307.6 | 193.3 | 193.3341 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.967 | 191 | 27 | 5 | 195 | 59.2862 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.37 | 0.025 | 250 | 302.8 | 4.7 | 307.5 | 193.5 | 193.5247 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2137 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9207 | 27.2 | 4.7
2008-10-20 | 16:21 | 2.378 | 0.025 | 250 | 302.8 | 4.6 | 307.4 | 193.4 | 193.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.033 | 191 | 27 | 5 | 195 | 59.2861 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9978 | 27.2 | 4.6
2008-10-20 | 16:21 | 2.38 | 0.025 | 250 | 302.8 | 4.6 | 307.4 | 193.6 | 193.6481 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9251 | 27.2 | 4.6
2008-10-20 | 16:21 | 2.38 | 0.025 | 250 | 302.8 | 4.5862 | 307.4 | 193.3 | 193.3294 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9519 | 27.2 | 4.6
2008-10-20 | 16:21 | 2.38 | 0.025 | 250 | 302.9 | 4.5 | 307.4 | 193.6 | 193.5589 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.029 | 27.2 | 4.5
2008-10-20 | 16:21 | 2.4 | 0.025 | 250 | 302.9 | 4.5 | 307.4 | 193.5 | 193.5201 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11.0878 | 27.2061 | 4.5
2008-10-20 | 16:21 | 2.517 | 0.0284 | 283.8532 | 302.9 | 4.4 | 307.3 | 193.6 | 193.6494 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0545 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9321 | 27.284 | 4.4
2008-10-20 | 16:21 | 3.325 | 0.0357 | 357.1643 | 334.7 | 4.3859 | 339 | 202.3 | 202.3462 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.056989 | 0.0009 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9613 | 27.2387 | 4.4
2008-10-20 | 16:21 | 3.523 | 0.034 | 340.4837 | 359.8 | 4.3 | 364.1 | 255.9 | 255.9433 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.063001 | 0.0009 | 0.0002 | 0 |  | n/a |  | 32.4555 | 55 | 1040 | 1004 | 11 | 27.2 | 4.3
2008-10-20 | 16:21 | 3.545 | 0.034 | 340 | 416 | 4.3 | 420.3 | 279.4 | 279.3605 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.071066 | 0.0009 | 0.0002 | 0 |  | n/a |  | 32.5445 | 55 | 1040 | 1004.138 | 10.9842 | 27.2 | 4.3
2008-10-20 | 16:21 | 3.56 | 0.0353 | 353.0115 | 445.5 | 4.4 | 449.9 | 285.2 | 285.207 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.075141 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.1377 | 10.907 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.58 | 0.036 | 360 | 448.3 | 4.4 | 452.7 | 286.8 | 286.7765 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.075514 | 0.0009 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.58 | 0.036 | 360 | 449.6 | 4.4 | 454 | 286.9 | 286.9 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2861 | 0.075728 | 0.000914 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9475 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.58 | 0.036 | 360 | 451 | 4.4 | 455.4 | 287.3 | 287.2595 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.076 | 0.001 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9753 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.593 | 0.036 | 360 | 451 | 4.4 | 455.4 | 287.2 | 287.1628 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.076 | 0.001 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.2019 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.1 | 4.4 | 455.5 | 287.4 | 287.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.076 | 0.001 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2792 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451 | 4.4 | 455.4 | 287.3 | 287.3463 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.076 | 0.001 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040.0657 | 1005 | 11.0691 | 27.2436 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451 | 4.4 | 455.4 | 287.6 | 287.5588 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.076 | 0.001 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11.0992 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.1 | 4.4 | 455.5 | 287.6 | 287.6 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.076 | 0.001014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9108 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.1 | 4.4 | 455.5 | 287.5 | 287.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.076014 | 0.0011 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004.862 | 10.988 | 27.2 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.1 | 4.4 | 455.5 | 287.8 | 287.8227 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0761 | 0.0011 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004.138 | 10.9347 | 27.2653 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451 | 4.4 | 455.4 | 287.4 | 287.4285 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0761 | 0.0011 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9425 | 27.3 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.1 | 4.4139 | 455.5 | 287.8 | 287.7588 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.0761 | 0.0011 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 11 | 27.2803 | 4.4
2008-10-20 | 16:21 | 3.602 | 0.036 | 360 | 451 | 4.5 | 455.5 | 287.6 | 287.6 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.285 | 0.0761 | 0.0011 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040.4688 | 1004.8498 | 11 | 27.203 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.036 | 360 | 451 | 4.5 | 455.5 | 287.7 | 287.7 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0761 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1004 | 11 | 27.2 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.0361 | 361.4083 | 450.9 | 4.5 | 455.4 | 287.7 | 287.7 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2136 | 0.0761 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1004 | 10.9487 | 27.2513 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.037 | 369.7417 | 450.9 | 4.5 | 455.4 | 287.8 | 287.7867 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0761 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.037 | 370 | 450.9 | 4.5 | 455.4 | 287.8 | 287.7736 | 0.1 | 13 | 957 | 952 | 941 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0761 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2941 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.037 | 370 | 450.9 | 4.5 | 455.4 | 287.6 | 287.6479 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.0762 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2171 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.037 | 370 | 450.8 | 4.5 | 455.3 | 287.6 | 287.6127 | 0.1 | 13.1 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0762 | 0.0012 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9599 | 27.2599 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.037 | 370 | 450.8 | 4.5 | 455.3 | 287.5 | 287.5434 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0762 | 0.001214 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9628 | 27.2628 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.0369 | 368.5917 | 450.7 | 4.5 | 455.2 | 287.7 | 287.663 | 0.1 | 13.1 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0762 | 0.0013 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.9 | 27.2147 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.036 | 360.2583 | 450.7 | 4.5 | 455.2 | 287.5 | 287.5469 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0762 | 0.0013 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.292 | 4.5
2008-10-20 | 16:21 | 3.605 | 0.036 | 360 | 450.7 | 4.5 | 455.2 | 287.6 | 287.6 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0762 | 0.0013 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2308 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 450.8 | 4.5 | 455.3 | 287.6 | 287.6 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.076214 | 0.0013 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.9 | 27.2464 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451 | 4.5 | 455.5 | 287.8 | 287.7666 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0763 | 0.0013 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.7 | 4.5 | 456.2 | 287.4 | 287.4 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0764 | 0.001314 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.9 | 27.2992 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.5 | 456.3 | 287.7 | 287.7425 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.9 | 27.2219 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.5 | 456.3 | 287.6 | 287.5915 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.2553 | 4.5
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.5 | 456.3 | 287.6 | 287.6436 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2675 | 4.5
2008-10-20 | 16:21 | 3.603 | 0.036 | 360 | 451.8 | 4.5 | 456.3 | 287.7 | 287.7 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.2097 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.036 | 360 | 451.8 | 4.4864 | 456.3 | 287.7 | 287.6776 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2136 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.9 | 27.287 | 4.5
2008-10-20 | 16:21 | 3.61 | 0.036 | 360 | 451.8 | 4.4 | 456.2 | 287.8 | 287.8257 | 0.1 | 13.1 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0764 | 0.0014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040.0697 | 1005 | 10.9642 | 27.2358 | 4.4
2008-10-20 | 16:21 | 3.602 | 0.036 | 360 | 451.8 | 4.3861 | 456.2 | 287.6 | 287.6265 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.076414 | 0.0014 | 0.000286 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9586 | 27.2414 | 4.4
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.3 | 456.1 | 287.8 | 287.8 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0765 | 0.0014 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.2862 | 456.1 | 287.8 | 287.8372 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0765 | 0.0014 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.2 | 456 | 288 | 287.9889 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0765 | 0.001414 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9731 | 27.3 | 4.2
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.7 | 4.2 | 455.9 | 287.8 | 287.8245 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.862 | 11 | 27.3 | 4.2
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.8 | 4.1 | 455.9 | 288 | 287.9731 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.1377 | 10.9725 | 27.3 | 4.1
2008-10-20 | 16:21 | 3.608 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.8 | 287.7941 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9047 | 27.3 | 4.1
2008-10-20 | 16:21 | 3.61 | 0.036 | 360 | 451.8 | 4.1 | 455.9 | 287.8 | 287.7741 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9819 | 27.3 | 4.1
2008-10-20 | 16:21 | 3.6 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.8 | 287.8147 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11 | 27.2409 | 4.1
2008-10-20 | 16:21 | 3.59 | 0.036 | 360 | 451.6 | 4.1 | 455.7 | 287.6 | 287.6215 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11 | 27.2361 | 4.1
2008-10-20 | 16:21 | 3.59 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.8 | 287.7638 | 0.1 | 13.1 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11 | 27.3 | 4.1
2008-10-20 | 16:21 | 3.59 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.6 | 287.6467 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11 | 27.3 | 4.1
2008-10-20 | 16:21 | 3.59 | 0.0361 | 361.395 | 451.8 | 4.1 | 455.9 | 287.5 | 287.5489 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 194.8 | 59.2138 | 0.0765 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1041 | 1004 | 10.9322 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.037 | 369.7983 | 451.8 | 4.1 | 455.9 | 287.6 | 287.5706 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 194.4 | 59.2 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1041 | 1004 | 10.9452 | 27.2548 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.037 | 370 | 451.8 | 4.1 | 455.9 | 287.8 | 287.7864 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 194 | 59.3 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9776 | 27.2224 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.0364 | 363.603 | 451.8 | 4.1 | 455.9 | 287.6 | 287.6 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 194.3 | 59.2862 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9003 | 27.2997 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.036 | 360 | 451.8 | 4.1 | 455.9 | 287.7 | 287.6864 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 194.7 | 59.2 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32.4555 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.6 | 287.5884 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.59 | 0.0361 | 361.3712 | 451.7 | 4.1 | 455.8 | 287.7 | 287.6746 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.286 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.596 | 0.037 | 369.7324 | 451.7 | 4.1 | 455.8 | 287.4 | 287.3522 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.6 | 0.0362 | 361.9601 | 451.7 | 4.1 | 455.8 | 287.5 | 287.4753 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.967 | 192 | 27 | 5 | 195 | 59.2862 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.6 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.7 | 287.6629 | 0.1 | 13.1 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.033 | 191 | 27 | 5 | 195 | 59.2865 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 3.714 | 0.036 | 360 | 451.7 | 4.1 | 455.8 | 287.7 | 287.6522 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0766 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:22 | 4.417 | 0.0404 | 403.6223 | 453.4 | 4.1 | 457.5 | 287.8 | 287.7715 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.076821 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9173 | 27.3 | 4.1
2008-10-20 | 16:22 | 4.721 | 0.0498 | 498.3902 | 490.5 | 4.1 | 494.6 | 324.7 | 324.6791 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.081718 | 0.0015 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.9945 | 27.3 | 4.1
2008-10-20 | 16:22 | 4.803 | 0.049 | 490.2602 | 547.6 | 4.1138 | 551.7 | 367.5 | 367.5443 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.089142 | 0.001514 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004.862 | 11 | 27.3 | 4.1
2008-10-20 | 16:22 | 4.81 | 0.049 | 490 | 590.2 | 4.3 | 594.5 | 378.4 | 378.4072 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.094452 | 0.0016 | 0.000214 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.138 | 10.951 | 27.3 | 4.3
2008-10-20 | 16:22 | 4.81 | 0.049 | 490 | 598.8 | 4.3 | 603.1 | 381.1 | 381.1231 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.967 | 191 | 27 | 5 | 195 | 59.3 | 0.0955 | 0.0016 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:22 | 4.81 | 0.049 | 490 | 598.9 | 4.3 | 603.2 | 381.6 | 381.5584 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0956 | 0.0016 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.8932 | 27.3 | 4.3
2008-10-20 | 16:22 | 4.81 | 0.049 | 490 | 599.1 | 4.3138 | 603.4 | 382 | 382.0346 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0956 | 0.0017 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.7388 | 27.3 | 4.3
2008-10-20 | 16:22 | 4.834 | 0.0489 | 488.6656 | 599.2 | 4.4 | 603.6 | 381.8 | 381.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2138 | 0.095614 | 0.0017 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.8157 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.834 | 0.048 | 480.3253 | 599.1 | 4.4 | 603.5 | 381.9 | 381.8823 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2862 | 0.0957 | 0.0017 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.935 | 27.265 | 4.4
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.2 | 4.4 | 603.6 | 382.2 | 382.193 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2139 | 0.0957 | 0.0017 | 0.0003 | 0 |  | n/a |  | 32.4555 | 55 | 1040 | 1004 | 11 | 27.2122 | 4.4
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.3 | 4.4 | 603.7 | 382.1 | 382.0531 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0957 | 0.0018 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 11 | 27.2894 | 4.4
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.4 | 4.4 | 603.8 | 382.2 | 382.1789 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2138 | 0.0957 | 0.0018 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.9334 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.3 | 4.4145 | 603.7 | 382.1 | 382.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2145 | 0.0957 | 0.0018 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.4 | 4.5 | 603.9 | 382.1 | 382.0821 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0958 | 0.0018 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.4 | 4.5 | 603.9 | 381.7 | 381.7271 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0958 | 0.0018 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.5 | 4.5 | 604 | 381.9 | 381.9164 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.214 | 0.0958 | 0.001814 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1003 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.5 | 4.6 | 604.1 | 382 | 381.9552 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0958 | 0.0019 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1003 | 10.9 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.6 | 4.6 | 604.2 | 381.9 | 381.913 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2862 | 0.0959 | 0.0019 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1003.138 | 10.9295 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.83 | 0.048 | 480 | 599.6 | 4.6 | 604.2 | 382.1 | 382.0534 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2862 | 0.0959 | 0.0019 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 11 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.838 | 0.048 | 480 | 599.5 | 4.6 | 604.1 | 382.2 | 382.1626 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0959 | 0.001914 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 11 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.5 | 4.6 | 604.1 | 382.2 | 382.1879 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0959 | 0.002 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.8769 | 27.2385 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.4 | 4.6 | 604 | 382.2 | 382.1563 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.096 | 0.002 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.8387 | 27.2387 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.5 | 4.6 | 604.1 | 382.1 | 382.1315 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.096 | 0.002 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.9159 | 27.2841 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.5 | 4.6 | 604.1 | 381.9 | 381.9248 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.096 | 0.002 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.9931 | 27.2069 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.6 | 4.6 | 604.2 | 382.3 | 382.2885 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.096 | 0.002 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9297 | 27.2703 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 599.8 | 4.6 | 604.4 | 382.1 | 382.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.096014 | 0.002014 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600 | 4.6 | 604.6 | 382.4 | 382.4417 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.286 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9247 | 27.2753 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.1 | 4.6 | 604.7 | 382.3 | 382.279 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 11 | 27.2019 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.1 | 4.6 | 604.7 | 382.2 | 382.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 11 | 27.2792 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.2 | 4.6 | 604.8 | 382.3 | 382.2707 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.9436 | 27.3 | 4.6
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.5 | 382.488 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.4 | 4.5 | 604.9 | 382.4 | 382.3623 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.9108 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.5 | 4.5 | 605 | 382.1 | 382.1486 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0961 | 0.0021 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1003 | 10.988 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.5 | 4.5 | 605 | 382.2 | 382.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0962 | 0.0021 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1003.138 | 10.9348 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.1 | 382.1258 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0962 | 0.0021 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.5 | 382.4752 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2141 | 0.0962 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.8609 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.84 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.2 | 382.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.286 | 0.0962 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.7065 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.848 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.4 | 382.3734 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0962 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.8478 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.3 | 382.2888 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0962 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.3 | 382.3472 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.0962 | 0.0022 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.8717 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.4861 | 604.8 | 382.1 | 382.0744 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.8 | 27.3 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.4 | 604.7 | 382.4 | 382.4104 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2862 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.8 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.4 | 604.7 | 382.6 | 382.5633 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2136 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.5445 | 55 | 1040 | 1005 | 10.8595 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.2 | 4.4 | 604.6 | 382.4 | 382.3545 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.2 | 4.4 | 604.6 | 382.5 | 382.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.3 | 4.4 | 604.7 | 382.3 | 382.3285 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.0963 | 0.0022 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.2 | 4.4 | 604.6 | 382.5 | 382.4898 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2141 | 0.0963 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.9688 | 27.2312 | 4.4
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.1 | 4.5 | 604.6 | 382.3 | 382.3257 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.214 | 0.0963 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 11 | 27.2 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.1 | 4.5 | 604.6 | 382.5 | 382.4739 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2138 | 0.0963 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.9306 | 27.2231 | 4.5
2008-10-20 | 16:22 | 4.85 | 0.048 | 480 | 600.1 | 4.5 | 604.6 | 382.3 | 382.3 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2138 | 0.0963 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.7006 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.85 | 0.048 | 480 | 600.1 | 4.5 | 604.6 | 382.5 | 382.4508 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0963 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.8551 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.845 | 0.048 | 480 | 600.1 | 4.5 | 604.6 | 382.3 | 382.324 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0964 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.7905 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.3 | 4.5 | 604.8 | 382.5 | 382.4574 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2862 | 0.0964 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.7639 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.5 | 4.5 | 605 | 382.3 | 382.2537 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0964 | 0.0023 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.5 | 4.4862 | 605 | 382.5 | 382.4598 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2138 | 0.0964 | 0.0024 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.5 | 4.4 | 604.9 | 382.5 | 382.5192 | 0.1 | 13 | 957 | 952 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0964 | 0.0024 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7729 | 27.3 | 4.4
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.5 | 4.4 | 604.9 | 382.3 | 382.3273 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0964 | 0.0024 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.863 | 27.3 | 4.4
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.5 | 4.4 | 604.9 | 382.5 | 382.4789 | 0.1 | 13 | 957 | 952 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.214 | 0.0964 | 0.0024 | 0.0003 | 0 |  | n/a |  | 32.4555 | 55 | 1040 | 1004 | 11.0463 | 27.3 | 4.4
2008-10-20 | 16:23 | 4.84 | 0.048 | 480 | 600.4 | 4.4 | 604.8 | 382.4 | 382.351 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2138 | 0.0964 | 0.0024 | 0.000286 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.8148 | 27.3 | 4.4
2008-10-20 | 16:23 | 4.883 | 0.048 | 480 | 600.3 | 4.3 | 604.6 | 382.6 | 382.6373 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0964 | 0.0024 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.8723 | 27.3 | 4.3
2008-10-20 | 16:23 | 5.127 | 0.0516 | 516.1469 | 600.4 | 4.3 | 604.7 | 382.7 | 382.6758 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.096414 | 0.0024 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:23 | 5.82 | 0.0598 | 597.551 | 635.1 | 4.3 | 639.4 | 400.5 | 400.5091 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.10062 | 0.0024 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.022 | 0.0581 | 580.5442 | 687.9 | 4.3 | 692.2 | 449.6 | 449.625 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.1064 | 0.0024 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8923 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.04 | 0.058 | 580 | 728.1 | 4.3138 | 732.4 | 470.8 | 470.7869 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.11132 | 0.002414 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.7378 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.04 | 0.058 | 580 | 741.7 | 4.4 | 746.1 | 475.2 | 475.2105 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2862 | 0.11285 | 0.0025 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.8166 | 27.2417 | 4.4
2008-10-20 | 16:23 | 6.04 | 0.0589 | 589.281 | 746.2 | 4.4 | 750.6 | 476.3 | 476.3009 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2138 | 0.11331 | 0.0025 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.2355 | 4.4
2008-10-20 | 16:23 | 6.04 | 0.06 | 600 | 747.2 | 4.4 | 751.6 | 476.7 | 476.7 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.1135 | 0.0025 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.04 | 0.06 | 600 | 747.7 | 4.4 | 752.1 | 476.6 | 476.6135 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.11351 | 0.002514 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.04 | 0.06 | 600 | 747.8 | 4.4 | 752.2 | 477 | 476.9705 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.1136 | 0.0026 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.7663 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.04 | 0.06 | 600 | 747.9 | 4.4 | 752.3 | 476.8 | 476.81 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.1136 | 0.0026 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8314 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.046 | 0.06 | 600 | 748.5 | 4.5 | 753 | 477 | 476.9634 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.1137 | 0.0026 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9792 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.6 | 4.5 | 753.1 | 476.8 | 476.8253 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2861 | 0.1137 | 0.0026 | 0.000214 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004.8614 | 10.9022 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.7 | 4.5 | 753.2 | 477.4 | 477.3753 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.1137 | 0.0026 | 0.0003 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004.1383 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.6 | 4.5 | 753.1 | 477.2 | 477.1557 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.1137 | 0.0026 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8471 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.5 | 4.5 | 753 | 477.3 | 477.3381 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.1137 | 0.0027 | 0.0003 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8302 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.5 | 4.5 | 753 | 477.3 | 477.3 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.1138 | 0.0027 | 0.0003 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.9074 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.5 | 4.5 | 753 | 477.3 | 477.3241 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.1138 | 0.0027 | 0.0003 | 0 |  | n/a |  | 32.5445 | 55 | 1040 | 1005 | 10.9846 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.4 | 4.5 | 752.9 | 477.2 | 477.1554 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2136 | 0.1138 | 0.0027 | 0.000286 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9382 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.06 | 0.06 | 600 | 748.6 | 4.5 | 753.1 | 477.2 | 477.1873 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.1138 | 0.0027 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8221 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.068 | 0.06 | 600 | 749.3 | 4.5 | 753.8 | 477.4 | 477.4489 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.1139 | 0.0027 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7323 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.07 | 0.06 | 600 | 749.4 | 4.5 | 753.9 | 477.1 | 477.0506 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.1139 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.8867 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.075 | 0.0592 | 592.1583 | 749.6 | 4.5 | 754.1 | 477.2 | 477.1885 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2862 | 0.1139 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.08 | 0.059 | 590 | 749.5 | 4.5 | 754 | 477 | 477.008 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2861 | 0.1139 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.8046 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.08 | 0.059 | 590 | 749.6 | 4.5 | 754.1 | 477.1 | 477.0644 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2862 | 0.11391 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004 | 10.7497 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.08 | 0.059 | 590 | 749.7 | 4.5 | 754.2 | 477 | 477 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.114 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004.138 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.08 | 0.059 | 590 | 750.2 | 4.4862 | 754.6 | 477.1 | 477.1043 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2138 | 0.11401 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.5
2008-10-20 | 16:23 | 6.074 | 0.059 | 590 | 750.5 | 4.4 | 754.9 | 477.1 | 477.1135 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040.9151 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.7 | 4.4 | 755.1 | 477 | 477.0126 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.033 | 189 | 27 | 5 | 195 | 59.2862 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.7 | 4.4 | 755.1 | 477 | 477.0175 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1004 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.7 | 4.4 | 755.1 | 476.9 | 476.9149 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.967 | 189 | 27 | 5 | 195 | 59.2138 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1004 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.7 | 4.4 | 755.1 | 477.3 | 477.2874 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.033 | 189 | 27 | 5 | 195 | 59.3 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.7 | 4.4 | 755.1 | 476.9 | 476.9258 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.8576 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 750.8 | 4.4 | 755.2 | 477.2 | 477.1639 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1141 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1041 | 1005 | 10.7804 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.4 | 755.7 | 477 | 476.9874 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2137 | 0.1142 | 0.0028 | 0.0002 | 0 |  | n/a |  | 32.4555 | 55 | 1041 | 1005 | 10.7032 | 27.3 | 4.4
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.3 | 755.6 | 477.1 | 477.0507 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1142 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040.0771 | 1005 | 10.848 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.1 | 4.3 | 755.4 | 477.2 | 477.1549 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1142 | 0.002814 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.3 | 755.6 | 477.3 | 477.2558 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1142 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.2 | 4.3 | 755.5 | 477.1 | 477.0557 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1142 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.8895 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.1 | 4.3 | 755.4 | 477.2 | 477.1743 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1142 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.7356 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.065 | 0.059 | 590 | 751.1 | 4.3 | 755.4 | 477.2 | 477.2199 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.214 | 0.1143 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.063 | 0.059 | 590 | 751.2 | 4.3 | 755.5 | 477.1 | 477.1133 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1143 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.774 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.2862 | 755.6 | 477.2 | 477.2168 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1143 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477.1 | 477.0751 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1143 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477.2 | 477.2375 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1143 | 0.0029 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1004.8614 | 10.7627 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.4 | 4.2 | 755.6 | 477.1 | 477.0921 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004 | 10.7917 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.5 | 4.2 | 755.7 | 477 | 477 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1004 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477 | 477 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.8995 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477.2 | 477.2341 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1004.8623 | 10.7452 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.5 | 4.2 | 755.7 | 477.1 | 477.1278 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.5445 | 55 | 1040 | 1004 | 10.7 | 27.3 | 4.2
2008-10-20 | 16:23 | 6.07 | 0.059 | 590 | 751.4 | 4.2 | 755.6 | 477.1 | 477.1 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.4555 | 55 | 1040 | 1004.1414 | 10.7635 | 27.3 | 4.2
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477.1 | 477.1367 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:24 | 6.065 | 0.059 | 590 | 751.3 | 4.2 | 755.5 | 477.1 | 477.1494 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1143 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:24 | 6.063 | 0.059 | 590 | 751.4 | 4.2 | 755.6 | 477.2 | 477.2249 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.544 | 55 | 1040 | 1005 | 10.7733 | 27.3 | 4.2
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.3 | 4.1861 | 755.5 | 477.1 | 477.0858 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 32 | 55 | 1040 | 1005 | 10.781 | 27.3 | 4.2
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.3 | 4.1 | 755.4 | 477.3 | 477.2529 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1040 | 1005 | 10.8646 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.4 | 4.1 | 755.5 | 477.1 | 477.1 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7102 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.4 | 4.1 | 755.5 | 477.3 | 477.273 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.07 | 0.059 | 590 | 751.3 | 4.1 | 755.4 | 477 | 476.9501 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7985 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.07 | 0.0603 | 602.9027 | 751.4 | 4.1 | 755.5 | 477.3 | 477.3158 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.8736 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.139 | 0.061 | 610 | 751.4 | 4.1 | 755.5 | 477.2 | 477.2252 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2141 | 0.1144 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8036 | 27.3 | 4.1
2008-10-20 | 16:24 | 6.615 | 0.0635 | 635.1515 | 753.5 | 4.1 | 757.6 | 477.4 | 477.3727 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.214 | 0.11441 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8808 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.1 | 0.0699 | 698.7807 | 791.7 | 4.1 | 795.8 | 501.6 | 501.5615 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1187 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.069 | 690.3133 | 844.5 | 4.1 | 848.6 | 547.7 | 547.7404 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1242 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.069 | 690 | 885.7 | 4.1 | 889.8 | 567.8 | 567.8471 | 0.1 | 13 | 958 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2139 | 0.12847 | 0.0031 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.069 | 690 | 894.9 | 4.1 | 899 | 571.1 | 571.13 | 0.1 | 13 | 958 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.12934 | 0.0031 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.0695 | 694.6011 | 897.2 | 4.1 | 901.3 | 572.2 | 572.2097 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2137 | 0.12961 | 0.0032 | 0.0002 | 0 |  | n/a |  | 32.5445 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.07 | 700 | 897.3 | 4.1 | 901.4 | 572.3 | 572.3286 | 0.1 | 13 | 958 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1297 | 0.0032 | 0.0002 | 0 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.07 | 700 | 897.3 | 4.1 | 901.4 | 572.8 | 572.7514 | 0.1 | 13 | 958 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1297 | 0.0032 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.23 | 0.07 | 700 | 897.4 | 4.1 | 901.5 | 572.7 | 572.7133 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1297 | 0.0032 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.235 | 0.07 | 700 | 897.5 | 4.1 | 901.6 | 572.7 | 572.7229 | 0.1 | 13 | 958 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.12971 | 0.0033 | 0.000214 | 1.4e-05 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.9 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.24 | 0.0706 | 706.2823 | 897.8 | 4.1 | 901.9 | 572.8 | 572.8 | 0.1 | 13 | 958 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2859 | 0.1298 | 0.0033 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.8476 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.24 | 0.071 | 710 | 897.8 | 4.1 | 901.9 | 572.9 | 572.943 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1298 | 0.0033 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7704 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.24 | 0.071 | 710 | 898.1 | 4.1 | 902.2 | 573 | 573 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1299 | 0.0034 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.707 | 27.3 | 4.1
2008-10-20 | 16:24 | 7.24 | 0.071 | 710 | 897.9 | 4.2 | 902.1 | 572.9 | 572.9446 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1299 | 0.0034 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7842 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.24 | 0.071 | 710 | 897.9 | 4.2 | 902.1 | 573.1 | 573.1079 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1299 | 0.0034 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.8614 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.245 | 0.071 | 710 | 897.7 | 4.2 | 901.9 | 573.1 | 573.1079 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.12991 | 0.0035 | 0.0003 | 0.0001 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.6 | 4.2 | 901.8 | 573.2 | 573.1898 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.13 | 0.0035 | 0.0003 | 0.0001 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.8685 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.6 | 4.2 | 901.8 | 572.9 | 572.9368 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.13 | 0.003514 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7142 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.6 | 4.2 | 901.8 | 573.3 | 573.3 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.13 | 0.0036 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6299 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.6 | 4.2 | 901.8 | 573.1 | 573.1316 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1301 | 0.0036 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7417 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.7 | 4.2 | 901.9 | 573.2 | 573.2182 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1301 | 0.0036 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.6 | 4.2 | 901.8 | 573.3 | 573.2643 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1301 | 0.003614 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8954 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.25 | 0.071 | 710 | 897.7 | 4.2 | 901.9 | 573.3 | 573.3074 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1301 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6637 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.251 | 0.071 | 710 | 897.7 | 4.2138 | 901.9 | 573.2 | 573.1623 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1302 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7678 | 27.3 | 4.2
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 897.8 | 4.3 | 902.1 | 573.1 | 573.1 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1302 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 898.4 | 4.3 | 902.7 | 573.3 | 573.3 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.13021 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 898.6 | 4.3 | 902.9 | 573.2 | 573.2467 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1303 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 898.6 | 4.3 | 902.9 | 573.3 | 573.2849 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1303 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7706 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 899.4 | 4.3 | 903.7 | 572.8 | 572.7641 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.13031 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 899.8 | 4.3 | 904.1 | 573.1 | 573.0789 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1304 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 899.9 | 4.3 | 904.2 | 572.9 | 572.9 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1304 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.268 | 0.071 | 710 | 900.1 | 4.3 | 904.4 | 573.1 | 573.1441 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.13041 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 900.6 | 4.3 | 904.9 | 573.1 | 573.1374 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1305 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 900.6 | 4.3 | 904.9 | 573.4 | 573.3632 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1305 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7554 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 901 | 4.3 | 905.3 | 573.3 | 573.2906 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.13051 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 902.4 | 4.3 | 906.7 | 573.1 | 573.1 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.1307 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 902.6 | 4.3 | 906.9 | 573.2 | 573.2 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1307 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7818 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 902.6 | 4.3 | 906.9 | 573.3 | 573.345 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1307 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7726 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 903.3 | 4.3 | 907.6 | 573.3 | 573.3037 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1307 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8727 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 904 | 4.3 | 908.3 | 573.2 | 573.2 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1308 | 0.0037 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7186 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.267 | 0.071 | 710 | 903.9 | 4.3 | 908.2 | 573.3 | 573.2506 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.003714 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8354 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 903.9 | 4.3 | 908.2 | 573 | 573.0271 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2141 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.81 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.26 | 0.071 | 710 | 903.9 | 4.314 | 908.2 | 573.3 | 573.2556 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 32.544 | 55 | 1041 | 1005 | 10.6776 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.268 | 0.071 | 710 | 903.8 | 4.4 | 908.2 | 573.3 | 573.3 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 32.456 | 55 | 1041 | 1005 | 10.6005 | 27.3 | 4.4
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 903.9 | 4.4 | 908.3 | 573.5 | 573.4595 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6767 | 27.3 | 4.4
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 903.9 | 4.4 | 908.3 | 573.1 | 573.1091 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:24 | 7.27 | 0.0716 | 716.0117 | 903.8 | 4.3861 | 908.2 | 573 | 573.0499 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:24 | 7.27 | 0.072 | 720 | 903.8 | 4.3138 | 908.1 | 573.2 | 573.2107 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.3
2008-10-20 | 16:24 | 7.27 | 0.0717 | 717.2789 | 904 | 4.4 | 908.4 | 573.2 | 573.2 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:24 | 7.27 | 0.071 | 710 | 904.1 | 4.4 | 908.5 | 573.2 | 573.2 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.27 | 0.071 | 710 | 903.9 | 4.4 | 908.3 | 573 | 572.9927 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7796 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.266 | 0.071 | 710 | 903.9 | 4.4 | 908.3 | 573.3 | 573.3266 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.883 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.9 | 4.4 | 908.3 | 573.2 | 573.2257 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8059 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.8 | 4.4 | 908.2 | 573.4 | 573.4346 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8713 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.7 | 4.4 | 908.1 | 573.4 | 573.413 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8031 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.6 | 4.4 | 908 | 573.4 | 573.3525 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.6 | 4.4 | 908 | 573.2 | 573.2281 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7029 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.6 | 4.4 | 908 | 573.4 | 573.3753 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.78 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.26 | 0.071 | 710 | 903.8 | 4.4 | 908.2 | 573.4 | 573.4221 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7428 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.326 | 0.071 | 710 | 903.7 | 4.4 | 908.1 | 573.3 | 573.3 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 7.791 | 0.0738 | 737.5191 | 903.6 | 4.4 | 908 | 573.5 | 573.4752 | 0.1 | 13.1 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1308 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 8.235 | 0.0809 | 808.7848 | 920.4 | 4.4 | 924.8 | 589.6 | 589.5753 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.13243 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 8.416 | 0.08 | 800.346 | 1008.5 | 4.4 | 1012.9 | 645 | 644.9626 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1408 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 8.45 | 0.08 | 800 | 1035.2 | 4.4 | 1039.6 | 664 | 664.0251 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14316 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 8.45 | 0.0806 | 806.0083 | 1043.3 | 4.4138 | 1047.7 | 668.2 | 668.2268 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.14394 | 0.0038 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.4
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1046.3 | 4.5 | 1050.8 | 669 | 669 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.14421 | 0.003814 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.5
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1046.8 | 4.5 | 1051.3 | 669.3 | 669.3038 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1443 | 0.0039 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.5
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1046.8 | 4.5139 | 1051.3 | 669.4 | 669.3812 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.1443 | 0.0039 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.5
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1046.9 | 4.6 | 1051.5 | 669.3 | 669.3121 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1444 | 0.0039 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.6
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1046.9 | 4.6 | 1051.5 | 669.5 | 669.5 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1444 | 0.004 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.6
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1047 | 4.6 | 1051.6 | 669.6 | 669.575 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14441 | 0.004 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.6
2008-10-20 | 16:25 | 8.45 | 0.0814 | 814.2917 | 1048.3 | 4.6 | 1052.9 | 669.6 | 669.6 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.14451 | 0.004 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5797 | 27.3 | 4.6
2008-10-20 | 16:25 | 8.45 | 0.0817 | 817.31 | 1049.9 | 4.6 | 1054.5 | 669.7 | 669.7 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.14471 | 0.004 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5749 | 27.3 | 4.6
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1051.2 | 4.7 | 1055.9 | 669.9 | 669.8702 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1448 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.081 | 810 | 1051.1 | 4.7 | 1055.8 | 669.7 | 669.7464 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1448 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.0818 | 817.7047 | 1051.3 | 4.7 | 1056 | 669.8 | 669.8 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1448 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.082 | 820 | 1051.5 | 4.7 | 1056.2 | 669.8 | 669.7743 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1449 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.0829 | 828.55 | 1051.4 | 4.7 | 1056.1 | 670 | 669.9514 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2861 | 0.1449 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.084 | 840 | 1051.6 | 4.7 | 1056.3 | 669.9 | 669.8533 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1449 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6993 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.0838 | 837.5696 | 1052.4 | 4.7 | 1057.1 | 669.8 | 669.8229 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.145 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6221 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.0821 | 820.692 | 1052.6 | 4.7 | 1057.3 | 670 | 670.0155 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.145 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6551 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.45 | 0.082 | 820 | 1052.5 | 4.7 | 1057.2 | 669.9 | 669.8723 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.145 | 0.0041 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.453 | 0.082 | 820 | 1052.5 | 4.7 | 1057.2 | 669.9 | 669.9 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.145 | 0.004114 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.0824 | 824.2667 | 1052.5 | 4.7 | 1057.2 | 669.7 | 669.7056 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.145 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1414 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052.3 | 4.7 | 1057 | 670.1 | 670.1251 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.145 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 669.9 | 669.9451 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.145 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 670.1 | 670.0808 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.145 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7363 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 669.8 | 669.8 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.145 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8906 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1051.8 | 4.7 | 1056.5 | 670 | 669.9567 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.14501 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7551 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1051.8 | 4.7 | 1056.5 | 670 | 670.0463 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1451 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 670.3 | 670.2732 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1451 | 0.0042 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6731 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052.2 | 4.7 | 1056.9 | 669.9 | 669.915 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.604 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052.1 | 4.7 | 1056.8 | 670.1 | 670.0742 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6812 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.46 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 669.9 | 669.9252 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.463 | 0.083 | 830 | 1051.9 | 4.7 | 1056.6 | 669.9 | 669.9215 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.286 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1051.9 | 4.7 | 1056.6 | 670.2 | 670.1533 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7253 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1051.9 | 4.7 | 1056.6 | 669.9 | 669.9 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8798 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.1 | 4.7 | 1056.8 | 670.1 | 670.0643 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7664 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.3 | 4.7 | 1057 | 669.9 | 669.8707 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1451 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.5 | 4.7 | 1057.2 | 670.1 | 670.1478 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1452 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.7 | 4.7 | 1057.4 | 670 | 669.9528 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1452 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.6 | 4.7 | 1057.3 | 670 | 669.9868 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2853 | 0.1452 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.851 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.2 | 4.7 | 1056.9 | 669.8 | 669.8105 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1452 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7944 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.2 | 4.7 | 1056.9 | 669.8 | 669.8253 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2859 | 0.1452 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8586 | 10.7 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.4 | 4.7 | 1057.1 | 669.9 | 669.8513 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1452 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6782 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.2 | 4.7 | 1056.9 | 670 | 669.9672 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.14521 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4468 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052 | 4.7 | 1056.7 | 670.1 | 670.0882 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7079 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.1 | 4.7 | 1056.8 | 669.9 | 669.9 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8225 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1052.1 | 4.7 | 1056.8 | 670.2 | 670.1563 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.684 | 27.3 | 4.7
2008-10-20 | 16:25 | 8.47 | 0.083 | 830 | 1051.8 | 4.8 | 1056.6 | 670 | 669.963 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005.1383 | 10.6068 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1051.8 | 4.8 | 1056.6 | 669.9 | 669.9 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8111 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1051.8 | 4.8 | 1056.6 | 670 | 669.9819 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7574 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.1 | 4.8 | 1056.9 | 670.3 | 670.2628 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6247 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.2 | 4.8 | 1057 | 669.9 | 669.8909 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2857 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.2 | 4.8 | 1057 | 669.8 | 669.8491 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.2 | 4.8 | 1057 | 669.9 | 669.9 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.004514 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.4 | 4.8 | 1057.2 | 669.9 | 669.9384 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1453 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.2 | 4.8 | 1057 | 670.1 | 670.0891 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.1453 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.47 | 0.083 | 830 | 1052.2 | 4.8138 | 1057 | 669.9 | 669.9 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.14531 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.7 | 27.3 | 4.8
2008-10-20 | 16:26 | 8.479 | 0.083 | 830 | 1052.3 | 4.9 | 1057.2 | 670 | 670 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.1454 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 4.9
2008-10-20 | 16:26 | 8.545 | 0.0861 | 861.3719 | 1052.4 | 4.9 | 1057.3 | 670 | 669.9711 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2863 | 0.1454 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6579 | 27.3 | 4.9
2008-10-20 | 16:26 | 9.321 | 0.0948 | 947.6335 | 1054.3 | 4.9 | 1059.2 | 675.5 | 675.4531 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.14558 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6193 | 27.3 | 4.9
2008-10-20 | 16:26 | 9.57 | 0.0931 | 931.1997 | 1129.6 | 4.9 | 1134.5 | 724.5 | 724.5328 | 0.1 | 13 | 958 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1521 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6965 | 27.3 | 4.9
2008-10-20 | 16:26 | 9.57 | 0.093 | 930 | 1167.9 | 4.9136 | 1172.9 | 756.6 | 756.6195 | 0.1 | 13 | 958 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.15526 | 0.0046 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6264 | 27.3 | 4.9
2008-10-20 | 16:26 | 9.573 | 0.093 | 930 | 1189.9 | 5 | 1194.9 | 764.2 | 764.2296 | 0.18 | 13 | 958 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.15707 | 0.004614 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6508 | 27.3 | 5
2008-10-20 | 16:26 | 9.58 | 0.093 | 930 | 1196.3 | 5.1 | 1201.4 | 765.6 | 765.5755 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2859 | 0.1576 | 0.0047 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 5.1
2008-10-20 | 16:26 | 9.58 | 0.0933 | 932.7703 | 1196.7 | 5.214 | 1201.9 | 766 | 766.0498 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.286 | 0.15771 | 0.004714 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.1397 | 10.7 | 27.3 | 5.2
2008-10-20 | 16:26 | 9.588 | 0.094 | 940 | 1198.5 | 5.3138 | 1203.8 | 766.3 | 766.285 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.1579 | 0.0048 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.7 | 27.3 | 5.3
2008-10-20 | 16:26 | 9.608 | 0.094 | 940 | 1198.5 | 5.4 | 1203.9 | 766.5 | 766.487 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.1579 | 0.0048 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6408 | 27.3 | 5.4
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1198.2 | 5.4138 | 1203.6 | 766.2 | 766.2288 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1579 | 0.0048 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6361 | 27.3 | 5.4
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1198.4 | 5.5 | 1203.9 | 766.2 | 766.2448 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.158 | 0.0048 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1198.7 | 5.5 | 1204.2 | 766.2 | 766.2209 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.158 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1199.9 | 5.5 | 1205.4 | 766.6 | 766.6 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.15811 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1201.2 | 5.5 | 1206.7 | 766.5 | 766.5179 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1582 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7903 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1201.3 | 5.5 | 1206.8 | 766.6 | 766.5856 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1582 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8326 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.62 | 0.094 | 940 | 1201.5 | 5.5 | 1207 | 766.5 | 766.5 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1582 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6009 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.623 | 0.094 | 940 | 1201.5 | 5.5 | 1207 | 766.4 | 766.4288 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1582 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6769 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.7 | 5.5 | 1207.2 | 766.4 | 766.4247 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.15821 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202 | 5.5 | 1207.5 | 766.4 | 766.4341 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1583 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.6688 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202 | 5.5 | 1207.5 | 766.5 | 766.5142 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1583 | 0.0049 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6083 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202.2 | 5.5 | 1207.7 | 766.1 | 766.1296 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1583 | 0.004914 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6855 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202.2 | 5.5 | 1207.7 | 766.5 | 766.5 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1583 | 0.005 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6373 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202.2 | 5.5 | 1207.7 | 766.2 | 766.2 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1583 | 0.005 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7195 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1202.1 | 5.5 | 1207.6 | 766.5 | 766.5 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1583 | 0.005 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.9 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.9 | 5.5 | 1207.4 | 766.3 | 766.3378 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1584 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.8 | 5.5 | 1207.3 | 766.7 | 766.7369 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1584 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.9 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.6 | 5.5 | 1207.1 | 766.7 | 766.7144 | 0.2 | 13 | 956 | 952 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8029 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.6 | 5.5 | 1207.1 | 766.7 | 766.6748 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.63 | 0.094 | 940 | 1201.4 | 5.5 | 1206.9 | 766.7 | 766.7298 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.631 | 0.094 | 940 | 1201.7 | 5.5 | 1207.2 | 766.6 | 766.6206 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.639 | 0.094 | 940 | 1201.9 | 5.5 | 1207.4 | 766.8 | 766.79 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.64 | 0.094 | 940 | 1201.7 | 5.5 | 1207.2 | 766.7 | 766.6744 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.64 | 0.094 | 940 | 1201.7 | 5.5 | 1207.2 | 766.8 | 766.8238 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.644 | 0.094 | 940 | 1201.7 | 5.5 | 1207.2 | 766.5 | 766.5378 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.5 | 5.5 | 1207 | 766.7 | 766.6752 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1584 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.3 | 5.5 | 1206.8 | 766.5 | 766.4545 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6569 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.3 | 5.5141 | 1206.8 | 766.6 | 766.6075 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.5
2008-10-20 | 16:26 | 9.65 | 0.0941 | 940.7833 | 1201.5 | 5.6 | 1207.1 | 766.7 | 766.7076 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.286 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.0949 | 949.1167 | 1201.6 | 5.6 | 1207.2 | 766.6 | 766.6438 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.095 | 950 | 1201.5 | 5.6 | 1207.1 | 766.8 | 766.8261 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6514 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.095 | 950 | 1201.8 | 5.6 | 1207.4 | 766.6 | 766.6 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.0946 | 945.8868 | 1201.6 | 5.6 | 1207.2 | 766.4 | 766.3761 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7109 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.6 | 5.6 | 1207.2 | 766.5 | 766.4758 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8663 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.5 | 5.6 | 1207.1 | 766.7 | 766.7234 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8397 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.3 | 5.6 | 1206.9 | 766.8 | 766.7726 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8375 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.3 | 5.6 | 1206.9 | 766.8 | 766.8293 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2857 | 0.1585 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8703 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.65 | 0.094 | 940 | 1201.2 | 5.6 | 1206.8 | 766.6 | 766.6185 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1585 | 0.005214 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7159 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.648 | 0.094 | 940 | 1201.4 | 5.6 | 1207 | 766.9 | 766.9493 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.631 | 0.0943 | 942.5267 | 1201.7 | 5.6 | 1207.3 | 766.7 | 766.7306 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6536 | 27.3 | 5.6
2008-10-20 | 16:26 | 9.63 | 0.095 | 950 | 1201.9 | 5.6 | 1207.5 | 766.6 | 766.6209 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.63 | 0.095 | 950 | 1201.9 | 5.6 | 1207.5 | 766.8 | 766.776 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1388 | 10.6 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.634 | 0.095 | 950 | 1201.9 | 5.6 | 1207.5 | 766.7 | 766.6539 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8632 | 10.6 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.64 | 0.095 | 950 | 1202.1 | 5.6 | 1207.7 | 767 | 767 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7653 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.64 | 0.0946 | 945.8258 | 1202.1 | 5.6 | 1207.7 | 767 | 767 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.9 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.64 | 0.0942 | 942.425 | 1202 | 5.6 | 1207.6 | 766.8 | 766.7589 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.8812 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.64 | 0.095 | 950 | 1202 | 5.6 | 1207.6 | 766.5 | 766.5285 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7269 | 27.3 | 5.6
2008-10-20 | 16:27 | 9.64 | 0.095 | 950 | 1201.8 | 5.7 | 1207.5 | 766.8 | 766.7617 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.005314 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.7
2008-10-20 | 16:27 | 9.636 | 0.095 | 950 | 1201.8 | 5.7 | 1207.5 | 766.7 | 766.7182 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.7
2008-10-20 | 16:27 | 9.63 | 0.095 | 950 | 1201.5 | 5.7 | 1207.2 | 767.1 | 767.0721 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1586 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.7
2008-10-20 | 16:27 | 9.63 | 0.095 | 950 | 1201.4 | 5.7 | 1207.1 | 766.9 | 766.89 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1587 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.7
2008-10-20 | 16:27 | 9.63 | 0.097 | 969.7599 | 1201.5 | 5.7 | 1207.2 | 766.9 | 766.9 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1587 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.356 | 0.103 | 1030 | 1204.7 | 5.7 | 1210.4 | 771.6 | 771.5527 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.15895 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6503 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.725 | 0.103 | 1030 | 1254 | 5.7 | 1259.7 | 803.2 | 803.1993 | 0.2 | 13 | 958 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.16283 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6269 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.81 | 0.1077 | 1076.5368 | 1312.3 | 5.7 | 1318 | 849.3 | 849.3279 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1654 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.696 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.81 | 0.109 | 1090 | 1336.2 | 5.7 | 1341.9 | 859.3 | 859.3129 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.16907 | 0.0055 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6188 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.809 | 0.109 | 1090 | 1342.8 | 5.7 | 1348.5 | 861.9 | 861.8656 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.16961 | 0.0055 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.801 | 0.1088 | 1087.648 | 1345.5 | 5.7138 | 1351.2 | 862.6 | 862.6133 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1698 | 0.005514 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.7
2008-10-20 | 16:27 | 10.8 | 0.108 | 1080 | 1345.6 | 5.9 | 1351.5 | 863 | 863 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1699 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6127 | 27.3 | 5.9
2008-10-20 | 16:27 | 10.8 | 0.108 | 1080 | 1345.6 | 5.9 | 1351.5 | 862.7 | 862.7383 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1699 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6899 | 27.3 | 5.9
2008-10-20 | 16:27 | 10.787 | 0.108 | 1080 | 1345.9 | 6 | 1351.9 | 863 | 862.9586 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.286 | 0.16991 | 0.005614 | 0.000314 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.108 | 1080 | 1346.3 | 6 | 1352.3 | 862.8 | 862.8 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.17 | 0.0057 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.108 | 1080 | 1346.4 | 6 | 1352.4 | 863 | 863.013 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.17 | 0.0057 | 0.000386 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6792 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.108 | 1080 | 1346.8 | 6 | 1352.8 | 863.3 | 863.2549 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1701 | 0.0057 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6022 | 27.3 | 6
2008-10-20 | 16:27 | 10.763 | 0.108 | 1080 | 1347.7 | 6 | 1353.7 | 863.2 | 863.1583 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.17011 | 0.005714 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6755 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.108 | 1080 | 1348.5 | 6 | 1354.5 | 863.3 | 863.3263 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1702 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.1072 | 1072.2222 | 1348.8 | 6 | 1354.8 | 863.1 | 863.1268 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1702 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1349.2 | 6 | 1355.2 | 863.3 | 863.2749 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.17021 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.7145 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1349.3 | 6 | 1355.3 | 863.1 | 863.1255 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1703 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.8688 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1350 | 6 | 1356 | 863.3 | 863.3 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.17031 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7768 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1350.4 | 6 | 1356.4 | 863.4 | 863.4 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1704 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1397 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1351.6 | 6 | 1357.6 | 863.6 | 863.5622 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.17041 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1351.7 | 6 | 1357.7 | 863.4 | 863.4078 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1705 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1351.9 | 6.0138 | 1357.9 | 863.3 | 863.2632 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1705 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1351.5 | 6.1 | 1357.6 | 863.2 | 863.2357 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2863 | 0.1705 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7951 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.76 | 0.107 | 1070 | 1351.7 | 6.1 | 1357.8 | 863.3 | 863.3161 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1706 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8259 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.768 | 0.107 | 1070 | 1351.9 | 6.1 | 1358 | 863.4 | 863.387 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1706 | 0.0059 | 0.000314 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.6019 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.8 | 6.1 | 1357.9 | 863.3 | 863.2725 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2136 | 0.1706 | 0.0059 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.679 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.8 | 6.1 | 1357.9 | 863.5 | 863.5 | 0.2 | 13.1 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1706 | 0.0059 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6.1 | 1358 | 863 | 863.0286 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1706 | 0.005914 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.7 | 6 | 1357.7 | 862.6 | 862.5649 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.286 | 0.1706 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6894 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1352 | 6 | 1358 | 862.5 | 862.5 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1706 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6123 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.8 | 6 | 1357.8 | 862.7 | 862.6863 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1706 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6649 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 863.2 | 863.1794 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1706 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.7 | 6 | 1357.7 | 863.7 | 863.6838 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.17061 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6808 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 864 | 864 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1707 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6036 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 864 | 864.0124 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1707 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6736 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 864.1 | 864.1 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1707 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.6492 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 864.1 | 864.1 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.1707 | 0.006 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.628 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 864 | 863.9956 | 0.2 | 13.1 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1707 | 0.006014 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7103 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 863.9 | 863.9371 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1707 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8644 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1352.1 | 6.1 | 1358.2 | 863.9 | 863.9262 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1707 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1394 | 10.7817 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.9 | 6.0862 | 1358 | 863.8 | 863.8172 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1707 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.1
2008-10-20 | 16:27 | 10.77 | 0.107 | 1070 | 1351.7 | 6 | 1357.7 | 863.7 | 863.7377 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1707 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.1069 | 1069.3578 | 1352.1 | 6 | 1358.1 | 863.9 | 863.8853 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1707 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.1061 | 1061.0175 | 1352 | 6 | 1358 | 863.5 | 863.5269 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1708 | 0.0061 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6322 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.106 | 1060 | 1352 | 6 | 1358 | 863.5 | 863.5049 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1708 | 0.006114 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6453 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.106 | 1060 | 1352.2 | 6 | 1358.2 | 863.7 | 863.7 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1708 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6775 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.106 | 1060 | 1352.1 | 6 | 1358.1 | 863.5 | 863.4804 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1708 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6003 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.106 | 1060 | 1352.4 | 6 | 1358.4 | 863.5 | 863.5451 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2857 | 0.17081 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6769 | 27.3 | 6
2008-10-20 | 16:27 | 10.77 | 0.1061 | 1060.6147 | 1352.4 | 6 | 1358.4 | 863.5 | 863.547 | 0.2 | 13.1 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1709 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5918 | 27.3 | 6
2008-10-20 | 16:28 | 10.77 | 0.1069 | 1069.2727 | 1352.2 | 6 | 1358.2 | 863.5 | 863.5498 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1709 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5625 | 27.3 | 6
2008-10-20 | 16:28 | 10.774 | 0.107 | 1070 | 1352.2 | 6 | 1358.2 | 863.7 | 863.687 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1709 | 0.0062 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.1064 | 1064.387 | 1352.1 | 6 | 1358.1 | 863.7 | 863.7 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.106 | 1060 | 1352.2 | 6 | 1358.2 | 863.7 | 863.6885 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6373 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.1063 | 1062.5303 | 1352 | 6 | 1358 | 863.4 | 863.4106 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.107 | 1070 | 1352 | 6 | 1358 | 863.5 | 863.5363 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6171 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 863.5 | 863.5269 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6942 | 27.3 | 6
2008-10-20 | 16:28 | 10.78 | 0.107 | 1070 | 1351.9 | 6 | 1357.9 | 863.7 | 863.6558 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5571 | 27.3 | 6
2008-10-20 | 16:28 | 10.778 | 0.107 | 1070 | 1351.7 | 6.1 | 1357.8 | 863.7 | 863.7 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.1709 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6946 | 27.3 | 6.1
2008-10-20 | 16:28 | 10.779 | 0.107 | 1070 | 1351.7 | 6.1 | 1357.8 | 863.9 | 863.875 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2857 | 0.17091 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.8225 | 27.3 | 6.1
2008-10-20 | 16:28 | 10.901 | 0.1092 | 1092.4354 | 1351.9 | 6.1 | 1358 | 863.9 | 863.9488 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.171 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.603 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.623 | 0.1169 | 1168.8324 | 1373.3 | 6.1 | 1379.4 | 869.3 | 869.3273 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.17251 | 0.0063 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6802 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.86 | 0.1152 | 1152.1518 | 1423.6 | 6.1 | 1429.7 | 921.6 | 921.6366 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.176 | 0.006314 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.864 | 0.115 | 1150 | 1474.9 | 6.1 | 1481 | 951.6 | 951.5934 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1782 | 0.0064 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.631 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.87 | 0.1156 | 1155.59 | 1494.5 | 6.1 | 1500.6 | 958.7 | 958.741 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1809 | 0.0064 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5117 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.87 | 0.1164 | 1163.9286 | 1499.2 | 6.1 | 1505.3 | 959.9 | 959.8735 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.1813 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5888 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.87 | 0.117 | 1170 | 1500.2 | 6.2 | 1506.4 | 960.3 | 960.3495 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.18131 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.877 | 0.117 | 1170 | 1500.1 | 6.2 | 1506.3 | 960.6 | 960.6 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1814 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6431 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.117 | 1170 | 1500.6 | 6.2 | 1506.8 | 960.7 | 960.7391 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1814 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6797 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.117 | 1170 | 1501.2 | 6.2136 | 1507.5 | 960.8 | 960.784 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2136 | 0.18141 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6025 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.878 | 0.117 | 1170 | 1501.6 | 6.3 | 1507.9 | 960.8 | 960.7707 | 0.2 | 13 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2141 | 0.1815 | 0.006514 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8586 | 10.6745 | 27.3 | 6.3
2008-10-20 | 16:28 | 11.87 | 0.117 | 1170 | 1502 | 6.3 | 1508.3 | 960.9 | 960.9246 | 0.2 | 13 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1815 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1397 | 10.7 | 27.3 | 6.3
2008-10-20 | 16:28 | 11.87 | 0.117 | 1170 | 1502 | 6.2 | 1508.2 | 960.6 | 960.6 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6714 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.877 | 0.117 | 1170 | 1502.1 | 6.2 | 1508.3 | 960.8 | 960.7756 | 0.2 | 13.1 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1816 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.6057 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.117 | 1170 | 1502.3 | 6.2 | 1508.5 | 960.6 | 960.6 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2137 | 0.1816 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6833 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.1177 | 1177.2477 | 1503.1 | 6.2 | 1509.3 | 960.8 | 960.8259 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.18161 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.821 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503.1 | 6.2 | 1509.3 | 960.8 | 960.7868 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1817 | 0.006614 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8247 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503.1 | 6.2 | 1509.3 | 960.5 | 960.5357 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1817 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503 | 6.2 | 1509.2 | 960.7 | 960.7443 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1817 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503.4 | 6.2 | 1509.6 | 961 | 960.973 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1817 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503.2 | 6.2 | 1509.4 | 960.8 | 960.7562 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1817 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.118 | 1180 | 1503.6 | 6.2 | 1509.8 | 960.8 | 960.7873 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2141 | 0.1818 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.1186 | 1185.744 | 1503.6 | 6.2 | 1509.8 | 960.9 | 960.9 | 0.2 | 13.1 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.214 | 0.1818 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6992 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.119 | 1190 | 1503.9 | 6.2 | 1510.1 | 961 | 961.0374 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1818 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.622 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.119 | 1190 | 1505.5 | 6.2 | 1511.7 | 960.8 | 960.8419 | 0.2 | 13.1 | 950 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1819 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6552 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.119 | 1190 | 1505.3 | 6.2 | 1511.5 | 960.7 | 960.674 | 0.2 | 13 | 950 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1819 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.119 | 1190 | 1505.6 | 6.1862 | 1511.8 | 960.8 | 960.8273 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1819 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.719 | 27.3 | 6.2
2008-10-20 | 16:28 | 11.88 | 0.119 | 1190 | 1505.5 | 6.1 | 1511.6 | 960.5 | 960.5268 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1819 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.8733 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.887 | 0.119 | 1190 | 1505.7 | 6.0862 | 1511.8 | 961.1 | 961.0514 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1819 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7086 | 27.3 | 6.1
2008-10-20 | 16:28 | 11.91 | 0.119 | 1190 | 1505.7 | 6 | 1511.7 | 960.7 | 960.6667 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.182 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.641 | 27.3 | 6
2008-10-20 | 16:28 | 11.91 | 0.119 | 1190 | 1506 | 6 | 1512 | 960.9 | 960.8718 | 0.2 | 13 | 951 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.182 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8637 | 10.6819 | 27.3 | 6
2008-10-20 | 16:28 | 11.91 | 0.1189 | 1189.4157 | 1507.2 | 6 | 1513.2 | 960.3 | 960.3498 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2141 | 0.182 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6047 | 27.3 | 6
2008-10-20 | 16:28 | 11.91 | 0.1181 | 1181.0684 | 1507.3 | 6 | 1513.3 | 960.8 | 960.7688 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.286 | 0.1821 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.8177 | 27.3 | 6
2008-10-20 | 16:28 | 11.91 | 0.118 | 1180 | 1507.4 | 6 | 1513.4 | 960.6 | 960.5878 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1821 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.7507 | 27.3 | 6
2008-10-20 | 16:28 | 11.91 | 0.118 | 1180 | 1507.1 | 6 | 1513.1 | 960.6 | 960.6363 | 0.2 | 13 | 951 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6269 | 27.3 | 6
2008-10-20 | 16:28 | 11.909 | 0.118 | 1180 | 1507.1 | 6 | 1513.1 | 960.6 | 960.5939 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.892 | 0.118 | 1180 | 1506.9 | 6 | 1512.9 | 960.6 | 960.6455 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.1181 | 1180.799 | 1506.9 | 6 | 1512.9 | 960.8 | 960.7533 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1821 | 0.0069 | 0.000386 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7584 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.1189 | 1189.2094 | 1507.4 | 6 | 1513.4 | 960.5 | 960.5 | 0.2 | 13 | 950 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7644 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.119 | 1190 | 1507.4 | 6 | 1513.4 | 960.9 | 960.9183 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.119 | 1190 | 1507.7 | 6 | 1513.7 | 960.8 | 960.7536 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.1186 | 1186.1366 | 1508 | 6 | 1514 | 961.1 | 961.1109 | 0.2 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6331 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1508 | 6 | 1514 | 960.7 | 960.683 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.214 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1003.8603 | 10.6439 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.9 | 6 | 1513.9 | 960.4 | 960.4402 | 0.2 | 13 | 950 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.679 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.3 | 6 | 1513.3 | 960.6 | 960.6 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.602 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.3 | 6 | 1513.3 | 960.6 | 960.6 | 0.2 | 13 | 951 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6757 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.3 | 6 | 1513.3 | 960.5 | 960.5441 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.2 | 6 | 1513.2 | 960.7 | 960.6711 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1506.9 | 6 | 1512.9 | 960.8 | 960.7898 | 0.2 | 13 | 950 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1821 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6927 | 27.3 | 6
2008-10-20 | 16:28 | 11.89 | 0.118 | 1180 | 1507.1 | 6 | 1513.1 | 960.8 | 960.7725 | 0.2 | 13 | 950 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2136 | 0.18211 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8637 | 10.6155 | 27.3 | 6
2008-10-20 | 16:29 | 11.89 | 0.118 | 1180 | 1507.3 | 6 | 1513.3 | 961.1 | 961.0748 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6617 | 27.3 | 6
2008-10-20 | 16:29 | 11.898 | 0.118 | 1180 | 1507.6 | 6 | 1513.6 | 960.5 | 960.4874 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 11.91 | 0.118 | 1180 | 1507.6 | 6 | 1513.6 | 960.8 | 960.8409 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 11.909 | 0.118 | 1180 | 1508.6 | 6 | 1514.6 | 960.5 | 960.4545 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 11.901 | 0.118 | 1180 | 1508.7 | 6 | 1514.7 | 960.7 | 960.7434 | 0.2 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6296 | 27.3 | 6
2008-10-20 | 16:29 | 11.9 | 0.118 | 1180 | 1508.5 | 6 | 1514.5 | 960.4 | 960.4022 | 0.2 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6476 | 27.3 | 6
2008-10-20 | 16:29 | 11.9 | 0.118 | 1180 | 1508.2 | 6 | 1514.2 | 960.4 | 960.3754 | 0.2 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1822 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 11.9 | 0.118 | 1180 | 1508.2 | 6 | 1514.2 | 960.6 | 960.6051 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.006914 | 0.000314 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6981 | 27.3 | 6
2008-10-20 | 16:29 | 11.9 | 0.118 | 1180 | 1507.9 | 6 | 1513.9 | 960.8 | 960.8374 | 0.2 | 13 | 950 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1822 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6209 | 27.3 | 6
2008-10-20 | 16:29 | 11.901 | 0.118 | 1180 | 1507.8 | 6 | 1513.8 | 960.8 | 960.8257 | 0.2 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6562 | 27.3 | 6
2008-10-20 | 16:29 | 11.909 | 0.118 | 1180 | 1508.2 | 6 | 1514.2 | 960.3 | 960.3417 | 0.2 | 13 | 950 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6332 | 27.3 | 6
2008-10-20 | 16:29 | 11.623 | 0.1172 | 1171.875 | 1508.6 | 6 | 1514.6 | 960.4 | 960.3717 | 0.2 | 13 | 950 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1822 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5211 | 27.3 | 6
2008-10-20 | 16:29 | 11.158 | 0.1061 | 1060.5736 | 1486.2 | 6 | 1492.2 | 955.5 | 955.4953 | 0.2 | 13 | 949 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2136 | 0.1818 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6755 | 27.3 | 6
2008-10-20 | 16:29 | 10.87 | 0.1057 | 1057.2749 | 1461 | 6 | 1467 | 900.6 | 900.5796 | 0.2 | 13 | 948 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.17902 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1003.8588 | 10.5052 | 27.3 | 6
2008-10-20 | 16:29 | 10.868 | 0.106 | 1060 | 1399.7 | 6 | 1405.7 | 872.8 | 872.812 | 0.2 | 13 | 949 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.17482 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4843 | 27.3 | 6
2008-10-20 | 16:29 | 10.86 | 0.106 | 1060 | 1370.2 | 6 | 1376.2 | 866.4 | 866.4219 | 0.3 | 13 | 950 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.17277 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.86 | 0.106 | 1060 | 1364.4 | 6 | 1370.4 | 864.2 | 864.1893 | 0.3 | 13 | 950 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.17229 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.845 | 0.106 | 1060 | 1361.5 | 6 | 1367.5 | 863.8 | 863.8376 | 0.28 | 13 | 951 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1721 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.84 | 0.106 | 1060 | 1361 | 6 | 1367 | 863.6 | 863.6298 | 0.2 | 13 | 951 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1721 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.836 | 0.106 | 1060 | 1360.9 | 6 | 1366.9 | 863.4 | 863.4121 | 0.2 | 13 | 951 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.1721 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.828 | 0.106 | 1060 | 1360.4 | 6 | 1366.4 | 863.6 | 863.5636 | 0.2 | 13 | 951 | 952 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.17209 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.82 | 0.106 | 1060 | 1357.5 | 6 | 1363.5 | 863.5 | 863.5 | 0.2 | 13 | 951 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.82 | 0.106 | 1060 | 1357.2 | 6 | 1363.2 | 863.3 | 863.2518 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.007 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.813 | 0.106 | 1060 | 1357.2 | 6 | 1363.2 | 863 | 863.0135 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2141 | 0.1718 | 0.007 | 0.000386 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6363 | 27.3 | 6
2008-10-20 | 16:29 | 10.81 | 0.106 | 1060 | 1357.2 | 6 | 1363.2 | 863.3 | 863.3 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 10.802 | 0.106 | 1060 | 1357.2 | 6 | 1363.2 | 863.2 | 863.213 | 0.2 | 13 | 952 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.4 | 6 | 1363.4 | 863.5 | 863.4846 | 0.2 | 13 | 952 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.632 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863.3 | 863.2849 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5093 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357 | 6 | 1363 | 863.2 | 863.2364 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4451 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.4 | 863.3702 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.007 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5995 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357 | 6 | 1363 | 863.2 | 863.1642 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5231 | 27.3 | 6
2008-10-20 | 16:29 | 10.786 | 0.106 | 1060 | 1357.2 | 6 | 1363.2 | 863.1 | 863.1 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5541 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 863.1 | 863.1 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 863.2 | 863.182 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863.3 | 863.2865 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1356.8 | 6 | 1362.8 | 863.4 | 863.3564 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.214 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1356.9 | 6 | 1362.9 | 863.3 | 863.3 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.6 | 863.5618 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.5 | 863.524 | 0.2 | 13 | 953 | 952 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1356.9 | 6 | 1362.9 | 863.2 | 863.1741 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6715 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 863.2 | 863.2258 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6514 | 27.3 | 6
2008-10-20 | 16:29 | 10.78 | 0.106 | 1060 | 1357.5 | 6 | 1363.5 | 863.2 | 863.2288 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.786 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863.3 | 863.3245 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.597 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863 | 863.0403 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5198 | 27.3 | 6
2008-10-20 | 16:29 | 10.788 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 862.8 | 862.8441 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1414 | 10.5573 | 27.3 | 6
2008-10-20 | 16:29 | 10.781 | 0.106 | 1060 | 1356.9 | 6 | 1362.9 | 863 | 862.9746 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6345 | 27.3 | 6
2008-10-20 | 16:29 | 10.789 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.2 | 863.1739 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6883 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.3 | 863.2832 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2136 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6111 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.4 | 863.376 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4679 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863 | 863.0196 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4864 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1356.9 | 6 | 1362.9 | 863.3 | 863.3022 | 0.2 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863.2 | 863.2493 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 863.3 | 863.3116 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.3 | 6 | 1363.3 | 863.5 | 863.4874 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.5483 | 27.3 | 6
2008-10-20 | 16:29 | 10.79 | 0.106 | 1060 | 1357.1 | 6 | 1363.1 | 863.3 | 863.2745 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5287 | 27.3 | 6
2008-10-20 | 16:29 | 10.794 | 0.106 | 1060 | 1357 | 6 | 1363 | 863.4 | 863.421 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5942 | 27.3 | 6
2008-10-20 | 16:29 | 10.796 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.3 | 863.3478 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5172 | 27.3 | 6
2008-10-20 | 16:29 | 10.779 | 0.106 | 1060 | 1356.7 | 6 | 1362.7 | 863.6 | 863.6 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4397 | 27.3 | 6
2008-10-20 | 16:29 | 10.771 | 0.106 | 1060 | 1356.5 | 6 | 1362.5 | 863.3 | 863.2772 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5134 | 27.3 | 6
2008-10-20 | 16:29 | 10.77 | 0.106 | 1060 | 1356.7 | 6.1 | 1362.8 | 863.3 | 863.3279 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.685 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.77 | 0.106 | 1060 | 1356.6 | 6.1 | 1362.7 | 863.3 | 863.2974 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6079 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.77 | 0.106 | 1060 | 1356.6 | 6.1 | 1362.7 | 863.4 | 863.4 | 0.2 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.17179 | 0.0069 | 0.000314 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.77 | 0.106 | 1060 | 1356.6 | 6.1 | 1362.7 | 863 | 863.0492 | 0.2 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.771 | 0.106 | 1060 | 1356.8 | 6.1 | 1362.9 | 863.3 | 863.3 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2136 | 0.17171 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.779 | 0.106 | 1060 | 1356.9 | 6.1 | 1363 | 863.1 | 863.1287 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1718 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.78 | 0.106 | 1060 | 1357.1 | 6.1 | 1363.2 | 863.4 | 863.3529 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.78 | 0.106 | 1060 | 1356.9 | 6.1 | 1363 | 863.1 | 863.0693 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4895 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.78 | 0.106 | 1060 | 1356.9 | 6.1 | 1363 | 863.2 | 863.1739 | 0.2 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1718 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4972 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.78 | 0.106 | 1060 | 1356.9 | 6.1138 | 1363 | 863.4 | 863.3508 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1718 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6904 | 27.3 | 6.1
2008-10-20 | 16:30 | 10.78 | 0.106 | 1060 | 1356.6 | 6.2 | 1362.8 | 863.2 | 863.2 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6133 | 27.3 | 6.2
2008-10-20 | 16:30 | 10.78 | 0.1062 | 1062.1093 | 1356.4 | 6.2 | 1362.6 | 863.1 | 863.1 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4722 | 27.3 | 6.2
2008-10-20 | 16:30 | 10.78 | 0.107 | 1070 | 1356 | 6.2 | 1362.2 | 863.1 | 863.1094 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4822 | 27.3 | 6.2
2008-10-20 | 16:30 | 10.764 | 0.107 | 1070 | 1356 | 6.2 | 1362.2 | 863.2 | 863.2 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 10.367 | 0.1013 | 1012.9593 | 1356.2 | 6.2 | 1362.4 | 863.3 | 863.3489 | 0.2 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1717 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.744 | 0.0963 | 962.5615 | 1332.6 | 6.2 | 1338.8 | 844.3 | 844.2927 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.16999 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4548 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.719 | 0.094 | 940 | 1246.3 | 6.2 | 1252.5 | 785.8 | 785.84 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.16342 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.4995 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.702 | 0.094 | 940 | 1218 | 6.2 | 1224.2 | 771.7 | 771.7063 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.16114 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.686 | 0.094 | 940 | 1212.3 | 6.2 | 1218.5 | 768.4 | 768.4117 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1607 | 0.0069 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1211.4 | 6.2 | 1217.6 | 767.6 | 767.5827 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1606 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1210.7 | 6.2 | 1216.9 | 767.2 | 767.2498 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2863 | 0.1606 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1210.5 | 6.2 | 1216.7 | 766.9 | 766.8722 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.16059 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1209.5 | 6.2 | 1215.7 | 767.1 | 767.0739 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1605 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1383 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1209.2 | 6.2 | 1215.4 | 766.7 | 766.7092 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1604 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.68 | 0.094 | 940 | 1209.1 | 6.2 | 1215.3 | 767 | 767.0427 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1604 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.675 | 0.094 | 940 | 1209.2 | 6.2 | 1215.4 | 766.9 | 766.9289 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1604 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.67 | 0.094 | 940 | 1209.3 | 6.2 | 1215.5 | 767 | 767.015 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1604 | 0.0068 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6211 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.67 | 0.094 | 940 | 1209 | 6.2 | 1215.2 | 766.9 | 766.9253 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.1604 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6982 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.671 | 0.094 | 940 | 1209.3 | 6.2 | 1215.5 | 766.7 | 766.7251 | 0.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1603 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4744 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.679 | 0.094 | 940 | 1209.4 | 6.2 | 1215.6 | 766.7 | 766.7181 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1603 | 0.0067 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1003.138 | 10.5583 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.666 | 0.094 | 940 | 1208.9 | 6.2 | 1215.1 | 766.9 | 766.8808 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1603 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6697 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.66 | 0.094 | 940 | 1208.9 | 6.2 | 1215.1 | 767.1 | 767.0788 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.16029 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.656 | 0.094 | 940 | 1209.1 | 6.2 | 1215.3 | 766.6 | 766.5549 | 0.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2861 | 0.1602 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1003.8614 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.65 | 0.094 | 940 | 1209.1 | 6.2 | 1215.3 | 766.7 | 766.7 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1602 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.65 | 0.0944 | 943.8192 | 1208.9 | 6.2 | 1215.1 | 766.6 | 766.6124 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1602 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.639 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.65 | 0.0948 | 947.9384 | 1208.9 | 6.2 | 1215.1 | 766.9 | 766.8514 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1602 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.65 | 0.094 | 940 | 1209 | 6.2 | 1215.2 | 766.6 | 766.5571 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1602 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7 | 27.3 | 6.2
2008-10-20 | 16:30 | 9.65 | 0.094 | 940 | 1206.4 | 6.1 | 1212.5 | 766.6 | 766.5631 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2141 | 0.16 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6296 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.65 | 0.094 | 940 | 1206.5 | 6.1 | 1212.6 | 766.5 | 766.5136 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.16 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.5524 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.65 | 0.0946 | 945.5887 | 1206.4 | 6.1 | 1212.5 | 766.3 | 766.3264 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.16 | 0.0066 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5248 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.65 | 0.095 | 950 | 1206.2 | 6.1 | 1212.3 | 766.2 | 766.2365 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.16 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.596 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.65 | 0.095 | 950 | 1206.2 | 6.1 | 1212.3 | 766.5 | 766.5094 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.15999 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1040.599 | 1004 | 10.4415 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.65 | 0.0949 | 949.3785 | 1205.9 | 6.1 | 1212 | 766.6 | 766.6446 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2861 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.5692 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.645 | 0.0941 | 941.3075 | 1205.9 | 6.1 | 1212 | 766.3 | 766.3121 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1383 | 10.6664 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.0947 | 947.0184 | 1205.7 | 6.1 | 1211.8 | 766.5 | 766.4628 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.7 | 6.1 | 1211.8 | 766.4 | 766.4 | 0.2 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.5 | 6.1 | 1211.6 | 766.7 | 766.6992 | 0.2 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.5 | 6.1 | 1211.6 | 766.6 | 766.6269 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6422 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.4 | 6.1 | 1211.5 | 766.7 | 766.6808 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1599 | 0.0065 | 0.0004 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6806 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.2 | 6.1 | 1211.3 | 766.7 | 766.7 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.286 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6035 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.3 | 6.0862 | 1211.4 | 766.7 | 766.7 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5263 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.2 | 6 | 1211.2 | 766.8 | 766.8 | 0.2 | 13.1 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6019 | 27.3 | 6
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.1 | 6 | 1211.1 | 766.7 | 766.7372 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6438 | 27.3 | 6
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205.3 | 6 | 1211.3 | 766.6 | 766.5748 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5052 | 27.3 | 6
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205 | 6 | 1211 | 766.4 | 766.4 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5824 | 27.3 | 6
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1205 | 6 | 1211 | 766.7 | 766.6598 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:30 | 9.636 | 0.095 | 950 | 1204.9 | 6.0138 | 1210.9 | 766.6 | 766.6135 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1599 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:30 | 9.632 | 0.095 | 950 | 1205 | 6.1 | 1211.1 | 766.7 | 766.7383 | 0.2 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1204.9 | 6.0862 | 1211 | 766.5 | 766.5444 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.64 | 0.095 | 950 | 1204.9 | 6.1 | 1211 | 766.8 | 766.7611 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8603 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:30 | 9.633 | 0.095 | 950 | 1205 | 6.1 | 1211.1 | 766.8 | 766.8232 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3 | 6.1
2008-10-20 | 16:31 | 9.63 | 0.095 | 950 | 1205 | 6.1 | 1211.1 | 766.7 | 766.7463 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6227 | 27.3 | 6.1
2008-10-20 | 16:31 | 9.63 | 0.095 | 950 | 1204.9 | 6.1 | 1211 | 766.8 | 766.7893 | 0.2 | 13.1 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6998 | 27.3 | 6.1
2008-10-20 | 16:31 | 9.632 | 0.095 | 950 | 1204.8 | 6.1 | 1210.9 | 766.7 | 766.736 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.623 | 27.3 | 6.1
2008-10-20 | 16:31 | 9.64 | 0.0946 | 946.0967 | 1205 | 6.0862 | 1211.1 | 766.6 | 766.5818 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3542 | 6.1
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1204.6 | 6 | 1210.6 | 766.6 | 766.6415 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.6314 | 27.3686 | 6
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1204.3 | 6 | 1210.3 | 766.9 | 766.9 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6914 | 27.3 | 6
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1204.1 | 6 | 1210.1 | 766.5 | 766.4714 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6142 | 27.3 | 6
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1203.9 | 6 | 1209.9 | 766.8 | 766.8098 | 0.2 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1204.1 | 6 | 1210.1 | 766.6 | 766.6434 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5599 | 27.3 | 6
2008-10-20 | 16:31 | 9.64 | 0.094 | 940 | 1204.1 | 6 | 1210.1 | 766.4 | 766.4298 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5173 | 27.3 | 6
2008-10-20 | 16:31 | 9.64 | 0.0922 | 922.0464 | 1204.2 | 6 | 1210.2 | 766.3 | 766.3393 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1598 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.5944 | 27.3 | 6
2008-10-20 | 16:31 | 8.878 | 0.0849 | 849.4431 | 1199 | 6 | 1205 | 756.9 | 756.9358 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.15939 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1377 | 10.8148 | 27.3 | 6
2008-10-20 | 16:31 | 8.55 | 0.0841 | 841.3721 | 1130 | 6 | 1136 | 718 | 717.9961 | 0.2 | 13.1 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.15366 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.7537 | 27.3 | 6
2008-10-20 | 16:31 | 8.539 | 0.084 | 840 | 1082.4 | 6 | 1088.4 | 679.9 | 679.8793 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14956 | 0.0065 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6259 | 27.3 | 6
2008-10-20 | 16:31 | 8.52 | 0.084 | 840 | 1064 | 6 | 1070 | 672.8 | 672.8286 | 0.2 | 13 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14831 | 0.0064 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6968 | 27.3 | 6
2008-10-20 | 16:31 | 8.52 | 0.084 | 840 | 1057.9 | 6 | 1063.9 | 671 | 670.9809 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1474 | 0.0064 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6196 | 27.3 | 6
2008-10-20 | 16:31 | 8.52 | 0.084 | 840 | 1056.5 | 6 | 1062.5 | 670.6 | 670.6266 | 0.2 | 13 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1473 | 0.0064 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6576 | 27.3 | 6
2008-10-20 | 16:31 | 8.499 | 0.084 | 840 | 1056.6 | 6 | 1062.6 | 670.3 | 670.3433 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1473 | 0.0064 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.6653 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1056.4 | 6 | 1062.4 | 670.4 | 670.4463 | 0.2 | 13 | 952 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1472 | 0.0064 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5881 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1056.3 | 6 | 1062.3 | 670.1 | 670.0547 | 0.2 | 13 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1472 | 0.006386 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.511 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.9 | 6 | 1061.9 | 670.2 | 670.1761 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1472 | 0.0063 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4338 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.8 | 6 | 1061.8 | 670.1 | 670.1166 | 0.2 | 13 | 952 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1472 | 0.0063 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4868 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.5 | 6 | 1061.5 | 670 | 670.0291 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1471 | 0.0063 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.6 | 6 | 1061.6 | 669.8 | 669.7891 | 0.2 | 13.1 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1471 | 0.0063 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8614 | 10.6 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.7 | 6 | 1061.7 | 669.3 | 669.26 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1471 | 0.0063 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4504 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.084 | 840 | 1055.5 | 6 | 1061.5 | 669.8 | 669.77 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1471 | 0.0062 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5555 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.0835 | 834.5238 | 1055.6 | 6 | 1061.6 | 669.7 | 669.7187 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.147 | 0.0062 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.1363 | 10.6711 | 27.3 | 6
2008-10-20 | 16:31 | 8.49 | 0.0826 | 826.192 | 1055.4 | 5.9 | 1061.3 | 670 | 670 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.147 | 0.0062 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.49 | 0.082 | 820 | 1055.4 | 5.9 | 1061.3 | 669.7 | 669.7253 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.147 | 0.0062 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.476 | 0.082 | 820 | 1055.5 | 5.9 | 1061.4 | 670.2 | 670.1536 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.147 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.3 | 5.9 | 1061.2 | 669.9 | 669.8519 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.147 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.8603 | 10.562 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.4 | 5.9 | 1061.3 | 670.1 | 670.0645 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14699 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4848 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.6 | 5.9 | 1061.5 | 669.9 | 669.9371 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1469 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.4076 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.5 | 5.9 | 1061.4 | 669.8 | 669.7721 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1469 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.138 | 10.539 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.2 | 5.9 | 1061.1 | 669.9 | 669.8889 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.1469 | 0.0061 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1055.1 | 5.9 | 1061 | 669.7 | 669.7261 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1469 | 0.006086 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5762 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.082 | 820 | 1054.8 | 5.9 | 1060.7 | 669.9 | 669.8561 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.14689 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.0827 | 826.9856 | 1054.8 | 5.9 | 1060.7 | 669.7 | 669.7444 | 0.2 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1468 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.5 | 5.9 | 1060.4 | 670 | 669.9592 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1468 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5554 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.5 | 5.9 | 1060.4 | 669.6 | 669.5636 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.1468 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.3 | 5.9 | 1060.2 | 670 | 669.9974 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1468 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.2 | 5.9 | 1060.1 | 669.9 | 669.9267 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054 | 5.9 | 1059.9 | 670 | 669.9886 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5359 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.2 | 5.9 | 1060.1 | 669.9 | 669.9376 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5826 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054.1 | 5.9 | 1060 | 670 | 670.0375 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6816 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1054 | 5.9 | 1059.9 | 670.1 | 670.0781 | 0.2 | 13.1 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6044 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.9 | 5.9 | 1059.8 | 669.8 | 669.8294 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.862 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.8 | 5.9 | 1059.7 | 670.1 | 670.0589 | 0.2 | 13.1 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040.9393 | 1004 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.7 | 5.9 | 1059.6 | 669.6 | 669.5733 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.5 | 5.9 | 1059.4 | 669.8 | 669.8419 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.6 | 5.9 | 1059.5 | 669.6 | 669.6136 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1467 | 0.006 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.4 | 5.9 | 1059.3 | 669.9 | 669.8738 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4826 | 27.3 | 5.9
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.2 | 5.8 | 1059 | 669.7 | 669.7 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4718 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.47 | 0.083 | 830.1898 | 1053.3 | 5.8 | 1059.1 | 669.9 | 669.9 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.587 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.475 | 0.0839 | 838.8179 | 1053.4 | 5.8 | 1059.2 | 669.8 | 669.7541 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5098 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.476 | 0.0833 | 832.9204 | 1053.2 | 5.8 | 1059 | 670 | 669.9576 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5672 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.2 | 5.8 | 1059 | 669.7 | 669.6744 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5558 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.5 | 5.8 | 1059.3 | 669.9 | 669.857 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5425 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.5 | 5.8 | 1059.3 | 669.7 | 669.7273 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2864 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6966 | 27.3 | 5.8
2008-10-20 | 16:31 | 8.47 | 0.083 | 830 | 1053.5 | 5.8 | 1059.3 | 669.6 | 669.6178 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6241 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1053.4 | 5.8 | 1059.2 | 669.7 | 669.6874 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1414 | 10.6 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1053.2 | 5.8 | 1059 | 669.6 | 669.6127 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5391 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1053.1 | 5.8 | 1058.9 | 669.7 | 669.7 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4154 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1052.8 | 5.8 | 1058.6 | 669.4 | 669.4373 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5698 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1052.9 | 5.8 | 1058.7 | 669.7 | 669.6807 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1052.9 | 5.7861 | 1058.7 | 669.6 | 669.5735 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2139 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1053.1 | 5.7138 | 1058.8 | 669.8 | 669.7736 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1466 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5673 | 27.3 | 5.7
2008-10-20 | 16:32 | 8.47 | 0.083 | 830 | 1053.1 | 5.8 | 1058.9 | 669.7 | 669.7131 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1465 | 0.0059 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.413 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.472 | 0.083 | 830 | 1053.3 | 5.8 | 1059.1 | 669.6 | 669.6482 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1465 | 0.005886 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.4 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.48 | 0.083 | 830 | 1053.4 | 5.7862 | 1059.2 | 669.7 | 669.7 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1465 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4958 | 27.3 | 5.8
2008-10-20 | 16:32 | 8.48 | 0.083 | 830 | 1053.4 | 5.7 | 1059.1 | 669.6 | 669.5574 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1465 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5498 | 27.3 | 5.7
2008-10-20 | 16:32 | 8.282 | 0.0828 | 828.1583 | 1053.3 | 5.7 | 1059 | 669.8 | 669.7726 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1465 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8586 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.721 | 0.072 | 719.825 | 1047.6 | 5.7 | 1053.3 | 667.9 | 667.9273 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.14598 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.33 | 0.0707 | 706.965 | 1007.7 | 5.7 | 1013.4 | 618.8 | 618.7676 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.14241 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.326 | 0.071 | 710 | 945.4 | 5.7 | 951.1 | 584.9 | 584.8994 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.13656 | 0.0058 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.31 | 0.0714 | 713.8091 | 912.4 | 5.7 | 918.1 | 576.2 | 576.1671 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.13334 | 0.005786 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1377 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.31 | 0.0718 | 717.9163 | 907.9 | 5.6861 | 913.6 | 574.6 | 574.5549 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1329 | 0.0057 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8614 | 10.4 | 27.3 | 5.7
2008-10-20 | 16:32 | 7.303 | 0.071 | 710 | 907.5 | 5.6 | 913.1 | 574.2 | 574.2279 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2863 | 0.1328 | 0.0057 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 5.6
2008-10-20 | 16:32 | 7.3 | 0.071 | 710 | 906.8 | 5.5 | 912.3 | 573.7 | 573.7 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1328 | 0.0057 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.485 | 27.3 | 5.5
2008-10-20 | 16:32 | 7.3 | 0.071 | 710 | 906.7 | 5.5 | 912.2 | 573.8 | 573.8436 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1327 | 0.0057 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5606 | 27.3 | 5.5
2008-10-20 | 16:32 | 7.298 | 0.071 | 710 | 906.7 | 5.5 | 912.2 | 573.7 | 573.7269 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1327 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4063 | 27.3 | 5.5
2008-10-20 | 16:32 | 7.29 | 0.071 | 710 | 906.7 | 5.5 | 912.2 | 573.6 | 573.6424 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1327 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 5.5
2008-10-20 | 16:32 | 7.29 | 0.071 | 710 | 906.7 | 5.5 | 912.2 | 573.6 | 573.6 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.13269 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8603 | 10.5023 | 27.3 | 5.5
2008-10-20 | 16:32 | 7.283 | 0.071 | 710 | 906.7 | 5.4 | 912.1 | 573.5 | 573.5131 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1326 | 0.0056 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5433 | 27.3 | 5.4
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.8 | 5.4 | 912.2 | 573.6 | 573.6 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1326 | 0.0055 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4111 | 27.3 | 5.4
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.7 | 5.4 | 912.1 | 573.5 | 573.4632 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1326 | 0.0055 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5652 | 27.3 | 5.4
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.9 | 5.4 | 912.3 | 573.5 | 573.475 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.13259 | 0.0055 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5.4
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.8 | 5.3 | 912.1 | 573.3 | 573.3 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1325 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.272 | 0.071 | 710 | 906.7 | 5.3 | 912 | 573.3 | 573.3 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1325 | 0.0054 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5863 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.277 | 0.071 | 710 | 906.6 | 5.3 | 911.9 | 573.3 | 573.3 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1325 | 0.005386 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5092 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.7 | 5.3 | 912 | 573.4 | 573.3755 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1324 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4316 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.28 | 0.071 | 710 | 906.7 | 5.3 | 912 | 573.2 | 573.2269 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1324 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1377 | 10.4 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.278 | 0.071 | 710 | 906.4 | 5.3 | 911.7 | 573.3 | 573.2815 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2141 | 0.1324 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 906.5 | 5.3 | 911.8 | 573.3 | 573.3 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.286 | 0.1324 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4003 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 906.4 | 5.2 | 911.6 | 573.1 | 573.1264 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.1324 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5546 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 904.7 | 5.2 | 909.9 | 573.2 | 573.2 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1322 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6545 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 904.6 | 5.2 | 909.8 | 573.2 | 573.1735 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1322 | 0.0053 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6367 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 904.6 | 5.2 | 909.8 | 573.4 | 573.3549 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1322 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 904.6 | 5.2 | 909.8 | 573.3 | 573.3493 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1321 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.071 | 710 | 904.6 | 5.2 | 909.8 | 573.2 | 573.1757 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1321 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4367 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.262 | 0.0712 | 711.7446 | 904.7 | 5.2 | 909.9 | 573.3 | 573.3 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1321 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.26 | 0.072 | 720 | 904.5 | 5.2 | 909.7 | 573.3 | 573.3 | 0.2 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.1321 | 0.0052 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4176 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.26 | 0.072 | 720 | 904.3 | 5.2 | 909.5 | 573.1 | 573.1271 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.1321 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4948 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.264 | 0.0713 | 713.0558 | 904.2 | 5.2 | 909.4 | 572.8 | 572.7565 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.132 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.572 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.0721 | 720.8509 | 904.3 | 5.2 | 909.5 | 573.1 | 573.0527 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.132 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.073 | 730 | 904.5 | 5.2 | 909.7 | 572.9 | 572.9468 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.132 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5474 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.073 | 730 | 904.4 | 5.2 | 909.6 | 573.1 | 573.1 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.132 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4069 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.263 | 0.073 | 730 | 904.3 | 5.2 | 909.5 | 573.1 | 573.1 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.132 | 0.0051 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8614 | 10.5613 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.265 | 0.073 | 730 | 904.6 | 5.2 | 909.8 | 573.1 | 573.0588 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.132 | 0.005086 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4844 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.0723 | 722.9651 | 904.6 | 5.2 | 909.8 | 572.9 | 572.9 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.132 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.47 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.5 | 5.2 | 909.7 | 573.3 | 573.2668 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.132 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5878 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.2 | 5.2 | 909.4 | 572.9 | 572.8881 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.132 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5106 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.2 | 5.2 | 909.4 | 573.1 | 573.1497 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.1319 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5666 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.4 | 5.2 | 909.6 | 573.1 | 573.1127 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.214 | 0.1319 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.5 | 5.2 | 909.7 | 573.1 | 573.1 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1319 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.5 | 5.2 | 909.7 | 573.1 | 573.1255 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2136 | 0.1319 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.5 | 5.2138 | 909.7 | 572.9 | 572.9442 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1319 | 0.005 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.4501 | 27.3 | 5.2
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.8 | 5.3 | 910.1 | 573.1 | 573.1182 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.504 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.8 | 5.3 | 910.1 | 572.9 | 572.8717 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:32 | 7.27 | 0.072 | 720 | 904.8 | 5.3 | 910.1 | 573.2 | 573.1901 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6061 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.7 | 5.3 | 910 | 572.9 | 572.9259 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6838 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.7 | 5.3 | 910 | 573.3 | 573.3374 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5165 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.5 | 5.3 | 909.8 | 573 | 573 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1319 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.4768 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.5 | 5.3 | 909.8 | 573.2 | 573.1686 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5688 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.4 | 5.3 | 909.7 | 572.9 | 572.8525 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.1397 | 10.4145 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.1 | 5.3 | 909.4 | 573 | 573.0397 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3301 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.3 | 5.3 | 909.6 | 573.2 | 573.2269 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4883 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.1 | 5.3 | 909.4 | 573.1 | 573.1127 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6758 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.2 | 5.3 | 909.5 | 573.1 | 573.0876 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.2 | 5.3 | 909.5 | 572.9 | 572.9 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.27 | 0.072 | 720 | 904.3 | 5.3 | 909.6 | 573 | 573.0489 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.1318 | 0.0049 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:33 | 7.277 | 0.072 | 720 | 904.2 | 5.3 | 909.5 | 572.9 | 572.9265 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1318 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.3
2008-10-20 | 16:33 | 6.948 | 0.072 | 720 | 904.2 | 5.3 | 909.5 | 572.7 | 572.7289 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.1318 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5899 | 27.3 | 5.3
2008-10-20 | 16:33 | 6.459 | 0.063 | 629.9655 | 899.2 | 5.3 | 904.5 | 570.8 | 570.8091 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1313 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5126 | 27.3 | 5.3
2008-10-20 | 16:33 | 6.1 | 0.0595 | 595.2571 | 872.8 | 5.2859 | 878.1 | 516.8 | 516.7936 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1286 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5646 | 27.3 | 5.3
2008-10-20 | 16:33 | 6.1 | 0.06 | 600 | 777.2 | 5.2 | 782.4 | 487 | 487.0173 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.286 | 0.11872 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5.2
2008-10-20 | 16:33 | 6.1 | 0.0598 | 598.3264 | 758.3 | 5.2 | 763.5 | 479.6 | 479.5605 | 0.2 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.11685 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5622 | 27.3 | 5.2
2008-10-20 | 16:33 | 6.1 | 0.059 | 590 | 755.9 | 5.2 | 761.1 | 478.1 | 478.0774 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1164 | 0.0048 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4079 | 27.3 | 5.2
2008-10-20 | 16:33 | 6.1 | 0.059 | 590 | 755.3 | 5.2 | 760.5 | 477.4 | 477.4167 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.1163 | 0.0047 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5466 | 27.3 | 5.2
2008-10-20 | 16:33 | 6.1 | 0.059 | 590 | 755.2 | 5.1 | 760.3 | 477.5 | 477.526 | 0.2 | 13 | 953 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.1163 | 0.0047 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.1 | 0.0595 | 595.2744 | 755.1 | 5.1 | 760.2 | 477 | 476.9797 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1163 | 0.0047 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5448 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.1 | 0.06 | 600 | 755 | 5.1 | 760.1 | 477.4 | 477.3774 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1162 | 0.0047 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.083 | 0.06 | 600 | 754.4 | 5.1 | 759.5 | 477.2 | 477.2 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.1161 | 0.0046 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754.3 | 5.1 | 759.4 | 477.2 | 477.2 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1161 | 0.0046 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5183 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754.2 | 5.0859 | 759.3 | 476.9 | 476.9427 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1161 | 0.0046 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8586 | 10.5274 | 27.3 | 5.1
2008-10-20 | 16:33 | 6.083 | 0.06 | 600 | 754.4 | 5 | 759.4 | 477.1 | 477.0746 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1161 | 0.0046 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4267 | 27.3 | 5
2008-10-20 | 16:33 | 6.09 | 0.06 | 600 | 754.3 | 5 | 759.3 | 477.1 | 477.1 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.116 | 0.004586 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5807 | 27.3 | 5
2008-10-20 | 16:33 | 6.09 | 0.06 | 600 | 754.2 | 5 | 759.2 | 477 | 477.0228 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.116 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.081 | 0.06 | 600 | 754.3 | 5 | 759.3 | 476.8 | 476.8056 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.116 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6445 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754.1 | 5 | 759.1 | 476.9 | 476.9035 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.11599 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6349 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754 | 5 | 759 | 476.9 | 476.875 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1159 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4016 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754 | 5 | 759 | 476.6 | 476.6252 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1159 | 0.0045 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5532 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754.1 | 5 | 759.1 | 477.2 | 477.1869 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1159 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 754 | 5 | 759 | 476.9 | 476.9478 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1159 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.08 | 0.06 | 600 | 753.9 | 5 | 758.9 | 477 | 477 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1159 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.059 | 0.06 | 600 | 753.8 | 5 | 758.8 | 476.8 | 476.7603 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.286 | 0.1158 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 753.5 | 5 | 758.5 | 477 | 476.9739 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1158 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 753 | 5 | 758 | 476.9 | 476.8959 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1158 | 0.0044 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 5
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 752.3 | 4.9862 | 757.3 | 476.7 | 476.7472 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.11569 | 0.004386 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5831 | 27.3 | 5
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 752 | 4.9 | 756.9 | 476.9 | 476.8625 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.1156 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5059 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 751.9 | 4.9 | 756.8 | 476.8 | 476.81 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2137 | 0.1156 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5713 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.05 | 0.06 | 600 | 752.1 | 4.9 | 757 | 477 | 476.9623 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1156 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5031 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.055 | 0.0592 | 591.7152 | 752.1 | 4.9 | 757 | 476.8 | 476.8243 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.1156 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4512 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.07 | 0.059 | 590 | 752.1 | 4.9 | 757 | 477 | 476.9787 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.1156 | 0.0043 | 0.0003 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.09 | 0.059 | 590 | 752.1 | 4.9 | 757 | 476.8 | 476.7808 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1156 | 0.0043 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.09 | 0.059 | 590 | 752.1 | 4.9 | 757 | 477.1 | 477.1269 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.11559 | 0.0043 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5428 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.09 | 0.059 | 590 | 752 | 4.9 | 756.9 | 477 | 477.0398 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.286 | 0.1155 | 0.0043 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5687 | 27.3 | 4.9
2008-10-20 | 16:33 | 6.09 | 0.059 | 590 | 752.1 | 4.8 | 756.9 | 476.9 | 476.9451 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0043 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6885 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.085 | 0.059 | 590 | 752 | 4.8 | 756.8 | 477 | 477.0207 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6113 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.08 | 0.059 | 590 | 752 | 4.8 | 756.8 | 477.1 | 477.1 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4681 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.08 | 0.059 | 590 | 751.9 | 4.8 | 756.7 | 477 | 477.0259 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.08 | 0.059 | 590 | 752 | 4.8 | 756.8 | 477 | 477.0237 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4608 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.08 | 0.059 | 590 | 752.2 | 4.8 | 757 | 477.2 | 477.1577 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.6921 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.067 | 0.059 | 590 | 752.1 | 4.8 | 756.9 | 477 | 477.0127 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6254 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 751.9 | 4.8 | 756.7 | 477.1 | 477.1219 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 751.9 | 4.8 | 756.7 | 476.9 | 476.9289 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 751.9 | 4.8 | 756.7 | 477.1 | 477.1436 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 751.8 | 4.8 | 756.6 | 476.9 | 476.9374 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 751.9 | 4.7862 | 756.7 | 476.8 | 476.837 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4804 | 27.3 | 4.8
2008-10-20 | 16:33 | 6.06 | 0.059 | 590 | 752 | 4.7 | 756.7 | 476.8 | 476.8076 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 4.7
2008-10-20 | 16:34 | 6.06 | 0.059 | 590 | 752 | 4.7 | 756.7 | 477 | 477.0459 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8614 | 10.4299 | 27.3 | 4.7
2008-10-20 | 16:34 | 6.063 | 0.059 | 590 | 752.1 | 4.7 | 756.8 | 477.1 | 477.1 | 0.2 | 13.1 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5843 | 27.3 | 4.7
2008-10-20 | 16:34 | 6.068 | 0.059 | 590 | 752.1 | 4.6862 | 756.8 | 476.7 | 476.7003 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1155 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3921 | 27.3 | 4.7
2008-10-20 | 16:34 | 6.06 | 0.059 | 590 | 752.1 | 4.6 | 756.7 | 477 | 477.0486 | 0.2 | 13.1 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.11549 | 0.0042 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4394 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.068 | 0.059 | 590 | 752 | 4.6 | 756.6 | 476.8 | 476.8285 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.6 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.07 | 0.059 | 590 | 752 | 4.6 | 756.6 | 477.1 | 477.0777 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.07 | 0.059 | 590 | 752 | 4.6 | 756.6 | 476.8 | 476.7503 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.07 | 0.059 | 590 | 751.8 | 4.6 | 756.4 | 476.8 | 476.7864 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5448 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.068 | 0.059 | 590 | 751.9 | 4.6 | 756.5 | 477 | 476.9611 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5324 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.06 | 0.059 | 590 | 752 | 4.6 | 756.6 | 477 | 477.0479 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5809 | 27.3 | 4.6
2008-10-20 | 16:34 | 6.068 | 0.059 | 590 | 752 | 4.6 | 756.6 | 477.1 | 477.1157 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4265 | 27.3 | 4.6
2008-10-20 | 16:34 | 5.828 | 0.059 | 590 | 751.9 | 4.5861 | 756.5 | 476.9 | 476.9 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.1154 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5918 | 27.3 | 4.6
2008-10-20 | 16:34 | 5.298 | 0.0492 | 492.1184 | 743.6 | 4.5 | 748.1 | 466 | 466.0348 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2137 | 0.11451 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5766 | 27.3 | 4.5
2008-10-20 | 16:34 | 4.9 | 0.0483 | 482.9942 | 702.7 | 4.5 | 707.2 | 427.3 | 427.3419 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.1099 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4366 | 27.3 | 4.5
2008-10-20 | 16:34 | 4.898 | 0.049 | 490 | 664.9 | 4.5 | 669.4 | 391.9 | 391.9312 | 0.2 | 13.1 | 952 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.10594 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.591 | 27.3 | 4.5
2008-10-20 | 16:34 | 4.89 | 0.0487 | 486.8224 | 611 | 4.5 | 615.5 | 384.9 | 384.9223 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.099272 | 0.0041 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.5
2008-10-20 | 16:34 | 4.89 | 0.048 | 480 | 608.6 | 4.4 | 613 | 383.2 | 383.1866 | 0.2 | 13 | 953 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0989 | 0.004 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6498 | 27.3 | 4.4
2008-10-20 | 16:34 | 4.883 | 0.048 | 480 | 606.9 | 4.3 | 611.2 | 382.5 | 382.4764 | 0.2 | 13 | 953 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.0987 | 0.004 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.673 | 27.3 | 4.3
2008-10-20 | 16:34 | 4.88 | 0.048 | 480 | 606.8 | 4.3 | 611.1 | 382.6 | 382.5932 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098686 | 0.004 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.3
2008-10-20 | 16:34 | 4.88 | 0.0473 | 473.4696 | 606.7 | 4.3 | 611 | 382.3 | 382.2702 | 0.2 | 13 | 953 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0986 | 0.004 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.3
2008-10-20 | 16:34 | 4.88 | 0.047 | 470 | 606.6 | 4.3 | 610.9 | 382.4 | 382.3729 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.0986 | 0.0039 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.3
2008-10-20 | 16:34 | 4.88 | 0.047 | 470 | 606 | 4.2 | 610.2 | 382.5 | 382.5 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.098486 | 0.0039 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.88 | 0.0472 | 471.5372 | 605.3 | 4.2 | 609.5 | 382.3 | 382.3292 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0984 | 0.0039 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.873 | 0.048 | 479.8017 | 605.3 | 4.2 | 609.5 | 382.2 | 382.2361 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0983 | 0.003886 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.87 | 0.048 | 480 | 605.3 | 4.2 | 609.5 | 382.1 | 382.1098 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0983 | 0.0038 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.864 | 0.0473 | 473.4486 | 605.2 | 4.2 | 609.4 | 382.2 | 382.2 | 0.2 | 13.1 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0983 | 0.0038 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5119 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.85 | 0.0475 | 475.0599 | 605.1 | 4.1857 | 609.3 | 382.2 | 382.2 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098286 | 0.0038 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8569 | 10.4422 | 27.3 | 4.2
2008-10-20 | 16:34 | 4.85 | 0.0473 | 473.1939 | 604.6 | 4.1 | 608.7 | 382.1 | 382.0958 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.098186 | 0.0038 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5963 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.85 | 0.046 | 460 | 603.7 | 4.1 | 607.8 | 382 | 382.0257 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.0981 | 0.0038 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.1397 | 10.4482 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.85 | 0.046 | 460 | 603.6 | 4.1 | 607.7 | 382.2 | 382.1739 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098086 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8623 | 10.3468 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.85 | 0.0468 | 468.1818 | 603.5 | 4.1 | 607.6 | 382 | 382.0365 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3304 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.85 | 0.047 | 470 | 603.6 | 4.1 | 607.7 | 382.1 | 382.0729 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4151 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.848 | 0.047 | 470 | 603.5 | 4.1 | 607.6 | 382.2 | 382.2 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5696 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.4 | 4.0862 | 607.5 | 381.9 | 381.9374 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.538 | 27.3 | 4.1
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.5 | 4 | 607.5 | 382 | 382 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5392 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.4 | 4 | 607.4 | 382.1 | 382.1 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.845 | 0.047 | 470 | 603.3 | 4 | 607.3 | 382.2 | 382.2247 | 0.2 | 13.1 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.847 | 0.047 | 470 | 603.3 | 4 | 607.3 | 382.1 | 382.072 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.098 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4585 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.3 | 4 | 607.3 | 382.2 | 382.177 | 0.2 | 13.1 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0979 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4958 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.2 | 4 | 607.2 | 382 | 382.0263 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0979 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.2 | 4 | 607.2 | 382.1 | 382.1472 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0979 | 0.0037 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603 | 4 | 607 | 382 | 382.0249 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0979 | 0.003686 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 602.8 | 4 | 606.8 | 382.1 | 382.0515 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2139 | 0.0979 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 602.8 | 4 | 606.8 | 382 | 382.0136 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0979 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 602.8 | 4 | 606.8 | 382.1 | 382.1 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2136 | 0.0979 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5779 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 602.9 | 4 | 606.9 | 382 | 382.0262 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0979 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4236 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603 | 4 | 607 | 381.9 | 381.9052 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0979 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603 | 4 | 607 | 382.1 | 382.0863 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.097886 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 4
2008-10-20 | 16:34 | 4.84 | 0.047 | 470 | 603.1 | 4 | 607.1 | 381.9 | 381.9255 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0978 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4397 | 27.3 | 4
2008-10-20 | 16:34 | 4.843 | 0.047 | 470 | 603.1 | 3.9 | 607 | 382.1 | 382.1218 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0978 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.594 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.85 | 0.047 | 470 | 603.2 | 3.9 | 607.1 | 382 | 382.013 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.0978 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.85 | 0.047 | 470 | 603 | 3.9 | 606.9 | 382.2 | 382.2487 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0978 | 0.0036 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4974 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.834 | 0.047 | 470 | 602.9 | 3.9 | 606.8 | 382 | 381.9693 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2139 | 0.0978 | 0.003586 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4568 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 602.9 | 3.9 | 606.8 | 381.9 | 381.8758 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 602.8 | 3.9 | 606.7 | 382 | 382.0401 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 602.8 | 3.9 | 606.7 | 382 | 382 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4812 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 602.8 | 3.9 | 606.7 | 382.2 | 382.1744 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 603 | 3.9 | 606.9 | 381.9 | 381.9465 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8623 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 603 | 3.9 | 606.9 | 382 | 381.9862 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0978 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:34 | 4.83 | 0.047 | 470 | 603 | 3.8 | 606.8 | 381.9 | 381.9128 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.097786 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.83 | 0.047 | 470 | 603 | 3.9 | 606.9 | 382.1 | 382.1 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:35 | 4.83 | 0.047 | 470 | 603.1 | 3.9 | 607 | 382.1 | 382.0731 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:35 | 4.831 | 0.0475 | 475.1195 | 603.1 | 3.9 | 607 | 382 | 381.9886 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2861 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 603 | 3.9 | 606.9 | 382 | 381.9811 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.1368 | 10.4 | 27.3 | 3.9
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.9 | 3.8 | 606.7 | 382.3 | 382.2844 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.8 | 3.8 | 606.6 | 382.1 | 382.0526 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3681 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.9 | 3.8 | 606.7 | 382.2 | 382.2279 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3091 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.8 | 3.8 | 606.6 | 382.3 | 382.3 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3863 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.8 | 3.8 | 606.6 | 382.2 | 382.2136 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.527 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.84 | 0.048 | 480 | 602.7 | 3.8 | 606.5 | 382 | 382.0402 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2859 | 0.0977 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5187 | 27.3 | 3.8
2008-10-20 | 16:35 | 4.5 | 0.046 | 460.423 | 602.6 | 3.8 | 606.4 | 381.9 | 381.8545 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.097686 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8603 | 10.4357 | 27.3 | 3.8
2008-10-20 | 16:35 | 3.949 | 0.0373 | 372.6337 | 572.5 | 3.7 | 576.2 | 376.3 | 376.3467 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.093972 | 0.0035 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5903 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.636 | 0.0362 | 361.9099 | 515.8 | 3.7 | 519.5 | 323.8 | 323.7943 | 0.2 | 13 | 954 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.086786 | 0.003486 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.6 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.58 | 0.036 | 360 | 466.1 | 3.7 | 469.8 | 295.1 | 295.0694 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.080245 | 0.0034 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5505 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.581 | 0.036 | 360 | 461.5 | 3.7 | 465.2 | 289.9 | 289.9382 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.079572 | 0.0034 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4734 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.59 | 0.036 | 360 | 458.8 | 3.7 | 462.5 | 288.7 | 288.7244 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.079172 | 0.0033 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.598 | 0.036 | 360 | 457.3 | 3.6862 | 461 | 288.2 | 288.1873 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.078986 | 0.0033 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4 | 27.3 | 3.7
2008-10-20 | 16:35 | 3.6 | 0.036 | 360 | 456.9 | 3.6 | 460.5 | 288.2 | 288.2046 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0788 | 0.0032 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.6 | 0.036 | 360 | 456.9 | 3.6 | 460.5 | 288 | 287.9539 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.078786 | 0.003186 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4354 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.6 | 0.0353 | 353.4056 | 456.8 | 3.6 | 460.4 | 288.1 | 288.0746 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0787 | 0.0031 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5126 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.6 | 0.035 | 350 | 456.9 | 3.6 | 460.5 | 287.9 | 287.8722 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0787 | 0.0031 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5897 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.6 | 0.035 | 350 | 456.8 | 3.6 | 460.4 | 287.7 | 287.7149 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0786 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3999 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.6 | 0.035 | 350 | 456 | 3.6 | 459.6 | 287.9 | 287.8945 | 0.2 | 13 | 955 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.078586 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3437 | 27.3 | 3.6
2008-10-20 | 16:35 | 3.606 | 0.035 | 350 | 456 | 3.5 | 459.5 | 287.7 | 287.7115 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0785 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3793 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.9 | 3.5 | 459.4 | 287.8 | 287.8248 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0784 | 0.003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3022 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.607 | 0.035 | 350 | 455.6 | 3.5 | 459.1 | 287.6 | 287.6108 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0784 | 0.0029 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3754 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.6 | 0.035 | 350 | 455.6 | 3.5 | 459.1 | 287.8 | 287.7715 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0784 | 0.0029 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5054 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.6 | 0.035 | 350 | 455.6 | 3.5 | 459.1 | 287.6 | 287.6251 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0783 | 0.0029 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5402 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.608 | 0.035 | 350 | 455.6 | 3.5 | 459.1 | 287.9 | 287.9154 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0783 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4073 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.6 | 3.5 | 459.1 | 287.8 | 287.8287 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0783 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8571 | 10.4844 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.5 | 3.5 | 459 | 287.9 | 287.8871 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0783 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5616 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.5 | 3.5 | 459 | 287.8 | 287.7541 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0783 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5225 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.5 | 3.5 | 459 | 287.7 | 287.7467 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.078286 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.4 | 3.5 | 458.9 | 287.9 | 287.8979 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2137 | 0.0782 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.4 | 3.5 | 458.9 | 287.7 | 287.727 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.0782 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5407 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 455.1 | 3.5 | 458.6 | 287.9 | 287.9261 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2137 | 0.078186 | 0.0028 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.7 | 3.5 | 458.2 | 287.8 | 287.8241 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.0781 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.8 | 3.5 | 458.3 | 288 | 288.0261 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0781 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5961 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.8 | 3.5 | 458.3 | 287.8 | 287.8251 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0781 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4418 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.8 | 3.5 | 458.3 | 287.8 | 287.8 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0781 | 0.002786 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.7 | 3.5 | 458.2 | 287.7 | 287.7277 | 0.2 | 13 | 956 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2859 | 0.0781 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4669 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.3 | 3.5 | 457.8 | 287.9 | 287.8724 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.214 | 0.078 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5787 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.2 | 3.5 | 457.7 | 287.8 | 287.7512 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.078 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4244 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454.2 | 3.5 | 457.7 | 287.8 | 287.7723 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.078 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5299 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.0355 | 354.9525 | 454.1 | 3.5 | 457.6 | 287.8 | 287.7888 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.078 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5157 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.0357 | 356.6693 | 454.2 | 3.5 | 457.7 | 287.8 | 287.8124 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.078 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4387 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.035 | 350 | 454 | 3.5 | 457.5 | 287.9 | 287.9292 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.078 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5931 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.617 | 0.035 | 350 | 454 | 3.5 | 457.5 | 287.8 | 287.8189 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.078 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5262 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.62 | 0.0358 | 358.352 | 454 | 3.5 | 457.5 | 288.1 | 288.0598 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.077986 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.249 | 3.5
2008-10-20 | 16:35 | 3.617 | 0.036 | 360 | 453.8 | 3.5 | 457.3 | 288.1 | 288.1 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4718 | 27.2282 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.7 | 3.5 | 457.2 | 288.5 | 288.4843 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.6 | 3.5 | 457.1 | 288.7 | 288.7 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.286 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.7 | 3.5 | 457.2 | 289.1 | 289.1409 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5183 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.7 | 3.5 | 457.2 | 289.2 | 289.1822 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4908 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.5 | 3.5 | 457 | 289.2 | 289.1742 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3418 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.4 | 3.5 | 456.9 | 289.2 | 289.1893 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.573 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.4 | 3.5 | 456.9 | 289.2 | 289.1727 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.5 | 3.5 | 457 | 289.2 | 289.155 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.0779 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4631 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.2 | 3.5 | 456.7 | 288.9 | 288.9206 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3914 | 27.3 | 3.5
2008-10-20 | 16:35 | 3.61 | 0.036 | 360 | 453.2 | 3.5 | 456.7 | 289.1 | 289.0747 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.7 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.61 | 0.036 | 360 | 453.3 | 3.5 | 456.8 | 288.8 | 288.8247 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4683 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.61 | 0.036 | 360 | 453.2 | 3.5 | 456.7 | 288.9 | 288.8585 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2141 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5088 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.613 | 0.036 | 360 | 453.1 | 3.5 | 456.6 | 288.7 | 288.7273 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5369 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.629 | 0.036 | 360 | 453.1 | 3.5 | 456.6 | 289 | 288.9604 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.63 | 0.036 | 360 | 453 | 3.5 | 456.5 | 288.9 | 288.918 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.617 | 0.036 | 360 | 452.9 | 3.5 | 456.4 | 289.1 | 289.0508 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.625 | 0.036 | 360 | 452.8 | 3.5 | 456.3 | 289.1 | 289.147 | 0.2 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4806 | 27.3 | 3.5
2008-10-20 | 16:36 | 3.64 | 0.036 | 360 | 452.8 | 3.6 | 456.4 | 289.2 | 289.2091 | 0.2 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2136 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 3.64 | 0.036 | 360 | 452.9 | 3.6 | 456.5 | 289.4 | 289.4 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0778 | 0.0027 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8623 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 3.64 | 0.0342 | 341.9199 | 452.9 | 3.6 | 456.5 | 289.2 | 289.2 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0778 | 0.002686 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.723 | 0.0292 | 291.8159 | 451.7 | 3.6 | 455.3 | 288.9 | 288.9408 | 0.2 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.07765 | 0.0026 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.429 | 0.0242 | 241.6067 | 411.2 | 3.6 | 414.8 | 244.3 | 244.2758 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.071971 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8586 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 324.2 | 3.6 | 327.8 | 207.1 | 207.0891 | 0.2 | 13 | 954 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.286 | 0.059307 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 307.6 | 3.6 | 311.2 | 197.5 | 197.5012 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0568 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 306.3 | 3.6 | 309.9 | 195.3 | 195.3232 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.056586 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 305.1 | 3.6 | 308.7 | 194.7 | 194.6795 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0564 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.6 | 308 | 194.3 | 194.3 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2139 | 0.0563 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.8614 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.6 | 308 | 194.2 | 194.1813 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0563 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.6 | 308 | 193.9 | 193.8537 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0562 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4003 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.6 | 308 | 193.9 | 193.8866 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0562 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4746 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.6 | 308 | 193.7 | 193.6621 | 0.2 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0562 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6377 | 27.3 | 3.6
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.5 | 307.9 | 193.8 | 193.7742 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.056186 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4053 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.5 | 307.9 | 193.5 | 193.5378 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.0561 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4753 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.4 | 3.5 | 307.9 | 193.7 | 193.6645 | 0.2 | 13 | 956 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0561 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4477 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.5 | 3.5 | 308 | 193.4 | 193.4391 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0561 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4587 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.5 | 3.5 | 308 | 193.5 | 193.5 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0561 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 304.5 | 3.5 | 308 | 193.4 | 193.3528 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.056086 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.0247 | 246.5116 | 304.4 | 3.5 | 307.9 | 193.3 | 193.3 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.056 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4775 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.026 | 260 | 303.7 | 3.5 | 307.2 | 193.3 | 193.2758 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.055986 | 0.0023 | 0.0002 | 8.6e-05 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5154 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.026 | 260 | 303.7 | 3.5 | 307.2 | 193.1 | 193.0743 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0559 | 0.002286 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.026 | 260 | 303.4 | 3.5 | 306.9 | 193.1 | 193.1282 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.055872 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.7 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.0254 | 253.5947 | 302.8 | 3.5 | 306.3 | 192.9 | 192.9185 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.63 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.025 | 250 | 302.8 | 3.4862 | 306.3 | 193.1 | 193.0731 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5056 | 27.3 | 3.5
2008-10-20 | 16:36 | 2.37 | 0.0253 | 252.9833 | 302.8 | 3.4 | 306.2 | 192.8 | 192.8411 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2141 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.0259 | 258.6403 | 302.7 | 3.4 | 306.1 | 193 | 193.0449 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4049 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.025 | 250.0344 | 302.7 | 3.4 | 306.1 | 192.8 | 192.8392 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6365 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.0234 | 233.9225 | 302.7 | 3.4 | 306.1 | 192.9 | 192.8977 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.023 | 230 | 302.7 | 3.4 | 306.1 | 192.8 | 192.8267 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0557 | 0.0022 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.023 | 230 | 302.7 | 3.4 | 306.1 | 192.8 | 192.7719 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0557 | 0.002186 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.669 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.0233 | 233.2426 | 302.8 | 3.4 | 306.2 | 192.8 | 192.7866 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.055686 | 0.0021 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4373 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 302.9 | 3.4 | 306.3 | 192.7 | 192.7 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.0021 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5942 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.7 | 192.7 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.0021 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.6 | 192.6 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.0021 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6426 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.7 | 192.6854 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2141 | 0.0556 | 0.0021 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4111 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303.1 | 3.4 | 306.5 | 192.6 | 192.613 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.214 | 0.0556 | 0.002086 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.1397 | 10.4 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.7 | 192.6808 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5015 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.6 | 192.6 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.7 | 192.6918 | 0.2 | 13 | 956 | 952 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5949 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303.1 | 3.4 | 306.5 | 192.6 | 192.6 | 0.2 | 13 | 956 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5178 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303 | 3.4 | 306.4 | 192.7 | 192.6717 | 0.2 | 13 | 956 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4406 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303.1 | 3.4 | 306.5 | 192.7 | 192.6921 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.0556 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4366 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.37 | 0.024 | 240 | 303.1 | 3.4 | 306.5 | 192.7 | 192.6748 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.055586 | 0.002 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.362 | 0.024 | 240 | 303.1 | 3.4 | 306.5 | 192.7 | 192.7 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0555 | 0.001986 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.36 | 0.024 | 240 | 303.2 | 3.4 | 306.6 | 192.5 | 192.5409 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2141 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3647 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.36 | 0.024 | 240 | 303.2 | 3.4 | 306.6 | 192.8 | 192.7619 | 0.2 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.214 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.357 | 0.024 | 240 | 303.2 | 3.4 | 306.6 | 192.7 | 192.6705 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3679 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.35 | 0.024 | 240 | 303.3 | 3.4 | 306.7 | 192.9 | 192.8706 | 0.2 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8623 | 10.5995 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.35 | 0.024 | 240 | 303.4 | 3.4 | 306.8 | 192.6 | 192.6361 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.523 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.358 | 0.024 | 240 | 303.5 | 3.4 | 306.9 | 192.8 | 192.7713 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2863 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4458 | 27.3 | 3.4
2008-10-20 | 16:36 | 2.366 | 0.024 | 240 | 303.5 | 3.4 | 306.9 | 192.8 | 192.8177 | 0.2 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.4
2008-10-20 | 16:37 | 2.37 | 0.024 | 240 | 303.6 | 3.4 | 307 | 192.8 | 192.8244 | 0.2 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2138 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.4
2008-10-20 | 16:37 | 2.37 | 0.024 | 240 | 303.6 | 3.5 | 307.1 | 192.8 | 192.8132 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2853 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.369 | 0.024 | 240 | 303.6 | 3.5 | 307.1 | 192.7 | 192.7076 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.36 | 0.024 | 240 | 303.6 | 3.5 | 307.1 | 192.8 | 192.8285 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.0555 | 0.0019 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.3599 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.36 | 0.024 | 240 | 303.6 | 3.5 | 307.1 | 192.7 | 192.7271 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3519 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.366 | 0.024 | 240 | 303.6 | 3.5 | 307.1 | 192.8 | 192.8228 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5836 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.37 | 0.0232 | 231.7672 | 303.6 | 3.5 | 307.1 | 192.7 | 192.7132 | 0.2 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5283 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.37 | 0.023 | 230 | 303.7 | 3.5 | 307.2 | 192.9 | 192.855 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4511 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.37 | 0.023 | 230 | 303.6 | 3.5 | 307.1 | 192.9 | 192.8805 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2861 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.5
2008-10-20 | 16:37 | 2.37 | 0.023 | 230 | 303.7 | 3.5 | 307.2 | 193 | 193.0217 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0555 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8617 | 10.4033 | 27.3 | 3.5
2008-10-20 | 16:37 | 1.697 | 0.0217 | 217.4682 | 301.7 | 3.5 | 305.2 | 192.8 | 192.8 | 0.2 | 13 | 955 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.055183 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4805 | 27.3 | 3.5
2008-10-20 | 16:37 | 1.305 | 0.0133 | 132.7942 | 262.6 | 3.4862 | 266.1 | 169.2 | 169.1838 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.049038 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4423 | 27.3 | 3.5
2008-10-20 | 16:37 | 1.18 | 0.013 | 130 | 181.7 | 3.4 | 185.1 | 117.1 | 117.0545 | 0.2 | 13 | 954 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.035741 | 0.0018 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4698 | 27.3 | 3.4
2008-10-20 | 16:37 | 1.18 | 0.0124 | 123.6154 | 161.1 | 3.4 | 164.5 | 100.9 | 100.8748 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.032086 | 0.0017 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5759 | 27.3 | 3.4
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 160.5 | 3.3859 | 163.9 | 98.5 | 98.4989 | 0.2 | 13 | 954 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2141 | 0.031986 | 0.0017 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4215 | 27.3 | 3.4
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 158.3 | 3.3 | 161.6 | 97.7 | 97.7279 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.031586 | 0.0017 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3.3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 157 | 3.3 | 160.3 | 97.6 | 97.6402 | 0.2 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0313 | 0.0017 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4873 | 27.3 | 3.3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 156.7 | 3.2 | 159.9 | 97.4 | 97.4 | 0.2 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0312 | 0.001686 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5584 | 27.3 | 3.2
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 156.6 | 3.2 | 159.8 | 97.3 | 97.324 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0312 | 0.0016 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4043 | 27.3 | 3.2
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 156.6 | 3.2 | 159.8 | 97.1 | 97.1255 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0311 | 0.001586 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 3.2
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 156.5 | 3.2 | 159.7 | 97.1 | 97.064 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0311 | 0.0015 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5039 | 27.3 | 3.2
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3.1 | 158.5 | 96.9 | 96.9279 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0309 | 0.0015 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5421 | 27.3 | 3.1
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3.0862 | 158.5 | 97.1 | 97.133 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0308 | 0.001486 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.4121 | 27.3 | 3.1
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3 | 158.4 | 96.9 | 96.9389 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0308 | 0.0014 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.5678 | 27.3 | 3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3 | 158.4 | 97 | 96.9875 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0308 | 0.0014 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3 | 158.4 | 96.9 | 96.937 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.0308 | 0.0014 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5617 | 27.3 | 3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3 | 158.4 | 96.8 | 96.8 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.286 | 0.030786 | 0.0014 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4845 | 27.3 | 3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.4 | 3 | 158.4 | 96.9 | 96.8931 | 0.1 | 13.1 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0307 | 0.0014 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4073 | 27.3 | 3
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 155.3 | 3 | 158.3 | 96.8 | 96.8366 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0307 | 0.0013 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3
2008-10-20 | 16:37 | 1.174 | 0.012 | 120 | 154.9 | 3 | 157.9 | 96.9 | 96.8558 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.030586 | 0.0013 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3058 | 27.3 | 3
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 154 | 3 | 157 | 96.7 | 96.7208 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.030486 | 0.0013 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2486 | 27.3 | 3
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.7 | 2.9862 | 156.7 | 96.9 | 96.8708 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0304 | 0.0013 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 3
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.9 | 156.7 | 96.7 | 96.7 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0304 | 0.0013 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.9
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.8864 | 156.7 | 96.9 | 96.8648 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.030386 | 0.001286 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5119 | 27.3 | 2.9
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.8 | 156.6 | 96.7 | 96.7255 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0303 | 0.0012 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.5006 | 27.3 | 2.8
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.8 | 156.6 | 96.6 | 96.6 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.0303 | 0.0012 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 2.8
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.9 | 2.8 | 156.7 | 96.7 | 96.7279 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.0012 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 2.8
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.9 | 2.7 | 156.6 | 96.6 | 96.6132 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.0012 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3647 | 27.3 | 2.7
2008-10-20 | 16:37 | 1.171 | 0.012 | 120 | 153.9 | 2.6862 | 156.6 | 96.8 | 96.7773 | 0.1 | 13.1 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.001186 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.7
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 153.9 | 2.6 | 156.5 | 96.7 | 96.6751 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0303 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.7 | 96.7069 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0303 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.5 | 96.5175 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.8 | 96.7754 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5013 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.5862 | 156.6 | 96.7 | 96.7132 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4885 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.179 | 0.012 | 120 | 154 | 2.5 | 156.5 | 96.8 | 96.7648 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2102 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 154 | 2.5 | 156.5 | 96.6 | 96.6287 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2864 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.1363 | 10.3644 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.178 | 0.012 | 120 | 154 | 2.5 | 156.5 | 96.8 | 96.7848 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2859 | 0.0302 | 0.0011 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8586 | 10.5184 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.8 | 96.8262 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0302 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5276 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.8 | 96.8 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0302 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 97 | 96.9691 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0302 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 96.9 | 96.9114 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0302 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.6 | 156.6 | 97 | 97.0267 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.030186 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.5862 | 156.6 | 96.9 | 96.9 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2138 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.3774 | 27.3 | 2.6
2008-10-20 | 16:37 | 1.18 | 0.012 | 120 | 154 | 2.5 | 156.5 | 97.1 | 97.0752 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3002 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.175 | 0.012 | 120 | 154 | 2.5 | 156.5 | 96.9 | 96.9285 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.377 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.9 | 2.5 | 156.4 | 97.1 | 97.0645 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3458 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.9 | 2.5 | 156.4 | 97 | 97 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8586 | 10.3314 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.9 | 2.5 | 156.4 | 97 | 96.9895 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.214 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.8603 | 10.3914 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.5 | 156.3 | 97.1 | 97.0644 | 0.1 | 13.1 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.3142 | 27.3 | 2.5
2008-10-20 | 16:37 | 1.17 | 0.012 | 120 | 153.8 | 2.5 | 156.3 | 97.1 | 97.0745 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2862 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.363 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 97.1 | 97.089 | 0.1 | 13.1 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 97.1 | 97.0719 | 0.1 | 13 | 956 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0301 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.6 | 2.5 | 156.1 | 97.1 | 97.1113 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.03 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.6 | 2.5 | 156.1 | 97 | 97 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.03 | 0.001 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.6 | 2.5 | 156.1 | 97.2 | 97.1717 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.6 | 2.5 | 156.1 | 96.9 | 96.8507 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2864 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 97 | 97.0451 | 0.1 | 13 | 957 | 952 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 96.8 | 96.8262 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 96.9 | 96.882 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.012 | 120 | 153.7 | 2.5 | 156.2 | 96.8 | 96.8284 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.17 | 0.0119 | 118.5928 | 153.7 | 2.5 | 156.2 | 96.9 | 96.8735 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2861 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 1.113 | 0.011 | 110.2664 | 153.7 | 2.5 | 156.2 | 96.8 | 96.8 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 0.624 | 0.0021 | 21.3406 | 150.8 | 2.5 | 153.3 | 96.7 | 96.7261 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.029475 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 0.07 | 0.0006 | 6.4023 | 102.2 | 2.5 | 104.7 | 53.7 | 53.7384 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2862 | 0.020713 | 0.0009 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 2.5
2008-10-20 | 16:38 | 0.02 | 0.001 | 10 | 44.2 | 2.362 | 46.6 | 14.3 | 14.2517 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.014448 | 0.0009 | 0.000186 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4207 | 27.3 | 2.4
2008-10-20 | 16:38 | 0.02 | 0.0007 | 7.0774 | 12.2 | 1.5 | 13.7 | 3.8 | 3.798 | 0.1 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.003293 | 0.000886 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4976 | 27.3 | 1.5
2008-10-20 | 16:38 | 0.02 | 0 | 0 | 4.3 | 1.4 | 5.7 | 1.2 | 1.1581 | 0.1 | 13 | 955 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.001671 | 0.0008 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3497 | 27.3 | 1.4
2008-10-20 | 16:38 | 0.02 | 0 | 0 | 3.4 | 1.3859 | 4.7 | 0.7 | 0.7 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.001486 | 0.0008 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4048 | 27.3 | 1.4
2008-10-20 | 16:38 | 0.02 | -0.0008 | -7.9409 | 2.9 | 1.3 | 4.2 | 0.3 | 0.2525 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.214 | 0.0013 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4408 | 27.3 | 1.3
2008-10-20 | 16:38 | 0.017 | -0.001 | -10 | 2.7 | 1.1862 | 3.8 | 0.3 | 0.2922 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2862 | 0.0012 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.307 | 27.3 | 1.2
2008-10-20 | 16:38 | 0.01 | -0.0005 | -5.2127 | 2.4 | 0.9758 | 3.4 | 0.1 | 0.1247 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0012 | 0.0007 | 8.6e-05 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3842 | 27.3 | 1
2008-10-20 | 16:38 | 0.01 | 0 | 0 | 2.4 | 0.1 | 2.5 | 0 | 0.0376 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0011 | 0.000686 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 0.1
2008-10-20 | 16:38 | 0.002 | 0 | 0 | 2.4 | 0.1 | 2.5 | 0 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0011 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 0.1
2008-10-20 | 16:38 | 0 | 0 | 0 | 2.4 | 0.1 | 2.5 | 0 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0011 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3842 | 27.3 | 0.1
2008-10-20 | 16:38 | 0.005 | 0 | 0 | 2.5 | 0 | 2.5 | 0 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.001 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.307 | 27.3 | 0
2008-10-20 | 16:38 | 0.01 | -0.0006 | -6.4701 | 2.4 | 0 | 2.4 | -0.2 | 0 | 0.1 | 13 | 955 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.001 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.138 | 10.3701 | 27.3 | 0
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.4 | 0 | 2.4 | -0.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2138 | 0.001 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 0
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.4 | 0 | 2.4 | -0.2 | 0 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.001 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3755 | 27.3 | 0
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.4 | -0.014143 | 2.4 | -0.1 | 0 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2141 | 0.001 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3017 | 27.3 | 0
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.2 | -0.1 | 2.2 | -0.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3789 | 27.3 | -0.1
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.1 | -0.1 | 2.1 | -0.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0009 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3438 | 27.3 | -0.1
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.3 | 0 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | -0.2
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.3 | 0 | 0.1 | 13 | 956 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3106 | 27.3 | -0.2
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0008 | 0.000386 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3878 | 27.3 | -0.2
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.3 | 0 | 0.1 | 13 | 956 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0008 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.335 | 27.3 | -0.2
2008-10-20 | 16:38 | 0.01 | -0.001 | -10 | 2.3 | -0.2138 | 2.3 | -0.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0008 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3422 | 27.3 | -0.2
2008-10-20 | 16:38 | 0.01 | -0.0002 | -2.0721 | 2.3 | -0.3 | 2.3 | -0.4 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0008 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.4 | 27.3 | -0.3
2008-10-20 | 16:38 | 0.01 | 0 | 0 | 2.3 | -0.3 | 2.3 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2864 | 0.0007 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.3
2008-10-20 | 16:38 | 0.01 | 0 | 0 | 2.3 | -0.3 | 2.3 | -0.4 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0007 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.2263 | -0.3
2008-10-20 | 16:38 | 0.01 | 0 | 0 | 2.3 | -0.3 | 2.3 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.286 | 0.0007 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3491 | 27.2509 | -0.3
2008-10-20 | 16:38 | 0.01 | 0 | 0 | 2.3 | -0.3 | 2.3 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0007 | 0.000286 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.2719 | -0.3
2008-10-20 | 16:38 | 0.004 | 0 | 0 | 2.3 | -0.3 | 2.3 | -0.1 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0007 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.2053 | -0.3
2008-10-20 | 16:38 | 0 | -0.0008 | -7.8619 | 2.3 | -0.3138 | 2.3 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0007 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.2825 | -0.3
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.000686 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3596 | 27.2404 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.2 | 0 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0006 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1039.6364 | 1004 | 10.4 | 27.2367 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0006 | 0.000186 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.4 | 27.2863 | -0.4
2008-10-20 | 16:38 | 0.008 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.4 | 27.2092 | -0.4
2008-10-20 | 16:38 | 0.004 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3317 | 27.2683 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3457 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.214 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3771 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.1 | 0 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1039.3627 | 1004 | 10.3 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0006 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0005 | 8.6e-05 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3545 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.4 | -0.4 | 2.4 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.4 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.391 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.4 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 944 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.3 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.862 | 10.3139 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.4 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2864 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3633 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.4 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.4 | 2.5 | -0.4 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3823 | 27.3 | -0.4
2008-10-20 | 16:38 | 0 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3051 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | -0.4 | 0 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.214 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.448 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2138 | 0.0005 | 0 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5469 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 188 | 27 | 5 | 195 | 59.2862 | 0.0005 | -1.4e-05 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3963 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2861 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.1386 | 10.3191 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2138 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3354 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.4 | 0 | 0.1 | 13 | 958 | 952 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.4 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.6 | -0.5 | 2.6 | -0.4 | 0 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2859 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4668 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.7 | -0.5 | 2.7 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4562 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.7 | -0.5 | 2.7 | -0.5 | 0 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0004 | -0.000114 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.8 | -0.5 | 2.8 | -0.3 | 0 | 0.1 | 13 | 957 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2864 | 0.0004 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.8 | -0.5 | 2.8 | -0.5 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2862 | 0.0004 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | -0.5
2008-10-20 | 16:39 | 0 | -0.001 | -10 | 2.7 | -0.5 | 2.7 | -0.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.000386 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3471 | 27.3 | -0.5
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-10-20
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.12
CO2 Span(%) | 11.9
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 1504
NO2 Span(ppm) | 255
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 957.0
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 16:18
Test End Time | 16:39
Test Duration (s) | 1266
Average Ambient Temperature (deg C) | -35
Average Ambient Pressure (mbar) | 988
Average Relative Humidity (%) | 18.27
Average Absolute Humidity (grains/lb dry air) | 0.247
Average Kh Factor | 0.833
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 10/20/2008 16:17:39.370 - None Found
Warnings:
0X0000 - 10/20/2008 16:17:39.370 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 123
Brake Specific Dropout Limit Count | 1266
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-10-20 | 16:18 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
Cerion
light w/generator

## THC Lin
Test Date:
Test Time:
10/20/2008
16:17:39
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
10/20/2008
16:17:39
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
10/20/2008
16:17:39
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
10/20/2008
16:17:39
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
10/20/2008
16:17:39
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
10/20/2008
16:17:39
NDUV Serial No:
208
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
11.99 | Rm = | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555
| C.F.= | 0.098698 | 0.19768 | 0.29695 | 0.39651 | 0.49636 | 0.5965 | 0.69693 | 0.79766 | 0.89868
| gas = | 1.1834 | 2.3702 | 3.5605 | 4.7542 | 5.9514 | 7.1521 | 8.3562 | 9.564 | 10.7752
Ro for CO2= | 0.8795