## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-710C | GDU-002 | 2008-01-01 | 2009-01-01
0 | 0 | 0 | 0
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 50568
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.04 | Scott gas | ALM0315016 | ±1% | 2010-01-29 | 93.12
200 ppm CO | 200.6
6% CO2 | 6
300 ppm NO | 298.1
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300.5
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 50568 |  |  | CO - AMBII Serial Number: |  |  | 50568 |  | HC - AMBII Serial Number: |  |  | 50568 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-05-08 |  | 6 | X | 0.1 | 3 |  |  | 200.6 |  | 50 | 3 |  | 31.04 |  | 4 | 3 |  | 298.1 |  | 15 | 3 |  | 93.12 | 100 | 5 | 2
Test Time: | 4:07
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -10 | -10 |  |  | 0 | 1000.2 | 1000.2 |  |  | 0 | -1.3 | -1.3 |  |  | 0 | 0 | 0
90 | 105 | 1 | 0.6 | 0.596 | 0.6 | 0.004 | 0.67 |  | 20.06 | 24 | 4 | 19.94 |  | 3.104 | 1000.4 | 997.3 | 32129.51 | Fail | 29.81 | 27.2 | -2.6 | -8.72 |  | 9.312 | 9.2 | -0.11 | -1.18
150 | 165 | 2 | 1.2 | 1.193 | 1.169 | -0.024 | -2.01 |  | 40.12 | 50 | 10 | 24.93 |  | 6.208 | 1000.4 | 994.2 | 16014.82 | Fail | 59.62 | 56.8 | -2.8 | -4.7 |  | 18.624 | 18.5 | -0.12 | -0.64
210 | 225 | 3 | 1.8 | 1.791 | 1.755 | -0.036 | -2.01 |  | 60.18 | 70 | 10 | 16.62 |  | 9.312 | 1000.9 | 991.6 | 10648.63 | Fail | 89.43 | 88.3 | -1.1 | -1.23 |  | 27.936 | 27.8 | -0.14 | -0.5
270 | 285 | 4 | 2.4 | 2.39 | 2.34 | -0.05 | -2.09 |  | 80.24 | 90 | 10 | 12.46 |  | 12.416 | 1000.3 | 987.9 | 7956.67 | Fail | 119.24 | 117.6 | -1.6 | -1.34 |  | 37.248 | 37.01 | -0.24 | -0.64
330 | 345 | 5 | 3 | 2.989 | 2.95 | -0.039 | -1.3 |  | 100.3 | 110 | 10 | 9.97 |  | 15.52 | 1000.3 | 984.8 | 6345.36 | Fail | 149.05 | 149.7 | 0.6 | 0.4 |  | 46.56 | 46.3 | -0.26 | -0.56
390 | 405 | 6 | 3.6 | 3.59 | 3.55 | -0.04 | -1.11 |  | 120.36 | 130 | 10 | 8.31 |  | 18.624 | 1001 | 982.4 | 5274.91 | Fail | 178.86 | 179.8 | 0.9 | 0.5 |  | 55.872 | 55.5 | -0.37 | -0.66
450 | 465 | 7 | 4.2 | 4.191 | 4.167 | -0.024 | -0.57 |  | 140.42 | 142 | 2 | 1.42 |  | 21.728 | 1000.8 | 979.1 | 4506.17 | Fail | 208.67 | 209.9 | 1.2 | 0.58 |  | 65.184 | 64.8 | -0.38 | -0.58
510 | 525 | 8 | 4.8 | 4.793 | 4.762 | -0.031 | -0.65 |  | 160.48 | 167 | 7 | 4.36 |  | 24.832 | 1001 | 976.2 | 3931.22 | Fail | 238.48 | 240.3 | 1.8 | 0.75 |  | 74.496 | 74.19 | -0.31 | -0.42
570 | 585 | 9 | 5.4 | 5.396 | 5.381 | -0.015 | -0.28 |  | 180.54 | 181 | 0 | 0 |  | 27.936 | 1001.2 | 973.3 | 3484.03 | Fail | 268.29 | 269.3 | 1 | 0.37 |  | 83.808 | 83.33 | -0.48 | -0.57
630 | 645 | 10 | 6 | 6 | 6 | 0 | 0 |  | 200.6 | 197 | -4 | -1.99 |  | 31.04 | 1001 | 970 | 3125 | Fail | 298.1 | 300.4 | 2.3 | 0.77 |  | 93.12 | 92.72 | -0.4 | -0.43
690 | 705 | 9 | 5.4 | 5.396 | 5.37 | -0.026 | -0.48 |  | 180.54 | 187 | 6 | 3.32 |  | 27.936 | 1001.1 | 973.2 | 3483.68 | Fail | 268.29 | 271.2 | 2.9 | 1.08 |  | 83.808 | 83.49 | -0.32 | -0.38
750 | 765 | 8 | 4.8 | 4.793 | 4.77 | -0.023 | -0.48 |  | 160.48 | 166 | 6 | 3.74 |  | 24.832 | 1001.5 | 976.7 | 3933.23 | Fail | 238.48 | 242.3 | 3.8 | 1.59 |  | 74.496 | 74.16 | -0.34 | -0.46
810 | 825 | 7 | 4.2 | 4.191 | 4.16 | -0.031 | -0.74 |  | 140.42 | 149 | 9 | 6.41 |  | 21.728 | 1000.8 | 979.1 | 4506.17 | Fail | 208.67 | 211.7 | 3 | 1.44 |  | 65.184 | 64.83 | -0.35 | -0.54
870 | 885 | 6 | 3.6 | 3.59 | 3.56 | -0.03 | -0.84 |  | 120.36 | 122 | 2 | 1.66 |  | 18.624 | 1001.3 | 982.7 | 5276.52 | Fail | 178.86 | 180.9 | 2 | 1.12 |  | 55.872 | 55.31 | -0.56 | -1
930 | 945 | 5 | 3 | 2.989 | 2.95 | -0.039 | -1.3 |  | 100.3 | 106 | 6 | 5.98 |  | 15.52 | 1001.1 | 985.6 | 6350.52 | Fail | 149.05 | 150.8 | 1.8 | 1.21 |  | 46.56 | 45.96 | -0.6 | -1.29
990 | 1005 | 4 | 2.4 | 2.39 | 2.345 | -0.045 | -1.88 |  | 80.24 | 90 | 10 | 12.46 |  | 12.416 | 1000.9 | 988.5 | 7961.5 | Fail | 119.24 | 121.8 | 2.6 | 2.18 |  | 37.248 | 36.7 | -0.55 | -1.48
1050 | 1065 | 3 | 1.8 | 1.791 | 1.759 | -0.032 | -1.79 |  | 60.18 | 70 | 10 | 16.62 |  | 9.312 | 1001 | 991.7 | 10649.7 | Fail | 89.43 | 91.9 | 2.5 | 2.8 |  | 27.936 | 27.49 | -0.45 | -1.61
1110 | 1125 | 2 | 1.2 | 1.193 | 1.172 | -0.021 | -1.76 |  | 40.12 | 48 | 8 | 19.94 |  | 6.208 | 1000.9 | 994.7 | 16022.87 | Fail | 59.62 | 62.6 | 3 | 5.03 |  | 18.624 | 18.1 | -0.52 | -2.79
1170 | 1185 | 1 | 0.6 | 0.596 | 0.595 | -0.001 | -0.17 |  | 20.06 | 20 | 0 | 0 |  | 3.104 | 1001 | 997.9 | 32148.84 | Fail | 29.81 | 33.5 | 3.7 | 12.41 |  | 9.312 | 8.89 | -0.42 | -4.51
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -10 | -10 |  |  | 0 | 1000.7 | 1000.7 |  |  | 0 | 2.5 | 2.5 |  |  | 0 | -0.5 | -0.5
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.017 | 0.06 |  |  |  | Intercept | 3 | 2 | Fail |  | Intercept | 1001 | 0 | Fail |  | Intercept | 0.1 | 3 |  |  | Intercept | -0.29 | 0.93
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 1.02 | 0.99 - 1.01 | Fail |  | Slope | 0.02 | 0.99 - 1.01 | Fail |  | Slope | 1.01 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0161 | 0.06 |  |  |  | SEE | 6.2671 | 2.006 | Fail |  | SEE | 0.2951 | 0.3104 |  |  | SEE | 1.9116 | 2.981 |  |  | SEE | 0.1614 | 0.9312
|  |  | r2 | 0.9999 | 0.998 |  |  |  | r2 | 0.9909 | 0.998 | Fail |  | r2 | 0.3255 | 0.998 | Fail |  | r2 | 0.9996 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-08 | 4:07 | 0.01 | 0 | 0 | -0.3 | 0.9 | 0.9 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2267 | 0.0011 | 0.001127 | 0 | 0 |  | n/a |  | 30 | 53 | 1035.5764 | 1000 | 29.5 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -0.3 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 13 | 951 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -0.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5321 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -0.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.6 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -0.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.6 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -0.8 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3189 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -1 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.000827 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2996 | 25.4498 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -1.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190.5 | 25 | 5 | 195 | 59.2734 | 0.0008 | 0.001027 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3706 | 25.5 | 0.9
2008-05-08 | 4:08 | 0.01 | 0 | 0 | -1.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2266 | 0.000727 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2911 | 25.4911 | 0.9
2008-05-08 | 4:08 | 0.006 | 0 | 0 | -1.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2117 | 25.4117 | 0.9
2008-05-08 | 4:08 | 0 | -0.0005 | -5.2296 | -1.4 | 0.9 | 0.9 | 0.1 | 0.1 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.3 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2678 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.9 | 0.9 | 0.1 | 0.072 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2529 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2267 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2267 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2061 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.005 | -0.001 | -10 | -1.3 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2855 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.01 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2734 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2349 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.008 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.274 | 0.0007 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1037 | 25.4241 | 0.9
2008-05-08 | 4:08 | 0.008 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.000927 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8105 | 25.5 | 0.9
2008-05-08 | 4:08 | 0.003 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.049 | 25.5 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0007 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.1621 | 25.4379 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2266 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1169 | 25.4415 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9788 | 25.5 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9032 | 25.4992 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2211 | 25.4197 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4196 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4219 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.004 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3 | 25.4 | 0.9
2008-05-08 | 4:08 | 0.008 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3 | 25.4 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.9 | 0.9 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9896 | 25.4776 | 0.9
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8266 | 0.8 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0094 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2636 | 25.4841 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5816 | 25.4046 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0005 | 25.4749 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.4 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.000827 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8334 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8733 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.3 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7145 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.2 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.000526 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1332 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.2 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 953 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.145 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.2 | 0.8 | 0.8 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29 | 25.5 | 0.8
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.2 | 0.7 | 0.7 | 0 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.99 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.1 | 0.7 | 0.7 | 0 | 0 | 0 | 12.9 | 953 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.000727 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.91 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1.1 | 0.7 | 0.7 | 0 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8304 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1 | 0.7 | 0.7 | -0.1 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | -0.001 | -10 | -1 | 0.7 | 0.7 | -0.1 | 0 | 0 | 12.9 | 952 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | -0.0006 | -6.4419 | -1 | 0.7 | 0.7 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8079 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | 0 | 0 | -1 | 0.7 | 0.7 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8874 | 25.5 | 0.7
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.6266 | 0.6 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7664 | 25.5 | 0.6
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.6 | 0.6 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8381 | 25.5 | 0.6
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.6 | 0.6 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0508 | 25.5 | 0.6
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.6 | 0.6 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1897 | 25.5 | 0.6
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.6 | 0.6 | -0.1 | 0 | 0 | 12.9 | 953 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0305 | 25.5 | 0.6
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.1 | 0.5266 | 0.5 | 0 | 0 | 0 | 12.9 | 952 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8716 | 25.5 | 0.5
2008-05-08 | 4:08 | 0 | 0 | -0.041632 | -1.2 | 0.5 | 0.5 | 0 | 0 | 0 | 12.9 | 952 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8873 | 25.5 | 0.5
2008-05-08 | 4:08 | 0 | -0.0008 | -8.368 | -1.2 | 0.5 | 0.5 | 0 | 0 | 0 | 12.9 | 953 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9538 | 25.5 | 0.5
2008-05-08 | 4:08 | 0 | -0.0003 | -3.3 | -1.3 | 0.5 | 0.5 | 0 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2267 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8025 | 25.5 | 0.5
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.3 | 0.5 | 0.5 | 0 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.882 | 25.5 | 0.5
2008-05-08 | 4:08 | 0 | 0 | 0 | -1.3 | 0.5 | 0.5 | 0 | 0 | 0 | 12.9 | 953 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8386 | 25.5 | 0.5
2008-05-08 | 4:08 | 0.059 | 0.0014 | 14.4288 | -1.3 | 0.5 | 0.5 | -0.1 | 0 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7591 | 25.5 | 0.5
2008-05-08 | 4:09 | 0.397 | 0.0068 | 68.0177 | -0.3 | 0.42675 | 0.4 | 0.1 | 0.084 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2267 | 0.00062 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7204 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.57 | 0.002 | 19.6696 | 5.7 | 0.4 | 6.1 | 2.8 | 2.8232 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.001874 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7999 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.58 | 0.0017 | 16.6667 | 22.3 | 0.4 | 22.7 | 6.7 | 6.6723 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.00514 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.58 | 0.002 | 20 | 25.7 | 0.4 | 26.1 | 8.3 | 8.3172 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188 | 25 | 5 | 195 | 59.2 | 0.0058 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8588 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.58 | 0.002 | 20 | 26 | 0.4 | 26.4 | 8.9 | 8.9381 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8618 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.588 | 0.002 | 20 | 26.3 | 0.4 | 26.7 | 9.1 | 9.1268 | 0 | 12.9 | 954 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5176 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.3 | 0.4 | 26.7 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5971 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.3 | 0.4 | 26.7 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5234 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.4 | 0.4 | 26.8 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.4 | 0.4 | 26.8 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0059 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8711 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2734 | 0.005973 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9849 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.000674 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9057 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.3 | 0 | 12.9 | 954 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8261 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.3 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.6396 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.3 | 0 | 12.9 | 954 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.5987 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.3 | 9.2537 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.3 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.5 | 0.4 | 26.9 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.6 | 0.4 | 27 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.6 | 0.4 | 27 | 9.3 | 9.3 | 0 | 12.9 | 954 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7335 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.6 | 0.4 | 27 | 9.3 | 9.3 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7699 | 25.5301 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.6 | 0.4 | 27 | 9.3 | 9.269 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.7094 | 25.5906 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.7 | 0.4 | 27.1 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7892 | 25.5108 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.8 | 0.4 | 27.2 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9373 | 25.5686 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.9 | 0.4 | 27.3 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9038 | 25.5519 | 0.4
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 26.9 | 0.4 | 27.3 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8276 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.595 | 0.002 | 20 | 26.9 | 0.4 | 27.3 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006073 | 0.000673 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.907 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27 | 0.4 | 27.4 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.9861 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27 | 0.4 | 27.4 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.006027 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.0658 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.1 | 0.4 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9183 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.1 | 0.4 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2266 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7748 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.1 | 0.4 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2267 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.0028 | 28.2931 | 27.1 | 0.4 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9369 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 189.5 | 25 | 5 | 195 | 59.2266 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.9851 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.3 | 0.4 | 27.7 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0556 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.3 | 0.4 | 27.7 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7325 | 28.8986 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2267 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 28.8189 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.0022 | 21.7152 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7599 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7196 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2266 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.799 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.644 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.0061 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7154 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 953 | 962 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006027 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.0027 | 26.5945 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8169 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.4 | 27.6 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8963 | 25.5 | 0.4
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.3266 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8242 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8549 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2267 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8657 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.6 | 0.003 | 30 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8286 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.596 | 0.003 | 30 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9874 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.003 | 30 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7325 | 28.7804 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.003 | 30 | 27.1 | 0.3 | 27.4 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7527 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.0028 | 28.418 | 27.1 | 0.3 | 27.4 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.7358 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.002 | 20.0916 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 190.5 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.6232 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7822 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.59 | 0.002 | 20 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.592 | 0.002 | 20 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.75 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.6 | 0.002 | 20 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 954 | 962 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2266 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7295 | 25.5 | 0.3
2008-05-08 | 4:09 | 0.83 | 0.003 | 30.2367 | 27.2 | 0.3 | 27.5 | 9.2 | 9.2 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.006 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 28.8 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.06 | 0.0079 | 79.2163 | 34.3 | 0.3 | 34.6 | 9.5 | 9.4945 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.006834 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.15 | 0.0055 | 55.4 | 42.3 | 0.3 | 42.6 | 13 | 13.0075 | 0 | 12.9 | 954 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.008999 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0716 | 25.4321 | 0.3
2008-05-08 | 4:10 | 1.15 | 0.005 | 50 | 51.9 | 0.3 | 52.2 | 16.3 | 16.3172 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.010861 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0581 | 25.4473 | 0.3
2008-05-08 | 4:10 | 1.15 | 0.005 | 50 | 54.6 | 0.3 | 54.9 | 17.7 | 17.7139 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.011347 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.15 | 0.005 | 50 | 55.2 | 0.3 | 55.5 | 18.2 | 18.2302 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.3 | 0.0115 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.15 | 0.005 | 50 | 55.3 | 0.3 | 55.6 | 18.4 | 18.351 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.0115 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.157 | 0.005 | 50 | 55.6 | 0.3 | 55.9 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8348 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.7 | 0.3 | 56 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.3 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8896 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.7 | 0.3 | 56 | 18.4 | 18.4 | 0 | 13 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9515 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.7 | 0.3 | 56 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 28.7963 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.8 | 0.3 | 56.1 | 18.4 | 18.4 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191 | 25 | 5 | 195 | 59.2266 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 28.7168 | 25.5 | 0.3
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.8 | 0.2266 | 56 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 191.5 | 25 | 5 | 195 | 59.2734 | 0.0116 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 28.8247 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.9 | 0.2 | 56.1 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 28.9 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.9 | 0.2 | 56.1 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 28.9216 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 55.9 | 0.2 | 56.1 | 18.4 | 18.4 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 29 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 56.3 | 0.2 | 56.5 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.3 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 56.3 | 0.2 | 56.5 | 18.5 | 18.4506 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0601 | 25.4399 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 56.3 | 0.2 | 56.5 | 18.4 | 18.4265 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0604 | 25.4396 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 56.3 | 0.2 | 56.5 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0117 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.962 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.165 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.803 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.1267 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9553 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2266 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1145 | 25.4427 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1631 | 25.4369 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.7341 | 29.1 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.266 | 29.1 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27.515 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3261 | 25.4246 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.4 | 0.2 | 56.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1258 | 25.4548 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.5 | 0.2 | 56.7 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.9343 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.5 | 0.2 | 56.7 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.6 | 0.2 | 56.8 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 29 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.6 | 0.2 | 56.8 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27.485 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 29.0727 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.6 | 0.2 | 56.8 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 999 | 29.1 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.5 | 0.2 | 56.7 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27.5145 | 192 | 25 | 5 | 195 | 59.2266 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.5 | 0.2 | 56.7 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0777 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.6 | 0.2 | 56.8 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.9186 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.7 | 0.2 | 56.9 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27.485 | 192 | 25 | 5 | 195 | 59.2734 | 0.0118 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1097 | 25.4301 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 56.9 | 0.2 | 57.1 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.011873 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1011 | 25.4494 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0291 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 27.515 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0914 | 25.4914 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2732 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0119 | 25.4119 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29 | 25.4675 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0468 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.164 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0476 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.16 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.163 | 0.005 | 50 | 57 | 0.2 | 57.2 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 946 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9 | 25.5 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57.1 | 0.2 | 57.3 | 18.5 | 18.5 | 0 | 12.9 | 954 | 961 | 945 | 19 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0119 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1598 | 25.4351 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57.2 | 0.2 | 57.4 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.011973 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1223 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.17 | 0.005 | 50 | 57.2 | 0.2 | 57.4 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9718 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.177 | 0.005 | 50 | 57.3 | 0.2 | 57.5 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1863 | 25.4034 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.0042 | 41.8948 | 57.3 | 0.2 | 57.5 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8686 | 25.4828 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.004 | 40 | 57.3 | 0.2 | 57.5 | 18.5 | 18.5 | 0 | 13 | 953 | 960 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1096 | 25.4381 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.004 | 40 | 57.3 | 0.2 | 57.5 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1756 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.0043 | 43.1376 | 57.3 | 0.2 | 57.5 | 18.5 | 18.5 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.734 | 29.1283 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.4 | 0.2 | 57.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.266 | 29.5939 | 25.4009 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.4 | 0.2 | 57.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0374 | 25.4804 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.4 | 0.2 | 57.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0794 | 25.4402 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.4 | 0.2 | 57.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 29.0823 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.4 | 0.2 | 57.6 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9561 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.005 | 50 | 57.5 | 0.2 | 57.7 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 19.5 | 0.29 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1947 | 25.4 | 0.2
2008-05-08 | 4:10 | 1.18 | 0.0057 | 57.0774 | 57.5 | 0.2 | 57.7 | 18.5 | 18.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.012 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0448 | 25.4776 | 0.2
2008-05-08 | 4:11 | 1.52 | 0.0099 | 98.7094 | 58.6 | 0.2 | 58.8 | 18.6 | 18.5514 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.01222 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9428 | 25.5 | 0.2
2008-05-08 | 4:11 | 1.72 | 0.0075 | 75.3435 | 65 | 0.2 | 65.2 | 20.7 | 20.7486 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0135 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9 | 25.4633 | 0.2
2008-05-08 | 4:11 | 1.733 | 0.0064 | 63.5405 | 80.2 | 0.2 | 80.4 | 24.8 | 24.8196 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2 | 0.016361 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9486 | 25.4 | 0.2
2008-05-08 | 4:11 | 1.74 | 0.006 | 60 | 84.8 | 0.2 | 85 | 26.8 | 26.7636 | 0 | 13 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.017173 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1871 | 25.4 | 0.2
2008-05-08 | 4:11 | 1.74 | 0.006 | 60 | 86 | 0.2 | 86.2 | 27.4 | 27.3603 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.017447 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0496 | 25.4752 | 0.2
2008-05-08 | 4:11 | 1.74 | 0.006 | 60 | 86.6 | 0.2 | 86.8 | 27.6 | 27.6389 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.017573 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0544 | 25.4456 | 0.2
2008-05-08 | 4:11 | 1.74 | 0.006 | 60 | 87 | 0.2 | 87.2 | 27.7 | 27.7 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0176 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.1339 | 25.4 | 0.2
2008-05-08 | 4:11 | 1.74 | 0.006 | 60 | 87.1 | 0.2 | 87.3 | 27.7 | 27.7302 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0176 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.2
2008-05-08 | 4:11 | 1.743 | 0.006 | 60 | 87.5 | 0.2734 | 87.7 | 27.8 | 27.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0177 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 87.7 | 0.3 | 88 | 27.8 | 27.8 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.017773 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0436 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9633 | 25.4316 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.1111 | 25.4889 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1906 | 25.4094 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0599 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1487 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1545 | 25.4291 | 0.3
2008-05-08 | 4:11 | 1.758 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8429 | 25.4914 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.006 | 60 | 88 | 0.3 | 88.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.24 | 25.412 | 0.3
2008-05-08 | 4:11 | 1.755 | 0.006 | 60 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1649 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.006 | 60 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0178 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1469 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.0065 | 64.7575 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2264 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3121 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4711 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3049 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0666 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.1 | 0.3 | 88.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.0196 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7325 | 29.4068 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5658 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.6 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4336 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2215 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3019 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4608 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3204 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2781 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.756 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4182 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4983 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 88.2 | 0.3 | 88.5 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2669 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 88.3 | 0.3 | 88.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 88.3 | 0.3 | 88.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0179 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 88.6 | 0.3 | 88.9 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.017973 | 0.000627 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1836 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89 | 0.3 | 89.3 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.018 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.1041 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.1 | 0.3 | 89.4 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.018 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3259 | 25.4 | 0.3
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.2 | 0.3734 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.018 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4546 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3638 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1138 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1933 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1272 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.76 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1523 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.757 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2318 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.3 | 0.4 | 89.7 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2888 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2093 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2292 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.3 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.007 | 70 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2323 | 25.4 | 0.4
2008-05-08 | 4:11 | 1.75 | 0.0076 | 76.3666 | 89.2 | 0.4 | 89.6 | 27.8 | 27.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0181 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2471 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.091 | 0.0109 | 108.5209 | 89.6 | 0.4 | 90 | 27.8 | 27.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2734 | 0.018173 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.2734 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.308 | 0.0086 | 86.1699 | 98.8 | 0.4 | 99.2 | 29.7 | 29.7264 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.019833 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 29.2124 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.008 | 80 | 108.6 | 0.4 | 109 | 33.9 | 33.8575 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.02172 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3711 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.342 | 0.0085 | 84.6864 | 114.4 | 0.4 | 114.8 | 35.8 | 35.8143 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.022773 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3349 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 115.9 | 0.4 | 116.3 | 36.7 | 36.6625 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2267 | 0.023073 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.5 | 0.4 | 116.9 | 36.9 | 36.9 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2266 | 0.0231 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3241 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.6 | 0.4 | 117 | 36.9 | 36.9397 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.023173 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4073 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.6 | 0.4 | 117 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5663 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.6 | 0.4 | 117 | 37 | 37.0352 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.351 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.6 | 0.4 | 117 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.7 | 0.4 | 117.1 | 37 | 37.0294 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.7 | 0.4 | 117.1 | 37.1 | 37.053 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2011 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.8 | 0.4 | 117.2 | 37.1 | 37.1 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2805 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.8 | 0.4 | 117.2 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 116.9 | 0.4 | 117.3 | 37 | 37 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3791 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.4 | 117.4 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0232 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4428 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.4 | 117.4 | 37 | 37 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2043 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.4 | 117.4 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2779 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.3 | 117.3 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.415 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.3 | 117.3 | 37 | 37.0352 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.537 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.35 | 0.009 | 90 | 117 | 0.3 | 117.3 | 37.1 | 37.1 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.6 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.349 | 0.009 | 90 | 117.1 | 0.3 | 117.4 | 37.1 | 37.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.6 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.1 | 0.3 | 117.4 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.5245 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.1 | 0.3 | 117.4 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3903 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.2 | 0.3 | 117.5 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2657 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.1 | 0.3 | 117.4 | 37.1 | 37.0631 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2562 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.2 | 0.3 | 117.5 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0233 | 0.000773 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5739 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.2 | 0.3 | 117.5 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0233 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.527 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.2 | 0.3 | 117.5 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.2 | 0.3 | 117.5 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0233 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.436 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.4 | 0.3 | 117.7 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0234 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3114 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.4 | 0.3 | 117.7 | 37 | 37.0146 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.3909 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.4 | 0.3 | 117.7 | 37.1 | 37.1 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2601 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.5 | 0.3 | 117.8 | 37 | 37.0053 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0234 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3984 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.5 | 0.3 | 117.8 | 37 | 37 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0234 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5705 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.5 | 0.3 | 117.8 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2264 | 0.0234 | 0.000726 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.491 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.6 | 0.3 | 117.9 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4115 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.7 | 0.3 | 118 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.468 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.7 | 0.3 | 118 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4051 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.7 | 0.3 | 118 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3 | 25.4 | 0.3
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.8 | 0.4 | 118.2 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2936 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.8 | 0.4 | 118.2 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.2144 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.8 | 0.4 | 118.2 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.396 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.8 | 0.4 | 118.2 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2263 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.365 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.8 | 0.4 | 118.2 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2979 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.9 | 0.4 | 118.3 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5921 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 117.9 | 0.4 | 118.3 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4331 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118 | 0.4 | 118.4 | 37 | 37 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5251 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118 | 0.4 | 118.4 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.023427 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.4317 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118 | 0.4 | 118.4 | 37 | 37 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118 | 0.4 | 118.4 | 37.1 | 37.1 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2043 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118 | 0.4 | 118.4 | 37.1 | 37.1 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4428 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.1 | 0.4 | 118.5 | 37.1 | 37.1 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3186 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.1 | 0.4 | 118.5 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.1 | 0.4 | 118.5 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.2 | 0.4 | 118.6 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.2 | 0.4 | 118.6 | 37 | 37 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2782 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.009 | 90 | 118.2 | 0.4 | 118.6 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4156 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.34 | 0.0087 | 86.9193 | 118.3 | 0.4 | 118.7 | 37 | 37 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2266 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.3877 | 25.4 | 0.4
2008-05-08 | 4:12 | 2.609 | 0.0085 | 85.2459 | 118.3 | 0.4 | 118.7 | 37 | 37 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0234 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.816 | 0.0118 | 118.0328 | 125.5 | 0.4 | 125.9 | 37.5 | 37.4548 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.024748 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.923 | 0.0112 | 112.1075 | 138.7 | 0.4 | 139.1 | 40.9 | 40.9405 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.02714 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 144.1 | 0.4 | 144.5 | 44.3 | 44.2592 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.028107 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2555 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 147.9 | 0.4 | 148.3 | 45.6 | 45.5776 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.02882 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.265 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 148.8 | 0.4 | 149.2 | 46 | 46.0291 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.028973 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 148.9 | 0.4 | 149.3 | 46.2 | 46.1656 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.734 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149 | 0.5 | 149.5 | 46.2 | 46.2 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.2675 | 29.2733 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149 | 0.5 | 149.5 | 46.3 | 46.3 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4588 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149 | 0.5 | 149.5 | 46.3 | 46.3 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5027 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149 | 0.5 | 149.5 | 46.3 | 46.3 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2523 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.1 | 0.5 | 149.6 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.934 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.1 | 0.5 | 149.6 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1124 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.1 | 0.5 | 149.6 | 46.3 | 46.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.000627 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.1406 | 1001 | 29.3511 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.2 | 0.5 | 149.7 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4701 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.4 | 0.5734 | 150 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.1302 | 1001 | 29.4094 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.6 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.4889 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.5684 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.4122 | 1000 | 29.4562 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3548 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.6 | 150.4 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.5 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.6 | 150.4 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.5 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3028 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2455 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.275 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2178 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0291 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.5358 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029027 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4731 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3967 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.6 | 150.4 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2372 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.6 | 150.4 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3219 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.6 | 150.4 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3595 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.3 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.6 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.3 | 25.4 | 0.6
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.8 | 0.5266 | 150.3 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.5355 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3672 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.2702 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2759 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.4 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.2665 | 29.2471 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.256 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2645 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2298 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3886 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2522 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3068 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4329 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.3 | 46.3 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 29.4626 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.3 | 46.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2241 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.2 | 46.2456 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2715 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.2 | 46.2302 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.266 | 29.4019 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.4697 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.2 | 46.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3813 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.2213 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2968 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.7 | 0.5 | 150.2 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4 | 25.4 | 0.5
2008-05-08 | 4:13 | 2.95 | 0.011 | 110 | 149.6 | 0.5 | 150.1 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3926 | 25.4 | 0.5
2008-05-08 | 4:13 | 3.021 | 0.011 | 110 | 149.5 | 0.5 | 150 | 46.2 | 46.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.029 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3131 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.339 | 0.011 | 110 | 149.6 | 0.4 | 150 | 46.3 | 46.2557 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.029073 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4991 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.53 | 0.0134 | 133.838 | 166.6 | 0.4 | 167 | 48.7 | 48.7315 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.032007 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.6 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.53 | 0.014 | 140 | 170.4 | 0.4 | 170.8 | 52.7 | 52.6821 | 0 | 13 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.032667 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.5241 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.531 | 0.014 | 140 | 175.5 | 0.4 | 175.9 | 54.5 | 54.5007 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.033573 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2952 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.539 | 0.0137 | 137.0546 | 176.2 | 0.4 | 176.6 | 55.1 | 55.0858 | 0 | 13 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.033673 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2157 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 176.5 | 0.4 | 176.9 | 55.3 | 55.3 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0337 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 176.7 | 0.4 | 177.1 | 55.4 | 55.4 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.033773 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.3301 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 177.1 | 0.4 | 177.5 | 55.5 | 55.4516 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0338 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.5 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 177.4 | 0.4 | 177.8 | 55.5 | 55.5 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2266 | 0.033873 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4963 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 177.7 | 0.4 | 178.1 | 55.5 | 55.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.033973 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3365 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 177.9 | 0.4 | 178.3 | 55.5 | 55.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.034 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2386 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 178 | 0.4 | 178.4 | 55.5 | 55.5 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.034 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.3634 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 178.7 | 0.4 | 179.1 | 55.5 | 55.5 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0341 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.4983 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 178.9 | 0.4 | 179.3 | 55.5 | 55.5 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.034173 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.1008 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 178.9 | 0.4 | 179.3 | 55.5 | 55.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.338 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 179 | 0.4 | 179.4 | 55.5 | 55.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.5177 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.54 | 0.013 | 130 | 179 | 0.4 | 179.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.6 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.548 | 0.013 | 130 | 179.1 | 0.4 | 179.5 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4935 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.2 | 0.4 | 179.6 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0158 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.2 | 0.4 | 179.6 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1537 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.547 | 0.013 | 130 | 179.2 | 0.4 | 179.6 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3127 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.541 | 0.013 | 130 | 179.3 | 0.4 | 179.7 | 55.5 | 55.5313 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.3283 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.2 | 0.4 | 179.6 | 55.6 | 55.5726 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0342 | 0.000673 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.3 | 0.4 | 179.7 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.3 | 0.4 | 179.7 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0342 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.3 | 0.4 | 179.7 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0342 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.3 | 0.4 | 179.7 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.034273 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.2667 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.6 | 0.4 | 180 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0343 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.7 | 0.4 | 180.1 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.034327 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7325 | 29.6154 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4949 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3305 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2733 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3151 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.4 | 180.2 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0332 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.9 | 0.4 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0922 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.9 | 0.4 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2254 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3103 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4695 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3715 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0345 | 0.000773 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2563 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2692 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 180 | 0.4 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.4
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.9 | 0.47325 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7329 | 29.3766 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.9 | 0.5 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2412 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3997 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.0129 | 128.9385 | 179.9 | 0.5 | 180.4 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2407 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.0121 | 120.5139 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 29.1409 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.012 | 120 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0345 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2543 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.0126 | 125.925 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0344 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4461 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0344 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2074 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.55 | 0.013 | 130 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0344 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:14 | 3.755 | 0.0133 | 132.8181 | 179.8 | 0.5 | 180.3 | 55.5 | 55.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0344 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.024 | 0.0158 | 158.4372 | 180.3 | 0.5 | 180.8 | 55.8 | 55.8264 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.034547 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.0144 | 144.2185 | 199.3 | 0.5 | 199.8 | 58.6 | 58.5784 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.037781 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2464 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 202.7 | 0.5 | 203.2 | 62.2 | 62.2358 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.03832 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.4843 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 207.4 | 0.5 | 207.9 | 63.8 | 63.8151 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.039093 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2766 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 208.1 | 0.5 | 208.6 | 64.4 | 64.4 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0392 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 208.1 | 0.5 | 208.6 | 64.6 | 64.6 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0392 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 208.1 | 0.5 | 208.6 | 64.7 | 64.7 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2741 | 0.0392 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 208.2 | 0.5 | 208.7 | 64.7 | 64.7 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0392 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 208.2 | 0.5 | 208.7 | 64.7 | 64.7302 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0392 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.344 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.161 | 0.014 | 140 | 208.2 | 0.5 | 208.7 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.039273 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.297 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.169 | 0.014 | 140 | 208.2 | 0.5 | 208.7 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0393 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2931 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 208.5 | 0.5 | 209 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0393 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4894 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.2 | 0.5 | 209.7 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4099 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.3 | 0.5 | 209.8 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3305 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.3 | 0.5 | 209.8 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.349 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.3 | 0.6 | 209.9 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2266 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3762 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.162 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1377 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 209.3 | 0.6 | 209.9 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3677 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.163 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5257 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.171 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0394 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 29.573 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.179 | 0.014 | 140 | 209.4 | 0.6 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7335 | 29.1755 | 25.4 | 0.6
2008-05-08 | 4:15 | 4.18 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.174 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2316 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3768 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3046 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4644 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3766 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2588 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2417 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.5 | 0.5 | 210 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4003 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.6 | 0.5 | 210.1 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.8329 | 1001 | 29.4797 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.7 | 0.5 | 210.2 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0395 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3223 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.8 | 0.5 | 210.3 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.039573 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 209.9 | 0.5 | 210.4 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.014 | 140 | 210 | 0.5 | 210.5 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.17 | 0.0148 | 147.9066 | 210 | 0.5 | 210.5 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.2772 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.167 | 0.015 | 150 | 210 | 0.5 | 210.5 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2433 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 210.1 | 0.5 | 210.6 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.0148 | 147.5021 | 210 | 0.5 | 210.5 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2626 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.1 | 0.5 | 210.6 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.3 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5806 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.2 | 0.5 | 210.7 | 64.8 | 64.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3761 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.2 | 0.5 | 210.7 | 64.8 | 64.8381 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2461 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.2 | 0.5 | 210.7 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1333 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.3 | 0.5 | 210.8 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0396 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0655 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.4 | 0.5 | 210.9 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.039673 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.4627 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.5 | 0.5 | 211 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0397 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2838 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.7 | 0.5 | 211.2 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0397 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.014 | 140 | 210.9 | 0.5 | 211.4 | 64.9 | 64.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2927 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.0146 | 145.9549 | 211 | 0.5 | 211.5 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.46 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.14 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.039773 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1695 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.9 | 64.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.9 | 64.8619 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2569 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3841 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211 | 0.5 | 211.5 | 64.8 | 64.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2251 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211.1 | 0.5 | 211.6 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.16 | 0.015 | 150 | 211.1 | 0.5 | 211.6 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2464 | 25.4 | 0.5
2008-05-08 | 4:15 | 4.481 | 0.0154 | 154.4854 | 211.1 | 0.5 | 211.6 | 64.8 | 64.8323 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0398 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.683 | 0.0188 | 187.9582 | 213.5 | 0.5 | 214 | 65.3 | 65.3158 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.040167 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2946 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0182 | 182.4578 | 230.2 | 0.5 | 230.7 | 69.1 | 69.106 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.04294 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2151 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.018 | 180 | 233.8 | 0.5 | 234.3 | 72.2 | 72.2175 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0434 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3288 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.018 | 180 | 236 | 0.5 | 236.5 | 73.5 | 73.4807 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.043873 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3123 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0175 | 174.66 | 237.6 | 0.5 | 238.1 | 73.9 | 73.8569 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.044173 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2929 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0161 | 161.069 | 237.8 | 0.5 | 238.3 | 74 | 74.0319 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0442 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.5923 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0169 | 169.4865 | 238.3 | 0.5 | 238.8 | 74.1 | 74.1 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2266 | 0.0443 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3531 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 238.9 | 0.5 | 239.4 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0444 | 0.000873 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3617 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 238.9 | 0.5 | 239.4 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0444 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3175 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 239.5 | 0.5 | 240 | 74.2 | 74.2 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.044474 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2416 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 239.7 | 0.5 | 240.2 | 74.2 | 74.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0445 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3995 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 239.8 | 0.5 | 240.3 | 74.2 | 74.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2405 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.766 | 0.017 | 170 | 239.9 | 0.5 | 240.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.0162 | 162.4488 | 240 | 0.5 | 240.5 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2769 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3823 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.3025 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2229 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2567 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.764 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2638 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2155 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2944 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 29.2256 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.1 | 74.1347 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.1 | 0.5 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2734 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2521 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.5701 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3119 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.097 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0162 | 162.396 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1236 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3898 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7325 | 29.3101 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3694 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.1 | 74.1 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3022 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2284 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2921 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0446 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2126 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.044673 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.017 | 170 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2461 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.0164 | 164.2144 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2492 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.3 | 0.5 | 240.8 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1157 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.76 | 0.016 | 160 | 240.3 | 0.4 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.3542 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.768 | 0.0161 | 160.6994 | 240.3 | 0.4 | 240.7 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3358 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.0169 | 169.0258 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3436 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.017 | 170 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.4 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.017 | 170 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.396 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.017 | 170 | 240.2 | 0.4 | 240.6 | 74.2 | 74.1709 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.236 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.017 | 170 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3232 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.0169 | 169.3083 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3589 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.0161 | 160.975 | 240.3 | 0.4 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2266 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2177 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.3 | 0.4 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.9 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.3 | 0.4 | 240.7 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.059 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.159 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.4 | 240.6 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.4734 | 240.7 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0873 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.1 | 74.1 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0716 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.2 | 0.5 | 240.7 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2305 | 25.4 | 0.5
2008-05-08 | 4:16 | 4.77 | 0.016 | 160 | 240.3 | 0.5 | 240.8 | 74.1 | 74.1 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.2675 | 29.3892 | 25.4 | 0.5
2008-05-08 | 4:17 | 4.997 | 0.016 | 160 | 240.3 | 0.5 | 240.8 | 74 | 74 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0447 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4741 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.258 | 0.0182 | 182.4872 | 242.1 | 0.5 | 242.6 | 74.2 | 74.2026 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.044993 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.5 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.39 | 0.0184 | 184.2461 | 256.1 | 0.5 | 256.6 | 77.8 | 77.7703 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0464 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3671 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.389 | 0.018 | 180 | 264.3 | 0.5 | 264.8 | 80.9 | 80.8519 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.048567 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.1 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.381 | 0.018 | 180 | 267.2 | 0.5 | 267.7 | 82.5 | 82.5172 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.049 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.1 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.388 | 0.018 | 180 | 267.6 | 0.5 | 268.1 | 83 | 83.0269 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0491 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2432 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.39 | 0.018 | 180 | 268.2 | 0.5 | 268.7 | 83.2 | 83.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.049173 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.2489 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.386 | 0.018 | 180 | 268.5 | 0.5 | 269 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0492 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 268.5 | 0.5 | 269 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0492 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.381 | 0.018 | 180 | 268.6 | 0.5 | 269.1 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0492 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.389 | 0.018 | 180 | 269 | 0.5 | 269.5 | 83.3 | 83.3261 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.049273 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.383 | 0.018 | 180 | 269.1 | 0.5 | 269.6 | 83.4 | 83.3602 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3452 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.1 | 0.5 | 269.6 | 83.3 | 83.3 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.4721 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.1 | 0.5 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.3853 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.1 | 0.5 | 269.6 | 83.3 | 83.3 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2272 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269 | 0.5 | 269.5 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0677 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269 | 0.5 | 269.5 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 29.1831 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269 | 0.5 | 269.5 | 83.3 | 83.3 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.3495 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269 | 0.5 | 269.5 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.5 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0493 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.266 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.5 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0494 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.5
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4266 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1137 | 25.4432 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.045 | 25.4775 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.4 | 269.7 | 83.3 | 83.3 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0494 | 0.000973 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4 | 269.6 | 83.4 | 83.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.2 | 0.4 | 269.6 | 83.4 | 83.4 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1601 | 25.4199 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.4 | 269.7 | 83.4 | 83.4 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0013 | 25.4994 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.0181 | 180.8374 | 269.3 | 0.4 | 269.7 | 83.4 | 83.4 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0789 | 25.5 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.0189 | 188.9675 | 269.3 | 0.4 | 269.7 | 83.4 | 83.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2266 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1584 | 25.5 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.019 | 190 | 269.4 | 0.4 | 269.8 | 83.4 | 83.4 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1242 | 25.5 | 0.4
2008-05-08 | 4:17 | 5.383 | 0.0184 | 184.057 | 269.4 | 0.4 | 269.8 | 83.4 | 83.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0693 | 25.4827 | 0.4
2008-05-08 | 4:17 | 5.389 | 0.018 | 180 | 269.4 | 0.4 | 269.8 | 83.4 | 83.3622 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.3873 | 25.4032 | 0.4
2008-05-08 | 4:17 | 5.381 | 0.018 | 180 | 269.4 | 0.4 | 269.8 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1712 | 25.4763 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.4 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.1557 | 25.4443 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.4 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.2 | 25.4 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.4 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.734 | 29.1416 | 25.4146 | 0.4
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.3 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.8241 | 25.494 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.3 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0494 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0943 | 25.4264 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.3 | 0.3 | 269.6 | 83.3 | 83.3 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 29.0938 | 25.4 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.4 | 0.3 | 269.7 | 83.3 | 83.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29.0651 | 25.4 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.4 | 0.3 | 269.7 | 83.3 | 83.2696 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.188 | 25.4 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.4 | 0.3 | 269.7 | 83.2 | 83.1719 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25.5 | 5 | 195 | 59.2734 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1085 | 25.4 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.4 | 0.3 | 269.7 | 83.1 | 83.1 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.3 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9584 | 25.4708 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.5 | 0.3 | 269.8 | 83.1 | 83.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0006 | 25.4497 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.4 | 0.3 | 269.7 | 83.2 | 83.1643 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.4299 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.5 | 0.3 | 269.8 | 83.2 | 83.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0814 | 25.5 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.5 | 0.3 | 269.8 | 83.2 | 83.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9223 | 25.5 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.018 | 180 | 269.5 | 0.3 | 269.8 | 83.2 | 83.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1049 | 25.5 | 0.3
2008-05-08 | 4:17 | 5.378 | 0.018 | 180 | 269.5 | 0.3 | 269.8 | 83.2 | 83.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1045 | 25.5 | 0.3
2008-05-08 | 4:17 | 5.37 | 0.018 | 180 | 269.6 | 0.3 | 269.9 | 83.2 | 83.2 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0817 | 25.4728 | 0.3
2008-05-08 | 4:17 | 5.37 | 0.0178 | 177.7833 | 269.6 | 0.3 | 269.9 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2801 | 25.4066 | 0.3
2008-05-08 | 4:17 | 5.37 | 0.017 | 170 | 269.5 | 0.3 | 269.8 | 83.2 | 83.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2268 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0427 | 25.4858 | 0.3
2008-05-08 | 4:17 | 5.37 | 0.017 | 170 | 269.6 | 0.3 | 269.9 | 83.2 | 83.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2618 | 25.4346 | 0.3
2008-05-08 | 4:17 | 5.37 | 0.017 | 170 | 269.7 | 0.3 | 270 | 83.3 | 83.2603 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.2649 | 25.445 | 0.3
2008-05-08 | 4:17 | 5.372 | 0.017 | 170 | 269.6 | 0.3 | 269.9 | 83.2 | 83.2 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.4755 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.017 | 170 | 269.6 | 0.3 | 269.9 | 83.2 | 83.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0919 | 25.4041 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.017 | 170 | 269.7 | 0.3 | 270 | 83.2 | 83.2493 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.933 | 25.4835 | 0.3
2008-05-08 | 4:17 | 5.38 | 0.017 | 170 | 269.7 | 0.3 | 270 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.623 | 0.017 | 170 | 269.7 | 0.3 | 270 | 83.3 | 83.3 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0495 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9843 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.878 | 0.0207 | 207.4357 | 274.1 | 0.3 | 274.4 | 83.6 | 83.574 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.050161 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.97 | 0.0214 | 214.344 | 280.7 | 0.3 | 281 | 87 | 87.0478 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0512 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1013 | 25.4987 | 0.3
2008-05-08 | 4:18 | 5.971 | 0.021 | 210 | 290.6 | 0.3 | 290.9 | 90.5 | 90.4681 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.052767 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1808 | 25.4192 | 0.3
2008-05-08 | 4:18 | 5.979 | 0.021 | 210 | 297 | 0.3 | 297.3 | 91.9 | 91.8822 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.05382 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0796 | 25.4602 | 0.3
2008-05-08 | 4:18 | 5.987 | 0.021 | 210 | 297.6 | 0.3 | 297.9 | 92.3 | 92.3297 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0539 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9604 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 297.9 | 0.3 | 298.2 | 92.5 | 92.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0539 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9191 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 298.7 | 0.3 | 299 | 92.6 | 92.6 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2266 | 0.0541 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9986 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 298.8 | 0.3 | 299.1 | 92.6 | 92.6 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0541 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1554 | 25.4223 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 299 | 0.3 | 299.3 | 92.6 | 92.6 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0541 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9706 | 25.4573 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 299.1 | 0.3 | 299.4 | 92.6 | 92.6 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2265 | 0.0541 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.874 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 299.2 | 0.3 | 299.5 | 92.6 | 92.6 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0541 | 0.001073 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.0164 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.021 | 210 | 299.1 | 0.3 | 299.4 | 92.6 | 92.6 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0541 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0959 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.0203 | 202.8207 | 299.1 | 0.3 | 299.4 | 92.6 | 92.6 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0542 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0246 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.02 | 200 | 299.2 | 0.3 | 299.5 | 92.6 | 92.6 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0542 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.02 | 200 | 299.3 | 0.3 | 299.6 | 92.7 | 92.7 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2734 | 0.0542 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9661 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.02 | 200 | 299.4 | 0.3 | 299.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0542 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8725 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.0199 | 199.0525 | 299.8 | 0.3 | 300.1 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7135 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.0182 | 182.132 | 299.9 | 0.3 | 300.2 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.918 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.0187 | 187.3631 | 300 | 0.3 | 300.3 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.992 | 0.019 | 190 | 300.1 | 0.3 | 300.4 | 92.7 | 92.7 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.1 | 0.3 | 300.4 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2733 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.011 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.1 | 0.3 | 300.4 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0905 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.993 | 0.019 | 190 | 300.2 | 0.3 | 300.5 | 92.7 | 92.7 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9605 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.99 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9492 | 25.5 | 0.3
2008-05-08 | 4:18 | 5.994 | 0.019 | 190 | 300.2 | 0.3 | 300.5 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.2 | 0.3 | 300.5 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9833 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8243 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0691 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 954 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1073 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.019 | 190 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2734 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9944 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.0196 | 195.609 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2734 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9149 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.0196 | 196.2365 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9645 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.0192 | 192.1732 | 300.3 | 0.3 | 300.6 | 92.7 | 92.7274 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.044 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.3 | 0.3 | 300.6 | 92.8 | 92.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.3 | 0.3 | 300.6 | 92.8 | 92.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.8 | 92.8 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0543 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9763 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.054373 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9415 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0998 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2734 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0204 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0383 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0477 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.7 | 92.7284 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2267 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8077 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.8 | 92.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2266 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8769 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.8 | 92.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9564 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.8 | 92.8308 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9284 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8303 | 25.5151 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.9 | 92.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.988 | 25.594 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.526 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.4 | 0.3 | 300.7 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 26 | 5 | 195 | 59.3 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0377 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2733 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.2762 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.3 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9426 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.5 | 0.3 | 300.8 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0544 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.5 | 0.3
2008-05-08 | 4:18 | 6 | 0.02 | 200 | 300.8 | 0.3 | 301.1 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.054473 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8604 | 25.5 | 0.3
2008-05-08 | 4:19 | 6 | 0.02 | 200 | 301.6 | 0.3 | 301.9 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.054573 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0094 | 25.5 | 0.3
2008-05-08 | 4:19 | 6 | 0.02 | 200 | 301.7 | 0.3 | 302 | 92.9 | 92.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0546 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0892 | 25.5 | 0.3
2008-05-08 | 4:19 | 6 | 0.02 | 200 | 301.8 | 0.3 | 302.1 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2734 | 0.0546 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9625 | 25.5688 | 0.3
2008-05-08 | 4:19 | 6 | 0.02 | 200 | 301.8 | 0.3 | 302.1 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.3 | 0.0546 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.5518 | 0.3
2008-05-08 | 4:19 | 6 | 0.0196 | 195.6883 | 301.8 | 0.3 | 302.1 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0546 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9829 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.639 | 0.018 | 180 | 301.8 | 0.3 | 302.1 | 92.9 | 92.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0546 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1859 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.443 | 0.018 | 180 | 296.9 | 0.3 | 297.2 | 92 | 92.0183 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.053839 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0269 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.0187 | 187.4005 | 279.6 | 0.3 | 279.9 | 87.6 | 87.5686 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.05116 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.868 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.019 | 190 | 274.9 | 0.3 | 275.2 | 85 | 85.0325 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.050427 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.5454 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.019 | 190 | 274.4 | 0.3 | 274.7 | 84 | 84.0271 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.050327 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8249 | 25.5751 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.019 | 190 | 274.2 | 0.3 | 274.5 | 83.7 | 83.6833 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0503 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9086 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.019 | 189.6075 | 273.6 | 0.3 | 273.9 | 83.6 | 83.6 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2264 | 0.0502 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7361 | 29.0676 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.384 | 0.0181 | 181.0751 | 272.7 | 0.3 | 273 | 83.6 | 83.6 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.050027 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1628 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 272.5 | 0.3 | 272.8 | 83.6 | 83.6 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2732 | 0.05 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.268 | 29.0732 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 272.4 | 0.3 | 272.7 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.05 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8778 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 272.4 | 0.3 | 272.7 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.3 | 0.05 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8032 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 272.3 | 0.3 | 272.6 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.05 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.962 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.8 | 0.3 | 272.1 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8791 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.6 | 0.3 | 271.9 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8399 | 25.5399 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.6 | 0.3 | 271.9 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9193 | 25.5807 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.6 | 0.3 | 271.9 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.9988 | 25.5012 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.5 | 0.3 | 271.8 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.9217 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.6 | 0.3 | 271.9 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9 | 25.5577 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.5 | 0.3 | 271.8 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.8782 | 1000 | 28.8628 | 25.5628 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.5 | 0.3 | 271.8 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2734 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.8332 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.5 | 0.3 | 271.8 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2737 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.9921 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.378 | 0.018 | 180 | 271.4 | 0.3 | 271.7 | 83.5 | 83.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0499 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 29 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.018 | 180 | 271.4 | 0.3 | 271.7 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.274 | 0.049826 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8909 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.018 | 180 | 271.3 | 0.3 | 271.6 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8339 | 25.5339 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.018 | 180 | 271.3 | 0.3 | 271.6 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2267 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9138 | 25.5862 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.018 | 180 | 271.3 | 0.3 | 271.6 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2738 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9932 | 25.5068 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.018 | 180 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9273 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.0186 | 185.5046 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9314 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.266 | 29.0105 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0903 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2735 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5697 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 29.0017 | 25.6 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9 | 25.6 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9082 | 25.6 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.5 | 83.5 | 0 | 13 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9877 | 25.6 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29 | 25.5329 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.5 | 83.5 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0463 | 25.5463 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.4 | 83.4 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1 | 25.5744 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.4 | 83.4 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0838 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.37 | 0.019 | 190 | 271.1 | 0.3 | 271.4 | 83.4 | 83.4411 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8455 | 25.5 | 0.3
2008-05-08 | 4:19 | 5.379 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.8643 | 25.5643 | 0.3
2008-05-08 | 4:19 | 5.394 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8126 | 25.5563 | 0.3
2008-05-08 | 4:19 | 5.4 | 0.019 | 190 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8152 | 25.523 | 0.3
2008-05-08 | 4:19 | 5.396 | 0.0183 | 182.9904 | 271.2 | 0.3 | 271.5 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.1955 | 25.6 | 0.3
2008-05-08 | 4:19 | 5.39 | 0.018 | 180 | 271.2 | 0.2266 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 29.036 | 25.6 | 0.2
2008-05-08 | 4:19 | 5.39 | 0.0184 | 183.6891 | 271.2 | 0.2 | 271.4 | 83.5 | 83.5 | 0 | 13 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8771 | 25.6 | 0.2
2008-05-08 | 4:19 | 5.381 | 0.019 | 190 | 271.2 | 0.1 | 271.3 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8817 | 25.5591 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.019 | 190 | 271.3 | 0.1 | 271.4 | 83.5 | 83.5 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9593 | 25.5203 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.019 | 190 | 271.3 | 0.1 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8005 | 25.5998 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.019 | 190 | 271.3 | 0.1 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9581 | 25.6 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.019 | 190 | 271.3 | 0.1 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8243 | 25.6 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.0186 | 186.3109 | 271.3 | 0.1 | 271.4 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8519 | 25.6 | 0.1
2008-05-08 | 4:19 | 5.38 | 0.018 | 180 | 271.2 | 0.1 | 271.3 | 83.5 | 83.5 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0826 | 25.6 | 0.1
2008-05-08 | 4:19 | 5.358 | 0.0179 | 179.1372 | 271.2 | 0.1 | 271.3 | 83.5 | 83.5 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.003 | 25.6 | 0.1
2008-05-08 | 4:20 | 5.055 | 0.0145 | 144.6247 | 270.9 | 0.1 | 271 | 83.5 | 83.5 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0498 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.8473 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.799 | 0.0162 | 161.5602 | 255.5 | 0.1 | 255.6 | 81.7 | 81.7238 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.047335 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.5442 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.0175 | 175.4047 | 248 | 0.1 | 248.1 | 77.5 | 77.4513 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.046106 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.5352 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.018 | 180 | 245.5 | 0.1 | 245.6 | 75.4 | 75.382 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.045707 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.018 | 180 | 243.8 | 0.1 | 243.9 | 74.6 | 74.6424 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.045427 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.018 | 179.7491 | 243.4 | 0.1 | 243.5 | 74.4 | 74.3698 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0454 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.0938 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.0171 | 171.0986 | 242.6 | 0.1 | 242.7 | 74.3 | 74.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.045227 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 29.0942 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.017 | 170 | 242.6 | 0.1 | 242.7 | 74.3 | 74.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0647 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.017 | 170 | 242.6 | 0.1 | 242.7 | 74.3 | 74.2697 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1527 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.0166 | 166.222 | 242.6 | 0.1 | 242.7 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.3 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8349 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.6 | 0.1 | 242.7 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0805 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.1006 | 25.5503 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29 | 25.5295 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.991 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9117 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.853 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7737 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7061 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7855 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9299 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9556 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 13 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2267 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9238 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.0161 | 160.6401 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0032 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.0169 | 168.9526 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 13 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2266 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0826 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.017 | 170 | 242.5 | 0.1 | 242.6 | 74.2 | 74.2 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2266 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9145 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.778 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 26 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8419 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.9992 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.3 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8403 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.1717 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.7408 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.266 | 28.8156 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.1 | 74.1488 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9821 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.4 | 0.1 | 242.5 | 74.1 | 74.1 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9025 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.3 | 0.1 | 242.4 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.3 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.7462 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.3 | 0.1 | 242.4 | 74.1 | 74.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8691 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170 | 242.3 | 0.1 | 242.4 | 74.1 | 74.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.9642 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.0163 | 162.9364 | 242.3 | 0.1 | 242.4 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8849 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8061 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.1755 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8738 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.3378 | 1001 | 28.8465 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25.5 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.774 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.3 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2734 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.3 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9006 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2 | 0.0452 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9101 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.78 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2734 | 0.045127 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7093 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.775 | 0.016 | 160 | 242.2 | 0.1 | 242.3 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2734 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7888 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.3 | 0.1 | 242.4 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2734 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.2 | 0.1 | 242.3 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8477 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.2 | 0.1 | 242.3 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2733 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.2 | 0.1 | 242.3 | 74.2 | 74.1701 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.3 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.913 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.2 | 0.1 | 242.3 | 74.1 | 74.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0712 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.1 | 0.1 | 242.2 | 74.1 | 74.1 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.3 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9695 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.016 | 160 | 242.1 | 0.1 | 242.2 | 74.1 | 74.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2733 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8104 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.017 | 170.4835 | 242 | 0.1 | 242.1 | 74.1 | 74.1356 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7485 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.0173 | 172.6218 | 242.1 | 0.1 | 242.2 | 74.2 | 74.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8963 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.77 | 0.0158 | 157.9773 | 242 | 0.1 | 242.1 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2266 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.817 | 25.6 | 0.1
2008-05-08 | 4:20 | 4.391 | 0.015 | 149.7005 | 241.9 | 0.1 | 242 | 74.2 | 74.2 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0451 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.204 | 0.0133 | 133.0616 | 225.8 | 0.1 | 225.9 | 72.1 | 72.0775 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.042439 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7582 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.19 | 0.0137 | 137.0224 | 217 | 0.1 | 217.1 | 68 | 67.9759 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.041033 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7212 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.19 | 0.014 | 140 | 214 | 0.1 | 214.1 | 66 | 65.9751 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.04048 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.19 | 0.014 | 140 | 213.2 | 0.1 | 213.3 | 65.2 | 65.2456 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.040327 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.19 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 65 | 64.9724 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0402 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7404 | 25.6596 | 0.1
2008-05-08 | 4:21 | 4.19 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0402 | 0.0011 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7387 | 25.6613 | 0.1
2008-05-08 | 4:21 | 4.179 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.9 | 64.8694 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0402 | 0.001026 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8178 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.166 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8978 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.8 | 64.8 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8434 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.153 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8307 | 25.6153 | 0.1
2008-05-08 | 4:21 | 4.15 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9889 | 25.6945 | 0.1
2008-05-08 | 4:21 | 4.154 | 0.0147 | 146.7833 | 212.3 | 0.1 | 212.4 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.3 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9258 | 25.6258 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.0145 | 144.6748 | 212.2 | 0.1 | 212.3 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 26 | 5 | 195 | 59.2266 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8462 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.9 | 64.9276 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 26 | 5 | 195 | 59.2266 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8663 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 65 | 65 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.2734 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.9874 | 25.6 | 0.1
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212.1 | 0 | 212.1 | 65 | 65 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2266 | 0.0402 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.908 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212 | 0 | 212 | 65 | 65 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.3 | 0.040127 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8286 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 212 | 0 | 212 | 65 | 65 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.0145 | 145.3713 | 211.9 | 0 | 211.9 | 65 | 65 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.8 | 25.6301 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.9 | 0 | 211.9 | 65 | 65 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7906 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 65 | 64.9711 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.711 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 149.9001 | 211.9 | 0 | 211.9 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.7 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.0142 | 141.5737 | 211.8 | 0 | 211.8 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7479 | 25.6521 | 0
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 211.9 | 0 | 211.9 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.7727 | 25.6273 | 0
2008-05-08 | 4:21 | 4.16 | 0.014 | 140 | 211.9 | 0 | 211.9 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.6934 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.0144 | 143.6891 | 211.8 | 0 | 211.8 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.6144 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25.5 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 64.9 | 64.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8244 | 25.6551 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.9 | 64.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 29.0514 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 64.9 | 64.8702 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8888 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6505 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6623 | 25.6623 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25.5 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8251 | 25.6583 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 26 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9573 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 26 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 952 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 26 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.7 | 0 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9191 | 25.6 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.9221 | 25.639 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 150 | 211.8 | 0 | 211.8 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.7633 | 25.7 | 0
2008-05-08 | 4:21 | 4.16 | 0.015 | 149.95 | 211.8 | -0.1 | 211.8 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6043 | 25.7 | -0.1
2008-05-08 | 4:21 | 4.16 | 0.0142 | 141.6097 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6772 | 25.6228 | -0.1
2008-05-08 | 4:21 | 4.164 | 0.014 | 140 | 211.8 | -0.1 | 211.8 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.2675 | 28.7566 | 25.6566 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.0145 | 145.391 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8721 | 25.7 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.015 | 150 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9694 | 25.6847 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.015 | 150 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8114 | 25.6057 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.015 | 150 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6741 | -0.1
2008-05-08 | 4:21 | 4.165 | 0.015 | 150 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 26 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.5197 | 1001 | 28.6928 | 25.7 | -0.1
2008-05-08 | 4:21 | 4.164 | 0.015 | 150 | 211.6 | -0.1 | 211.6 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.3 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6661 | 25.7 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.015 | 150 | 211.6 | -0.1 | 211.6 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8249 | 25.6875 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.015 | 150 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.9838 | 25.6081 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.0148 | 147.9789 | 211.7 | -0.1 | 211.7 | 64.8 | 64.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8578 | 25.6711 | -0.1
2008-05-08 | 4:21 | 4.157 | 0.014 | 140 | 211.7 | -0.1 | 211.7 | 64.8 | 64.7745 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7494 | 25.6494 | -0.1
2008-05-08 | 4:21 | 4.155 | 0.014 | 140 | 211.7 | -0.1 | 211.7 | 64.7 | 64.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7301 | 25.6 | -0.1
2008-05-08 | 4:21 | 4.164 | 0.014 | 140 | 211.7 | -0.1 | 211.7 | 64.7 | 64.7 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.014 | 140 | 211.7 | -0.1 | 211.7 | 64.7 | 64.7 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | -0.1
2008-05-08 | 4:21 | 4.17 | 0.014 | 140 | 211.7 | 0 | 211.7 | 64.7 | 64.7 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6683 | 0
2008-05-08 | 4:21 | 4.17 | 0.0142 | 142 | 211.6 | 0 | 211.6 | 64.7 | 64.7 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7523 | 25.7 | 0
2008-05-08 | 4:21 | 3.963 | 0.015 | 149.9335 | 211.6 | 0 | 211.6 | 64.7 | 64.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0401 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5914 | 25.7 | 0
2008-05-08 | 4:22 | 3.71 | 0.0133 | 133.3084 | 210.8 | 0 | 210.8 | 64.4 | 64.4468 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.039953 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3195 | 25.7 | 0
2008-05-08 | 4:22 | 3.57 | 0.0144 | 143.704 | 194.3 | 0 | 194.3 | 61.5 | 61.4552 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.03716 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5568 | 25.7 | 0
2008-05-08 | 4:22 | 3.57 | 0.0145 | 144.7024 | 186.5 | 0 | 186.5 | 58.2 | 58.1513 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.035833 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6652 | 25.6348 | 0
2008-05-08 | 4:22 | 3.57 | 0.014 | 140 | 184.5 | -0.0734 | 184.5 | 56.3 | 56.2844 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0355 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7447 | 25.6 | -0.1
2008-05-08 | 4:22 | 3.57 | 0.014 | 140 | 184 | -0.1 | 184 | 55.7 | 55.7494 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0354 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6241 | -0.1
2008-05-08 | 4:22 | 3.57 | 0.014 | 140 | 183.9 | -0.1 | 183.9 | 55.5 | 55.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0354 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7965 | 25.7 | -0.1
2008-05-08 | 4:22 | 3.565 | 0.014 | 140 | 183.8 | -0.1 | 183.8 | 55.5 | 55.4709 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2266 | 0.0354 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7171 | 25.7 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.0133 | 133.1067 | 183.7 | -0.1 | 183.7 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0353 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.638 | 25.7 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.6 | -0.1 | 182.6 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.035127 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7659 | 25.6585 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.5 | -0.1 | 182.5 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.9972 | 1001 | 28.9578 | 25.6 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.5 | -0.1 | 182.5 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.4 | -0.1 | 182.4 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.8 | 25.6 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.4 | -0.1 | 182.4 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7407 | 25.6593 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.584 | 25.7 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.2 | 0 | 182.2 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4724 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.2 | 0 | 182.2 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0351 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.7902 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.1 | 0 | 182.1 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6463 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182.1 | 0 | 182.1 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.2665 | 28.7126 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 182 | 0 | 182 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6929 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 181.9 | 0 | 181.9 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.515 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 181.8 | 0 | 181.8 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.594 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 181.8 | 0 | 181.8 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.035 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6 | 25.7 | 0
2008-05-08 | 4:22 | 3.552 | 0.013 | 130 | 181.6 | 0 | 181.6 | 55.4 | 55.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.034927 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7065 | 25.6467 | 0
2008-05-08 | 4:22 | 3.55 | 0.013 | 130 | 181.1 | 0 | 181.1 | 55.3 | 55.3312 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7674 | 25.6 | 0
2008-05-08 | 4:22 | 3.555 | 0.013 | 130 | 181 | 0 | 181 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.6 | 0
2008-05-08 | 4:22 | 3.56 | 0.0123 | 123.3667 | 181 | 0 | 181 | 55.3 | 55.33 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.6 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.4 | 55.37 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6292 | 25.6708 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.4 | 55.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6502 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.4 | 55.3865 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.3 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7296 | 25.6704 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.791 | 25.6 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0349 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7116 | 25.6 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.034827 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4967 | 25.6678 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 181 | 0 | 181 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2264 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5878 | 25.653 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7213 | 25.6262 | 0
2008-05-08 | 4:22 | 3.56 | 0.0122 | 121.9262 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5179 | 25.694 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 129.9518 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.756 | 25.6147 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6058 | 25.6 | 0
2008-05-08 | 4:22 | 3.56 | 0.0123 | 123.3888 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.6324 | 25.6441 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.2675 | 28.5533 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5232 | 25.6589 | 0
2008-05-08 | 4:22 | 3.56 | 0.012 | 120 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2735 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6795 | 25.6205 | 0
2008-05-08 | 4:22 | 3.56 | 0.0125 | 125.2976 | 180.9 | 0 | 180.9 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.6002 | 25.6998 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25.5 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.5207 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.8 | 0 | 180.8 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6173 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.7 | 0 | 180.7 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.738 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7653 | 25.6827 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6064 | 25.6032 | 0
2008-05-08 | 4:22 | 3.558 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.7325 | 28.7524 | 25.6 | 0
2008-05-08 | 4:22 | 3.55 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.8556 | 25.6 | 0
2008-05-08 | 4:22 | 3.55 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.7254 | 25.6349 | 0
2008-05-08 | 4:22 | 3.55 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.457 | 25.7 | 0
2008-05-08 | 4:22 | 3.555 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.7737 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6539 | 25.7 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | 0 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7051 | 25.6475 | 0
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | -0.1 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6724 | 25.6319 | -0.1
2008-05-08 | 4:22 | 3.56 | 0.013 | 130 | 180.6 | -0.1 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7 | -0.1
2008-05-08 | 4:22 | 3.225 | 0.013 | 130 | 180.6 | -0.1 | 180.6 | 55.3 | 55.3 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0348 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7 | -0.1
2008-05-08 | 4:23 | 3.001 | 0.0114 | 113.5774 | 173.7 | -0.1 | 173.7 | 54.8 | 54.7741 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2266 | 0.033572 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6095 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.0117 | 116.6752 | 157.6 | -0.1 | 157.6 | 50.6 | 50.6293 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.03078 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6507 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 154.3 | -0.1 | 154.3 | 47.8 | 47.8033 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0302 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6576 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 154.2 | -0.1 | 154.2 | 46.6 | 46.643 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0302 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 153 | -0.1 | 153 | 46.3 | 46.277 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.03 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.8 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 151.8 | -0.1 | 151.8 | 46.1 | 46.1 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.029727 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6659 | 25.7 | -0.1
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 151.6 | -0.2 | 151.6 | 46 | 46.0488 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0297 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.5077 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 151.5 | -0.2 | 151.5 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0297 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7335 | 28.4508 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.012 | 120 | 151.5 | -0.2 | 151.5 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0297 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5948 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.0117 | 116.7177 | 151.5 | -0.2 | 151.5 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0297 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5155 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.942 | 0.011 | 110 | 151.4 | -0.2 | 151.4 | 46 | 46 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0297 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6278 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 151 | -0.2 | 151 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.4837 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 151 | -0.2 | 151 | 46 | 46 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3129 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5784 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.5 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.94 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5195 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.945 | 0.011 | 110 | 150.8 | -0.2 | 150.8 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2737 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5989 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.2 | 150.8 | 46 | 46 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5426 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.2 | 150.9 | 45.9 | 45.9 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.611 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.2 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7509 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.2 | 150.8 | 46 | 46 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5124 | 25.7 | -0.2
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5752 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5454 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5339 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6133 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.0106 | 106.4726 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6927 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.009 | 90 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5559 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.009 | 90 | 150.8 | -0.3 | 150.8 | 46 | 45.965 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.5 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.0096 | 96.4892 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9289 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4693 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.0105 | 104.8207 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 13 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2733 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4395 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7582 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2266 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5241 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.3 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.3 | 150.9 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4831 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.7 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 46 | 46 | 0 | 12.9 | 952 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5689 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.3 | 150.9 | 45.9 | 45.9 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5245 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.3 | 150.9 | 45.9 | 45.9 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.596 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 150.9 | -0.3 | 150.9 | 45.9 | 45.9 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5167 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151 | -0.3 | 151 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5626 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151 | -0.3 | 151 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.6 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.1 | -0.3 | 151.1 | 45.9 | 45.93 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.5571 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.1 | -0.3 | 151.1 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.3992 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.2 | -0.3 | 151.2 | 46 | 46 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2266 | 0.0295 | 0.0009 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3198 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.2 | -0.3 | 151.2 | 46 | 46 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25.5 | 5 | 195 | 59.2 | 0.0295 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.4786 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.3 | -0.3 | 151.3 | 46 | 46 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.3 | -0.3 | 151.3 | 46 | 45.9852 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5462 | 25.7179 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.3 | -0.3 | 151.3 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0295 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3069 | 25.7977 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.4 | -0.3 | 151.4 | 45.9 | 45.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.029573 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5312 | 25.7229 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.4 | -0.3 | 151.4 | 45.9 | 45.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.6564 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.011 | 110 | 151.4 | -0.3 | 151.4 | 45.9 | 45.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2735 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5926 | 25.7 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.0105 | 105.2206 | 151.4 | -0.3 | 151.4 | 46 | 46 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3849 | 25.7151 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.01 | 100 | 151.5 | -0.3 | 151.5 | 46 | 46 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3061 | 25.7939 | -0.3
2008-05-08 | 4:23 | 2.95 | 0.01 | 100 | 151.4 | -0.3 | 151.4 | 46 | 45.9719 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7263 | -0.3
2008-05-08 | 4:23 | 2.603 | 0.01 | 100 | 151.5 | -0.3 | 151.5 | 46 | 45.9502 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0296 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3533 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.381 | 0.0083 | 83.2124 | 142.4 | -0.3 | 142.4 | 44.7 | 44.7049 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.027899 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.0087 | 86.7608 | 127.8 | -0.3 | 127.8 | 40.8 | 40.7636 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.025286 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.3759 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 124.5 | -0.3 | 124.5 | 38.1 | 38.0935 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.02468 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.2171 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 124 | -0.3 | 124 | 37.2 | 37.2401 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.3 | 0.0246 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2706 | 25.6294 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 124 | -0.3 | 124 | 36.9 | 36.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0246 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.65 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 123.7 | -0.3 | 123.7 | 36.8 | 36.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.024527 | 0.0008 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3295 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 123.2 | -0.3 | 123.2 | 36.8 | 36.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.024427 | 0.000727 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3911 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 123.2 | -0.3 | 123.2 | 36.8 | 36.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0244 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3118 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 123.1 | -0.3 | 123.1 | 36.8 | 36.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0244 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3675 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 122.9 | -0.3 | 122.9 | 36.8 | 36.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0244 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.4467 | 25.7467 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.8 | 36.8 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.474 | 25.774 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3886 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7335 | 28.2298 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2644 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.343 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3438 | 25.7 | -0.3
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.3 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.723 | -0.3
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.2266 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3945 | 25.7982 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.7 | -0.2 | 122.7 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1547 | 25.7182 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.2 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2837 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.2 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0243 | 0.0007 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.5219 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.2 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2737 | 0.024226 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.6398 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.8 | -0.2 | 122.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0242 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4017 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.7 | -0.2 | 122.7 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0242 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.3214 | 25.7786 | -0.2
2008-05-08 | 4:24 | 2.36 | 0.009 | 90 | 122.4 | -0.2 | 122.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.274 | 0.0242 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7399 | 28.4742 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.355 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0242 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.5249 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3831 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3037 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7244 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.2665 | 28.2453 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.2 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2341 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.1266 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 122.3 | -0.1 | 122.3 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0241 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.9 | -0.1 | 121.9 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1549 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.9 | -0.1 | 121.9 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2037 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.8 | -0.1 | 121.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2265 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3309 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.9 | -0.1 | 121.9 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3902 | 25.7098 | -0.1
2008-05-08 | 4:24 | 2.345 | 0.009 | 90 | 121.9 | -0.1 | 121.9 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.3102 | 25.7898 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.8 | -0.1 | 121.8 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.3 | 25.7307 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.6 | -0.1 | 121.6 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.024 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.3487 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.1 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3719 | 25.7 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.1 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7361 | 28.3 | 25.7074 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.1 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2266 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.9031 | 1001 | 28.3 | 25.7866 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.1 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2733 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1678 | 25.8 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.1 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2367 | 25.7544 | -0.1
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3502 | 25.7 | -0.2
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2086 | 25.7043 | -0.2
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3673 | 25.7837 | -0.2
2008-05-08 | 4:24 | 2.34 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2746 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.345 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.347 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2436 | 25.7782 | -0.2
2008-05-08 | 4:24 | 2.342 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7011 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.4 | 25.7805 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2802 | 25.8 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.009 | 90 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2737 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.278 | 25.761 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.0082 | 82.0083 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.382 | 25.718 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.008 | 80 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.302 | 25.798 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.008 | 80 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4549 | 25.7226 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.008 | 80 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3295 | 25.7568 | -0.2
2008-05-08 | 4:24 | 2.35 | 0.0082 | 81.6426 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7638 | -0.2
2008-05-08 | 4:24 | 2.192 | 0.009 | 89.6554 | 121.4 | -0.2 | 121.4 | 36.7 | 36.7 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0239 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7155 | -0.2
2008-05-08 | 4:25 | 1.959 | 0.0066 | 66.1148 | 121 | -0.2 | 121 | 36.6 | 36.6231 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.023827 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7947 | -0.2
2008-05-08 | 4:25 | 1.79 | 0.0073 | 73.0285 | 105 | -0.2 | 105 | 34.1 | 34.1441 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.021162 | 0.0006 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1258 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.008 | 80 | 97.8 | -0.2 | 97.8 | 30.2 | 30.2039 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.019513 | 0.000527 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2607 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.0077 | 76.8147 | 95.3 | -0.2 | 95.3 | 28.6 | 28.5838 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.019127 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.8923 | 1000.2665 | 28.3671 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 94.7 | -0.2 | 94.7 | 27.8 | 27.8404 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.019027 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.7325 | 28.3 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 94.7 | -0.2 | 94.7 | 27.7 | 27.6705 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.019 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 93.6 | -0.2 | 93.6 | 27.6 | 27.6 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.018727 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.4414 | 25.7293 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 93.1 | -0.2 | 93.1 | 27.6 | 27.5687 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0187 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.249 | 25.7 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 93.1 | -0.2 | 93.1 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0187 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.1194 | 25.7298 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 93 | -0.2 | 93 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.3632 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.77 | 0.007 | 70 | 93 | -0.2 | 93 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.0457 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.764 | 0.007 | 70 | 92.9 | -0.2 | 92.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.2717 | 25.8 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.9 | -0.2 | 92.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.3531 | 25.7531 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.9 | -0.2 | 92.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.3261 | 25.7261 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.0065 | 65.3091 | 92.8 | -0.2 | 92.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.794 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.8 | -0.2 | 92.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.7145 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.8 | -0.2 | 92.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3352 | 25.7648 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.7 | -0.2 | 92.7 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0186 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1674 | 25.7558 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.6 | -0.2 | 92.6 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.018527 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0235 | 25.7 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.5 | -0.2 | 92.5 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1053 | 25.7026 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.5 | -0.2 | 92.5 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2645 | 25.7822 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.006 | 60 | 92.5 | -0.2 | 92.5 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2267 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.115 | 25.7383 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.0061 | 61.2562 | 92.5 | -0.2 | 92.5 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1231 | 25.7 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 69.5757 | 92.4 | -0.2 | 92.4 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2796 | 25.7 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.3 | -0.2 | 92.3 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2262 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2002 | 25.7 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.2 | -0.2 | 92.2 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0426 | 25.7787 | -0.2
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.2 | -0.1266 | 92.2 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28 | 25.7417 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.2 | -0.1 | 92.2 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0379 | 25.7379 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92.1 | -0.1 | 92.1 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0185 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1172 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.754 | 0.007 | 70 | 92 | -0.1 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1967 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.75 | 0.007 | 70 | 92 | -0.1 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3521 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.753 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92 | -0.1 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2266 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3316 | 25.7658 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1717 | 25.7141 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0129 | 25.7936 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.2636 | 28.2188 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92 | -0.1 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92 | -0.1 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2265 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.2645 | 28.289 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2096 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2734 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2697 | 25.7303 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2265 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1527 | 25.7491 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.4 | 27.4372 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.3 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.436 | 1001 | 27.9715 | 25.8 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 27.9397 | 25.7921 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | -0.1 | 91.9 | 27.5 | 27.5 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3365 | 25.7127 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.8 | -0.1 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2667 | 25.7 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.8 | -0.1 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.154 | 25.7 | -0.1
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.8 | -0.026653 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2267 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1 | 25.7 | 0
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1048 | 25.7048 | 0
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.7325 | 28.1841 | 25.7841 | 0
2008-05-08 | 4:25 | 1.759 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2734 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.073 | 25.8 | 0
2008-05-08 | 4:25 | 1.75 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1286 | 25.8 | 0
2008-05-08 | 4:25 | 1.75 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2778 | 25.8 | 0
2008-05-08 | 4:25 | 1.75 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2016 | 25.7984 | 0
2008-05-08 | 4:25 | 1.75 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.5 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.281 | 25.719 | 0
2008-05-08 | 4:25 | 1.753 | 0.007 | 70 | 91.8 | 0 | 91.8 | 27.5 | 27.4753 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2733 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2396 | 25.7 | 0
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 91.9 | 0 | 91.9 | 27.4 | 27.4481 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2266 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.1205 | 25.7 | 0
2008-05-08 | 4:25 | 1.76 | 0.007 | 70 | 92 | 0 | 92 | 27.5 | 27.5 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.0574 | 25.7 | 0
2008-05-08 | 4:25 | 1.76 | 0.0068 | 68.4831 | 92 | 0 | 92 | 27.5 | 27.5 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2734 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2955 | 25.7 | 0
2008-05-08 | 4:25 | 1.76 | 0.006 | 60.4575 | 92.1 | 0 | 92.1 | 27.4 | 27.4 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.0668 | 25.7 | 0
2008-05-08 | 4:26 | 1.565 | 0.006 | 60 | 92.2 | 0 | 92.2 | 27.4 | 27.4 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2266 | 0.0184 | 0.0005 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.0572 | 25.7 | 0
2008-05-08 | 4:26 | 1.296 | 0.0054 | 53.8282 | 90.1 | 0 | 90.1 | 27.4 | 27.4 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.018033 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1367 | 25.7 | 0
2008-05-08 | 4:26 | 1.199 | 0.0055 | 54.6318 | 73.3 | 0 | 73.3 | 24.4 | 24.4091 | 0 | 13 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2734 | 0.014859 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.2 | 25.716 | 0
2008-05-08 | 4:26 | 1.19 | 0.0057 | 56.9199 | 66.3 | 0 | 66.3 | 20.8 | 20.771 | 0 | 12.9 | 952 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.013533 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2 | 25.7955 | 0
2008-05-08 | 4:26 | 1.19 | 0.0051 | 51.3257 | 64.8 | 0 | 64.8 | 19 | 19.0086 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.013253 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2748 | 25.7252 | 0
2008-05-08 | 4:26 | 1.19 | 0.006 | 59.509 | 63.6 | 0 | 63.6 | 18.5 | 18.468 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.013027 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.3 | 25.7 | 0
2008-05-08 | 4:26 | 1.185 | 0.006 | 60 | 63.3 | 0 | 63.3 | 18.3 | 18.2686 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.3 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.006 | 60 | 63.2 | 0 | 63.2 | 18.2 | 18.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2871 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.0055 | 55.3682 | 63.2 | 0 | 63.2 | 18.2 | 18.2 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.2645 | 28.2077 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63.2 | 0 | 63.2 | 18.2 | 18.2 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63.2 | 0 | 63.2 | 18.2 | 18.2 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3021 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63.1 | 0 | 63.1 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0129 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3697 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63 | 0 | 63 | 18.1 | 18.1448 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63.1 | 0 | 63.1 | 18.1 | 18.1251 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63.1 | 0 | 63.1 | 18.2 | 18.171 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.0943 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.005 | 50 | 63 | 0 | 63 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.1438 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.0048 | 47.6392 | 63 | 0 | 63 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2727 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.0031 | 31.0141 | 63 | 0 | 63 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2067 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.9 | 0 | 62.9 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2856 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.9 | 0 | 62.9 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3 | 25.7 | 0
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.9 | -0.1 | 62.9 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.1654 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.8 | -0.1 | 62.8 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.0486 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.8 | -0.1 | 62.8 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2734 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1036.6276 | 1001 | 28.1927 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.003 | 30 | 62.8 | -0.1 | 62.8 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2267 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.034 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.0038 | 37.8908 | 62.8 | -0.1 | 62.8 | 18.1 | 18.1 | 0 | 13 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 27.9377 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.004 | 40 | 62.8 | -0.1 | 62.8 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2266 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1087 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.004 | 40 | 62.7 | -0.1 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2267 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.3364 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.0043 | 43.0137 | 62.7 | -0.1 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0128 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1011 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.0049 | 48.8861 | 62.7 | -0.1 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2598 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.18 | 0.0041 | 40.5109 | 62.7 | -0.1 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.1814 | 25.7 | -0.1
2008-05-08 | 4:26 | 1.175 | 0.004 | 40 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.177 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.0046 | 46.2041 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2468 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.0074 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.1537 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3124 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.2932 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.16 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2734 | 0.0127 | 0.0004 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4775 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001.734 | 28.5 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2343 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.0976 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2266 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2881 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2087 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.6 | 0 | 62.6 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2267 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1296 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 13 | 952 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1499 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.266 | 28.2294 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.3176 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.005 | 50 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.4763 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.0034 | 34.2185 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.3647 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.003 | 30 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7467 | 0
2008-05-08 | 4:26 | 1.17 | 0.003 | 30 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2742 | 25.8 | 0
2008-05-08 | 4:26 | 1.17 | 0.003 | 30 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7942 | 0
2008-05-08 | 4:26 | 1.17 | 0.0031 | 31.4019 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2 | 25.7148 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 39.5057 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3291 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.000227 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2681 | 25.744 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1466 | 25.7767 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2733 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.2675 | 28.2975 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.218 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2615 | 25.7 | 0
2008-05-08 | 4:26 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3 | 25.7 | 0
2008-05-08 | 4:27 | 1.17 | 0.004 | 40 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25 | 5 | 195 | 59.2266 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2594 | 25.7 | 0
2008-05-08 | 4:27 | 1.014 | 0.0028 | 27.831 | 62.7 | 0 | 62.7 | 18.1 | 18.1 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1006 | 25.7 | 0
2008-05-08 | 4:27 | 0.643 | 0.002 | 20 | 60.3 | 0 | 60.3 | 17.5 | 17.4712 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.012086 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.0215 | 25.7 | 0
2008-05-08 | 4:27 | 0.592 | 0.002 | 20 | 43.9 | 0 | 43.9 | 13.6 | 13.6484 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2266 | 0.008932 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1163 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 38.7 | 0 | 38.7 | 10.8 | 10.7855 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.007886 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2379 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 36.6 | 0 | 36.6 | 9.5 | 9.5077 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0075 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3172 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 36.1 | 0 | 36.1 | 9.1 | 9.0703 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.007353 | 0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.3966 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.9 | 0 | 34.9 | 9 | 9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.007127 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.4 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.5 | 0 | 34.5 | 8.9 | 8.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.007 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.1782 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.4 | 0 | 34.4 | 8.9 | 8.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.007 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1045 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.4 | 0 | 34.4 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.007 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2716 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.4 | 0.1 | 34.5 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2266 | 0.007 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.113 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34.4 | 0.1 | 34.5 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.007 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.173 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 34 | 0.1 | 34.1 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0069 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3571 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.8 | 0.1 | 33.9 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.006827 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3413 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.8 | 0.1 | 33.9 | 8.9 | 8.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2731 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0334 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.8 | 0.1 | 33.9 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2716 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.8 | 0.1 | 33.9 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 28.3698 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2524 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1575 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2755 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0068 | 2.7e-05 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0374 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0068 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2007 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0068 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1154 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.595 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0068 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0019 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.7 | 0.1 | 33.8 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0068 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2847 | 25.7 | 0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.6 | 0 | 33.6 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.006727 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.0463 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.6 | 0 | 33.6 | 8.8 | 8.8352 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1919 | 25.636 | 0
2008-05-08 | 4:27 | 0.592 | 0.002 | 20 | 33.6 | 0 | 33.6 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2133 | 25.6433 | 0
2008-05-08 | 4:27 | 0.596 | 0.002 | 20 | 33.6 | 0 | 33.6 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1902 | 25.7 | 0
2008-05-08 | 4:27 | 0.595 | 0.002 | 20 | 33.6 | 0 | 33.6 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4986 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.5 | 0 | 33.5 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.3 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4186 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.5 | 0 | 33.5 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2267 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1563 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2733 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.121 | 25.7 | 0
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2266 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3197 | 25.7 | 0
2008-05-08 | 4:27 | 0.596 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3991 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2266 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.6345 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.5268 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.4 | 0 | 33.4 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3627 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.5 | 0 | 33.5 | 8.9 | 8.9 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.25 | 25.7 | 0
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.5 | 0 | 33.5 | 8.9 | 8.9 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0117 | 25.7 | 0
2008-05-08 | 4:27 | 0.594 | 0.002 | 20 | 33.5 | -0.1 | 33.5 | 8.8 | 8.8291 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2264 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 33.5 | -0.1 | 33.5 | 8.9 | 8.8504 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0067 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3549 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.593 | 0.002 | 20 | 33.4 | -0.1 | 33.4 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0066 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3657 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.4 | -0.1 | 33.4 | 8.8 | 8.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0066 | 0 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3137 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 33.4 | -0.1 | 33.4 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0066 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3932 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.592 | 0.002 | 20 | 33.4 | -0.1 | 33.4 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0066 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1824 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.596 | 0.002 | 20 | 33.3 | -0.1 | 33.3 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0066 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1519 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 32.7 | -0.1 | 32.7 | 8.8 | 8.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.006527 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1377 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0309 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2267 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2701 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.6 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0915 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.592 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0489 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2266 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1849 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3922 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3129 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.1 | -0.1 | 32.1 | 8.8 | 8.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2082 | 25.7 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.1 | -0.1 | 32.1 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.3 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.6747 | -0.1
2008-05-08 | 4:27 | 0.59 | 0.002 | 20 | 32.2 | -0.1 | 32.2 | 8.8 | 8.8 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188 | 25 | 5 | 195 | 59.2734 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0952 | 25.6 | -0.1
2008-05-08 | 4:28 | 0.354 | -0.0035 | -35.2353 | 32.1 | -0.1 | 32.1 | 8.8 | 8.8 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25.5 | 5 | 195 | 59.2 | 0.006327 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0159 | 25.6 | -0.1
2008-05-08 | 4:28 | 0.044 | -0.002 | -19.7496 | 20.3 | -0.1734 | 20.3 | 7.8 | 7.8012 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25.5 | 5 | 195 | 59.2 | 0.004719 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1906 | 25.6635 | -0.2
2008-05-08 | 4:28 | 0.001 | 0 | 0 | 10.2 | -0.2 | 10.2 | 3.5 | 3.5163 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.001952 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2141 | 25.7 | -0.2
2008-05-08 | 4:28 | 0.009 | 0 | 0 | 6.7 | -0.2 | 6.7 | 1 | 0.997 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.00128 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1222 | 25.7 | -0.2
2008-05-08 | 4:28 | 0.01 | 0 | 0 | 5.3 | -0.2 | 5.3 | 0 | 0.0411 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 188.5 | 25 | 5 | 195 | 59.2 | 0.001 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2 | 25.7 | -0.2
2008-05-08 | 4:28 | 0.004 | 0 | 0 | 4.2 | -0.3 | 4.2 | -0.2 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2267 | 0.0008 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 4.1 | -0.3 | 4.1 | -0.3 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.000727 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0788 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 4.1 | -0.3 | 4.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.0007 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.12 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 4 | -0.3 | 4 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0007 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2612 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 3.7 | -0.3 | 3.7 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0007 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1024 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 3.3 | -0.3 | 3.3 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0006 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.334 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 3.1 | -0.3 | 3.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3425 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 3.1 | -0.3 | 3.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.263 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | 0 | 0 | 3.1 | -0.3 | 3.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2164 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.0004 | -4.2333 | 3.1 | -0.3 | 3.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2959 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 3 | -0.3 | 3 | -0.4 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1495 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.008 | -0.001 | -10 | 3 | -0.3 | 3 | -0.4 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.001 | -10 | 2.9 | -0.3 | 2.9 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2266 | 0.000427 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.1 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.006 | -0.001 | -10 | 2.8 | -0.3 | 2.8 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1403 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.8 | -0.3 | 2.8 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2733 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3786 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2734 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3483 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2698 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0304 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.3464 | 25.6307 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.4026 | 25.6 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.2157 | 25.6281 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 189.5 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.0149 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.1737 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000 | 28.0013 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.7 | -0.3 | 2.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0828 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.6 | -0.3 | 2.6 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.2497 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.6 | -0.2 | 2.6 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.734 | 28.1 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.6 | -0.2 | 2.6 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.5 | -0.2 | 2.5 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2888 | 25.6371 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.5 | -0.2 | 2.5 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.6 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.5 | -0.2 | 2.5 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2733 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.2665 | 28.3779 | 25.6221 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.5 | -0.2 | 2.5 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3014 | 25.6986 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.4 | -0.2 | 2.4 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3809 | 25.6191 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.4 | -0.2 | 2.4 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2189 | 25.6604 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.4 | -0.2 | 2.4 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2 | 0.000326 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.6602 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.6192 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1 | 25.6986 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1778 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 26 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2572 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25.5 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1897 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 2 | -0.2 | 2 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190 | 25 | 5 | 195 | 59.2734 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0486 | 25.7 | -0.2
2008-05-08 | 4:28 | 0.004 | -0.001 | -10 | 2 | -0.2 | 2 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 190.5 | 25 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.734 | 28.2867 | 25.7 | -0.2
2008-05-08 | 4:28 | 0.008 | -0.001 | -10 | 1.9 | -0.2 | 1.9 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.3 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.001 | -10 | 1.9 | -0.2 | 1.9 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.2639 | 28.1366 | 25.7 | -0.2
2008-05-08 | 4:28 | 0 | -0.0013 | -12.5754 | 1.9 | -0.3 | 1.9 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.000227 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0678 | 25.7 | -0.3
2008-05-08 | 4:28 | 0 | -0.002 | -20 | 1.9 | -0.3 | 1.9 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 947 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.2133 | 25.6867 | -0.3
2008-05-08 | 4:28 | 0 | -0.002 | -20 | 1.8 | -0.3 | 1.8 | -0.5 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.2665 | 28.2927 | 25.6073 | -0.3
2008-05-08 | 4:28 | 0.004 | -0.002 | -20 | 1.7 | -0.3 | 1.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.6721 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.002 | -20 | 1.7 | -0.3 | 1.7 | -0.5 | 0 | 0 | 12.9 | 954 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2485 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.2 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1901 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1104 | 25.7 | -0.3
2008-05-08 | 4:28 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2734 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.0309 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2734 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1456 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 13 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2162 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.002 | -20 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2266 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.0018 | -17.5084 | 1.6 | -0.3 | 1.6 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.001 | -10 | 1.7 | -0.3 | 1.7 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 191.5 | 26 | 5 | 195 | 59.2734 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1321 | 25.6339 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.001 | -10 | 1.7 | -0.3 | 1.7 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2266 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2 | 25.6455 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.001 | -10 | 1.7 | -0.3 | 1.7 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2737 | 0.0002 | -0.000174 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2249 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.01 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.009 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.7 | -0.3
2008-05-08 | 4:29 | 0.001 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3628 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.7325 | 28.2734 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.122 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1986 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.8 | -0.3 | 1.8 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1192 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001.266 | 28.1602 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.3 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2396 | 25.6604 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.319 | 25.619 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3984 | 25.6984 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.4 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 13 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2266 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2284 | 25.6428 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 1.9 | -0.3 | 1.9 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1367 | 25.6367 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.4 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.216 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2955 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.5 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1502 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2734 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2086 | 25.6457 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2663 | 25.6337 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.3 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.2131 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2 | -0.3 | 2 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.2925 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.3 | 2.1 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1561 | 25.7 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.3 | 2.1 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2263 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.0487 | 25.6487 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.3 | 2.1 | -0.5 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.0306 | 25.6 | -0.3
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.2266 | 2.1 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.1193 | 25.6097 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1002 | 28.2789 | 25.6895 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.3 | 25.6311 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 25.5 | 5 | 195 | 59.3 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.2516 | 25.6 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1001 | 28.1722 | 25.6278 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.5 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1036 | 1000.266 | 28.0929 | 25.6929 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.0136 | 25.6136 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.5 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000 | 28.1978 | 25.6659 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1000.734 | 28.3 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.4 | 0 | 0 | 12.9 | 954 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.2255 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.3 | -0.2 | 2.3 | -0.4 | 0 | 0 | 13 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2733 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1001 | 28.0126 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28 | 192 | 26 | 5 | 195 | 59.2733 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.2509 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 13 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28.515 | 192 | 26 | 5 | 195 | 59.2735 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.1123 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 29 | 192 | 26 | 5 | 195 | 59.2267 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.084 | 25.658 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.2 | 2.1 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 29 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.1781 | 25.6 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.1 | -0.2 | 2.1 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 29 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.104 | 25.6 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 29 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1003 | 28.3421 | 25.6 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28.485 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.3399 | 25.6601 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 962 | 945 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28.515 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28.1819 | 25.7 | -0.2
2008-05-08 | 4:29 | 0 | -0.001 | -10 | 2.2 | -0.2 | 2.2 | -0.4 | 0 | 0 | 12.9 | 953 | 961 | 946 | 20 | 0.3 | 0.01 | 984 | -34 | 0 | 28.4855 | 192 | 26 | 5 | 195 | 59.2 | 0.0002 | -0.0003 | 0 | 0 |  | n/a |  | 30 | 53 | 1037 | 1002 | 28 | 25.7 | -0.2
Summary Information:
Post Processor DLL Version | 6.81
Status: | No Processing Errors Detected
Test Date | 2008-05-08
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 50568
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 325
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 298
NO2 Span(ppm) | 507.8
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.00
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 2
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 4:07
Test End Time | 4:29
Test Duration (s) | 1312
Average Ambient Temperature (deg C) | -34
Average Ambient Pressure (mbar) | 984
Average Relative Humidity (%) | 19.875
Average Absolute Humidity (grains/lb dry air) | 0.299
Average Kh Factor | 0.886
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Percent Time in NTE Zone | 0
Total Number of NTE Windows | 0
Total Number of non-consecutive NTE Windows | 0
| Avg. NTE Emissions | Maximum NTE Emissions | Std. Dev.
CO2 (g/kW-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/08/2008 03:56:30.978 - None Found
0X0000 - 05/08/2008 04:01:34.948 - None Found
0X0000 - 05/08/2008 04:05:14.234 - None Found
0X0000 - 05/08/2008 04:07:29.120 - None Found
Warnings:
0X0000 - 05/08/2008 03:56:30.978 - None Found
0X0000 - 05/08/2008 04:01:34.948 - None Found
0X0000 - 05/08/2008 04:05:14.234 - None Found
0X0000 - 05/08/2008 04:07:29.120 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 173
Brake Specific Dropout Limit Count | 1312
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-05-08
Time | 4:02
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
CO(ppm) | 9.375 | -10 | 19.375
CO2(%) | 0 | 0 | 0
NO(ppm) | 0.79375 | 0.28 | 0.51375
NO2(ppm) | 1.2 | -0.1 | 1.3
THC(ppmC) | -0.7 | -0.1 | -0.6
Span
InfoVer | 1
Date | 2008-05-08
Time | 4:05
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 200 | 6 | 20.9 | 325 | x^0: 298; x^1: 507.8; x^2: 0; x^3: 93
Gas | Previous | Current | Difference
CO(ppm) | 191.875 | 196 | -4.125
CO2(%) | 6 | 6.0007 | -0.000667
NO(ppm) | 296.1062 | 297.9133 | -1.8071
THC(ppmC) | 92.8063 | 93.08 | -0.27375
Test Information:
SEMTECH_DATA_FILE | 2008-05-08 | 3:56 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## THC Lin
Test Date:
Test Time:
5/8/2008
4:07:29
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
5/8/2008
4:07:29
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
5/8/2008
4:07:29
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
5/8/2008
4:07:29
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
5/8/2008
4:07:29
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
5/8/2008
4:07:29
NDUV Serial No:
208
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6 | Rm = | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277
| C.F.= | 0.099349 | 0.19884 | 0.29848 | 0.39826 | 0.49819 | 0.59826 | 0.69847 | 0.79884 | 0.89935
| gas = | 0.59609 | 1.193 | 1.7909 | 2.3896 | 2.9891 | 3.5895 | 4.1908 | 4.793 | 5.3961
Ro for CO2= | 0.8795