## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 66
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 28261
NDUV Serial Number: | 131
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | CO2 - AMBII Serial Number: |  |  | 28261 |  |  | CO - AMBII Serial Number: |  |  | 28261 |  | HC - AMBII Serial Number: |  |  | 28261 |  | NO - NDUV Serial Number: |  |  | 131 |  | THC - FID Serial Number: |  |  | 66
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-12-02 |  | 11.99 | X | 0.1 | 3 |  |  | 1196 |  | 50 | 3 |  | 319 |  | 4 | 3 |  | 1504 |  | 15 | 3 |  | 957 | 1000 | 5 | 2
Test Time: | 12:09
Test Software: | 2.018 66
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 6 | 6 |  |  | 0 | 1014 | 1014 |  |  | 0 | 0.9 | 0.9 |  |  | 0 | 0 | 0
90 | 105 | 1 | 1.199 | 1.183 | 1.128 | -0.055 | -4.65 |  | 119.6 | 132 | 12 | 10.03 |  | 31.9 | 1014 | 982.1 | 3078.68 | Fail | 150.4 | 150.3 | -0.1 | -0.07 |  | 95.7 | 96.5 | 0.8 | 0.84
150 | 165 | 2 | 2.398 | 2.37 | 2.296 | -0.074 | -3.12 |  | 239.2 | 254 | 15 | 6.27 |  | 63.8 | 1014.1 | 950.3 | 1489.5 | Fail | 300.8 | 302.4 | 1.6 | 0.53 |  | 191.4 | 192.6 | 1.2 | 0.63
210 | 225 | 3 | 3.597 | 3.56 | 3.464 | -0.096 | -2.7 |  | 358.8 | 372 | 13 | 3.62 |  | 95.7 | 1014 | 918.3 | 959.56 | Fail | 451.2 | 450.6 | -0.6 | -0.13 |  | 287.1 | 286.4 | -0.7 | -0.24
270 | 285 | 4 | 4.796 | 4.754 | 4.654 | -0.1 | -2.1 |  | 478.4 | 485 | 7 | 1.46 |  | 127.6 | 1014.5 | 886.9 | 695.06 | Fail | 601.6 | 600.6 | -1 | -0.17 |  | 382.8 | 380.3 | -2.5 | -0.65
330 | 345 | 5 | 5.995 | 5.951 | 5.852 | -0.099 | -1.66 |  | 598 | 606 | 8 | 1.34 |  | 159.5 | 1014.6 | 855.1 | 536.11 | Fail | 752 | 750.9 | -1.1 | -0.15 |  | 478.5 | 474.5 | -4 | -0.84
390 | 405 | 6 | 7.194 | 7.152 | 7.081 | -0.071 | -0.99 |  | 717.6 | 726 | 8 | 1.11 |  | 191.4 | 1014.5 | 823.1 | 430.04 | Fail | 902.4 | 902.1 | -0.3 | -0.03 |  | 574.2 | 570 | -4.2 | -0.73
450 | 465 | 7 | 8.393 | 8.356 | 8.298 | -0.058 | -0.69 |  | 837.2 | 841 | 4 | 0.48 |  | 223.3 | 1014.7 | 791.4 | 354.41 | Fail | 1052.8 | 1052.5 | -0.3 | -0.03 |  | 669.9 | 666.5 | -3.4 | -0.51
510 | 525 | 8 | 9.592 | 9.564 | 9.487 | -0.077 | -0.81 |  | 956.8 | 950 | -7 | -0.73 |  | 255.2 | 1014.3 | 759.1 | 297.45 | Fail | 1203.2 | 1202.5 | -0.7 | -0.06 |  | 765.6 | 762.7 | -2.9 | -0.38
570 | 585 | 9 | 10.791 | 10.775 | 10.681 | -0.094 | -0.87 |  | 1076.4 | 1072 | -4 | -0.37 |  | 287.1 | 1014.6 | 727.5 | 253.4 | Fail | 1353.6 | 1351.9 | -1.7 | -0.13 |  | 861.3 | 859.3 | -2 | -0.23
630 | 645 | 10 | 11.99 | 11.99 | 11.865 | -0.125 | -1.04 |  | 1196 | 1190 | -6 | -0.5 |  | 319 | 1014.9 | 695.9 | 218.15 | Fail | 1504 | 1502.6 | -1.4 | -0.09 |  | 957 | 957.2 | 0.2 | 0.02
690 | 705 | 9 | 10.791 | 10.775 | 10.689 | -0.086 | -0.8 |  | 1076.4 | 1069 | -7 | -0.65 |  | 287.1 | 1014 | 726.9 | 253.19 | Fail | 1353.6 | 1357.4 | 3.8 | 0.28 |  | 861.3 | 859.2 | -2.1 | -0.24
750 | 765 | 8 | 9.592 | 9.564 | 9.491 | -0.073 | -0.76 |  | 956.8 | 954 | -3 | -0.31 |  | 255.2 | 1014.4 | 759.2 | 297.49 | Fail | 1203.2 | 1208.8 | 5.6 | 0.47 |  | 765.6 | 762.7 | -2.9 | -0.38
810 | 825 | 7 | 8.393 | 8.356 | 8.3 | -0.056 | -0.67 |  | 837.2 | 837 | 0 | 0 |  | 223.3 | 1014 | 790.7 | 354.1 | Fail | 1052.8 | 1058 | 5.2 | 0.49 |  | 669.9 | 666.3 | -3.6 | -0.54
870 | 885 | 6 | 7.194 | 7.152 | 7.089 | -0.063 | -0.88 |  | 717.6 | 719 | 1 | 0.14 |  | 191.4 | 1014 | 822.6 | 429.78 | Fail | 902.4 | 907.6 | 5.2 | 0.58 |  | 574.2 | 570 | -4.2 | -0.73
930 | 945 | 5 | 5.995 | 5.951 | 5.863 | -0.088 | -1.48 |  | 598 | 594 | -4 | -0.67 |  | 159.5 | 1013.9 | 854.4 | 535.67 | Fail | 752 | 757 | 5 | 0.66 |  | 478.5 | 474.6 | -3.9 | -0.82
990 | 1005 | 4 | 4.796 | 4.754 | 4.661 | -0.093 | -1.96 |  | 478.4 | 491 | 13 | 2.72 |  | 127.6 | 1014 | 886.4 | 694.67 | Fail | 601.6 | 605 | 3.4 | 0.57 |  | 382.8 | 380 | -2.8 | -0.73
1050 | 1065 | 3 | 3.597 | 3.56 | 3.465 | -0.095 | -2.67 |  | 358.8 | 370 | 11 | 3.07 |  | 95.7 | 1013.4 | 917.7 | 958.93 | Fail | 451.2 | 455 | 3.8 | 0.84 |  | 287.1 | 286.1 | -1 | -0.35
1110 | 1125 | 2 | 2.398 | 2.37 | 2.281 | -0.089 | -3.76 |  | 239.2 | 242 | 3 | 1.25 |  | 63.8 | 1013.6 | 949.8 | 1488.71 | Fail | 300.8 | 306.7 | 5.9 | 1.96 |  | 191.4 | 192.1 | 0.7 | 0.37
1170 | 1185 | 1 | 1.199 | 1.183 | 1.123 | -0.06 | -5.07 |  | 119.6 | 122 | 2 | 1.67 |  | 31.9 | 1013.8 | 981.9 | 3078.06 | Fail | 150.4 | 154.5 | 4.1 | 2.73 |  | 95.7 | 95.8 | 0.1 | 0.1
1230 | 1245 | 0 | 0 | 0 | 0.003 | 0.003 |  |  | 0 | -10 | -10 |  |  | 0 | 1013 | 1013 |  |  | 0 | 1 | 1 |  |  | 0 | -0.3 | -0.3
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.049 | 0.12 |  |  |  | Intercept | 9 | 12 |  |  | Intercept | 1014 | 3 | Fail |  | Intercept | 2.2 | 15 |  |  | Intercept | -0.2 | 9.6
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.99 | 0.99 - 1.01 |  |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0267 | 0.1199 |  |  |  | SEE | 6.842 | 11.96 |  |  | SEE | 0.3453 | 3.19 |  |  | SEE | 2.7496 | 15.04 |  |  | SEE | 1.5718 | 9.572
|  |  | r2 | 1 | 0.998 |  |  |  | r2 | 0.9997 | 0.998 |  |  | r2 | 0.4473 | 0.998 | Fail |  | r2 | 1 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 2 | 0.6 | 2.6 | 0.3 | 0.3 | 0.1 | 13.1 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0015 | 0.0011 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.6 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 2 | 0.6 | 2.6 | 0.2 | 0.2 | 0.1 | 13.1 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.001471 | 0.0011 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.5047 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.9 | 0.6 | 2.5 | 0.2 | 0.2 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0014 | 0.0011 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.5 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.8 | 0.6 | 2.4 | 0.2 | 0.2 | 0.1 | 13.1 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.0014 | 0.001071 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.5535 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0.0013 | 12.8524 | 1.8 | 0.6 | 2.4 | 0.2 | 0.2 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8288 | 0.001371 | 0.001 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013.2883 | 11.5359 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0.002 | 20 | 1.8 | 0.6 | 2.4 | 0.1 | 0.1087 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.0013 | 0.001 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013.7123 | 11.4 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0.0014 | 14.3946 | 1.8 | 0.6 | 2.4 | 0.1 | 0.1 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 27.659 | 192 | 28 | 5 | 195.5 | 59.8 | 0.0013 | 0.001 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.4 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.7 | 0.57036 | 2.3 | 0.1 | 0.1089 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 27 | 192 | 28 | 5 | 195.8 | 59.8 | 0.00127 | 0.00097 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.4677 | 27.7 | 0.6
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.7 | 0.5 | 2.2 | 0.1 | 0.1 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 27.341 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0012 | 0.0009 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.5463 | 27.7 | 0.5
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.6 | 0.5 | 2.1 | 0.1 | 0.1 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.0012 | 0.0009 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.6 | 27.7 | 0.5
2008-12-02 | 12:09 | 0.01 | 0 | 0 | 1.6 | 0.5 | 2.1 | 0.1 | 0.1 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8 | 0.0012 | 0.0009 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050 | 1013.7083 | 11.5932 | 27.7 | 0.5
2008-12-02 | 12:09 | 0.009 | 0 | 0 | 1.6 | 0.5 | 2.1 | 0.1 | 0.1 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.001171 | 0.0009 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1050.3552 | 1014 | 11.436 | 27.7 | 0.5
2008-12-02 | 12:09 | 0.001 | 0 | 0 | 1.6 | 0.5 | 2.1 | 0.1 | 0.1 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0011 | 0.000871 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4606 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.5 | 0.5 | 2 | 0.1 | 0.1 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0011 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.5 | 1.9 | 0.1 | 0.1 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.7117 | 11.4824 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.5 | 1.9 | 0.1 | 0.1 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4041 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.5 | 1.9 | 0 | 0 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3253 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.5 | 1.9 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8296 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4601 | 27.7 | 0.5
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.4 | 1.8 | 0 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.568 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.4 | 1.8 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8293 | 0.001 | 0.0008 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.2931 | 11.4895 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.4 | 0.4 | 1.8 | 0 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.000971 | 0.000771 | -2.9e-05 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4114 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | 0 | 0 | 1.3 | 0.4 | 1.7 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0009 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4675 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | -0.0006 | -6.2694 | 1.3 | 0.4 | 1.7 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0009 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | -0.001 | -10 | 1.3 | 0.4 | 1.7 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0009 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 27.7 | 0.4
2008-12-02 | 12:09 | 0 | -0.001 | -10 | 1.3 | 0.37111 | 1.7 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0009 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 27.7 | 0.4
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.2 | 0.3 | 1.5 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.0009 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 27.7 | 0.3
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.2 | 0.27123 | 1.5 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.379 | 27.7 | 0.3
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.2 | 0.2 | 1.4 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8286 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.1 | 0.2 | 1.3 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3176 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.1 | 0.2 | 1.3 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.8294 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3956 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.1 | 0.2 | 1.3 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4746 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1.1 | 0.2 | 1.3 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1 | 0.2 | 1.2 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1 | 0.2 | 1.2 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1 | 0.2 | 1.2 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.000771 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | -0.001 | -10 | 1 | 0.2 | 1.2 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0007 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | 0.0002 | 1.5197 | 0.9 | 0.2 | 1.1 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0007 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | 0.0029 | 28.7652 | 0.9 | 0.2 | 1.1 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8288 | 0.0007 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4753 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | 0.0021 | 20.6418 | 0.9 | 0.2 | 1.1 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0007 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4033 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | 0.002 | 20 | 0.8 | 0.17036 | 1 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0007 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4819 | 27.7 | 0.2
2008-12-02 | 12:10 | 0 | 0.002 | 20 | 0.8 | 0.1 | 0.9 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0007 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0.002 | 20 | 0.8 | 0.1 | 0.9 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0.002 | 20 | 0.8 | 0.1 | 0.9 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4826 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0.0019 | 18.6869 | 0.8 | 0.1 | 0.9 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.2929 | 11.4045 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0.001 | 10.2694 | 0.8 | 0.1 | 0.9 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4745 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0.0002 | 2.2249 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8289 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.2883 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.001 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.7129 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.7055 | 11.4539 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.7069 | 11.3753 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.829 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0006 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.7 | 0.1 | 0.8 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.000571 | 0.000471 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3778 | 27.7 | 0.1
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0005 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4826 | 27.7 | 0
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0005 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4041 | 27.7 | 0
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0005 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 27.7 | 0
2008-12-02 | 12:10 | 0.01 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8704 | 0.0005 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.2964 | 11.4532 | 27.7 | 0
2008-12-02 | 12:10 | 0.01 | 0.0034 | 33.617 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0005 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4365 | 27.7 | 0
2008-12-02 | 12:10 | 0.803 | 0.0131 | 130.8007 | 11.4 | 0 | 11.4 | 0 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.002317 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051.2534 | 1014 | 11.3103 | 27.7 | 0
2008-12-02 | 12:10 | 1.099 | 0.0139 | 139.3186 | 86 | 0 | 86 | 0 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0146 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3886 | 27.7 | 0
2008-12-02 | 12:10 | 1.14 | 0.0132 | 132.3205 | 142.8 | 0 | 142.8 | 44.7 | 44.6752 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.023703 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4673 | 27.7 | 0
2008-12-02 | 12:10 | 1.14 | 0.013 | 130 | 147 | 0.029293 | 147.1 | 84.7 | 84.748 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.024346 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 0
2008-12-02 | 12:10 | 1.14 | 0.013 | 130 | 150 | 0.12892 | 150.2 | 93.8 | 93.7889 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.024829 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.451 | 27.7 | 0.1
2008-12-02 | 12:10 | 1.14 | 0.013 | 130 | 151.2 | 0.2 | 151.4 | 95.3 | 95.3122 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.025 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.3061 | 27.7 | 0.2
2008-12-02 | 12:10 | 1.132 | 0.0131 | 130.7923 | 151.4 | 0.22883 | 151.6 | 95.8 | 95.7937 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8712 | 0.025 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.7117 | 11.4632 | 27.7 | 0.2
2008-12-02 | 12:10 | 1.13 | 0.0139 | 139.1326 | 151.5 | 0.32877 | 151.8 | 95.9 | 95.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4398 | 27.7 | 0.3
2008-12-02 | 12:10 | 1.13 | 0.0132 | 132.3615 | 151.5 | 0.42872 | 151.9 | 96.1 | 96.1 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8287 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4388 | 27.7 | 0.4
2008-12-02 | 12:10 | 1.13 | 0.013 | 130 | 151.5 | 0.52964 | 152 | 96.1 | 96.1 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8704 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 0.5
2008-12-02 | 12:10 | 1.13 | 0.013 | 130 | 151.5 | 0.62945 | 152.1 | 96.3 | 96.2931 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0251 | 0.0004 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 0.6
2008-12-02 | 12:10 | 1.13 | 0.0132 | 132.452 | 151.5 | 1.6293 | 153.1 | 96.3 | 96.3 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0251 | 0.0004 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 1.6
2008-12-02 | 12:10 | 1.13 | 0.014 | 140 | 151.5 | 2.0612 | 153.6 | 96.4 | 96.3946 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0251 | 0.0004 | 2.9e-05 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 2.1
2008-12-02 | 12:10 | 1.13 | 0.014 | 140 | 151.5 | 2.7293 | 154.2 | 96.4 | 96.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.0251 | 0.0004 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 2.7
2008-12-02 | 12:10 | 1.13 | 0.0132 | 132.2968 | 151.5 | 3.7291 | 155.2 | 96.4 | 96.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8291 | 0.0251 | 0.0004 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 3.7
2008-12-02 | 12:10 | 1.127 | 0.013 | 130 | 151.4 | 4.0889 | 155.5 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8711 | 0.0251 | 0.000429 | 0.000129 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 4.1
2008-12-02 | 12:10 | 1.121 | 0.013 | 130 | 150.7 | 5.1306 | 155.8 | 96.5 | 96.5263 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8712 | 0.025 | 0.0005 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 5.1
2008-12-02 | 12:10 | 1.129 | 0.013 | 130 | 150.5 | 5.8 | 156.3 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8712 | 0.025 | 0.000529 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4542 | 27.7 | 5.8
2008-12-02 | 12:10 | 1.13 | 0.013 | 130 | 149.9 | 5.9287 | 155.8 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.024971 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4244 | 27.7 | 5.9
2008-12-02 | 12:10 | 1.124 | 0.013 | 130 | 149.8 | 6 | 155.8 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0249 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6
2008-12-02 | 12:10 | 1.12 | 0.013 | 130 | 149.7 | 6.0294 | 155.7 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0249 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6
2008-12-02 | 12:10 | 1.12 | 0.013 | 130 | 149.7 | 6.1 | 155.8 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0249 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4399 | 27.7 | 6.1
2008-12-02 | 12:10 | 1.12 | 0.013 | 130 | 149.7 | 6.1292 | 155.8 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.8292 | 0.025 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 6.1
2008-12-02 | 12:10 | 1.12 | 0.013 | 130 | 149.7 | 7.2579 | 157 | 96.6 | 96.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.025 | 0.0006 | 0.000329 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.3
2008-12-02 | 12:11 | 1.12 | 0.013 | 130 | 149.7 | 10.8929 | 160.6 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.025 | 0.0006 | 0.0005 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 10.9
2008-12-02 | 12:11 | 1.126 | 0.013 | 130 | 149.6 | 11.6289 | 161.2 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.025 | 0.0006 | 0.0005 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4743 | 27.7 | 11.6
2008-12-02 | 12:11 | 1.13 | 0.013 | 130 | 149.6 | 11.7 | 161.3 | 96.5 | 96.5 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.025 | 0.0006 | 0.0005 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 11.7
2008-12-02 | 12:11 | 1.128 | 0.013 | 130 | 149.6 | 11.6712 | 161.3 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.025 | 0.0006 | 0.0005 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 11.7
2008-12-02 | 12:11 | 1.121 | 0.013 | 130 | 149.7 | 11.0245 | 160.7 | 96.5 | 96.5 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.025029 | 0.0007 | 0.000471 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.49 | 27.7 | 11
2008-12-02 | 12:11 | 1.129 | 0.0132 | 132.4147 | 150.1 | 9.3016 | 159.4 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.02513 | 0.00073 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.2984 | 11.4119 | 27.7 | 9.3
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.1 | 8.6 | 158.7 | 96.5 | 96.469 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8296 | 0.0252 | 0.0008 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.467 | 27.7 | 8.6
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.1 | 8.5 | 158.6 | 96.5 | 96.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0008 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 8.5
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.2 | 8.1362 | 158.3 | 96.5 | 96.5 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0008 | 0.000371 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4757 | 27.7 | 8.1
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.2 | 7.5 | 157.7 | 96.5 | 96.5 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.5
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.2 | 7.5 | 157.7 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.829 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.5
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.3 | 7.4 | 157.7 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.4
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.3 | 7.3 | 157.6 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.3
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.4 | 7.2711 | 157.6 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8711 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.417 | 27.7 | 7.3
2008-12-02 | 12:11 | 1.13 | 0.014 | 140 | 150.4 | 7.2 | 157.6 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8712 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4955 | 27.7 | 7.2
2008-12-02 | 12:11 | 1.13 | 0.0134 | 133.9498 | 150.4 | 7.2 | 157.6 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8287 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4258 | 27.7 | 7.2
2008-12-02 | 12:11 | 1.129 | 0.013 | 130 | 150.4 | 7.1 | 157.5 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 196 | 59.8 | 0.0252 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7.1
2008-12-02 | 12:11 | 1.121 | 0.0123 | 122.856 | 150.4 | 7.1 | 157.5 | 96.5 | 96.5 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4313 | 27.7 | 7.1
2008-12-02 | 12:11 | 1.12 | 0.0101 | 100.7083 | 150.4 | 7.0707 | 157.5 | 96.5 | 96.5 | 0.1 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8293 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4902 | 27.7 | 7.1
2008-12-02 | 12:11 | 1.12 | 0.0109 | 109.0417 | 150.4 | 7 | 157.4 | 96.5 | 96.5 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8292 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4122 | 27.7 | 7
2008-12-02 | 12:11 | 1.12 | 0.011 | 110 | 150.5 | 7 | 157.5 | 96.5 | 96.5 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8705 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4668 | 27.7 | 7
2008-12-02 | 12:11 | 1.12 | 0.0116 | 116.0351 | 150.5 | 7 | 157.5 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4544 | 27.7 | 7
2008-12-02 | 12:11 | 1.12 | 0.012 | 120 | 150.5 | 6.9709 | 157.5 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.7 | 7
2008-12-02 | 12:11 | 1.12 | 0.012 | 120 | 150.5 | 6.9 | 157.4 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4027 | 27.7 | 6.9
2008-12-02 | 12:11 | 1.12 | 0.012 | 120 | 150.5 | 6.8 | 157.3 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8288 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4812 | 27.7 | 6.8
2008-12-02 | 12:11 | 1.12 | 0.012 | 120 | 150.5 | 6.7 | 157.2 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8714 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.7
2008-12-02 | 12:11 | 1.12 | 0.012 | 120 | 150.5 | 6.7 | 157.2 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8713 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.7
2008-12-02 | 12:11 | 1.12 | 0.0126 | 126.0418 | 150.5 | 6.7 | 157.2 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.7
2008-12-02 | 12:11 | 1.12 | 0.013 | 130 | 150.5 | 6.6 | 157.1 | 96.5 | 96.5 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8706 | 0.0253 | 0.0008 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.6
2008-12-02 | 12:11 | 1.12 | 0.013 | 130 | 150.5 | 6.5 | 157 | 96.5 | 96.498 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8707 | 0.0253 | 0.000829 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.5
2008-12-02 | 12:11 | 1.12 | 0.013 | 130 | 150.6 | 6.5 | 157.1 | 96.5 | 96.5 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0253 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.7 | 6.5
2008-12-02 | 12:11 | 1.126 | 0.013 | 130 | 150.6 | 6.4708 | 157.1 | 96.4 | 96.411 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8708 | 0.0253 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4688 | 27.7 | 6.5
2008-12-02 | 12:11 | 1.13 | 0.013 | 130 | 150.6 | 6.3709 | 157 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0253 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4097 | 27.7 | 6.4
2008-12-02 | 12:11 | 1.13 | 0.0136 | 136.0705 | 150.6 | 6.3 | 156.9 | 96.5 | 96.5 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0253 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4883 | 27.7 | 6.3
2008-12-02 | 12:11 | 1.144 | 0.014 | 140 | 150.6 | 6.3 | 156.9 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8288 | 0.0254 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4331 | 27.7 | 6.3
2008-12-02 | 12:11 | 1.316 | 0.0171 | 170.5246 | 150.9 | 6.3 | 157.2 | 96.5 | 96.5 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.025458 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.2877 | 11.4 | 27.7 | 6.3
2008-12-02 | 12:11 | 2.044 | 0.0268 | 267.9784 | 186.5 | 6.3 | 192.8 | 96.5 | 96.5 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.030777 | 0.0009 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.376 | 27.7 | 6.3
2008-12-02 | 12:11 | 2.26 | 0.0243 | 243.02 | 264.9 | 6.9632 | 271.8 | 100.2 | 100.248 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.042089 | 0.0009 | 0.00033 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.3052 | 27.7 | 7
2008-12-02 | 12:11 | 2.264 | 0.0248 | 247.6038 | 292.7 | 8.5 | 301.2 | 161.4 | 161.3651 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.045888 | 0.0009 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4619 | 27.7 | 8.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 294.6 | 8.5 | 303.1 | 186.5 | 186.4692 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8708 | 0.0462 | 0.0009 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 8.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 295.8 | 8.5 | 304.3 | 191 | 190.9706 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.046329 | 0.000929 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 8.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 297 | 8.5 | 305.5 | 191.9 | 191.875 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0466 | 0.001 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4833 | 27.7 | 8.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 297.1 | 8.6 | 305.7 | 192.2 | 192.2127 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0466 | 0.001 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4047 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.264 | 0.025 | 250 | 297.3 | 8.6 | 305.9 | 192.2 | 192.2061 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8709 | 0.0467 | 0.001029 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4739 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.26 | 0.0258 | 257.6038 | 297.4 | 8.7 | 306.1 | 192.3 | 192.2939 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8289 | 0.0467 | 0.0011 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5525 | 27.7 | 8.7
2008-12-02 | 12:11 | 2.262 | 0.026 | 260 | 297.5 | 8.7 | 306.2 | 192.3 | 192.3 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0467 | 0.0011 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5689 | 27.7 | 8.7
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 297.6 | 8.6 | 306.2 | 192.5 | 192.478 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0468 | 0.001129 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 297.8 | 8.6296 | 306.5 | 192.4 | 192.4087 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0469 | 0.0012 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7036 | 11.5 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 298 | 8.7 | 306.7 | 192.5 | 192.4919 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8294 | 0.0469 | 0.0012 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 8.7
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 298.1 | 8.7 | 306.8 | 192.6 | 192.5951 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.047 | 0.001229 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5905 | 27.7 | 8.7
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 298.2 | 8.6 | 306.8 | 192.5 | 192.5031 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.1 | 59.8 | 0.047029 | 0.0013 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6524 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.277 | 0.026 | 260 | 298.3 | 8.6 | 306.9 | 192.7 | 192.6862 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.0471 | 0.0013 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5047 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.297 | 0.026 | 260 | 298.5 | 8.6 | 307.1 | 192.6 | 192.6033 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8293 | 0.047129 | 0.0013 | 0.0004 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6613 | 27.7 | 8.6
2008-12-02 | 12:11 | 2.273 | 0.026 | 260 | 299.2 | 8.2382 | 307.5 | 192.7 | 192.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0472 | 0.0013 | 0.000371 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.7 | 8.2
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 299.3 | 7.5 | 306.8 | 192.6 | 192.6062 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0472 | 0.0013 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7381 | 27.7 | 7.5
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 299.7 | 7.4712 | 307.2 | 192.6 | 192.6 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.047258 | 0.001329 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7834 | 27.7 | 7.5
2008-12-02 | 12:11 | 2.27 | 0.026 | 260 | 300.7 | 7.141 | 307.9 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.047429 | 0.0014 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7123 | 11.705 | 27.7 | 7.1
2008-12-02 | 12:11 | 2.27 | 0.0254 | 254.1974 | 301 | 6.4714 | 307.5 | 192.6 | 192.5909 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0475 | 0.0014 | 0.0003 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.7 | 6.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 301.1 | 4.5 | 305.6 | 192.6 | 192.6047 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0475 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5954 | 27.7 | 4.5
2008-12-02 | 12:11 | 2.27 | 0.025 | 250 | 301.2 | 4.5 | 305.7 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0475 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5617 | 27.7 | 4.5
2008-12-02 | 12:12 | 2.27 | 0.0251 | 250.5932 | 301.3 | 4.5 | 305.8 | 192.6 | 192.6073 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0475 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.7 | 4.5
2008-12-02 | 12:12 | 2.293 | 0.0259 | 258.9474 | 301.5 | 4.5 | 306 | 192.7 | 192.6906 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.047529 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.2917 | 11.7 | 27.7 | 4.5
2008-12-02 | 12:12 | 2.31 | 0.026 | 260 | 301.8 | 4.4705 | 306.3 | 192.6 | 192.6058 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0476 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7664 | 27.7 | 4.5
2008-12-02 | 12:12 | 2.305 | 0.0254 | 254.2114 | 301.8 | 4.4 | 306.2 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0476 | 0.0014 | 0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7549 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 301.9 | 4.4 | 306.3 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0476 | 0.0014 | 0.000171 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.6526 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 301.9 | 4.4 | 306.3 | 192.6 | 192.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7111 | 11.5023 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 302 | 4.4 | 306.4 | 192.6 | 192.5939 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5805 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 302.1 | 4.4 | 306.5 | 192.5 | 192.5087 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5408 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 302.1 | 4.4 | 306.5 | 192.6 | 192.5939 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8287 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5758 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.292 | 0.025 | 250 | 302.2 | 4.4 | 306.6 | 192.6 | 192.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6672 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.3 | 4.4 | 306.7 | 192.5 | 192.5028 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0476 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5101 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.4 | 4.4 | 306.8 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.5 | 4.4 | 306.9 | 192.6 | 192.5985 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5519 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.6 | 4.4 | 307 | 192.6 | 192.6 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5693 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.5 | 4.4 | 306.9 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5185 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.025 | 250 | 302.6 | 4.4 | 307 | 192.6 | 192.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6754 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.0258 | 258.2864 | 302.6 | 4.4 | 307 | 192.5 | 192.5101 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.027 | 270 | 302.6 | 4.3712 | 307 | 192.5 | 192.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6101 | 27.7 | 4.4
2008-12-02 | 12:12 | 2.3 | 0.0269 | 269.4093 | 302.7 | 4.3 | 307 | 192.6 | 192.5934 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0477 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.547 | 27.7 | 4.3
2008-12-02 | 12:12 | 2.294 | 0.0261 | 261.0899 | 302.7 | 4.3 | 307 | 192.5 | 192.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.698 | 27.7 | 4.3
2008-12-02 | 12:12 | 2.29 | 0.026 | 260 | 302.8 | 4.2706 | 307.1 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6198 | 27.7 | 4.3
2008-12-02 | 12:12 | 2.288 | 0.026 | 260 | 302.8 | 4.2 | 307 | 192.5 | 192.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4821 | 27.7 | 4.2
2008-12-02 | 12:12 | 2.28 | 0.026 | 260 | 302.9 | 4.2 | 307.1 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4377 | 27.7 | 4.2
2008-12-02 | 12:12 | 2.28 | 0.026 | 260 | 303 | 4.2 | 307.2 | 192.6 | 192.5969 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5324 | 27.7 | 4.2
2008-12-02 | 12:12 | 2.28 | 0.026 | 260 | 303.1 | 4.1 | 307.2 | 192.6 | 192.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6896 | 27.7 | 4.1
2008-12-02 | 12:12 | 2.28 | 0.026 | 260 | 303.1 | 4.1 | 307.2 | 192.5 | 192.5098 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.2909 | 11.7 | 27.7 | 4.1
2008-12-02 | 12:12 | 2.28 | 0.026 | 260 | 303.1 | 4 | 307.1 | 192.4 | 192.4083 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7111 | 11.5968 | 27.7 | 4
2008-12-02 | 12:12 | 2.282 | 0.026 | 260 | 303.1 | 4 | 307.1 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4695 | 27.7 | 4
2008-12-02 | 12:12 | 2.29 | 0.026 | 260 | 303.1 | 4 | 307.1 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.409 | 27.7 | 4
2008-12-02 | 12:12 | 2.29 | 0.026 | 260 | 303.1 | 4 | 307.1 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8714 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4876 | 27.7 | 4
2008-12-02 | 12:12 | 2.29 | 0.0259 | 259.4016 | 303.1 | 3.9 | 307 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0477 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.9
2008-12-02 | 12:12 | 2.29 | 0.0251 | 250.7285 | 303.1 | 3.8 | 306.9 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.047729 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.7069 | 11.4553 | 27.7 | 3.8
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 303.2 | 3.7708 | 307 | 192.4 | 192.4055 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0478 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4233 | 27.7 | 3.8
2008-12-02 | 12:12 | 2.29 | 0.025 | 250 | 303.2 | 3.7 | 306.9 | 192.5 | 192.493 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0478 | 0.001329 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5036 | 27.7 | 3.7
2008-12-02 | 12:12 | 2.323 | 0.025 | 250 | 303.2 | 3.7 | 306.9 | 192.4 | 192.4098 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.0478 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.66 | 27.7 | 3.7
2008-12-02 | 12:12 | 2.638 | 0.0281 | 281.0201 | 304.9 | 3.7 | 308.6 | 192.5 | 192.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.048033 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.5826 | 27.7 | 3.7
2008-12-02 | 12:12 | 3.211 | 0.039 | 390 | 347.3 | 3.7 | 351 | 192.5 | 192.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.052269 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.7111 | 11.4625 | 27.7 | 3.7
2008-12-02 | 12:12 | 3.437 | 0.039 | 390 | 425.6 | 3.7 | 429.3 | 198 | 198.027 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.063305 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.416 | 27.7 | 3.7
2008-12-02 | 12:12 | 3.45 | 0.039 | 390 | 444 | 3.7288 | 447.7 | 256.1 | 256.1298 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8288 | 0.065458 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4946 | 27.7 | 3.7
2008-12-02 | 12:12 | 3.45 | 0.0384 | 384.4629 | 447 | 3.8 | 450.8 | 281.1 | 281.1371 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0658 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.6452 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.038 | 380 | 448.1 | 3.8 | 451.9 | 284.7 | 284.7395 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.066 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.546 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.038 | 380 | 448.3 | 3.8 | 452.1 | 285.7 | 285.6749 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.066 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.4 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.0379 | 378.8529 | 448.4 | 3.8 | 452.2 | 285.8 | 285.8 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8293 | 0.066029 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.0362 | 362.2278 | 448.7 | 3.8 | 452.5 | 285.8 | 285.8 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0661 | 0.001529 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 448.8 | 3.8 | 452.6 | 286 | 285.9953 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.0662 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5315 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 448.9 | 3.8 | 452.7 | 286.1 | 286.1 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195.7 | 59.8 | 0.0662 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5555 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 448.9 | 3.8 | 452.7 | 286.2 | 286.2 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.829 | 0.0662 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 449 | 3.8 | 452.8 | 286.3 | 286.2915 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8288 | 0.0662 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5016 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 448.9 | 3.8286 | 452.7 | 286.4 | 286.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8714 | 0.0662 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5799 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.45 | 0.036 | 360 | 449 | 3.9 | 452.9 | 286.3 | 286.3263 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5414 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.454 | 0.0375 | 374.7241 | 449.2 | 3.9 | 453.1 | 286.5 | 286.4905 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.3 | 3.9 | 453.2 | 286.4 | 286.4069 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.3 | 3.9 | 453.2 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8707 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.4 | 3.9 | 453.3 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6447 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.5 | 3.9 | 453.4 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0663 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.6 | 3.8707 | 453.5 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0663 | 0.001629 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6398 | 27.7 | 3.9
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.7 | 3.8 | 453.5 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0663 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.6 | 3.8 | 453.4 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8711 | 0.066329 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.8 | 3.8 | 453.6 | 286.5 | 286.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8288 | 0.0664 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7117 | 11.5 | 27.7 | 3.8
2008-12-02 | 12:12 | 3.46 | 0.038 | 380 | 449.9 | 3.8 | 453.7 | 286.5 | 286.5 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0664 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.8
2008-12-02 | 12:13 | 3.46 | 0.038 | 380 | 450 | 3.8 | 453.8 | 286.4 | 286.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8287 | 0.0664 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4772 | 27.7 | 3.8
2008-12-02 | 12:13 | 3.46 | 0.038 | 380 | 450 | 3.8 | 453.8 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8704 | 0.0664 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4013 | 27.7 | 3.8
2008-12-02 | 12:13 | 3.46 | 0.038 | 380 | 450.2 | 3.8 | 454 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.4798 | 27.7 | 3.8
2008-12-02 | 12:13 | 3.462 | 0.038 | 380 | 450.3 | 3.8 | 454.1 | 286.3 | 286.306 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4416 | 27.7 | 3.8
2008-12-02 | 12:13 | 3.469 | 0.0376 | 376.0984 | 450.3 | 3.7 | 454 | 286.4 | 286.3969 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.437 | 27.7 | 3.7
2008-12-02 | 12:13 | 3.461 | 0.037 | 370 | 450.4 | 3.7 | 454.1 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 196 | 59.871 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.7
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.4 | 3.7 | 454.1 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.7
2008-12-02 | 12:13 | 3.466 | 0.037 | 370 | 450.4 | 3.6709 | 454.1 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8 | 0.0665 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4278 | 27.7 | 3.7
2008-12-02 | 12:13 | 3.47 | 0.0377 | 377.0909 | 450.4 | 3.6 | 454 | 286.3 | 286.2927 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0665 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3491 | 27.7 | 3.6
2008-12-02 | 12:13 | 3.47 | 0.0374 | 374.4521 | 450.5 | 3.6 | 454.1 | 286.2 | 286.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3595 | 27.7 | 3.6
2008-12-02 | 12:13 | 3.47 | 0.037 | 370 | 450.5 | 3.6 | 454.1 | 286.4 | 286.3911 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8288 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4917 | 27.7 | 3.6
2008-12-02 | 12:13 | 3.47 | 0.037 | 370 | 450.5 | 3.6 | 454.1 | 286.3 | 286.3197 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7129 | 11.4131 | 27.7 | 3.6
2008-12-02 | 12:13 | 3.463 | 0.037 | 370 | 450.6 | 3.5 | 454.1 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8296 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4654 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.6 | 3.5 | 454.1 | 286.3 | 286.3049 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.7 | 3.5 | 454.2 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.0375 | 375.4379 | 450.7 | 3.5 | 454.2 | 286.4 | 286.3923 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.0376 | 376.2219 | 450.9 | 3.5 | 454.4 | 286.5 | 286.4949 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.9 | 3.5 | 454.4 | 286.4 | 286.4047 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6164 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.9 | 3.5 | 454.4 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0666 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6266 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 451 | 3.5 | 454.5 | 286.5 | 286.5 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 450.9 | 3.5 | 454.4 | 286.5 | 286.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.0376 | 375.5451 | 451 | 3.5 | 454.5 | 286.5 | 286.5 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.038 | 380 | 451 | 3.5 | 454.5 | 286.4 | 286.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.038 | 380 | 450.9 | 3.5 | 454.4 | 286.5 | 286.4932 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195.4 | 59.8 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.0379 | 378.8989 | 451.1 | 3.4707 | 454.5 | 286.5 | 286.495 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195.8 | 59.9 | 0.0667 | 0.0018 | 0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.2931 | 11.492 | 27.7 | 3.5
2008-12-02 | 12:13 | 3.46 | 0.0362 | 362.4651 | 451 | 3.4 | 454.4 | 286.5 | 286.5 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 196 | 59.8292 | 0.0667 | 0.0018 | 7.1e-05 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4134 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.036 | 360 | 451 | 3.4 | 454.4 | 286.4 | 286.4032 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.829 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4651 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.036 | 360 | 451.1 | 3.4 | 454.5 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4563 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.036 | 360 | 451 | 3.4 | 454.4 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8291 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4222 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.0362 | 362.4186 | 451.1 | 3.4 | 454.5 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5008 | 27.7008 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 451.2 | 3.4 | 454.6 | 286.4 | 286.3998 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5793 | 27.7793 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 451.2 | 3.4 | 454.6 | 286.4 | 286.3909 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4842 | 27.7421 | 3.4
2008-12-02 | 12:13 | 3.46 | 0.037 | 370 | 451.2 | 3.4 | 454.6 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0667 | 0.0018 | 0 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4364 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.458 | 0.037 | 370 | 451.2 | 3.4 | 454.6 | 286.4 | 286.4 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0667 | 0.0018 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.53 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.45 | 0.037 | 370 | 451.2 | 3.4 | 454.6 | 286.4 | 286.4 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8708 | 0.0667 | 0.0018 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.687 | 27.7 | 3.4
2008-12-02 | 12:13 | 3.45 | 0.0389 | 388.8105 | 451.2 | 3.4 | 454.6 | 286.5 | 286.4983 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0667 | 0.0018 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5558 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.23 | 0.0461 | 460.5646 | 455 | 3.4 | 458.4 | 286.4 | 286.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.871 | 0.067135 | 0.0018 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2903 | 11.5 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.599 | 0.0469 | 468.748 | 534.7 | 3.4 | 538.1 | 286.4 | 286.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.076224 | 0.001829 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4708 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.651 | 0.047 | 470 | 591.7 | 3.4 | 595.1 | 330.9 | 330.9403 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.082358 | 0.0019 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2909 | 11.4078 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.67 | 0.047 | 470 | 595.3 | 3.4 | 598.7 | 368.5 | 368.4985 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.9 | 0.0827 | 0.0019 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4863 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.67 | 0.047 | 470 | 596.8 | 3.4 | 600.2 | 377.6 | 377.5589 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0829 | 0.00193 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4352 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.67 | 0.0476 | 476.199 | 597.3 | 3.4 | 600.7 | 379.2 | 379.1516 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.083 | 0.002 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.67 | 0.05 | 500 | 598.7 | 3.4 | 602.1 | 379.7 | 379.6928 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8287 | 0.083129 | 0.002 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4219 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.67 | 0.05 | 500 | 598.8 | 3.4 | 602.2 | 379.8 | 379.7909 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.0832 | 0.002 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4995 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.666 | 0.05 | 500 | 598.8 | 3.4 | 602.2 | 380 | 379.9852 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0832 | 0.002 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.421 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.66 | 0.0489 | 489.1901 | 598.9 | 3.4 | 602.3 | 380.1 | 380.093 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0832 | 0.0021 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4575 | 27.7 | 3.4
2008-12-02 | 12:13 | 4.66 | 0.048 | 480 | 598.9 | 3.5 | 602.4 | 380 | 380.0077 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8708 | 0.083229 | 0.0021 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7361 | 3.5
2008-12-02 | 12:13 | 4.66 | 0.048 | 480 | 598.9 | 3.5293 | 602.5 | 380.1 | 380.1 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0833 | 0.0021 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4853 | 27.7853 | 3.5
2008-12-02 | 12:13 | 4.66 | 0.048 | 480 | 599.1 | 3.6 | 602.7 | 380.1 | 380.1 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.083329 | 0.002129 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4068 | 27.7068 | 3.6
2008-12-02 | 12:13 | 4.66 | 0.048 | 480 | 599.2 | 3.6 | 602.8 | 380.1 | 380.1 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0834 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4718 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.66 | 0.0494 | 494.1118 | 599.3 | 3.6 | 602.9 | 380.1 | 380.1 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.0834 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2883 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.662 | 0.0494 | 494.4511 | 599.2 | 3.6 | 602.8 | 380.1 | 380.1 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0834 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.669 | 0.049 | 490 | 599.3 | 3.6 | 602.9 | 380.1 | 380.1111 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.083429 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.661 | 0.049 | 490 | 599.5 | 3.6 | 603.1 | 380.2 | 380.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0835 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7036 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.66 | 0.049 | 490 | 599.6 | 3.6 | 603.2 | 380.2 | 380.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.0835 | 0.0022 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.666 | 0.049 | 490 | 599.8 | 3.6 | 603.4 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.0835 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.67 | 0.049 | 490 | 599.8 | 3.6 | 603.4 | 380.2 | 380.2077 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.0835 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:13 | 4.67 | 0.049 | 490 | 599.9 | 3.6 | 603.5 | 380.3 | 380.2872 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0835 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5003 | 27.7002 | 3.6
2008-12-02 | 12:13 | 4.67 | 0.049 | 490 | 599.9 | 3.6 | 603.5 | 380.2 | 380.204 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6574 | 27.7787 | 3.6
2008-12-02 | 12:13 | 4.67 | 0.049 | 490 | 599.9 | 3.6 | 603.5 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5856 | 27.8 | 3.6
2008-12-02 | 12:14 | 4.67 | 0.049 | 490 | 600 | 3.6 | 603.6 | 380.3 | 380.3042 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.7 | 28 | 5 | 195 | 59.9 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7642 | 3.6
2008-12-02 | 12:14 | 4.67 | 0.049 | 490 | 600.1 | 3.6 | 603.7 | 380.4 | 380.3956 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.8 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.666 | 0.0483 | 482.9599 | 600.1 | 3.6 | 603.7 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.8 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.2 | 3.6 | 603.8 | 380.4 | 380.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.8 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.2 | 3.6 | 603.8 | 380.3 | 380.3068 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.4 | 59.8296 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.45 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.3 | 3.6 | 603.9 | 380.2 | 380.2073 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.7 | 59.8 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.4 | 3.6 | 604 | 380.3 | 380.3 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.9 | 59.8293 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4071 | 27.7071 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.4 | 3.6 | 604 | 380.4 | 380.3945 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.5 | 59.8708 | 0.0836 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4857 | 27.7857 | 3.6
2008-12-02 | 12:14 | 4.66 | 0.048 | 480 | 600.4 | 3.6 | 604 | 380.3 | 380.3 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.3 | 28 | 5 | 195.1 | 59.829 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5642 | 27.8 | 3.6
2008-12-02 | 12:14 | 4.658 | 0.048 | 480 | 600.4 | 3.6 | 604 | 380.4 | 380.3939 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2883 | 11.5573 | 27.7573 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.048 | 480 | 600.4 | 3.6 | 604 | 380.3 | 380.3031 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7091 | 11.4574 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.0482 | 482.0628 | 600.6 | 3.6 | 604.2 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3003 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.8 | 3.6 | 604.4 | 380.3 | 380.3 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4568 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.8 | 3.6 | 604.4 | 380.2 | 380.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.83 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6131 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.8 | 3.6 | 604.4 | 380.3 | 380.3177 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8704 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5599 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.8 | 3.6 | 604.4 | 380.3 | 380.3028 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.328 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.9 | 3.6 | 604.5 | 380.3 | 380.3071 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4852 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.8 | 3.6 | 604.4 | 380.3 | 380.3252 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8708 | 0.0837 | 0.0023 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4288 | 27.7 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 600.9 | 3.6 | 604.5 | 380.3 | 380.3062 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0837 | 0.002329 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4496 | 27.7496 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 601 | 3.6 | 604.6 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0837 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7071 | 11.5 | 27.8 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 601.1 | 3.5709 | 604.7 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7932 | 3.6
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 601.2 | 3.5 | 604.7 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7147 | 3.5
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 601.2 | 3.5 | 604.7 | 380.3 | 380.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5639 | 27.7639 | 3.5
2008-12-02 | 12:14 | 4.65 | 0.049 | 490 | 601.2 | 3.5 | 604.7 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5576 | 27.7576 | 3.5
2008-12-02 | 12:14 | 4.657 | 0.049 | 490 | 601.3 | 3.5 | 604.8 | 380.4 | 380.3931 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2871 | 11.479 | 27.7 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.2 | 3.5 | 604.7 | 380.3 | 380.3069 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4005 | 27.7 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.2 | 3.5 | 604.7 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5561 | 27.7781 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.4 | 3.5 | 604.9 | 380.4 | 380.3927 | 0.1 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0838 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5434 | 27.7434 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.5 | 3.5 | 605 | 380.5 | 380.4733 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.083829 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.6 | 3.5 | 605.1 | 380.3 | 380.3096 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0839 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.6 | 3.4 | 605 | 380.5 | 380.494 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0839 | 0.0024 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.4
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.6 | 3.4 | 605 | 380.3 | 380.3108 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8711 | 0.0839 | 0.002429 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4292 | 27.7708 | 3.4
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.6 | 3.4 | 605 | 380.4 | 380.3939 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8712 | 0.0839 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4493 | 27.8 | 3.4
2008-12-02 | 12:14 | 4.66 | 0.049 | 490 | 601.7 | 3.4 | 605.1 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0839 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.8 | 3.4
2008-12-02 | 12:14 | 4.704 | 0.049 | 490 | 601.7 | 3.4 | 605.1 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0839 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.8 | 3.4
2008-12-02 | 12:14 | 5.343 | 0.052 | 520.0411 | 606.9 | 3.4 | 610.3 | 380.4 | 380.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.084434 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.8 | 3.4
2008-12-02 | 12:14 | 5.775 | 0.0619 | 618.9482 | 665 | 3.4 | 668.4 | 380.4 | 380.4089 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.089442 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7365 | 3.4
2008-12-02 | 12:14 | 5.899 | 0.0603 | 602.767 | 733.6 | 3.4 | 737 | 399 | 399.0116 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.097105 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5841 | 27.742 | 3.4
2008-12-02 | 12:14 | 5.89 | 0.0607 | 607.0433 | 746 | 3.5 | 749.5 | 455.1 | 455.0688 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.8292 | 0.098258 | 0.0025 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.6589 | 27.8 | 3.5
2008-12-02 | 12:14 | 5.89 | 0.061 | 610 | 748.7 | 3.5 | 752.2 | 469.2 | 469.2264 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0986 | 0.0026 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5016 | 27.8 | 3.5
2008-12-02 | 12:14 | 5.89 | 0.0606 | 606.1817 | 748.8 | 3.5 | 752.3 | 472.6 | 472.6281 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8293 | 0.0986 | 0.0026 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.7223 | 3.5
2008-12-02 | 12:14 | 5.89 | 0.06 | 600 | 749 | 3.5 | 752.5 | 473.3 | 473.2936 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.098629 | 0.0026 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.7 | 3.5
2008-12-02 | 12:14 | 5.883 | 0.06 | 600 | 749.1 | 3.5289 | 752.6 | 473.7 | 473.6932 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0987 | 0.0026 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2889 | 11.5348 | 27.7348 | 3.5
2008-12-02 | 12:14 | 5.88 | 0.06 | 600 | 748.9 | 4.5287 | 753.5 | 474.1 | 474.0937 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8713 | 0.0987 | 0.002629 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2869 | 11.6133 | 27.8 | 4.5
2008-12-02 | 12:14 | 5.869 | 0.06 | 600 | 749 | 4.6 | 753.6 | 474.3 | 474.2937 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8288 | 0.0987 | 0.0027 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7123 | 11.6918 | 27.8 | 4.6
2008-12-02 | 12:14 | 5.85 | 0.0605 | 605.4839 | 749.1 | 4.6 | 753.7 | 474.3 | 474.3031 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.0988 | 0.0027 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7 | 27.7296 | 4.6
2008-12-02 | 12:14 | 5.85 | 0.061 | 610 | 749.2 | 4.6 | 753.8 | 474.3 | 474.2951 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8294 | 0.0988 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.749 | 27.749 | 4.6
2008-12-02 | 12:14 | 5.85 | 0.061 | 610 | 749.3 | 4.5 | 753.8 | 474.5 | 474.4914 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0988 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8 | 27.8 | 4.5
2008-12-02 | 12:14 | 5.85 | 0.061 | 610 | 749.4 | 4.5 | 753.9 | 474 | 473.9542 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0988 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7939 | 27.8 | 4.5
2008-12-02 | 12:14 | 5.85 | 0.061 | 610 | 749.5 | 4.5 | 754 | 474.3 | 474.2876 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.098829 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7154 | 27.8 | 4.5
2008-12-02 | 12:14 | 5.854 | 0.061 | 610 | 749.6 | 4.4 | 754 | 474.5 | 474.4789 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8293 | 0.0989 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7632 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.862 | 0.061 | 610 | 749.7 | 4.4 | 754.1 | 474.5 | 474.4959 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8709 | 0.0989 | 0.0028 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7583 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.0603 | 602.85 | 749.8 | 4.4 | 754.2 | 474.6 | 474.5969 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8289 | 0.0989 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.059 | 590 | 749.9 | 4.4 | 754.3 | 474.4 | 474.4074 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.098929 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.059 | 590 | 750 | 4.4 | 754.4 | 474.6 | 474.5957 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8712 | 0.099 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.059 | 590 | 750.2 | 4.4 | 754.6 | 474.4 | 474.4186 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8287 | 0.099 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7559 | 27.8 | 4.4
2008-12-02 | 12:14 | 5.866 | 0.059 | 590 | 750.2 | 4.4 | 754.6 | 474.5 | 474.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.099 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7656 | 27.7656 | 4.4
2008-12-02 | 12:14 | 5.862 | 0.0595 | 595.3382 | 750.2 | 4.4 | 754.6 | 474.6 | 474.5909 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195 | 59.8706 | 0.099 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7261 | 27.713 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.06 | 600 | 750.2 | 4.4 | 754.6 | 474.6 | 474.6 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.099 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8832 | 27.7916 | 4.4
2008-12-02 | 12:14 | 5.87 | 0.06 | 600 | 750.2 | 4.4 | 754.6 | 474.6 | 474.5527 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0991 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7598 | 27.8 | 4.4
2008-12-02 | 12:15 | 5.87 | 0.06 | 600 | 750.3 | 4.4 | 754.7 | 474.4 | 474.4096 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0991 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.797 | 27.8 | 4.4
2008-12-02 | 12:15 | 5.87 | 0.06 | 600 | 750.4 | 4.3707 | 754.8 | 474.4 | 474.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.0991 | 0.0029 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.846 | 27.8 | 4.4
2008-12-02 | 12:15 | 5.87 | 0.06 | 600 | 750.5 | 4.3 | 754.8 | 474.6 | 474.5875 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8291 | 0.0991 | 0.002929 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7057 | 27.8 | 4.3
2008-12-02 | 12:15 | 5.87 | 0.06 | 600 | 750.6 | 4.3 | 754.9 | 474.5 | 474.5042 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7843 | 27.8 | 4.3
2008-12-02 | 12:15 | 5.869 | 0.06 | 600 | 750.6 | 4.3 | 754.9 | 474.5 | 474.4965 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7372 | 27.7372 | 4.3
2008-12-02 | 12:15 | 5.861 | 0.06 | 600 | 750.6 | 4.3 | 754.9 | 474.5 | 474.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7414 | 27.7 | 4.3
2008-12-02 | 12:15 | 5.853 | 0.06 | 600 | 750.7 | 4.3 | 755 | 474.4 | 474.4045 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.78 | 27.72 | 4.3
2008-12-02 | 12:15 | 5.85 | 0.06 | 600 | 750.7 | 4.2 | 754.9 | 474.6 | 474.5946 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8704 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7015 | 27.7985 | 4.2
2008-12-02 | 12:15 | 5.85 | 0.06 | 600 | 750.7 | 4.1706 | 754.9 | 474.4 | 474.4112 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8706 | 0.0991 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7771 | 27.8 | 4.2
2008-12-02 | 12:15 | 5.85 | 0.06 | 600 | 750.7 | 4.1 | 754.8 | 474.2 | 474.2 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7444 | 27.7444 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.06 | 600 | 750.8 | 4.1 | 754.9 | 474.4 | 474.3895 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7083 | 11.7682 | 27.7341 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.0602 | 602.2417 | 751 | 4.1 | 755.1 | 474.4 | 474.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8747 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751 | 4.1 | 755.1 | 474.4 | 474.3757 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8707 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7176 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751 | 4.1 | 755.1 | 474.4 | 474.3893 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7697 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751 | 4.1 | 755.1 | 474.5 | 474.506 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8285 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8482 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.0615 | 615.2903 | 751 | 4.1 | 755.1 | 474.5 | 474.5 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8465 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.0616 | 616.4696 | 751 | 4.1 | 755.1 | 474.5 | 474.5 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8702 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7054 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751 | 4 | 755 | 474.7 | 474.686 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8296 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7838 | 27.8 | 4
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751.1 | 4 | 755.1 | 474.5 | 474.5201 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8624 | 27.8 | 4
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751.2 | 4.0293 | 755.3 | 474.5 | 474.5095 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8707 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7069 | 11.8592 | 27.8 | 4
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751.2 | 4 | 755.2 | 474.6 | 474.6 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751.2 | 4 | 755.2 | 474.5 | 474.5066 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.871 | 0.0992 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4
2008-12-02 | 12:15 | 5.85 | 0.061 | 610 | 751.1 | 4.1 | 755.2 | 474.6 | 474.594 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.099229 | 0.003 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:15 | 5.859 | 0.061 | 610 | 751.3 | 4 | 755.3 | 474.5 | 474.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8551 | 27.8 | 4
2008-12-02 | 12:15 | 5.86 | 0.061 | 610 | 751.4 | 4 | 755.4 | 474.6 | 474.5939 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8711 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8666 | 27.8 | 4
2008-12-02 | 12:15 | 5.865 | 0.061 | 610 | 751.3 | 4 | 755.3 | 474.5 | 474.5 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7885 | 27.8 | 4
2008-12-02 | 12:15 | 5.87 | 0.061 | 610 | 751.4 | 4 | 755.4 | 474.6 | 474.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7123 | 11.7092 | 27.8 | 4
2008-12-02 | 12:15 | 5.87 | 0.061 | 610 | 751.6 | 4 | 755.6 | 474.7 | 474.6894 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8388 | 27.8 | 4
2008-12-02 | 12:15 | 5.87 | 0.061 | 610 | 751.6 | 4 | 755.6 | 474.6 | 474.6027 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.9 | 27.8 | 4
2008-12-02 | 12:15 | 5.861 | 0.061 | 610 | 751.5 | 4 | 755.5 | 474.4 | 474.4064 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0993 | 0.0031 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8735 | 27.8 | 4
2008-12-02 | 12:15 | 5.86 | 0.061 | 610 | 751.7 | 4 | 755.7 | 474.4 | 474.4052 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.4 | 59.9 | 0.099329 | 0.003129 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7083 | 11.8 | 27.8 | 4
2008-12-02 | 12:15 | 5.86 | 0.061 | 610 | 751.8 | 4 | 755.8 | 474.6 | 474.5937 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8705 | 0.0994 | 0.0032 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8 | 27.8 | 4
2008-12-02 | 12:15 | 5.86 | 0.061 | 610 | 751.8 | 4 | 755.8 | 474.6 | 474.585 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.099429 | 0.0032 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8 | 27.8 | 4
2008-12-02 | 12:15 | 5.862 | 0.0605 | 604.5194 | 751.8 | 4 | 755.8 | 474.6 | 474.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8709 | 0.0995 | 0.0032 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8406 | 27.8 | 4
2008-12-02 | 12:15 | 5.887 | 0.06 | 600 | 751.8 | 4 | 755.8 | 474.7 | 474.7275 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0995 | 0.0032 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8809 | 27.8 | 4
2008-12-02 | 12:15 | 6.187 | 0.0629 | 628.5897 | 752.2 | 4 | 756.2 | 474.6 | 474.6085 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0995 | 0.0032 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8023 | 27.8 | 4
2008-12-02 | 12:15 | 6.768 | 0.0749 | 748.7167 | 789.6 | 4 | 793.6 | 474.5 | 474.5042 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8712 | 0.10295 | 0.003229 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2877 | 11.8763 | 27.8 | 4
2008-12-02 | 12:15 | 7.048 | 0.0725 | 724.5036 | 870.8 | 4 | 874.8 | 478.1 | 478.0823 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.11022 | 0.0033 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8452 | 27.8 | 4
2008-12-02 | 12:15 | 7.093 | 0.072 | 720 | 895.5 | 4 | 899.5 | 539.2 | 539.1533 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.11243 | 0.0033 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7668 | 27.8 | 4
2008-12-02 | 12:15 | 7.078 | 0.072 | 720 | 896.2 | 4 | 900.2 | 563.8 | 563.8208 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1125 | 0.0034 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.7113 | 27.8 | 4
2008-12-02 | 12:15 | 7.07 | 0.0723 | 723.487 | 896.3 | 4 | 900.3 | 568.1 | 568.1407 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1125 | 0.0034 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7904 | 27.8 | 4
2008-12-02 | 12:15 | 7.07 | 0.0728 | 728.0858 | 897.2 | 4 | 901.2 | 569.2 | 569.1831 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1126 | 0.0034 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.869 | 27.8 | 4
2008-12-02 | 12:15 | 7.07 | 0.072 | 720 | 897.3 | 4.1 | 901.4 | 569.4 | 569.4152 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1127 | 0.0034 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.9 | 27.8 | 4.1
2008-12-02 | 12:15 | 7.07 | 0.072 | 720 | 897.3 | 4.2 | 901.5 | 569.6 | 569.6 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1127 | 0.0034 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8482 | 27.8 | 4.2
2008-12-02 | 12:15 | 7.07 | 0.072 | 720 | 897.6 | 4.2291 | 901.8 | 569.8 | 569.7881 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.11273 | 0.003429 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7078 | 27.8 | 4.2
2008-12-02 | 12:15 | 7.068 | 0.072 | 720 | 897.9 | 4.3 | 902.2 | 569.7 | 569.7 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8289 | 0.1128 | 0.0035 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.866 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.072 | 720 | 898.8 | 4.3 | 903.1 | 569.8 | 569.7939 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1129 | 0.0035 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7766 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.072 | 720 | 900 | 4.3 | 904.3 | 569.8 | 569.8 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.113 | 0.0035 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7806 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.0719 | 719.1577 | 900.5 | 4.3 | 904.8 | 570 | 569.9899 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.113 | 0.0035 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.9 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.0703 | 702.6424 | 900.5 | 4.3 | 904.8 | 569.8 | 569.8095 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.1131 | 0.003529 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.9 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.07 | 700 | 900.8 | 4.3 | 905.1 | 569.8 | 569.8 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.11313 | 0.0036 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8241 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.07 | 700 | 901 | 4.3 | 905.3 | 570 | 569.9845 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.1132 | 0.0036 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.07 | 700 | 900.9 | 4.3 | 905.2 | 570 | 570 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.829 | 0.1132 | 0.0036 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.07 | 700 | 901 | 4.3 | 905.3 | 570 | 569.9975 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8293 | 0.1132 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.07 | 700.4505 | 901.2 | 4.3 | 905.5 | 569.9 | 569.9 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1132 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7091 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.0708 | 708.4956 | 901.2 | 4.3 | 905.5 | 570 | 570 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8289 | 0.11323 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.0717 | 716.9726 | 901.4 | 4.3 | 905.7 | 569.8 | 569.8063 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8712 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.072 | 720 | 901.4 | 4.3 | 905.7 | 569.6 | 569.6044 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8288 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8256 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.06 | 0.072 | 720 | 901.4 | 4.3 | 905.7 | 569.4 | 569.4043 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8713 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8929 | 27.8 | 4.3
2008-12-02 | 12:15 | 7.068 | 0.072 | 720 | 901.4 | 4.3 | 905.7 | 569.4 | 569.396 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8296 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7036 | 11.7347 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 901.5 | 4.3 | 905.8 | 569.2 | 569.2049 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.7 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 901.7 | 4.3 | 906 | 569.5 | 569.4892 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7399 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 901.7 | 4.3 | 906 | 569.6 | 569.5946 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8708 | 0.1133 | 0.0037 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7083 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 901.8 | 4.3 | 906.1 | 569.7 | 569.6902 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.829 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 901.9 | 4.3 | 906.2 | 569.8 | 569.8 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8753 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.07 | 0.072 | 720 | 902 | 4.3 | 906.3 | 569.9 | 569.8957 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8291 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.9 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.077 | 0.072 | 720 | 902.1 | 4.3 | 906.4 | 569.9 | 569.9427 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8289 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8349 | 27.8 | 4.3
2008-12-02 | 12:16 | 7.08 | 0.072 | 720 | 902 | 4.2 | 906.2 | 570 | 570.0465 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7223 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.072 | 720 | 902 | 4.2 | 906.2 | 569.9 | 569.9 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8793 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.0725 | 725.2227 | 902 | 4.2 | 906.2 | 569.9 | 569.9066 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6954 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.073 | 730 | 902 | 4.2 | 906.2 | 569.9 | 569.9 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8704 | 0.1134 | 0.0038 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2964 | 11.6935 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.0728 | 728.1005 | 902 | 4.2 | 906.2 | 569.9 | 569.9 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.11343 | 0.003829 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.072 | 720 | 902.1 | 4.2 | 906.3 | 570 | 570 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.2
2008-12-02 | 12:16 | 7.08 | 0.072 | 720 | 902.2 | 4.1 | 906.3 | 569.9 | 569.908 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.084 | 0.0734 | 733.9551 | 902.2 | 4.1 | 906.3 | 570 | 569.9968 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.09 | 0.0735 | 734.7623 | 902.1 | 4.1 | 906.2 | 570.1 | 570.1077 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.09 | 0.073 | 730 | 902.2 | 4.1 | 906.3 | 570.1 | 570.1064 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7822 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.09 | 0.073 | 730 | 902.2 | 4.1 | 906.3 | 570 | 570.0031 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.704 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.083 | 0.073 | 730 | 902.3 | 4.1 | 906.4 | 570.2 | 570.1773 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8712 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2877 | 11.7 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.073 | 730 | 902.2 | 4.1 | 906.3 | 570 | 570.0064 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8072 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.073 | 730 | 902.4 | 4.1 | 906.5 | 570.1 | 570.0954 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8296 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.8679 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.072 | 719.5877 | 902.3 | 4.1 | 906.4 | 570 | 570.0061 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7055 | 11.8106 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.4 | 4.1 | 906.5 | 570.1 | 570.1 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2931 | 11.8887 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.3 | 4.1 | 906.4 | 570 | 570.0028 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8291 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.8325 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.087 | 0.071 | 710 | 902.3 | 4.1 | 906.4 | 570.2 | 570.1955 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.7537 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.09 | 0.071 | 710 | 902.4 | 4.1 | 906.5 | 570.1 | 570.0957 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.8293 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7071 | 11.7 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.087 | 0.071 | 710 | 902.4 | 4.1 | 906.5 | 570.1 | 570.1201 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8291 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7091 | 11.7033 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.5 | 4.1 | 906.6 | 570.1 | 570.1102 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8289 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7819 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.5 | 4.1 | 906.6 | 569.8 | 569.8166 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7395 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.6 | 4.1 | 906.7 | 570 | 570 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.661 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.7 | 4.1 | 906.8 | 570.1 | 570.0932 | 0.1 | 13.1 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6174 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.5 | 4.1 | 906.6 | 570 | 570.0047 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6957 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.08 | 0.071 | 710 | 902.6 | 4.1 | 906.7 | 570.1 | 570.1 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.6255 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.082 | 0.071 | 710 | 902.6 | 4.1 | 906.7 | 569.9 | 569.9094 | 0.1 | 13.1 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8292 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.7062 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.112 | 0.071 | 710 | 902.6 | 4.1 | 906.7 | 570 | 569.9979 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.871 | 0.1135 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.705 | 27.8 | 4.1
2008-12-02 | 12:16 | 7.828 | 0.0732 | 731.9165 | 906.5 | 4.1 | 910.6 | 570 | 570 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.11385 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053.5284 | 1014 | 11.5203 | 27.8102 | 4.1
2008-12-02 | 12:16 | 8.234 | 0.083 | 829.9833 | 960.3 | 4.1 | 964.4 | 569.9 | 569.9062 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8291 | 0.11756 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6767 | 27.8883 | 4.1
2008-12-02 | 12:16 | 8.35 | 0.0822 | 821.657 | 1031.7 | 4.1 | 1035.8 | 581 | 580.9938 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.12414 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8328 | 4.1
2008-12-02 | 12:16 | 8.336 | 0.082 | 820 | 1043.3 | 4.1 | 1047.4 | 639.7 | 639.7485 | 0.1 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8287 | 0.12503 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2869 | 11.7 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.082 | 820 | 1045.7 | 4.1 | 1049.8 | 662.3 | 662.2885 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.12523 | 0.004 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6756 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.082 | 820 | 1045.7 | 4.1 | 1049.8 | 665 | 664.9679 | 0.1 | 13.2 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.1253 | 0.004 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.0822 | 821.8624 | 1046.1 | 4.1 | 1050.2 | 665.7 | 665.6975 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.12533 | 0.004 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.083 | 830 | 1048.2 | 4.1 | 1052.3 | 665.9 | 665.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8706 | 0.12553 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.66 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.083 | 830 | 1048.7 | 4.1 | 1052.8 | 666 | 665.9941 | 0.1 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8707 | 0.1256 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7385 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.083 | 830 | 1048.8 | 4.1 | 1052.9 | 666 | 666 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195.5 | 59.8 | 0.1256 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.083 | 830 | 1049.8 | 4.1294 | 1053.9 | 666.2 | 666.1876 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8706 | 0.12573 | 0.004129 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.1
2008-12-02 | 12:16 | 8.31 | 0.0833 | 833.3985 | 1050.3 | 4.2 | 1054.5 | 666.1 | 666.1 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8291 | 0.1258 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5778 | 27.8741 | 4.2
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1050.6 | 4.2289 | 1054.8 | 666.1 | 666.1 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8289 | 0.12583 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5527 | 27.9 | 4.2
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1050.9 | 4.3 | 1055.2 | 666.1 | 666.1 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8712 | 0.1259 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6625 | 27.8688 | 4.3
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1050.9 | 4.3 | 1055.2 | 666.3 | 666.2913 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8288 | 0.1259 | 0.004229 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8097 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1051.1 | 4.3 | 1055.4 | 666.3 | 666.3 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8287 | 0.1259 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.888 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1051.2 | 4.3 | 1055.5 | 666.3 | 666.276 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1259 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7666 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.31 | 0.084 | 840 | 1051.2 | 4.3 | 1055.5 | 666.3 | 666.2888 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8301 | 0.12593 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6547 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.302 | 0.084 | 840 | 1051.2 | 4.3 | 1055.5 | 666.3 | 666.3 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.126 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6477 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.307 | 0.084 | 840 | 1051.5 | 4.3 | 1055.8 | 666.4 | 666.3907 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.126 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.295 | 0.084 | 840 | 1051.7 | 4.3 | 1056 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.12603 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.287 | 0.084 | 840 | 1051.7 | 4.3 | 1056 | 666.3 | 666.3021 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.3 | 0.084 | 840 | 1051.8 | 4.3 | 1056.1 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8291 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.8 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.3 | 0.084 | 840 | 1051.8 | 4.3 | 1056.1 | 666.5 | 666.4955 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7834 | 27.8 | 4.3
2008-12-02 | 12:16 | 8.3 | 0.084 | 840 | 1051.9 | 4.3 | 1056.2 | 666.4 | 666.4 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8288 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7049 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1051.9 | 4.3 | 1056.2 | 666.4 | 666.4 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1051.9 | 4.3 | 1056.2 | 666.5 | 666.4911 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8287 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0847 | 846.6277 | 1052 | 4.3 | 1056.3 | 666.4 | 666.3749 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8704 | 0.1261 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.1 | 4.3 | 1056.4 | 666.5 | 666.4864 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8706 | 0.1262 | 0.004429 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7092 | 27.8092 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.3 | 4.3 | 1056.6 | 666.5 | 666.5 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8707 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2931 | 11.7877 | 27.8877 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.2 | 4.3 | 1056.5 | 666.4 | 666.4047 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195.8 | 59.8292 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6674 | 27.8337 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.2 | 4.3 | 1056.5 | 666.4 | 666.4 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.829 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6448 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.2 | 4.3 | 1056.5 | 666.4 | 666.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8293 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7233 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.3 | 4.3 | 1056.6 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7103 | 11.798 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 850 | 1052.4 | 4.3 | 1056.7 | 666.4 | 666.4 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.1262 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.6957 | 11.7195 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0847 | 846.6354 | 1052.5 | 4.3 | 1056.8 | 666.4 | 666.4 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.12623 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2883 | 11.582 | 27.859 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0842 | 841.8115 | 1052.6 | 4.3 | 1056.9 | 666.5 | 666.4913 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8712 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5375 | 27.8625 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.085 | 849.935 | 1052.6 | 4.3 | 1056.9 | 666.4 | 666.4023 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8713 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0842 | 841.7333 | 1052.5 | 4.3 | 1056.8 | 666.4 | 666.427 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0847 | 846.8394 | 1052.5 | 4.3 | 1056.8 | 666.5 | 666.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8293 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.745 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.0834 | 834.4226 | 1052.5 | 4.3 | 1056.8 | 666.4 | 666.4038 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8708 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6463 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.082 | 820 | 1052.6 | 4.3 | 1056.9 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8295 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5603 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.292 | 0.0822 | 821.8708 | 1052.8 | 4.3 | 1057.1 | 666.6 | 666.5959 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.1263 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7071 | 11.7 | 27.8086 | 4.3
2008-12-02 | 12:17 | 8.29 | 0.083 | 830 | 1052.8 | 4.3 | 1057.1 | 666.6 | 666.5934 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8709 | 0.1263 | 0.004529 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8868 | 4.3
2008-12-02 | 12:17 | 8.29 | 0.083 | 830 | 1052.9 | 4.2711 | 1057.2 | 666.5 | 666.5046 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8711 | 0.1263 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2889 | 11.7 | 27.8344 | 4.3
2008-12-02 | 12:17 | 8.294 | 0.083 | 830 | 1053 | 4.2 | 1057.2 | 666.4 | 666.4041 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.3 | 59.8288 | 0.1263 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6114 | 27.8443 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1053 | 4.2 | 1057.2 | 666.6 | 666.5776 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1263 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5685 | 27.8772 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.6 | 666.5934 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.12633 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7973 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.4 | 666.4135 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6403 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5416 | 27.8584 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.5 | 666.4894 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5738 | 27.8631 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.5 | 666.5109 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8708 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6845 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.083 | 830 | 1052.9 | 4.2 | 1057.1 | 666.6 | 666.5927 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.606 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.0833 | 833.3419 | 1053.1 | 4.2 | 1057.3 | 666.4 | 666.4065 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8707 | 0.1264 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6721 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.3 | 4.2 | 1057.5 | 666.5 | 666.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.1264 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2889 | 11.7 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.2 | 4.2 | 1057.4 | 666.5 | 666.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8288 | 0.1264 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.7 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.1 | 4.2288 | 1057.4 | 666.5 | 666.5 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.1264 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6837 | 27.8 | 4.2
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.3 | 4.3 | 1057.6 | 666.4 | 666.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.1264 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5267 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.5 | 4.3 | 1057.8 | 666.5 | 666.493 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1265 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6296 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.3 | 0.084 | 840 | 1053.5 | 4.3 | 1057.8 | 666.5 | 666.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.1265 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2945 | 11.6132 | 27.8 | 4.3
2008-12-02 | 12:17 | 8.998 | 0.0861 | 860.51 | 1053.5 | 4.3 | 1057.8 | 666.5 | 666.4537 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.1265 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7069 | 11.5223 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.397 | 0.095 | 949.5574 | 1110.7 | 4.3 | 1115 | 666.2 | 666.2056 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1308 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.52 | 0.0933 | 933.1639 | 1187.5 | 4.3 | 1191.8 | 679.7 | 679.7336 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.13546 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.52 | 0.093 | 930 | 1194.9 | 4.3 | 1199.2 | 738 | 738.0133 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.13696 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5421 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.52 | 0.093 | 930 | 1196.9 | 4.3 | 1201.2 | 757.6 | 757.5548 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.13713 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.52 | 0.0933 | 933.3247 | 1197.7 | 4.3 | 1202 | 760.9 | 760.9438 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8289 | 0.1372 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5149 | 27.8 | 4.3
2008-12-02 | 12:17 | 9.503 | 0.094 | 940 | 1197.8 | 4.6306 | 1202.5 | 762 | 761.9694 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1372 | 0.0048 | 2.9e-05 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5935 | 27.8 | 4.6
2008-12-02 | 12:17 | 9.5 | 0.094 | 940 | 1198.1 | 5.2 | 1203.3 | 762.3 | 762.2915 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.7 | 28 | 5 | 195 | 59.8288 | 0.1372 | 0.0048 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6719 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.0965 | 964.6205 | 1198.5 | 5.2 | 1203.7 | 762.4 | 762.4055 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.3 | 28 | 5 | 195 | 59.8298 | 0.1373 | 0.00483 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7 | 27.8505 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.097 | 970 | 1198.8 | 5.2 | 1204 | 762.5 | 762.5 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.7 | 28 | 5 | 195 | 59.8296 | 0.13733 | 0.0049 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.729 | 27.871 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.097 | 970 | 1199 | 5.2 | 1204.2 | 762.6 | 762.5946 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.3 | 28 | 5 | 195 | 59.8293 | 0.1374 | 0.0049 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7925 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.0967 | 966.6867 | 1199.1 | 5.2 | 1204.3 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1374 | 0.0049 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7139 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1199.5 | 5.2 | 1204.7 | 762.7 | 762.6936 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1375 | 0.005 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5711 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1199.6 | 5.2 | 1204.8 | 762.7 | 762.6957 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1375 | 0.005 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.543 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1199.8 | 5.2 | 1205 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.1375 | 0.005 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6217 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1200.1 | 5.2 | 1205.3 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.1376 | 0.0051 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6995 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.0955 | 955.1333 | 1200.3 | 5.2 | 1205.5 | 762.6 | 762.6031 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1377 | 0.0051 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5424 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.0953 | 953.3189 | 1200.3 | 5.2 | 1205.5 | 762.6 | 762.6046 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1377 | 0.0051 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5571 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1200.3 | 5.2 | 1205.5 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1377 | 0.0051 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6712 | 27.8356 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1200.2 | 5.2 | 1205.4 | 762.8 | 762.7983 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1377 | 0.0051 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7857 | 27.8857 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1200.7 | 5.2 | 1205.9 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.1377 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7071 | 27.8071 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1200.8 | 5.1707 | 1206 | 762.7 | 762.7 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.1377 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6287 | 27.8 | 5.2
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1201 | 5.1 | 1206.1 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8292 | 0.1377 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6498 | 27.8 | 5.1
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1201 | 5.1 | 1206.1 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8295 | 0.1377 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.7 | 27.8 | 5.1
2008-12-02 | 12:17 | 9.5 | 0.096 | 960 | 1201.1 | 5 | 1206.1 | 762.8 | 762.7934 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.686 | 27.8 | 5
2008-12-02 | 12:18 | 9.513 | 0.096 | 960 | 1201.2 | 5 | 1206.2 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8291 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2909 | 11.529 | 27.8 | 5
2008-12-02 | 12:18 | 9.52 | 0.096 | 960 | 1201.1 | 5 | 1206.1 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8713 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.52 | 0.0953 | 953.4121 | 1201.2 | 5 | 1206.2 | 762.8 | 762.7909 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5425 | 27.8 | 5
2008-12-02 | 12:18 | 9.511 | 0.095 | 950 | 1201.3 | 5 | 1206.3 | 762.7 | 762.7032 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6 | 27.8 | 5
2008-12-02 | 12:18 | 9.471 | 0.095 | 950 | 1201.3 | 5 | 1206.3 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.1378 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6 | 27.8 | 5
2008-12-02 | 12:18 | 9.478 | 0.095 | 950 | 1201.5 | 5 | 1206.5 | 762.7 | 762.7119 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8294 | 0.13783 | 0.0052 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7055 | 11.5219 | 27.8781 | 5
2008-12-02 | 12:18 | 9.473 | 0.095 | 950 | 1202 | 5 | 1207 | 762.8 | 762.8 | 0.1 | 13.1 | 916 | 953 | 942 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.1379 | 0.0053 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2931 | 11.6128 | 27.8436 | 5
2008-12-02 | 12:18 | 9.47 | 0.095 | 950 | 1202.1 | 5 | 1207.1 | 762.8 | 762.7923 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8708 | 0.13793 | 0.0053 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6302 | 27.8 | 5
2008-12-02 | 12:18 | 9.473 | 0.095 | 950 | 1202.3 | 5 | 1207.3 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.138 | 0.0053 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2903 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.48 | 0.095 | 950 | 1202.4 | 5 | 1207.4 | 762.6 | 762.61 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.138 | 0.0053 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7071 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.48 | 0.095 | 950 | 1202.5 | 5 | 1207.5 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.138 | 0.0053 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5708 | 27.8 | 5
2008-12-02 | 12:18 | 9.48 | 0.095 | 950 | 1202.6 | 5 | 1207.6 | 762.7 | 762.6942 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8289 | 0.138 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5511 | 27.8 | 5
2008-12-02 | 12:18 | 9.48 | 0.095 | 950 | 1202.7 | 5 | 1207.7 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.138 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.485 | 0.095 | 950 | 1202.6 | 5 | 1207.6 | 762.7 | 762.7062 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.13803 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7123 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.493 | 0.095 | 950 | 1202.8 | 5 | 1207.8 | 762.8 | 762.7953 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1381 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 9.5 | 0.095 | 950 | 1203 | 5 | 1208 | 762.7 | 762.7034 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1381 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8633 | 5
2008-12-02 | 12:18 | 9.5 | 0.095 | 950 | 1202.9 | 5 | 1207.9 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.1381 | 0.0054 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4584 | 27.9 | 5
2008-12-02 | 12:18 | 9.508 | 0.095 | 950 | 1203.2 | 5 | 1208.2 | 762.8 | 762.7936 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195.4 | 59.9 | 0.1381 | 0.005429 | 7.1e-05 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4198 | 27.8802 | 5
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.4 | 4.9 | 1208.3 | 762.7 | 762.7198 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4989 | 27.8011 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.4 | 4.9 | 1208.3 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5775 | 27.8 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.7 | 4.9 | 1208.6 | 762.7 | 762.7 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8293 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.544 | 27.8 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.4 | 4.9 | 1208.3 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8709 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5345 | 27.8 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.4 | 4.9 | 1208.3 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.587 | 27.8 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.3 | 4.9 | 1208.2 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5084 | 27.8 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.3 | 4.9 | 1208.2 | 762.6 | 762.6 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8701 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.3 | 4.8713 | 1208.2 | 762.5 | 762.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 4.9
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.4 | 4.8 | 1208.2 | 762.5 | 762.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4732 | 27.9 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.5 | 4.8 | 1208.3 | 762.6 | 762.6 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4051 | 27.8949 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.6 | 4.8 | 1208.4 | 762.5 | 762.5051 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8293 | 0.1382 | 0.0055 | 2.9e-05 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4841 | 27.8159 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.6 | 4.8 | 1208.4 | 762.6 | 762.5894 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8295 | 0.1382 | 0.0055 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4373 | 27.8627 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.6 | 4.8 | 1208.4 | 762.5 | 762.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.8589 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.6 | 4.8 | 1208.4 | 762.5 | 762.5 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8709 | 0.1382 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4195 | 27.8 | 4.8
2008-12-02 | 12:18 | 9.507 | 0.095 | 950 | 1203.8 | 4.8 | 1208.6 | 762.5 | 762.5 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8289 | 0.1383 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4983 | 27.8 | 4.8
2008-12-02 | 12:18 | 9.5 | 0.095 | 950 | 1203.8 | 4.8 | 1208.6 | 762.7 | 762.6893 | 0.1 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8288 | 0.1383 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 4.8
2008-12-02 | 12:18 | 9.51 | 0.095 | 950 | 1203.8 | 4.8 | 1208.6 | 762.6 | 762.6065 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.1383 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4447 | 27.8 | 4.8
2008-12-02 | 12:18 | 10.03 | 0.0968 | 967.9167 | 1204 | 4.8 | 1208.8 | 762.5 | 762.5033 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8287 | 0.1383 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4338 | 27.8 | 4.8
2008-12-02 | 12:18 | 10.504 | 0.107 | 1069.6797 | 1232.4 | 4.8 | 1237.2 | 762.5 | 762.5 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 196 | 59.9 | 0.14027 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5121 | 27.8 | 4.8
2008-12-02 | 12:18 | 10.7 | 0.1054 | 1053.6669 | 1316.9 | 4.8 | 1321.7 | 771.1 | 771.0667 | 0.1 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.1459 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5909 | 27.8 | 4.8
2008-12-02 | 12:18 | 10.7 | 0.105 | 1050 | 1343.4 | 4.8 | 1348.2 | 830.8 | 830.8488 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.14763 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8694 | 4.8
2008-12-02 | 12:18 | 10.699 | 0.105 | 1050 | 1345.9 | 4.9 | 1350.8 | 853 | 853.0316 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.14783 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5521 | 27.8521 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.105 | 1050 | 1346.7 | 4.9 | 1351.6 | 857.2 | 857.2362 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8705 | 0.1479 | 0.005629 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4206 | 27.8 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.1051 | 1051.4751 | 1346.8 | 4.9 | 1351.7 | 858.3 | 858.2502 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8707 | 0.14793 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2929 | 11.2151 | 27.8 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1347.1 | 4.9 | 1352 | 858.5 | 858.4957 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8291 | 0.148 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7091 | 11.4505 | 27.8 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1347.7 | 4.9 | 1352.6 | 858.7 | 858.689 | 0.1 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8289 | 0.1481 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.562 | 27.8 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1348.4 | 4.9288 | 1353.4 | 858.7 | 858.7 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.14813 | 0.0058 | 2.9e-05 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5595 | 27.8 | 4.9
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1348.6 | 5 | 1353.6 | 858.8 | 858.8 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8702 | 0.1482 | 0.0058 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.481 | 27.819 | 5
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1349.5 | 5 | 1354.5 | 858.8 | 858.8 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1482 | 0.0058 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4024 | 27.8976 | 5
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1349.8 | 5 | 1354.8 | 858.8 | 858.7981 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8706 | 0.1483 | 0.0058 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4761 | 27.9 | 5
2008-12-02 | 12:18 | 10.69 | 0.106 | 1060 | 1350 | 5 | 1355 | 858.9 | 858.9 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.14833 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2931 | 11.5546 | 27.8454 | 5
2008-12-02 | 12:18 | 10.695 | 0.106 | 1060 | 1350.1 | 5 | 1355.1 | 858.9 | 858.9 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1484 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5671 | 27.8329 | 5
2008-12-02 | 12:18 | 10.7 | 0.106 | 1060 | 1350.1 | 5 | 1355.1 | 858.9 | 858.9 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.871 | 0.1484 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8888 | 5
2008-12-02 | 12:18 | 10.7 | 0.106 | 1060 | 1350.4 | 5 | 1355.4 | 859 | 858.9949 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8707 | 0.1484 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8099 | 5
2008-12-02 | 12:18 | 10.7 | 0.106 | 1060 | 1350.2 | 5 | 1355.2 | 859.1 | 859.1 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.1484 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5687 | 27.8 | 5
2008-12-02 | 12:18 | 10.692 | 0.106 | 1060 | 1350.2 | 5 | 1355.2 | 859.2 | 859.1958 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.1484 | 0.0059 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5056 | 27.8472 | 5
2008-12-02 | 12:18 | 10.683 | 0.106 | 1060 | 1350.5 | 5 | 1355.5 | 859.1 | 859.1085 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.14843 | 0.005929 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4512 | 27.8744 | 5
2008-12-02 | 12:18 | 10.68 | 0.106 | 1060 | 1351.2 | 5 | 1356.2 | 859.1 | 859.1 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8712 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5964 | 27.8 | 5
2008-12-02 | 12:18 | 10.68 | 0.106 | 1060 | 1351.4 | 5 | 1356.4 | 859.2 | 859.2 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8713 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5174 | 27.8 | 5
2008-12-02 | 12:18 | 10.68 | 0.1065 | 1064.7373 | 1351.4 | 5 | 1356.4 | 859.2 | 859.2 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8 | 5
2008-12-02 | 12:18 | 10.68 | 0.107 | 1070 | 1351.5 | 5 | 1356.5 | 859.2 | 859.2 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8294 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4602 | 27.8398 | 5
2008-12-02 | 12:18 | 10.68 | 0.107 | 1070 | 1351.5 | 5 | 1356.5 | 859.2 | 859.2 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8707 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4182 | 27.8818 | 5
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070.1122 | 1351.3 | 5 | 1356.3 | 859.2 | 859.2 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4963 | 27.8037 | 5
2008-12-02 | 12:19 | 10.68 | 0.1078 | 1078.125 | 1351.3 | 5 | 1356.3 | 859.1 | 859.1079 | 0.1 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.1485 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8753 | 5
2008-12-02 | 12:19 | 10.68 | 0.108 | 1080 | 1351.4 | 5 | 1356.4 | 859.2 | 859.1948 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8709 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8461 | 5
2008-12-02 | 12:19 | 10.681 | 0.108 | 1080 | 1351.6 | 5 | 1356.6 | 859.2 | 859.2245 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8289 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7111 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.69 | 0.108 | 1080 | 1351.6 | 5 | 1356.6 | 859.4 | 859.3917 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8712 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2883 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.69 | 0.108 | 1080 | 1351.5 | 5 | 1356.5 | 859.4 | 859.4 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.69 | 0.108 | 1079.9168 | 1351.6 | 5 | 1356.6 | 859.4 | 859.4 | 0.1 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.695 | 0.1072 | 1071.5973 | 1351.7 | 5 | 1356.7 | 859.3 | 859.3034 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8296 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.691 | 0.107 | 1070 | 1351.7 | 5 | 1356.7 | 859.4 | 859.389 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1486 | 0.006 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 5
2008-12-02 | 12:19 | 10.67 | 0.107 | 1070 | 1352 | 5 | 1357 | 859.2 | 859.2021 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.1486 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8033 | 5
2008-12-02 | 12:19 | 10.67 | 0.107 | 1070 | 1352 | 4.9 | 1356.9 | 859.3 | 859.2903 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1486 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.882 | 4.9
2008-12-02 | 12:19 | 10.678 | 0.107 | 1070 | 1352 | 4.9 | 1356.9 | 859.3 | 859.3 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.8705 | 0.1486 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.1 | 4.9 | 1357 | 859.2 | 859.2064 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1486 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.2 | 4.9 | 1357.1 | 859.3 | 859.2948 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8824 | 4.9
2008-12-02 | 12:19 | 10.677 | 0.107 | 1070 | 1352.2 | 4.9 | 1357.1 | 859.3 | 859.3 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.804 | 4.9
2008-12-02 | 12:19 | 10.67 | 0.107 | 1070 | 1352.1 | 4.9 | 1357 | 859.3 | 859.3 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8 | 4.9
2008-12-02 | 12:19 | 10.67 | 0.107 | 1070 | 1352.1 | 4.9 | 1357 | 859.1 | 859.1064 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4469 | 27.8 | 4.9
2008-12-02 | 12:19 | 10.678 | 0.107 | 1070 | 1352 | 4.9 | 1356.9 | 859.3 | 859.2958 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.8 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352 | 4.9 | 1356.9 | 859.3 | 859.2909 | 0.1 | 13.1 | 914 | 953 | 942 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4102 | 27.8 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.1 | 4.9 | 1357 | 859.2 | 859.2 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4887 | 27.8 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.2 | 4.9 | 1357.1 | 859.2 | 859.2 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5672 | 27.8672 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.3 | 4.9 | 1357.2 | 859.2 | 859.2 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5542 | 27.9 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.5 | 4.9 | 1357.4 | 859.2 | 859.2044 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4757 | 27.8757 | 4.9
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.5 | 4.8 | 1357.3 | 859.3 | 859.3 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4 | 27.8028 | 4.8
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.5 | 4.8 | 1357.3 | 859.3 | 859.2512 | 0.1 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4 | 27.8807 | 4.8
2008-12-02 | 12:19 | 10.68 | 0.107 | 1070 | 1352.6 | 4.8 | 1357.4 | 859.2 | 859.2 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4595 | 27.9 | 4.8
2008-12-02 | 12:19 | 10.674 | 0.1083 | 1083.0915 | 1353 | 4.8 | 1357.8 | 859.2 | 859.2 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4617 | 27.9 | 4.8
2008-12-02 | 12:19 | 10.66 | 0.1085 | 1085.2038 | 1352.8 | 4.8 | 1357.6 | 859.1 | 859.1021 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4169 | 27.9 | 4.8
2008-12-02 | 12:19 | 10.66 | 0.1083 | 1082.9083 | 1353.1 | 4.8 | 1357.9 | 859 | 859.0102 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4954 | 27.9 | 4.8
2008-12-02 | 12:19 | 10.668 | 0.109 | 1090 | 1353.1 | 4.8 | 1357.9 | 859.1 | 859.0934 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8261 | 4.8
2008-12-02 | 12:19 | 10.657 | 0.109 | 1090 | 1352.9 | 4.8 | 1357.7 | 859.1 | 859.1 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8707 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.8524 | 4.8
2008-12-02 | 12:19 | 10.659 | 0.109 | 1090 | 1352.9 | 4.8 | 1357.7 | 859.1 | 859.1 | 0.1 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8708 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.869 | 4.8
2008-12-02 | 12:19 | 10.67 | 0.1084 | 1083.6464 | 1352.9 | 4.8 | 1357.7 | 859.1 | 859.1 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4905 | 27.8 | 4.8
2008-12-02 | 12:19 | 10.673 | 0.1075 | 1075.3207 | 1352.8 | 4.8 | 1357.6 | 859.1 | 859.1 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4124 | 27.8 | 4.8
2008-12-02 | 12:19 | 10.69 | 0.107 | 1070 | 1352.6 | 4.8 | 1357.4 | 859.1 | 859.1 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1487 | 0.0061 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.8664 | 4.8
2008-12-02 | 12:19 | 11.135 | 0.1088 | 1088.2143 | 1352.6 | 4.8 | 1357.4 | 859.3 | 859.2875 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.14873 | 0.006129 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4451 | 27.9 | 4.8
2008-12-02 | 12:19 | 11.643 | 0.1186 | 1185.6169 | 1382.9 | 4.8 | 1387.7 | 859.3 | 859.3251 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.15065 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5235 | 27.9 | 4.8
2008-12-02 | 12:19 | 11.87 | 0.1174 | 1173.577 | 1468.7 | 4.8 | 1473.5 | 866.4 | 866.4311 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.15593 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.598 | 27.9 | 4.8
2008-12-02 | 12:19 | 11.87 | 0.1176 | 1176.4706 | 1492.6 | 4.8 | 1497.4 | 928.5 | 928.4529 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8703 | 0.15736 | 0.00623 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5197 | 27.9 | 4.8
2008-12-02 | 12:19 | 11.87 | 0.118 | 1180 | 1497.4 | 4.8 | 1502.2 | 951.7 | 951.7135 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8704 | 0.1577 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 4.8
2008-12-02 | 12:19 | 11.87 | 0.1183 | 1182.8941 | 1497.9 | 5.8 | 1503.7 | 955.4 | 955.358 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.1577 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 5.8
2008-12-02 | 12:19 | 11.862 | 0.119 | 1190 | 1497.9 | 5.8 | 1503.7 | 956.6 | 956.579 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1577 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 5.8
2008-12-02 | 12:19 | 11.86 | 0.119 | 1190 | 1498.1 | 5.8 | 1503.9 | 956.6 | 956.6069 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1577 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 5.8
2008-12-02 | 12:19 | 11.86 | 0.119 | 1190 | 1498.3 | 5.8 | 1504.1 | 956.8 | 956.7959 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1577 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.646 | 27.827 | 5.8
2008-12-02 | 12:19 | 11.86 | 0.119 | 1190 | 1498.9 | 5.8 | 1504.7 | 956.8 | 956.8063 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1578 | 0.006329 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6484 | 27.8516 | 5.8
2008-12-02 | 12:19 | 11.861 | 0.119 | 1190 | 1499 | 5.8 | 1504.8 | 956.9 | 956.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.1578 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2889 | 11.5698 | 27.9 | 5.8
2008-12-02 | 12:19 | 11.87 | 0.119 | 1190 | 1499 | 5.8 | 1504.8 | 957 | 956.9969 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1578 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8913 | 5.8
2008-12-02 | 12:19 | 11.878 | 0.119 | 1190 | 1499 | 5.8 | 1504.8 | 957 | 957.0249 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1578 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.8128 | 5.8
2008-12-02 | 12:19 | 11.88 | 0.119 | 1190 | 1499.1 | 5.7702 | 1504.9 | 957.1 | 957.1 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8702 | 0.1578 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6315 | 27.8657 | 5.8
2008-12-02 | 12:19 | 11.88 | 0.119 | 1190 | 1499.4 | 5.7 | 1505.1 | 957.2 | 957.1944 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1579 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6115 | 27.9 | 5.7
2008-12-02 | 12:19 | 11.874 | 0.119 | 1190 | 1499.7 | 5.7 | 1505.4 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.1579 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5228 | 27.8772 | 5.7
2008-12-02 | 12:19 | 11.86 | 0.1195 | 1194.6487 | 1500.1 | 5.6707 | 1505.8 | 957.1 | 957.1072 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.158 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5987 | 27.8 | 5.7
2008-12-02 | 12:19 | 11.86 | 0.12 | 1200 | 1500.2 | 5.6 | 1505.8 | 957.2 | 957.1945 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.158 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5201 | 27.8 | 5.6
2008-12-02 | 12:19 | 11.86 | 0.12 | 1200 | 1500.5 | 5.6 | 1506.1 | 957.1 | 957.1067 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.158 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4416 | 27.8584 | 5.6
2008-12-02 | 12:19 | 11.86 | 0.12 | 1200 | 1500.9 | 5.5707 | 1506.4 | 957.3 | 957.3 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.158 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4369 | 27.9 | 5.6
2008-12-02 | 12:19 | 11.837 | 0.12 | 1200 | 1501 | 5.5 | 1506.5 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.158 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:19 | 11.813 | 0.12 | 1200 | 1501.3 | 5.5 | 1506.8 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8289 | 0.1581 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:19 | 11.82 | 0.12 | 1200 | 1501.6 | 5.5288 | 1507.1 | 957.4 | 957.3936 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8712 | 0.15813 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.428 | 27.9 | 5.5
2008-12-02 | 12:19 | 11.82 | 0.12 | 1200 | 1501.9 | 5.5 | 1507.4 | 957.3 | 957.3108 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.4507 | 27.9 | 5.5
2008-12-02 | 12:19 | 11.82 | 0.12 | 1200 | 1501.9 | 5.5 | 1507.4 | 957.2 | 957.2044 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8716 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.5294 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.839 | 0.12 | 1200 | 1502 | 5.5 | 1507.5 | 957.3 | 957.2934 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8296 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8921 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.1192 | 1192.0626 | 1502.1 | 5.5 | 1507.6 | 957.2 | 957.2023 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8135 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.2 | 5.5 | 1507.7 | 957.2 | 957.1891 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8292 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.6 | 27.8 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.1 | 5.5 | 1507.6 | 957.3 | 957.2935 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.871 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5568 | 27.8432 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.2 | 5.5 | 1507.7 | 957.2 | 957.2043 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 196 | 59.9 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.898 | 0.119 | 1190 | 1502.2 | 5.5 | 1507.7 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.91 | 0.119 | 1190 | 1502.2 | 5.5 | 1507.7 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1582 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.91 | 0.119 | 1190 | 1502.1 | 5.5 | 1507.6 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8288 | 0.15823 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5573 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.891 | 0.119 | 1190 | 1502.2 | 5.5 | 1507.7 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8288 | 0.1583 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5642 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.3 | 5.5 | 1507.8 | 957.3 | 957.2861 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1583 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4854 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.4 | 5.5 | 1507.9 | 957.2 | 957.2024 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4069 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.7 | 5.5 | 1508.2 | 957.2 | 957.2 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4717 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.8 | 5.5 | 1508.3 | 957.1 | 957.1081 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8293 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.7069 | 11.5 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.9 | 5.5 | 1508.4 | 957.2 | 957.1943 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.2905 | 11.4713 | 27.8713 | 5.5
2008-12-02 | 12:20 | 11.853 | 0.119 | 1190 | 1502.8 | 5.5 | 1508.3 | 957.2 | 957.2 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.829 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4071 | 27.8071 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1502.9 | 5.5 | 1508.4 | 957.2 | 957.1955 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4857 | 27.8857 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1502.7 | 5.5 | 1508.2 | 957.1 | 957.1061 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.1584 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4358 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.852 | 0.119 | 1190 | 1502.9 | 5.5 | 1508.4 | 957.1 | 957.1076 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1584 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.85 | 0.119 | 1190 | 1502.9 | 5.5 | 1508.4 | 957.3 | 957.2911 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.1584 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.855 | 0.119 | 1190 | 1502.9 | 5.5 | 1508.4 | 957.3 | 957.3 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.1584 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1503.1 | 5.5 | 1508.6 | 957.1 | 957.1 | 0.1 | 13.2 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.1584 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3217 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1503.1 | 5.5 | 1508.6 | 957.2 | 957.1934 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.1584 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3568 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1503.1 | 5.5 | 1508.6 | 957.3 | 957.2974 | 0.1 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8294 | 0.15843 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3646 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1503.2 | 5.5 | 1508.7 | 957.3 | 957.3 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8293 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3556 | 27.8861 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.119 | 1190 | 1503.4 | 5.5 | 1508.9 | 957.3 | 957.2902 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.6695 | 27.8076 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.1198 | 1197.914 | 1503.4 | 5.5 | 1508.9 | 957.4 | 957.3934 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195 | 59.9 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4165 | 27.8709 | 5.5
2008-12-02 | 12:20 | 11.863 | 0.12 | 1200 | 1503.5 | 5.5 | 1509 | 957.4 | 957.4 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8293 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.87 | 0.12 | 1200 | 1503.8 | 5.5 | 1509.3 | 957.5 | 957.498 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3547 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.87 | 0.12 | 1200 | 1503.7 | 5.5 | 1509.2 | 957.3 | 957.3 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4936 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.878 | 0.12 | 1200 | 1503.7 | 5.5 | 1509.2 | 957.3 | 957.3 | 0.1 | 13.2 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8712 | 0.1585 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4151 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.88 | 0.12 | 1200 | 1503.7 | 5.5 | 1509.2 | 957.3 | 957.3 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8287 | 0.1585 | 0.006829 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.88 | 0.1192 | 1192.1482 | 1503.7 | 5.5 | 1509.2 | 957.4 | 957.4 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8285 | 0.1585 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.88 | 0.119 | 1190 | 1503.8 | 5.5 | 1509.3 | 957.4 | 957.4 | 0.1 | 13.2 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8706 | 0.1585 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.877 | 0.1195 | 1194.5167 | 1504.1 | 5.5 | 1509.6 | 957.4 | 957.4 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8293 | 0.1586 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.86 | 0.1203 | 1202.8861 | 1504.1 | 5.5 | 1509.6 | 957.4 | 957.4 | 0.1 | 13.2 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.1586 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4776 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.868 | 0.1197 | 1196.6723 | 1504.2 | 5.471 | 1509.7 | 957.3 | 957.3042 | 0.1 | 13.2 | 913 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.1586 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.444 | 27.9 | 5.5
2008-12-02 | 12:20 | 11.202 | 0.1112 | 1111.7827 | 1497.3 | 5.4 | 1502.7 | 957.2 | 957.1938 | 0.1 | 13.2 | 913 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.15819 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.765 | 0.1078 | 1078.4194 | 1434.8 | 5.4 | 1440.2 | 957.1 | 957.1044 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8291 | 0.15442 | 0.0069 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7091 | 11.4 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.696 | 0.107 | 1070 | 1375.1 | 5.4 | 1380.5 | 927.3 | 927.3096 | 0.1 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.15088 | 0.006871 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1363.5 | 5.4 | 1368.9 | 877.6 | 877.5563 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.14997 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5402 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1359.8 | 5.4 | 1365.2 | 863.4 | 863.44 | 0.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.1498 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4546 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1359.7 | 5.4 | 1365.1 | 860.6 | 860.5616 | 0.1 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8287 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3268 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1359.5 | 5.4 | 1364.9 | 859.8 | 859.8094 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8296 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052.0695 | 1014 | 11.4052 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1359.5 | 5.4 | 1364.9 | 859.4 | 859.4073 | 0.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4841 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.71 | 0.107 | 1070 | 1359.5 | 5.4 | 1364.9 | 859.3 | 859.3 | 0.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8707 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4373 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.704 | 0.107 | 1070 | 1359.4 | 5.4 | 1364.8 | 859.1 | 859.1093 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.871 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4411 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.662 | 0.107 | 1070 | 1359.2 | 5.4 | 1364.6 | 859 | 859.0052 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.676 | 0.1069 | 1068.719 | 1359.2 | 5.4 | 1364.6 | 859 | 859.0317 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.686 | 0.106 | 1060.4545 | 1359.1 | 5.4 | 1364.5 | 859.2 | 859.1938 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8711 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.69 | 0.106 | 1060 | 1359 | 5.4 | 1364.4 | 859.1 | 859.1107 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8302 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.69 | 0.1066 | 1066.3162 | 1358.9 | 5.4 | 1364.3 | 859.2 | 859.2039 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4663 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.69 | 0.107 | 1070 | 1358.7 | 5.3713 | 1364.1 | 859.4 | 859.3936 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4122 | 27.9 | 5.4
2008-12-02 | 12:20 | 10.69 | 0.107 | 1070 | 1358.6 | 5.3 | 1363.9 | 859.4 | 859.4 | 0.1 | 13.1 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4908 | 27.9 | 5.3
2008-12-02 | 12:20 | 10.69 | 0.1069 | 1068.7199 | 1358.5 | 5.3 | 1363.8 | 859.4 | 859.4 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8294 | 0.1497 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4307 | 27.9 | 5.3
2008-12-02 | 12:20 | 10.69 | 0.106 | 1060.1289 | 1358.5 | 5.3 | 1363.8 | 859.4 | 859.4 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.14967 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4478 | 27.9 | 5.3
2008-12-02 | 12:20 | 10.69 | 0.106 | 1060 | 1358.5 | 5.2708 | 1363.7 | 859.4 | 859.4 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8708 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.3
2008-12-02 | 12:20 | 10.69 | 0.1073 | 1072.6262 | 1358.6 | 5.2 | 1363.8 | 859.4 | 859.4048 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8295 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5049 | 27.9 | 5.2
2008-12-02 | 12:20 | 10.693 | 0.108 | 1080 | 1358.6 | 5.2 | 1363.8 | 859.5 | 859.498 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8293 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.7071 | 11.5834 | 27.9 | 5.2
2008-12-02 | 12:20 | 10.709 | 0.108 | 1080 | 1358.4 | 5.2 | 1363.6 | 859.4 | 859.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8291 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9 | 5.2
2008-12-02 | 12:20 | 10.71 | 0.108 | 1080 | 1358.2 | 5.2 | 1363.4 | 859.4 | 859.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.8289 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5597 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.71 | 0.108 | 1080 | 1358.1 | 5.2 | 1363.3 | 859.4 | 859.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.71 | 0.1072 | 1072.2 | 1358 | 5.2 | 1363.2 | 859.4 | 859.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.71 | 0.107 | 1070 | 1357.9 | 5.2 | 1363.1 | 859.3 | 859.3071 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.424 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.71 | 0.107 | 1070 | 1357.6 | 5.2 | 1362.8 | 859.3 | 859.3 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.71 | 0.107 | 1070 | 1357.6 | 5.1708 | 1362.8 | 859.3 | 859.3 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5.2
2008-12-02 | 12:21 | 10.694 | 0.107 | 1070 | 1357.7 | 5.1 | 1362.8 | 859.2 | 859.2075 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4112 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.107 | 1070 | 1357.7 | 5.1 | 1362.8 | 859.4 | 859.3888 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4899 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.681 | 0.107 | 1070 | 1357.6 | 5.1 | 1362.7 | 859.2 | 859.2178 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4316 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.673 | 0.107 | 1070 | 1357.5 | 5.1 | 1362.6 | 859.3 | 859.2939 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.829 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.68 | 0.1066 | 1065.5546 | 1357.4 | 5.1 | 1362.5 | 859.3 | 859.3 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4254 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.68 | 0.106 | 1060 | 1357.4 | 5.1 | 1362.5 | 859.2 | 859.2044 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.68 | 0.1063 | 1062.7876 | 1357.4 | 5.1 | 1362.5 | 859.2 | 859.2089 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1496 | 0.0068 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.686 | 0.108 | 1079.5802 | 1357.3 | 5.1 | 1362.4 | 859.2 | 859.1909 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8298 | 0.1495 | 0.00677 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4389 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.1072 | 1072.1561 | 1357.2 | 5.1 | 1362.3 | 859.1 | 859.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8704 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.107 | 1070 | 1357.2 | 5.1 | 1362.3 | 859.1 | 859.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.418 | 27.882 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.1065 | 1065.4002 | 1357.1 | 5.1 | 1362.2 | 859.1 | 859.0789 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4964 | 27.8036 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.106 | 1060 | 1357.1 | 5.1 | 1362.2 | 859.1 | 859.0963 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8708 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4249 | 27.8751 | 5.1
2008-12-02 | 12:21 | 10.698 | 0.1063 | 1062.5612 | 1357.3 | 5.1 | 1362.4 | 859.1 | 859.0883 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4536 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.694 | 0.1079 | 1078.8744 | 1357.3 | 5.1 | 1362.4 | 859.1 | 859.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.108 | 1080 | 1357.1 | 5.1 | 1362.2 | 859.1 | 859.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4894 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.1074 | 1073.5768 | 1357 | 5.1 | 1362.1 | 859.2 | 859.1956 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8289 | 0.1495 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4109 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.107 | 1070 | 1357.1 | 5.1 | 1362.2 | 859.1 | 859.1044 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4677 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.69 | 0.107 | 1070 | 1357.1 | 5.1 | 1362.2 | 859 | 859.0042 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.698 | 0.107 | 1070 | 1357 | 5.1 | 1362.1 | 859 | 858.9923 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 5.1
2008-12-02 | 12:21 | 10.7 | 0.107 | 1070 | 1357 | 5 | 1362 | 858.7 | 858.7093 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4968 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.1078 | 1077.817 | 1357 | 5 | 1362 | 858.8 | 858.7927 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8294 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4183 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1357 | 5 | 1362 | 859 | 859 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8707 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1357 | 5 | 1362 | 858.9 | 858.9106 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4387 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1356.8 | 5 | 1361.8 | 858.7 | 858.7122 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8707 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4828 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1356.7 | 5 | 1361.7 | 858.7 | 858.7 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4045 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1356.8 | 5 | 1361.8 | 858.8 | 858.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8289 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4742 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.108 | 1080 | 1356.8 | 5 | 1361.8 | 858.9 | 858.898 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3944 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.1067 | 1067.431 | 1356.8 | 5 | 1361.8 | 859 | 858.9956 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3626 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.106 | 1060 | 1357 | 5 | 1362 | 859.1 | 859.1 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4902 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.106 | 1060 | 1357.1 | 5 | 1362.1 | 859.2 | 859.1775 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.412 | 27.9 | 5
2008-12-02 | 12:21 | 10.7 | 0.1044 | 1043.8449 | 1356.9 | 5 | 1361.9 | 859.2 | 859.1909 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8706 | 0.1494 | 0.0067 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5
2008-12-02 | 12:21 | 9.985 | 0.0935 | 935.4212 | 1349 | 5 | 1354 | 859.4 | 859.3971 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8292 | 0.1489 | 0.006671 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 5
2008-12-02 | 12:21 | 9.546 | 0.0961 | 961.0204 | 1262.7 | 4.9 | 1267.6 | 859.1 | 859.1341 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8708 | 0.14325 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4239 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.53 | 0.097 | 970 | 1223.3 | 4.9 | 1228.2 | 818.9 | 818.906 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.829 | 0.14068 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4977 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.53 | 0.0965 | 965.4348 | 1218.1 | 4.9 | 1223 | 775.2 | 775.2075 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8707 | 0.14024 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4193 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.53 | 0.096 | 960 | 1215.6 | 4.9 | 1220.5 | 766 | 766.0085 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.14 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.53 | 0.0959 | 958.7469 | 1215.4 | 4.9 | 1220.3 | 763.8 | 763.8297 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.14 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.524 | 0.0951 | 950.6103 | 1214.9 | 4.9 | 1219.8 | 763.3 | 763.3041 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.14 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4328 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.52 | 0.0958 | 957.7019 | 1214.3 | 4.9 | 1219.2 | 763.1 | 763.1029 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1399 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5898 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.52 | 0.0954 | 953.7316 | 1214 | 4.8704 | 1218.9 | 763 | 763.0066 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.1399 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5267 | 27.9 | 4.9
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1213.6 | 4.8 | 1218.4 | 763 | 762.9818 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.13987 | 0.0066 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4481 | 27.9 | 4.8
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1213.2 | 4.8 | 1218 | 762.9 | 762.9 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.13977 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4305 | 27.9 | 4.8
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1212.8 | 4.8 | 1217.6 | 762.9 | 762.9 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1397 | 0.0065 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4821 | 27.9 | 4.8
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1211.9 | 4.771 | 1216.6 | 763 | 763 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.1396 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3251 | 27.9 | 4.8
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1211.6 | 4.7 | 1216.3 | 762.9 | 762.9043 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.1396 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2929 | 11.4978 | 27.9 | 4.7
2008-12-02 | 12:21 | 9.52 | 0.095 | 950 | 1211.4 | 4.6709 | 1216 | 763 | 762.9917 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.1395 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5111 | 27.9 | 4.7
2008-12-02 | 12:21 | 9.519 | 0.095 | 950 | 1211.4 | 4.6 | 1216 | 763 | 762.9947 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8289 | 0.1395 | 0.0064 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.423 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.51 | 0.095 | 950 | 1211.1 | 4.6 | 1215.7 | 763 | 763 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8712 | 0.13947 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4984 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.51 | 0.095 | 950 | 1211 | 4.6 | 1215.6 | 762.9 | 762.9021 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4201 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.504 | 0.095 | 950 | 1210.8 | 4.6 | 1215.4 | 762.9 | 762.9 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8714 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4585 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.5 | 0.095 | 950 | 1210.6 | 4.6 | 1215.2 | 762.7 | 762.7068 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195.4 | 59.8704 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.463 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.503 | 0.095 | 950 | 1210.5 | 4.6 | 1215.1 | 762.9 | 762.8912 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8294 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.51 | 0.095 | 950 | 1210.5 | 4.6 | 1215.1 | 762.8 | 762.8 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.51 | 0.095 | 950 | 1210.5 | 4.6 | 1215.1 | 762.8 | 762.8 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4722 | 27.9 | 4.6
2008-12-02 | 12:21 | 9.51 | 0.095 | 950 | 1210.4 | 4.6 | 1215 | 762.8 | 762.7749 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.871 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.516 | 0.095 | 950 | 1210.3 | 4.6 | 1214.9 | 762.7 | 762.7 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8709 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.52 | 0.095 | 950 | 1210.3 | 4.6 | 1214.9 | 762.7 | 762.7064 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.1394 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4919 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.52 | 0.095 | 950 | 1210.2 | 4.6 | 1214.8 | 762.7 | 762.7 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4134 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.52 | 0.095 | 950 | 1210.1 | 4.6 | 1214.7 | 762.7 | 762.7 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.52 | 0.095 | 950 | 1209.8 | 4.6 | 1214.4 | 762.5 | 762.5045 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.52 | 0.095 | 950 | 1209.8 | 4.6 | 1214.4 | 762.6 | 762.5983 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8296 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.477 | 0.095 | 950 | 1209.8 | 4.5706 | 1214.4 | 762.5 | 762.5051 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4006 | 27.9 | 4.6
2008-12-02 | 12:22 | 9.454 | 0.0958 | 957.682 | 1208.9 | 4.5 | 1213.4 | 762.5 | 762.524 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4791 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.466 | 0.0954 | 953.8205 | 1208.8 | 4.5 | 1213.3 | 762.5 | 762.5053 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1393 | 0.0063 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.481 | 0.095 | 950 | 1208.8 | 4.5 | 1213.3 | 762.6 | 762.595 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1392 | 0.006271 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.489 | 0.095 | 950 | 1208.7 | 4.5 | 1213.2 | 762.7 | 762.7 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.49 | 0.0951 | 951.2927 | 1208.7 | 4.5 | 1213.2 | 762.7 | 762.7 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.49 | 0.096 | 959.633 | 1208.6 | 4.5 | 1213.1 | 762.7 | 762.6934 | 0.1 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2889 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.49 | 0.096 | 960 | 1208.6 | 4.5 | 1213.1 | 762.7 | 762.7 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.2883 | 11.4498 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.49 | 0.0954 | 954.0099 | 1208.5 | 4.5 | 1213 | 762.7 | 762.7092 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195.6 | 59.9 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4287 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.491 | 0.0954 | 954.2964 | 1208.4 | 4.5 | 1212.9 | 762.8 | 762.7967 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8287 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.5 | 0.096 | 960 | 1208.4 | 4.5 | 1212.9 | 762.9 | 762.8934 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.5 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.5 | 0.0959 | 958.7916 | 1208.4 | 4.5 | 1212.9 | 762.8 | 762.7931 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8301 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4358 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.5 | 0.0951 | 950.6813 | 1208.4 | 4.5 | 1212.9 | 762.7 | 762.707 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4427 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.3 | 4.5 | 1212.8 | 762.7 | 762.7 | 0.1 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4788 | 27.9 | 4.5
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.3 | 4.4 | 1212.7 | 762.8 | 762.7955 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.871 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7097 | 11.4002 | 27.9 | 4.4
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.3 | 4.4 | 1212.7 | 762.7 | 762.7 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.4
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.2 | 4.4 | 1212.6 | 762.6 | 762.6088 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4 | 27.9 | 4.4
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.2 | 4.4 | 1212.6 | 762.7 | 762.7 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8289 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.365 | 27.9 | 4.4
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.1 | 4.3 | 1212.4 | 762.6 | 762.6107 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8288 | 0.1392 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2724 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208 | 4.3 | 1212.3 | 762.7 | 762.6952 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.13917 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1154 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208 | 4.3 | 1212.3 | 762.6 | 762.6464 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8702 | 0.1391 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7016 | 11.3832 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.1 | 4.3 | 1212.4 | 762.7 | 762.7 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8704 | 0.1391 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4508 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.1 | 4.3 | 1212.4 | 762.7 | 762.7 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8707 | 0.1391 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1208.1 | 4.3 | 1212.4 | 762.8 | 762.8 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8292 | 0.1391 | 0.0062 | 0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1207.9 | 4.3 | 1212.2 | 762.8 | 762.7784 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.829 | 0.1391 | 0.0062 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1207.7 | 4.3 | 1212 | 762.9 | 762.8876 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1391 | 0.0062 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1207.6 | 4.3 | 1211.9 | 762.8 | 762.8041 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1391 | 0.0062 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4417 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.095 | 950 | 1207.5 | 4.3 | 1211.8 | 762.8 | 762.8 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8711 | 0.1391 | 0.0062 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4798 | 27.9 | 4.3
2008-12-02 | 12:22 | 9.5 | 0.0931 | 931.3613 | 1207.4 | 4.3 | 1211.7 | 762.7 | 762.7087 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1391 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4012 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.679 | 0.0873 | 873.0526 | 1207.4 | 4.3 | 1211.7 | 762.7 | 762.6937 | 0.1 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8 | 0.13907 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.43 | 0.0824 | 824.4696 | 1161.9 | 4.3 | 1166.2 | 761.4 | 761.4198 | 0.1 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8287 | 0.13579 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.7129 | 11.3441 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.367 | 0.082 | 820 | 1075.7 | 4.3 | 1080 | 728.3 | 728.3431 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8704 | 0.12943 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3344 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.29 | 0.082 | 820 | 1065.9 | 4.3 | 1070.2 | 680.7 | 680.702 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8294 | 0.12867 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.291 | 0.0829 | 828.6116 | 1062.2 | 4.3 | 1066.5 | 669.6 | 669.6277 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8707 | 0.1284 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 4.3
2008-12-02 | 12:22 | 8.299 | 0.084 | 840 | 1061.5 | 4.2 | 1065.7 | 667.5 | 667.4593 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1283 | 0.0061 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5397 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.084 | 840 | 1061.4 | 4.2 | 1065.6 | 666.9 | 666.8535 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.829 | 0.12827 | 0.006071 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4066 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.084 | 840 | 1061 | 4.2 | 1065.2 | 666.6 | 666.6045 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.1281 | 0.0059 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3075 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.084 | 840 | 1060.7 | 4.2 | 1064.9 | 666.4 | 666.4063 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.1281 | 0.0059 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.0834 | 833.9599 | 1060.5 | 4.2 | 1064.7 | 666.4 | 666.4137 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.128 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.083 | 830 | 1060.3 | 4.2 | 1064.5 | 666.9 | 666.8751 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8286 | 0.128 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4128 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.0827 | 827.3339 | 1060 | 4.2 | 1064.2 | 667.1 | 667.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1279 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3815 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.0823 | 822.5021 | 1060 | 4.2 | 1064.2 | 666.8 | 666.8249 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8704 | 0.1279 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4809 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.0839 | 839.1827 | 1060 | 4.2 | 1064.2 | 666.7 | 666.7 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8294 | 0.1279 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4021 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.296 | 0.084 | 840 | 1059.5 | 4.2 | 1063.7 | 666.5 | 666.5072 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1279 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.4764 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.293 | 0.084 | 840 | 1059.4 | 4.2 | 1063.6 | 666.5 | 666.5 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.3 | 59.8708 | 0.1279 | 0.0058 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.7083 | 11.5 | 27.9 | 4.2
2008-12-02 | 12:22 | 8.3 | 0.084 | 840 | 1059.4 | 4.1 | 1063.5 | 666.5 | 666.5 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.871 | 0.1279 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.3 | 0.0837 | 837.3476 | 1059.3 | 4.1 | 1063.4 | 666.3 | 666.3211 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.1278 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.3 | 0.0831 | 831.1641 | 1059.3 | 4.1 | 1063.4 | 666.5 | 666.4938 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8709 | 0.1278 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.3 | 0.0839 | 839.2482 | 1059.2 | 4.1 | 1063.3 | 666.4 | 666.4102 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8289 | 0.1278 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.296 | 0.0832 | 832.3395 | 1059.2 | 4.1 | 1063.3 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.1278 | 0.0057 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4525 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.29 | 0.083 | 830 | 1059 | 4.1 | 1063.1 | 666.3 | 666.3 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1277 | 0.005671 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.426 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.291 | 0.0834 | 834.2202 | 1058.6 | 4.1 | 1062.7 | 666.5 | 666.4953 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8287 | 0.1277 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.7129 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.299 | 0.084 | 840 | 1058.6 | 4.1 | 1062.7 | 666.3 | 666.3064 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1276 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4.1
2008-12-02 | 12:22 | 8.3 | 0.084 | 840 | 1058.6 | 4.1 | 1062.7 | 666.4 | 666.3927 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1276 | 0.0056 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5615 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058.6 | 4.1 | 1062.7 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1276 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.56 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058.6 | 4.1 | 1062.7 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.1276 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4817 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.0834 | 834.0248 | 1058.6 | 4.1 | 1062.7 | 666.3 | 666.3043 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1276 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4037 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.083 | 830 | 1058.5 | 4.1 | 1062.6 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.12757 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.0833 | 832.6311 | 1058.4 | 4.1 | 1062.5 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1275 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4541 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058.4 | 4.1 | 1062.5 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1275 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4349 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058.2 | 4.1 | 1062.3 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8288 | 0.1275 | 0.0055 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3221 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058.2 | 4.0713 | 1062.2 | 666.3 | 666.3066 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.1275 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4791 | 27.9 | 4.1
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1058 | 4 | 1062 | 666.4 | 666.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8704 | 0.1275 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4319 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1057.9 | 4 | 1061.9 | 666.2 | 666.2101 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8706 | 0.1275 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4465 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1057.8 | 4 | 1061.8 | 666.3 | 666.293 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8293 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1057.8 | 4 | 1061.8 | 666.3 | 666.2949 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4968 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1057.7 | 4 | 1061.7 | 666.3 | 666.3 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.829 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.418 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0832 | 832.3475 | 1057.7 | 4 | 1061.7 | 666.2 | 666.2062 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4606 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.083 | 830 | 1057.8 | 4 | 1061.8 | 666.2 | 666.2251 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.083 | 830 | 1057.8 | 4 | 1061.8 | 666.3 | 666.3 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8711 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0833 | 832.5125 | 1057.9 | 4 | 1061.9 | 666.5 | 666.4889 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8712 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0839 | 839.1554 | 1057.9 | 4 | 1061.9 | 666.3 | 666.3032 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4255 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0831 | 830.7095 | 1058 | 4 | 1062 | 666.2 | 666.2045 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8713 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0838 | 837.6781 | 1057.9 | 4 | 1061.9 | 666.3 | 666.2988 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8704 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4316 | 27.9 | 4
2008-12-02 | 12:23 | 8.305 | 0.084 | 840 | 1057.8 | 4 | 1061.8 | 666.2 | 666.2048 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4797 | 27.9 | 4
2008-12-02 | 12:23 | 8.319 | 0.084 | 840 | 1057.8 | 4 | 1061.8 | 666.4 | 666.3949 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3228 | 27.9 | 4
2008-12-02 | 12:23 | 8.311 | 0.084 | 840 | 1057.7 | 4 | 1061.7 | 666.4 | 666.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.1274 | 0.0054 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4342 | 27.9 | 4
2008-12-02 | 12:23 | 8.31 | 0.084 | 840 | 1057.6 | 4 | 1061.6 | 666.2 | 666.209 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1274 | 0.005371 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4544 | 27.9 | 4
2008-12-02 | 12:23 | 8.31 | 0.084 | 840 | 1057.6 | 4 | 1061.6 | 666.1 | 666.1119 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1273 | 0.0053 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 4
2008-12-02 | 12:23 | 8.31 | 0.084 | 840 | 1057.6 | 4 | 1061.6 | 666.2 | 666.1979 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.1273 | 0.0053 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3973 | 27.9 | 4
2008-12-02 | 12:23 | 8.31 | 0.084 | 840 | 1057.5 | 4 | 1061.5 | 666.1 | 666.1043 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.1273 | 0.0053 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3194 | 27.9 | 4
2008-12-02 | 12:23 | 8.309 | 0.084 | 840 | 1057.4 | 4 | 1061.4 | 666.2 | 666.1937 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.12727 | 0.005271 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4187 | 27.9 | 4
2008-12-02 | 12:23 | 8.301 | 0.084 | 840 | 1057.3 | 3.9 | 1061.2 | 666.1 | 666.1032 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8287 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 3.9
2008-12-02 | 12:23 | 8.315 | 0.0839 | 839.1814 | 1057.1 | 3.9 | 1061 | 666.1 | 666.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4834 | 27.9 | 3.9
2008-12-02 | 12:23 | 8.308 | 0.0831 | 830.7426 | 1057 | 4 | 1061 | 666.1 | 666.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8706 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4049 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.0838 | 837.6703 | 1056.9 | 3.9707 | 1060.9 | 666.1 | 666.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4736 | 27.9 | 4
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1056.9 | 3.9 | 1060.8 | 666.1 | 666.11 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4479 | 27.9 | 3.9
2008-12-02 | 12:23 | 8.3 | 0.084 | 840 | 1056.9 | 3.9 | 1060.8 | 666.2 | 666.1968 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3694 | 27.9 | 3.9
2008-12-02 | 12:23 | 8.3 | 0.0838 | 837.5271 | 1056.7 | 3.9 | 1060.6 | 666.1 | 666.1103 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.7117 | 11.3184 | 27.9 | 3.9
2008-12-02 | 12:23 | 8.292 | 0.0819 | 819.4148 | 1056.8 | 3.9 | 1060.7 | 666.1 | 666.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1272 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4754 | 27.9 | 3.9
2008-12-02 | 12:23 | 7.601 | 0.0709 | 709.1518 | 1050.6 | 3.9 | 1054.5 | 666.1 | 666.1061 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.12671 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4338 | 27.9 | 3.9
2008-12-02 | 12:23 | 7.116 | 0.0731 | 730.6946 | 964.9 | 3.9 | 968.8 | 666 | 666 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.11992 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3553 | 27.9 | 3.9
2008-12-02 | 12:23 | 7.11 | 0.0734 | 733.9783 | 918.7 | 3.9 | 922.6 | 621.7 | 621.6784 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.11604 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3232 | 27.9 | 3.9
2008-12-02 | 12:23 | 7.11 | 0.073 | 730 | 914.6 | 3.8702 | 918.4 | 581 | 581.0154 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8298 | 0.11567 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 3.9
2008-12-02 | 12:23 | 7.11 | 0.073 | 730 | 913 | 3.8 | 916.8 | 572.9 | 572.9444 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.11557 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 3.8
2008-12-02 | 12:23 | 7.11 | 0.073 | 730 | 912.6 | 3.8 | 916.4 | 571.1 | 571.1255 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1155 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4586 | 27.9 | 3.8
2008-12-02 | 12:23 | 7.086 | 0.073 | 730 | 912.3 | 3.7 | 916 | 570.7 | 570.7061 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 196 | 59.8707 | 0.1155 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.07 | 0.073 | 730 | 911.6 | 3.7 | 915.3 | 570.4 | 570.4126 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8708 | 0.1154 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4844 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.07 | 0.073 | 730 | 911.2 | 3.7 | 914.9 | 570.4 | 570.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.1154 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4058 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.071 | 0.0726 | 725.8882 | 911 | 3.7 | 914.7 | 570.3 | 570.3062 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.1154 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4722 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.079 | 0.072 | 720 | 910.9 | 3.7 | 914.6 | 570.3 | 570.3 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.11537 | 0.0052 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.072 | 720 | 910.8 | 3.7 | 914.5 | 570.3 | 570.2957 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1153 | 0.005171 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4407 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.086 | 0.072 | 720 | 910.2 | 3.7 | 913.9 | 570.2 | 570.2 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1153 | 0.0051 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3163 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.09 | 0.072 | 720 | 909.3 | 3.7 | 913 | 570.2 | 570.2 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1152 | 0.0051 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4733 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.087 | 0.072 | 720 | 908.8 | 3.7 | 912.5 | 570 | 570.0141 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1152 | 0.0051 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3704 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.072 | 720 | 908.6 | 3.7 | 912.3 | 570.1 | 570.093 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1151 | 0.0051 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3867 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.0718 | 717.5604 | 908.5 | 3.7 | 912.2 | 570 | 570 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.1151 | 0.0051 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4779 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.4 | 3.7 | 912.1 | 569.9 | 569.9075 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8708 | 0.11507 | 0.005071 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.3 | 3.7 | 912 | 569.8 | 569.8437 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8705 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.3 | 3.7 | 912 | 569.7 | 569.7 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4 | 27.9 | 3.7
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.3 | 3.6 | 911.9 | 569.7 | 569.7 | 0.1 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4719 | 27.9 | 3.6
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.3 | 3.6 | 911.9 | 569.7 | 569.7 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6147 | 27.9 | 3.6
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908.2 | 3.6 | 911.8 | 569.9 | 569.8877 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6932 | 27.9 | 3.6
2008-12-02 | 12:23 | 7.08 | 0.071 | 710 | 908 | 3.6 | 911.6 | 569.7 | 569.7044 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.115 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5567 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.08 | 0.071 | 710 | 908 | 3.6 | 911.6 | 569.8 | 569.7953 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.1149 | 0.005 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4498 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.084 | 0.0718 | 717.6284 | 907.9 | 3.6 | 911.5 | 569.8 | 569.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.1149 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4286 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.088 | 0.072 | 720 | 907.8 | 3.6 | 911.4 | 569.8 | 569.8108 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8293 | 0.1149 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4929 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.08 | 0.0724 | 724.3983 | 907.9 | 3.6 | 911.5 | 569.9 | 569.9 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195.6 | 59.8292 | 0.1149 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4144 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.08 | 0.073 | 730 | 907.7 | 3.6 | 911.3 | 569.9 | 569.9 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.829 | 0.1149 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4641 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.103 | 0.0729 | 729.2574 | 907.7 | 3.6 | 911.3 | 570 | 570 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8 | 0.1149 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5427 | 27.9 | 3.6
2008-12-02 | 12:24 | 7.11 | 0.0721 | 720.8186 | 907.6 | 3.6 | 911.2 | 570 | 570 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8709 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5789 | 27.9211 | 3.6
2008-12-02 | 12:24 | 7.11 | 0.0728 | 727.6525 | 907.6 | 3.6 | 911.2 | 570.1 | 570.0968 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5005 | 27.9995 | 3.6
2008-12-02 | 12:24 | 7.11 | 0.073 | 730 | 907.6 | 3.5713 | 911.1 | 569.9 | 569.9154 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9219 | 3.6
2008-12-02 | 12:24 | 7.108 | 0.073 | 730 | 907.5 | 3.5 | 911 | 570.1 | 570.0909 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8712 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5565 | 27.9565 | 3.5
2008-12-02 | 12:24 | 7.083 | 0.073 | 730 | 907.6 | 3.5 | 911.1 | 570.2 | 570.1978 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5301 | 27.965 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.0729 | 729.2988 | 907.6 | 3.5 | 911.1 | 570.2 | 570.1936 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8713 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.441 | 27.9137 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.0721 | 720.9516 | 907.6 | 3.5 | 911.1 | 569.9 | 569.9146 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6765 | 27.9922 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.0712 | 712.3818 | 907.6 | 3.5 | 911.1 | 570.1 | 570.1 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4881 | 27.9294 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.071 | 710 | 907.6 | 3.5 | 911.1 | 570.1 | 570.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8708 | 0.1148 | 0.0049 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.449 | 27.949 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.071 | 710 | 907.6 | 3.5 | 911.1 | 570.2 | 570.1944 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.871 | 0.1148 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9726 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.071 | 710 | 907.6 | 3.5 | 911.1 | 570 | 570.0063 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8293 | 0.1148 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4938 | 27.9062 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.0709 | 709.2923 | 907.5 | 3.5 | 911 | 570.1 | 570.1 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.1148 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4153 | 27.9847 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.0701 | 700.8677 | 907.3 | 3.5 | 910.8 | 570 | 570 | 0.1 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.1148 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5264 | 27.9368 | 3.5
2008-12-02 | 12:24 | 7.08 | 0.07 | 700 | 907.4 | 3.5 | 910.9 | 569.9 | 569.9071 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.1148 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5168 | 27.9 | 3.5
2008-12-02 | 12:24 | 7.078 | 0.07 | 700 | 907.4 | 3.5 | 910.9 | 570 | 569.9935 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8288 | 0.11477 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4401 | 27.92 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.07 | 700 | 907.3 | 3.5 | 910.8 | 570.1 | 570.0895 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5973 | 27.9987 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.07 | 700 | 907.2 | 3.5 | 910.7 | 570.1 | 570.1 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195.2 | 59.8296 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6772 | 27.9228 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.0701 | 701.0353 | 907 | 3.5 | 910.5 | 570 | 570.0035 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8294 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5887 | 27.9557 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.0709 | 709.061 | 907 | 3.5 | 910.5 | 569.9 | 569.9 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 907 | 3.5 | 910.5 | 570.1 | 570.0938 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9873 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 906.9 | 3.5 | 910.4 | 570 | 570.0096 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.829 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9088 | 3.5
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 906.9 | 3.4 | 910.3 | 570 | 570.0239 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8709 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6394 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 906.9 | 3.4 | 910.3 | 570.1 | 570.1 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8711 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6036 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.0711 | 710.6911 | 906.7 | 3.4 | 910.1 | 570.2 | 570.2 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.0719 | 719.0175 | 906.7 | 3.4 | 910.1 | 570.1 | 570.1037 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8288 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4948 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.072 | 720 | 906.7 | 3.4 | 910.1 | 570.1 | 570.1022 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8287 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4163 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.0714 | 713.9447 | 906.7 | 3.4 | 910.1 | 570 | 569.9974 | 0.1 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5242 | 27.9 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 906.7 | 3.4 | 910.1 | 569.9 | 569.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8294 | 0.1147 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5188 | 27.9406 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.071 | 710 | 906.8 | 3.4 | 910.2 | 569.8 | 569.8 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8293 | 0.11467 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4191 | 28 | 3.4
2008-12-02 | 12:24 | 7.07 | 0.0703 | 703.2556 | 906.8 | 3.4 | 910.2 | 569.7 | 569.7 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.1146 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4976 | 28 | 3.4
2008-12-02 | 12:24 | 6.393 | 0.062 | 619.9917 | 903.4 | 3.4 | 906.8 | 569.9 | 569.8794 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.11434 | 0.0048 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3.4
2008-12-02 | 12:24 | 5.886 | 0.0618 | 617.6059 | 827.2 | 3.4 | 830.6 | 569.9 | 569.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8293 | 0.10772 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4454 | 28 | 3.4
2008-12-02 | 12:24 | 5.812 | 0.062 | 620 | 766.4 | 3.4 | 769.8 | 543.6 | 543.5723 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.1 | 59.9 | 0.10217 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.4662 | 27.9669 | 3.4
2008-12-02 | 12:24 | 5.82 | 0.062 | 620 | 761.6 | 3.4 | 765 | 488.4 | 488.4319 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.10184 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.2889 | 11.6115 | 27.9115 | 3.4
2008-12-02 | 12:24 | 5.82 | 0.0618 | 617.6642 | 761.5 | 3.3713 | 764.8 | 478.1 | 478.1038 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8287 | 0.10167 | 0.0047 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6901 | 27.9901 | 3.4
2008-12-02 | 12:24 | 5.82 | 0.061 | 610 | 761.1 | 3.3 | 764.4 | 476.1 | 476.065 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8288 | 0.1016 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6314 | 28 | 3.3
2008-12-02 | 12:24 | 5.826 | 0.061 | 610 | 760.9 | 3.3 | 764.2 | 475.5 | 475.5089 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8287 | 0.1015 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5529 | 28 | 3.3
2008-12-02 | 12:24 | 5.834 | 0.061 | 610 | 760.6 | 3.3 | 763.9 | 475.3 | 475.3069 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8706 | 0.1015 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5254 | 27.9746 | 3.3
2008-12-02 | 12:24 | 5.84 | 0.061 | 610 | 760 | 3.3 | 763.3 | 474.9 | 474.9205 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1014 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9034 | 3.3
2008-12-02 | 12:24 | 5.839 | 0.061 | 610 | 759.9 | 3.3 | 763.2 | 474.9 | 474.9 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1014 | 0.0046 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9825 | 3.3
2008-12-02 | 12:24 | 5.831 | 0.061 | 610 | 759.7 | 3.3 | 763 | 475.1 | 475.0603 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.829 | 0.1014 | 0.004571 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6611 | 28 | 3.3
2008-12-02 | 12:24 | 5.83 | 0.061 | 610 | 759.5 | 3.3 | 762.8 | 475 | 474.9881 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1013 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6207 | 27.9604 | 3.3
2008-12-02 | 12:24 | 5.83 | 0.061 | 610 | 759.5 | 3.3 | 762.8 | 474.9 | 474.91 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.1013 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.918 | 3.3
2008-12-02 | 12:24 | 5.834 | 0.061 | 610 | 759.2 | 3.3 | 762.5 | 474.9 | 474.8734 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8711 | 0.1013 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.996 | 3.3
2008-12-02 | 12:24 | 5.842 | 0.061 | 610 | 759.1 | 3.3 | 762.4 | 474.8 | 474.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8712 | 0.10127 | 0.0045 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.925 | 3.3
2008-12-02 | 12:24 | 5.851 | 0.061 | 610 | 759 | 3.3 | 762.3 | 475 | 474.9935 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8712 | 0.1012 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9 | 3.3
2008-12-02 | 12:24 | 5.859 | 0.061 | 610 | 758.8 | 3.3 | 762.1 | 474.7 | 474.7057 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1012 | 0.0044 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.532 | 27.9 | 3.3
2008-12-02 | 12:24 | 5.868 | 0.061 | 610 | 758.6 | 3.3 | 761.9 | 474.8 | 474.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1011 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9 | 3.3
2008-12-02 | 12:24 | 5.864 | 0.061 | 610 | 758.6 | 3.2706 | 761.8 | 474.8 | 474.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.10107 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.9 | 3.3
2008-12-02 | 12:24 | 5.864 | 0.061 | 610 | 758.3 | 3.2 | 761.5 | 474.9 | 474.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.101 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6675 | 27.9675 | 3.2
2008-12-02 | 12:24 | 5.868 | 0.061 | 610 | 758.1 | 3.2 | 761.3 | 474.8 | 474.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.101 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6539 | 27.9539 | 3.2
2008-12-02 | 12:24 | 5.86 | 0.061 | 610 | 758.2 | 3.2 | 761.4 | 474.7 | 474.71 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.101 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5509 | 27.9246 | 3.2
2008-12-02 | 12:24 | 5.86 | 0.061 | 610 | 758.1 | 3.2 | 761.3 | 474.8 | 474.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.10097 | 0.0043 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4031 | 28 | 3.2
2008-12-02 | 12:24 | 5.845 | 0.0607 | 607.113 | 758 | 3.2 | 761.2 | 474.8 | 474.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1009 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4816 | 28 | 3.2
2008-12-02 | 12:25 | 5.858 | 0.0565 | 565.272 | 757.7 | 3.2 | 760.9 | 474.7 | 474.6641 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.1009 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.56 | 28 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.0575 | 574.5667 | 757.6 | 3.2 | 760.8 | 474.9 | 474.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1009 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6385 | 28 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.058 | 580 | 757.4 | 3.2 | 760.6 | 474.7 | 474.7 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.1009 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6659 | 28 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.058 | 580 | 757.4 | 3.2 | 760.6 | 474.7 | 474.7082 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1009 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5089 | 28 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.058 | 580 | 757.4 | 3.2 | 760.6 | 474.6 | 474.6035 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.10087 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.426 | 27.926 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.058 | 580 | 757.4 | 3.2 | 760.6 | 474.9 | 474.8899 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4526 | 27.9526 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.058 | 580 | 757.3 | 3.2 | 760.5 | 474.5 | 474.5154 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3.2
2008-12-02 | 12:25 | 5.87 | 0.0587 | 587.2811 | 757.2 | 3.2 | 760.4 | 474.6 | 474.6 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4809 | 28 | 3.2
2008-12-02 | 12:25 | 5.868 | 0.059 | 590 | 757.1 | 3.2 | 760.3 | 474.6 | 474.6 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3246 | 28 | 3.2
2008-12-02 | 12:25 | 5.86 | 0.059 | 590 | 757 | 3.2 | 760.2 | 474.7 | 474.6957 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4991 | 27.9336 | 3.2
2008-12-02 | 12:25 | 5.86 | 0.0592 | 592.4491 | 757.1 | 3.1711 | 760.2 | 474.7 | 474.6938 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 27.945 | 3.2
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 757.1 | 3.1 | 760.2 | 474.5 | 474.5263 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5765 | 28 | 3.1
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 757.1 | 3.1 | 760.2 | 474.6 | 474.6 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8714 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3.1
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 757.1 | 3.1 | 760.2 | 474.6 | 474.6 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5 | 28 | 3.1
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.9 | 3.1 | 760 | 474.7 | 474.7059 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5589 | 28 | 3.1
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.9 | 3.1 | 760 | 474.6 | 474.6 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5625 | 28 | 3.1
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.8 | 3 | 759.8 | 474.8 | 474.8 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1008 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.7 | 3 | 759.7 | 474.5 | 474.5045 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.1007 | 0.0042 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 3
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.7 | 2.971 | 759.6 | 474.6 | 474.6098 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.871 | 0.1007 | 0.004171 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9271 | 3
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.5 | 2.9 | 759.4 | 474.5 | 474.5196 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.1007 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 27.9514 | 2.9
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.3 | 2.9 | 759.2 | 474.5 | 474.5125 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8291 | 0.10067 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5897 | 27.9701 | 2.9
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.2 | 2.9 | 759.1 | 474.6 | 474.598 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8711 | 0.1006 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7748 | 27.9084 | 2.9
2008-12-02 | 12:25 | 5.86 | 0.06 | 600 | 756.2 | 2.9 | 759.1 | 474.6 | 474.5741 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.1006 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5393 | 27.9869 | 2.9
2008-12-02 | 12:25 | 5.867 | 0.06 | 600 | 756.2 | 2.9 | 759.1 | 474.7 | 474.6913 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1006 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.435 | 28 | 2.9
2008-12-02 | 12:25 | 5.87 | 0.06 | 600 | 756.1 | 2.9 | 759 | 474.7 | 474.7129 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1006 | 0.0041 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4435 | 28 | 2.9
2008-12-02 | 12:25 | 5.87 | 0.06 | 600 | 755.9 | 2.9 | 758.8 | 474.5 | 474.506 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.1006 | 0.00407 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5448 | 28 | 2.9
2008-12-02 | 12:25 | 5.87 | 0.06 | 600 | 755.9 | 2.9 | 758.8 | 474.4 | 474.4 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.1005 | 0.004 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6972 | 28 | 2.9
2008-12-02 | 12:25 | 5.869 | 0.06 | 600 | 755.8 | 2.9 | 758.7 | 474.7 | 474.6899 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.1005 | 0.004 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4617 | 28 | 2.9
2008-12-02 | 12:25 | 5.852 | 0.06 | 600 | 755.9 | 2.9 | 758.8 | 474.6 | 474.6029 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.1005 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4579 | 28 | 2.9
2008-12-02 | 12:25 | 5.85 | 0.06 | 600 | 755.8 | 2.9 | 758.7 | 474.6 | 474.6043 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.1005 | 0.0039 | 0 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.9
2008-12-02 | 12:25 | 5.85 | 0.06 | 600 | 755.9 | 2.9 | 758.8 | 474.5 | 474.5 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.9 | 0.1005 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.4851 | 28 | 2.9
2008-12-02 | 12:25 | 5.85 | 0.06 | 600 | 755.9 | 2.8 | 758.7 | 474.5 | 474.5 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.1004 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.2889 | 11.4066 | 28 | 2.8
2008-12-02 | 12:25 | 5.85 | 0.06 | 600 | 755.8 | 2.7712 | 758.5 | 474.5 | 474.5 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1004 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.2883 | 11.5437 | 27.9282 | 2.8
2008-12-02 | 12:25 | 5.799 | 0.06 | 600 | 755.8 | 2.7 | 758.5 | 474.5 | 474.4776 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.1004 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5497 | 27.9503 | 2.7
2008-12-02 | 12:25 | 5.064 | 0.0571 | 571.1111 | 754.6 | 2.7 | 757.3 | 474.4 | 474.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.10028 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5864 | 28 | 2.7
2008-12-02 | 12:25 | 4.773 | 0.0471 | 470.5188 | 693.6 | 2.7 | 696.3 | 474.4 | 474.4 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0944 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.7781 | 27.9927 | 2.7
2008-12-02 | 12:25 | 4.69 | 0.0479 | 478.887 | 650.9 | 2.7 | 653.6 | 452.6 | 452.6242 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.090175 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5424 | 27.9141 | 2.7
2008-12-02 | 12:25 | 4.686 | 0.048 | 480 | 611.2 | 2.7 | 613.9 | 397.3 | 397.3202 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8293 | 0.086141 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5 | 27.9642 | 2.7
2008-12-02 | 12:25 | 4.68 | 0.048 | 480 | 609.1 | 2.7 | 611.8 | 383.9 | 383.9195 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0859 | 0.0039 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5 | 28 | 2.7
2008-12-02 | 12:25 | 4.68 | 0.048 | 480 | 608.2 | 2.7 | 610.9 | 381.5 | 381.4614 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.829 | 0.085871 | 0.003871 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.5213 | 28 | 2.7
2008-12-02 | 12:25 | 4.68 | 0.048 | 480 | 607.7 | 2.7 | 610.4 | 380.8 | 380.8062 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0858 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.2929 | 11.5998 | 28 | 2.7
2008-12-02 | 12:25 | 4.68 | 0.048 | 480 | 607.6 | 2.7 | 610.3 | 380.6 | 380.6024 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0858 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5217 | 28 | 2.7
2008-12-02 | 12:25 | 4.646 | 0.048 | 480 | 607.4 | 2.7 | 610.1 | 380.6 | 380.6041 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.0857 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.7
2008-12-02 | 12:25 | 4.636 | 0.048 | 480 | 607.2 | 2.7 | 609.9 | 380.6 | 380.6061 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0857 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.7
2008-12-02 | 12:25 | 4.658 | 0.048 | 480 | 607.2 | 2.7 | 609.9 | 380.6 | 380.5741 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0857 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5275 | 28 | 2.7
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 607.2 | 2.6 | 609.8 | 380.4 | 380.4047 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0857 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6845 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 606.9 | 2.6 | 609.5 | 380.4 | 380.4 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8294 | 0.0856 | 0.0038 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6292 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 606.5 | 2.6 | 609.1 | 380.4 | 380.4 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.0856 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5509 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 606.3 | 2.6 | 608.9 | 380.4 | 380.4 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.555 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 606.1 | 2.6 | 608.7 | 380.3 | 380.305 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.871 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6875 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 606 | 2.6 | 608.6 | 380.4 | 380.3953 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5306 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 605.9 | 2.6 | 608.5 | 380.3 | 380.3031 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195 | 59.8291 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6264 | 28 | 2.6
2008-12-02 | 12:25 | 4.65 | 0.048 | 480 | 605.8 | 2.6 | 608.4 | 380.3 | 380.3232 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8711 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6584 | 28 | 2.6
2008-12-02 | 12:25 | 4.657 | 0.048 | 480 | 605.9 | 2.6 | 608.5 | 380.5 | 380.496 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0855 | 0.0037 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6201 | 28 | 2.6
2008-12-02 | 12:25 | 4.66 | 0.048 | 480 | 605.8 | 2.6 | 608.4 | 380.2 | 380.2 | 0.1 | 13.1 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0854 | 0.003671 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6987 | 28 | 2.6
2008-12-02 | 12:25 | 4.664 | 0.048 | 480 | 605.7 | 2.6 | 608.3 | 380.3 | 380.2926 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8713 | 0.0854 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5457 | 28 | 2.6
2008-12-02 | 12:25 | 4.67 | 0.048 | 480 | 605.8 | 2.6 | 608.4 | 380.2 | 380.2 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8704 | 0.08537 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5557 | 28 | 2.6
2008-12-02 | 12:25 | 4.67 | 0.048 | 480 | 605.8 | 2.6 | 608.4 | 380.2 | 380.2091 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0853 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.566 | 28 | 2.6
2008-12-02 | 12:25 | 4.67 | 0.048 | 480 | 605.7 | 2.6 | 608.3 | 380 | 380.0135 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8293 | 0.0853 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:25 | 4.67 | 0.048 | 480 | 605.6 | 2.6 | 608.2 | 380.1 | 380.1 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0853 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.048 | 480 | 605.7 | 2.6 | 608.3 | 380.2 | 380.1961 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8293 | 0.0853 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7089 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.048 | 480 | 605.6 | 2.6 | 608.2 | 380.1 | 380.1045 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8709 | 0.0853 | 0.0036 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7038 | 28 | 2.6
2008-12-02 | 12:26 | 4.668 | 0.0491 | 491.4484 | 605.5 | 2.6 | 608.1 | 380.3 | 380.2915 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8289 | 0.0853 | 0.003571 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5735 | 28 | 2.6
2008-12-02 | 12:26 | 4.658 | 0.05 | 500 | 605.5 | 2.6 | 608.1 | 380.4 | 380.4059 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0853 | 0.0035 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5046 | 28 | 2.6
2008-12-02 | 12:26 | 4.625 | 0.0498 | 497.8765 | 605.4 | 2.6 | 608 | 380.1 | 380.1253 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195.6 | 59.8288 | 0.0853 | 0.0035 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5834 | 28 | 2.6
2008-12-02 | 12:26 | 4.656 | 0.049 | 490 | 605.4 | 2.6 | 608 | 380.1 | 380.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.0853 | 0.0035 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.476 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.049 | 490 | 605.3 | 2.6 | 607.9 | 380.1 | 380.1115 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.08527 | 0.0035 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4811 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.049 | 490 | 605.2 | 2.6 | 607.8 | 380.1 | 380.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0852 | 0.003471 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5809 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.0484 | 484.238 | 605.1 | 2.6 | 607.7 | 380 | 380.0073 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5024 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.048 | 480 | 605.1 | 2.6 | 607.7 | 380.2 | 380.2 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.0484 | 484.3134 | 605.2 | 2.6 | 607.8 | 380.1 | 380.1 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.871 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6089 | 28 | 2.6
2008-12-02 | 12:26 | 4.67 | 0.05 | 500 | 605 | 2.6 | 607.6 | 380 | 380.0062 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6343 | 28 | 2.6
2008-12-02 | 12:26 | 4.664 | 0.05 | 500 | 605 | 2.6 | 607.6 | 380.2 | 380.1935 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.05 | 500 | 604.9 | 2.6 | 607.5 | 380.1 | 380.1085 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0851 | 0.0034 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.0494 | 494.4615 | 604.9 | 2.6 | 607.5 | 379.9 | 379.9073 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8712 | 0.085071 | 0.003371 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6368 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.049 | 490 | 604.8 | 2.6 | 607.4 | 380 | 380 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8288 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6063 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.049 | 490 | 604.7 | 2.6 | 607.3 | 379.9 | 379.9045 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8287 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.049 | 490 | 604.7 | 2.6 | 607.3 | 379.9 | 379.9209 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5069 | 28 | 2.6
2008-12-02 | 12:26 | 4.655 | 0.049 | 490 | 604.6 | 2.6 | 607.2 | 380 | 380 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6647 | 28 | 2.6
2008-12-02 | 12:26 | 4.65 | 0.049 | 490 | 604.6 | 2.6 | 607.2 | 380 | 380 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8707 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6391 | 28 | 2.6
2008-12-02 | 12:26 | 4.65 | 0.049 | 490 | 604.5 | 2.6 | 607.1 | 380.1 | 380.0944 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.085 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.2911 | 11.5213 | 28 | 2.6
2008-12-02 | 12:26 | 4.65 | 0.049 | 490 | 604.4 | 2.6 | 607 | 380 | 380.0024 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0849 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.2888 | 11.4534 | 28 | 2.6
2008-12-02 | 12:26 | 4.65 | 0.049 | 490 | 604.4 | 2.6 | 607 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.8293 | 0.0849 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6885 | 28 | 2.6
2008-12-02 | 12:26 | 4.657 | 0.049 | 490 | 604.5 | 2.6 | 607.1 | 380.2 | 380.1972 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0849 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.4757 | 28 | 2.6
2008-12-02 | 12:26 | 4.671 | 0.049 | 490 | 604.5 | 2.6 | 607.1 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.0849 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5066 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.049 | 490 | 604.5 | 2.6 | 607.1 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.0849 | 0.0033 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.049 | 490 | 604.5 | 2.6 | 607.1 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0849 | 0.003271 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.0486 | 486.094 | 604.5 | 2.6 | 607.1 | 380.1 | 380.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0849 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.048 | 480 | 604.5 | 2.6 | 607.1 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0849 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.533 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.048 | 480 | 604.4 | 2.6 | 607 | 380.3 | 380.2901 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0849 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5 | 28 | 2.6
2008-12-02 | 12:26 | 4.68 | 0.048 | 480 | 604.4 | 2.6 | 607 | 380.1 | 380.1096 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0849 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5242 | 28 | 2.6
2008-12-02 | 12:26 | 4.676 | 0.048 | 480 | 604.3 | 2.6 | 606.9 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0848 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5973 | 28 | 2.6
2008-12-02 | 12:26 | 4.668 | 0.048 | 480 | 604.4 | 2.6 | 607 | 380.1 | 380.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.0848 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5193 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.048 | 480 | 604.4 | 2.6 | 607 | 380 | 380.0032 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.0848 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5594 | 28 | 2.6
2008-12-02 | 12:26 | 4.66 | 0.0478 | 477.8057 | 604.5 | 2.6 | 607.1 | 380.2 | 380.1856 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8711 | 0.0848 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6381 | 28 | 2.6
2008-12-02 | 12:26 | 3.947 | 0.0462 | 462.0382 | 604.3 | 2.6 | 606.9 | 380.1 | 380.0537 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0848 | 0.0032 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7166 | 28 | 2.6
2008-12-02 | 12:26 | 3.573 | 0.0382 | 382.4204 | 587.9 | 2.5712 | 590.5 | 380 | 379.9625 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.083102 | 0.003171 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7951 | 28 | 2.6
2008-12-02 | 12:26 | 3.498 | 0.0377 | 377.1044 | 496.1 | 2.5 | 498.6 | 359.4 | 359.4409 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8702 | 0.073009 | 0.0031 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.7016 | 11.6528 | 28 | 2.5
2008-12-02 | 12:26 | 3.48 | 0.0374 | 374.3302 | 459.3 | 2.5 | 461.8 | 302.8 | 302.7904 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8713 | 0.068771 | 0.0031 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.6 | 28 | 2.5
2008-12-02 | 12:26 | 3.48 | 0.037 | 370 | 458.3 | 2.4706 | 460.8 | 290.5 | 290.4686 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.068671 | 0.003071 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.6 | 28 | 2.5
2008-12-02 | 12:26 | 3.48 | 0.037 | 370 | 457.4 | 2.4 | 459.8 | 287.5 | 287.4553 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0685 | 0.003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6181 | 28 | 2.4
2008-12-02 | 12:26 | 3.48 | 0.037 | 370 | 457.3 | 2.3 | 459.6 | 286.8 | 286.8154 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0685 | 0.003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.775 | 28 | 2.3
2008-12-02 | 12:26 | 3.474 | 0.037 | 370 | 457.3 | 2.3 | 459.6 | 286.6 | 286.6 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0685 | 0.003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.734 | 28 | 2.3
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 457 | 2.3 | 459.3 | 286.4 | 286.4098 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.0684 | 0.003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7 | 28 | 2.3
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.9 | 2.3 | 459.2 | 286.4 | 286.4 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0684 | 0.003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6771 | 28 | 2.3
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.7 | 2.3 | 459 | 286.3 | 286.3061 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8288 | 0.0684 | 0.002971 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5986 | 28 | 2.3
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.6 | 2.3 | 458.9 | 286.4 | 286.3968 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.87 | 0.06837 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.5202 | 28 | 2.3
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.4 | 2.2 | 458.6 | 286.3 | 286.3 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.6167 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.4 | 2.2 | 458.6 | 286.3 | 286.3 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.7 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.4 | 2.2 | 458.6 | 286.2 | 286.2068 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8294 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.2945 | 11.7 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.4 | 2.2 | 458.6 | 286.2 | 286.2155 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.7 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.4 | 2.2 | 458.6 | 286.2 | 286.2467 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8709 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.2 | 2.2 | 458.4 | 286.2 | 286.2096 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.0683 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5986 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456.1 | 2.2 | 458.3 | 286.2 | 286.2 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.068271 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5293 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 456 | 2.2 | 458.2 | 286.2 | 286.2 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.0682 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 455.9 | 2.2 | 458.1 | 286.1 | 286.1041 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0682 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2.2
2008-12-02 | 12:26 | 3.47 | 0.037 | 370 | 455.8 | 2.1712 | 457.9 | 286.2 | 286.2 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0682 | 0.0029 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2.2
2008-12-02 | 12:26 | 3.469 | 0.0374 | 373.8804 | 455.7 | 2.1 | 457.8 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0682 | 0.002871 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6429 | 28 | 2.1
2008-12-02 | 12:26 | 3.461 | 0.038 | 380 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8713 | 0.068171 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6783 | 28 | 2.1
2008-12-02 | 12:26 | 3.46 | 0.038 | 380 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8294 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5998 | 28 | 2.1
2008-12-02 | 12:27 | 3.46 | 0.038 | 380 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5213 | 28 | 2.1
2008-12-02 | 12:27 | 3.464 | 0.038 | 380 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8708 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.038 | 380 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5706 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.0376 | 376.3105 | 455.7 | 2.1 | 457.8 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6719 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455.6 | 2.1 | 457.7 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5149 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455.4 | 2.1 | 457.5 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0681 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455.3 | 2.1 | 457.4 | 286 | 286 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.068071 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455.3 | 2.1 | 457.4 | 286 | 286 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8288 | 0.068 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4721 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455.1 | 2.1 | 457.2 | 286.1 | 286.0911 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8713 | 0.068 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4065 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 454.9 | 2.1 | 457 | 286 | 286 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8704 | 0.068 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4849 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 454.9 | 2.1 | 457 | 286 | 286 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.068 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5634 | 28 | 2.1
2008-12-02 | 12:27 | 3.47 | 0.037 | 370 | 455 | 2 | 457 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.068 | 0.0028 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6416 | 28 | 2
2008-12-02 | 12:27 | 3.465 | 0.037 | 370 | 455 | 2 | 457 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8292 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6802 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 455 | 2 | 457 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6012 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 455 | 2 | 457 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.9 | 2 | 456.9 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8711 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5442 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.9 | 2 | 456.9 | 286.2 | 286.1969 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2883 | 11.5682 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.8 | 2 | 456.8 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6752 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.8 | 2 | 456.8 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.068 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5175 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.8 | 2 | 456.8 | 286.1 | 286.1066 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.9 | 0.06797 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7789 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.037 | 370 | 454.7 | 2 | 456.7 | 286 | 286.0089 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8036 | 28 | 2
2008-12-02 | 12:27 | 3.458 | 0.037 | 370 | 454.7 | 2 | 456.7 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8707 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6733 | 28 | 2
2008-12-02 | 12:27 | 3.434 | 0.037 | 370 | 454.7 | 2 | 456.7 | 286.1 | 286.1035 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5896 | 28 | 2
2008-12-02 | 12:27 | 3.43 | 0.037 | 370 | 454.7 | 2 | 456.7 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.871 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4328 | 28 | 2
2008-12-02 | 12:27 | 3.43 | 0.037 | 370 | 454.7 | 2 | 456.7 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5863 | 28 | 2
2008-12-02 | 12:27 | 3.43 | 0.037 | 370 | 454.6 | 2 | 456.6 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8291 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 2
2008-12-02 | 12:27 | 3.438 | 0.0364 | 364.4482 | 454.7 | 2 | 456.7 | 286.1 | 286.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.681 | 28 | 2
2008-12-02 | 12:27 | 3.47 | 0.036 | 360 | 454.6 | 2 | 456.6 | 286.2 | 286.198 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0679 | 0.0027 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6025 | 28 | 2
2008-12-02 | 12:27 | 3.47 | 0.036 | 360 | 454.6 | 2 | 456.6 | 286.2 | 286.2 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.676 | 28 | 2
2008-12-02 | 12:27 | 3.47 | 0.036 | 360 | 454.4 | 2 | 456.4 | 286.1 | 286.1 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8287 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6455 | 28 | 2
2008-12-02 | 12:27 | 3.464 | 0.036 | 360 | 454.4 | 2 | 456.4 | 286.1 | 286.0932 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8704 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.036 | 360 | 454.4 | 2 | 456.4 | 286 | 286.0088 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.036 | 360 | 454.5 | 2 | 456.5 | 286 | 286.0071 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8707 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.036 | 360 | 454.5 | 2 | 456.5 | 286 | 286 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8708 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 2
2008-12-02 | 12:27 | 3.46 | 0.036 | 360 | 454.4 | 2 | 456.4 | 286 | 286 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8295 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5532 | 28 | 2
2008-12-02 | 12:27 | 2.931 | 0.0353 | 352.7356 | 454.5 | 2 | 456.5 | 286 | 286 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8707 | 0.0678 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5252 | 28 | 2
2008-12-02 | 12:27 | 2.496 | 0.0244 | 244.3119 | 441.6 | 2 | 443.6 | 285.9 | 285.91 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.066258 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6068 | 28 | 2
2008-12-02 | 12:27 | 2.306 | 0.0244 | 244.2295 | 357.7 | 1.9711 | 359.7 | 280.5 | 280.5378 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.056929 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7643 | 28 | 2
2008-12-02 | 12:27 | 2.3 | 0.025 | 250 | 312.5 | 1.8712 | 314.3 | 219.2 | 219.155 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.050185 | 0.0026 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6178 | 28 | 1.9
2008-12-02 | 12:27 | 2.3 | 0.025 | 250 | 309.6 | 1.8 | 311.4 | 197.5 | 197.4794 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195 | 59.9 | 0.049771 | 0.002571 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7123 | 11.5392 | 28 | 1.8
2008-12-02 | 12:27 | 2.3 | 0.025 | 250 | 309.5 | 1.8 | 311.3 | 193.8 | 193.7874 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8287 | 0.0497 | 0.0025 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 1.8
2008-12-02 | 12:27 | 2.293 | 0.025 | 250 | 309.4 | 1.8 | 311.2 | 193 | 193.0396 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.04967 | 0.00247 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 1.8
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 309.2 | 1.8 | 311 | 192.8 | 192.797 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.049483 | 0.002312 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 1.8
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308.9 | 1.8 | 310.7 | 192.6 | 192.6 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0492 | 0.0021 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5469 | 28 | 1.8
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308.4 | 1.8 | 310.2 | 192.4 | 192.4041 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0491 | 0.0021 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5631 | 28 | 1.8
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308.2 | 1.7 | 309.9 | 192.4 | 192.4 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.049071 | 0.002 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308.2 | 1.7 | 309.9 | 192.4 | 192.4 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.049 | 0.002 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308.1 | 1.7 | 309.8 | 192.4 | 192.4 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8289 | 0.049 | 0.002 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 308 | 1.7 | 309.7 | 192.3 | 192.3 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.049 | 0.002 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 307.9 | 1.7 | 309.6 | 192.3 | 192.3 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.048971 | 0.002 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7239 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 307.9 | 1.7 | 309.6 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8298 | 0.0489 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7952 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 307.8 | 1.7 | 309.5 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8704 | 0.0489 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6394 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 307.7 | 1.7 | 309.4 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8294 | 0.0489 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6591 | 28 | 1.7
2008-12-02 | 12:27 | 2.29 | 0.025 | 250 | 307.7 | 1.7 | 309.4 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.8293 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.7
2008-12-02 | 12:27 | 2.279 | 0.025 | 250 | 307.7 | 1.6708 | 309.4 | 192.2 | 192.2 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8708 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7162 | 28 | 1.7
2008-12-02 | 12:27 | 2.27 | 0.025 | 250 | 307.7 | 1.6 | 309.3 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7947 | 28 | 1.6
2008-12-02 | 12:27 | 2.27 | 0.025 | 250 | 307.6 | 1.6 | 309.2 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8293 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7271 | 28 | 1.6
2008-12-02 | 12:27 | 2.27 | 0.025 | 250 | 307.4 | 1.6 | 309 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.803 | 28 | 1.6
2008-12-02 | 12:27 | 2.27 | 0.025 | 250 | 307.4 | 1.6 | 309 | 192.2 | 192.2 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8397 | 28 | 1.6
2008-12-02 | 12:27 | 2.277 | 0.025 | 250 | 307.3 | 1.6 | 308.9 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7086 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.025 | 250 | 307.2 | 1.6 | 308.8 | 192.1 | 192.1 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0488 | 0.0019 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7871 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.0257 | 257.0462 | 307.2 | 1.6 | 308.8 | 192.2 | 192.1951 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8287 | 0.048771 | 0.001871 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6689 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.0249 | 249.2003 | 307.1 | 1.6 | 308.7 | 192.2 | 192.2 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8294 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.644 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 307 | 1.6 | 308.6 | 192.2 | 192.2 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.9 | 1.6 | 308.5 | 192.1 | 192.1212 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.0241 | 240.5574 | 306.8 | 1.6 | 308.4 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8295 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.0249 | 248.6349 | 306.8 | 1.5707 | 308.4 | 192.2 | 192.2 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8293 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7576 | 28 | 1.6
2008-12-02 | 12:28 | 2.28 | 0.025 | 250 | 306.7 | 1.5 | 308.2 | 192.2 | 192.2 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8709 | 0.0487 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.0245 | 244.676 | 306.8 | 1.5 | 308.3 | 192.2 | 192.2 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7852 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.7 | 1.5 | 308.2 | 192.2 | 192.1891 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8288 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7067 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.8 | 1.5 | 308.3 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8288 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6282 | 28 | 1.5
2008-12-02 | 12:28 | 2.287 | 0.0241 | 240.5565 | 306.8 | 1.5 | 308.3 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195.5 | 59.8713 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 1.5
2008-12-02 | 12:28 | 2.285 | 0.0249 | 248.621 | 306.8 | 1.5 | 308.3 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8296 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.0243 | 242.9447 | 306.7 | 1.5 | 308.2 | 192 | 192.0019 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.7 | 1.5 | 308.2 | 192 | 192 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.7 | 1.5 | 308.2 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6641 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8707 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192 | 192.0071 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192 | 192 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7111 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.544 | 28 | 1.5
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8288 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6128 | 28 | 1.5
2008-12-02 | 12:28 | 2.281 | 0.024 | 240 | 306.6 | 1.5 | 308.1 | 192 | 192.0055 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0486 | 0.0018 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.289 | 0.024 | 240 | 306.5 | 1.5 | 308 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8296 | 0.0486 | 0.00177 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.29 | 0.024 | 240 | 306.5 | 1.5 | 308 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | 1.5
2008-12-02 | 12:28 | 2.29 | 0.024 | 240 | 306.5 | 1.5 | 308 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7703 | 28 | 1.5
2008-12-02 | 12:28 | 2.29 | 0.024 | 240 | 306.4 | 1.5 | 307.9 | 192.1 | 192.1 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8296 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 1.5
2008-12-02 | 12:28 | 2.288 | 0.024 | 240 | 306.3 | 1.5 | 307.8 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.871 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.773 | 28 | 1.5
2008-12-02 | 12:28 | 2.281 | 0.024 | 240 | 306.3 | 1.4 | 307.7 | 192.2 | 192.194 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7057 | 28 | 1.4
2008-12-02 | 12:28 | 2.289 | 0.024 | 240 | 306.2 | 1.4 | 307.6 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7841 | 28 | 1.4
2008-12-02 | 12:28 | 2.283 | 0.024 | 240 | 306.2 | 1.4 | 307.6 | 192.1 | 192.1 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7374 | 28 | 1.4
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.2 | 1.4 | 307.6 | 192.1 | 192.1 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8712 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7409 | 28 | 1.4
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.1 | 1.4 | 307.5 | 192 | 192.0044 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8288 | 0.0485 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7808 | 28 | 1.4
2008-12-02 | 12:28 | 2.28 | 0.024 | 240 | 306.2 | 1.4 | 307.6 | 192 | 192 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.04847 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.702 | 28 | 1.4
2008-12-02 | 12:28 | 2.272 | 0.024 | 240 | 306.2 | 1.4 | 307.6 | 192 | 192.0108 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8704 | 0.0484 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 1.4
2008-12-02 | 12:28 | 2.151 | 0.0213 | 212.8485 | 306.2 | 1.4 | 307.6 | 192 | 192 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8706 | 0.0484 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6451 | 28 | 1.4
2008-12-02 | 12:28 | 1.448 | 0.0102 | 101.66 | 280 | 1.4 | 281.4 | 192 | 191.9952 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.044861 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7069 | 11.6334 | 28 | 1.4
2008-12-02 | 12:28 | 1.16 | 0.0126 | 125.7177 | 232.6 | 1.4 | 234 | 190.9 | 190.9385 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0383 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7237 | 28 | 1.4
2008-12-02 | 12:28 | 1.141 | 0.013 | 130 | 164.5 | 1.4 | 165.9 | 134.6 | 134.6093 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.028252 | 0.0017 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.8807 | 28 | 1.4
2008-12-02 | 12:28 | 1.11 | 0.013 | 130 | 157.8 | 1.3707 | 159.2 | 102.7 | 102.713 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.027171 | 0.0016 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7624 | 28 | 1.4
2008-12-02 | 12:28 | 1.11 | 0.013 | 130 | 157 | 1.3 | 158.3 | 97.8 | 97.755 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8291 | 0.027071 | 0.0016 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7473 | 28 | 1.3
2008-12-02 | 12:28 | 1.11 | 0.013 | 130 | 156.9 | 1.3 | 158.2 | 96.7 | 96.7339 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8711 | 0.027 | 0.0016 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7488 | 28 | 1.3
2008-12-02 | 12:28 | 1.11 | 0.013 | 130 | 156.5 | 1.3 | 157.8 | 96.4 | 96.4042 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.026885 | 0.001514 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 1.3
2008-12-02 | 12:28 | 1.115 | 0.013 | 130 | 156 | 1.3 | 157.3 | 96.3 | 96.3032 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8713 | 0.026571 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2871 | 11.6 | 28 | 1.3
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155.5 | 1.2 | 156.7 | 96.2 | 96.2 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 1.2
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155.5 | 1.1706 | 156.7 | 96.1 | 96.1 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8294 | 0.0265 | 0.0012 | -0.000129 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6396 | 28 | 1.2
2008-12-02 | 12:28 | 1.12 | 0.0126 | 126.4967 | 155.4 | 1.1 | 156.5 | 96 | 96 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.718 | 28 | 1.1
2008-12-02 | 12:28 | 1.129 | 0.012 | 120 | 155.4 | 1.1 | 156.5 | 96 | 96 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8708 | 0.026471 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7963 | 28 | 1.1
2008-12-02 | 12:28 | 1.13 | 0.012 | 120 | 155.4 | 1.1 | 156.5 | 95.9 | 95.9 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.0264 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 1.1
2008-12-02 | 12:28 | 1.13 | 0.012 | 120 | 155.4 | 1.1 | 156.5 | 95.9 | 95.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8707 | 0.0264 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 1.1
2008-12-02 | 12:28 | 1.13 | 0.012 | 120 | 155.3 | 1.1 | 156.4 | 95.9 | 95.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8291 | 0.0264 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7363 | 28 | 1.1
2008-12-02 | 12:28 | 1.13 | 0.012 | 120 | 155.3 | 1.1 | 156.4 | 95.8 | 95.8063 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8711 | 0.0264 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6309 | 28 | 1.1
2008-12-02 | 12:28 | 1.13 | 0.0123 | 123.495 | 155.2 | 1.1 | 156.3 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8288 | 0.0263 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.8647 | 28 | 1.1
2008-12-02 | 12:28 | 1.121 | 0.013 | 130 | 155.1 | 1 | 156.1 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8288 | 0.0263 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8329 | 28 | 1
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155 | 0.97128 | 156 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8713 | 0.0263 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7543 | 28 | 1
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155 | 0.9 | 155.9 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.02627 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7242 | 28 | 0.9
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155 | 0.9 | 155.9 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 0.9
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155 | 0.9 | 155.9 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8293 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 0.9
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 155 | 0.9 | 155.9 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8708 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 0.9
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 154.9 | 0.9 | 155.8 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.8 | 28 | 0.9
2008-12-02 | 12:28 | 1.12 | 0.013 | 130 | 154.9 | 0.9 | 155.8 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.767 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.013 | 130 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8289 | 0.0262 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.61 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.013 | 130 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0262 | 0.001071 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7458 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.013 | 130 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7485 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.013 | 130 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8298 | 0.0262 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 0.9
2008-12-02 | 12:29 | 1.129 | 0.013 | 130 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.0262 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7036 | 11.6912 | 28 | 0.9
2008-12-02 | 12:29 | 1.13 | 0.013 | 129.5948 | 154.8 | 0.9 | 155.7 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8706 | 0.026171 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6128 | 28 | 0.9
2008-12-02 | 12:29 | 1.13 | 0.0121 | 121.4911 | 154.7 | 0.9 | 155.6 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.3 | 28 | 5 | 195 | 59.8 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5346 | 28 | 0.9
2008-12-02 | 12:29 | 1.13 | 0.012 | 120 | 154.6 | 0.9 | 155.5 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8708 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7083 | 11.5879 | 28 | 0.9
2008-12-02 | 12:29 | 1.128 | 0.012 | 120 | 154.6 | 0.9 | 155.5 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.829 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6322 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.6 | 0.9 | 155.5 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4031 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.0122 | 121.9092 | 154.6 | 0.9 | 155.5 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8709 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6387 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.013 | 129.5575 | 154.5 | 0.9 | 155.4 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8289 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.0122 | 121.5125 | 154.5 | 0.9 | 155.4 | 95.7 | 95.7 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8712 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6635 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.5 | 0.9 | 155.4 | 95.8 | 95.7728 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8288 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6149 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.5 | 0.8697 | 155.4 | 95.8 | 95.7911 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6934 | 28 | 0.9
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.4 | 0.8 | 155.2 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8294 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7718 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.4 | 0.8 | 155.2 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6492 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.4 | 0.8 | 155.2 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5861 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.8 | 95.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8295 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7857 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.8 | 95.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6287 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.7 | 95.7 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8291 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5359 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.4 | 0.8 | 155.2 | 95.8 | 95.7923 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5426 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.4 | 0.8 | 155.2 | 95.8 | 95.7891 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8712 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.7 | 95.7 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8713 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2871 | 11.522 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.7 | 95.7 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.3 | 0.8 | 155.1 | 95.7 | 95.7093 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4652 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.2 | 0.8 | 155 | 95.8 | 95.7954 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.3 | 28 | 5 | 195.2 | 59.8707 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4265 | 28 | 0.8
2008-12-02 | 12:29 | 1.12 | 0.012 | 120 | 154.2 | 0.8 | 155 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5835 | 28 | 0.8
2008-12-02 | 12:29 | 1.134 | 0.012 | 120 | 154.2 | 0.8 | 155 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5298 | 28 | 0.8
2008-12-02 | 12:29 | 1.14 | 0.012 | 120 | 154.2 | 0.8 | 155 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8293 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4513 | 28 | 0.8
2008-12-02 | 12:29 | 1.14 | 0.012 | 120 | 154.2 | 0.8 | 155 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8709 | 0.026 | 0.000971 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4271 | 28 | 0.8
2008-12-02 | 12:29 | 1.14 | 0.012 | 120 | 154.2 | 0.77111 | 155 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.026 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5055 | 28 | 0.8
2008-12-02 | 12:29 | 1.123 | 0.012 | 120 | 154.2 | 0.7 | 154.9 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8 | 0.026 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5841 | 28 | 0.7
2008-12-02 | 12:29 | 0.559 | 0.0089 | 88.8843 | 150 | 0.7 | 150.7 | 95.8 | 95.8 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.025971 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | 0.7
2008-12-02 | 12:29 | 0.149 | -0.0029 | -29.2991 | 106.6 | 0.7 | 107.3 | 95.8 | 95.7764 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8298 | 0.018508 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5592 | 28 | 0.7
2008-12-02 | 12:29 | 0.015 | -0.0012 | -12.2051 | 75.2 | 0.7 | 75.9 | 82.3 | 82.2751 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8704 | 0.013429 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5192 | 28 | 0.7
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 23.3 | 0.6 | 23.9 | 22.2 | 22.1957 | 0.1 | 13.2 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.004823 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5978 | 28 | 0.6
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 4.8 | 0.23676 | 5 | 4.6 | 4.5761 | 0.1 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0017 | 0.0009 | -0.000229 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | 0.2
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 3.9 | -0.4 | 3.9 | 1.2 | 1.1578 | 0.1 | 13.2 | 905 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.001571 | 0.0009 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2917 | 11.6 | 28 | -0.4
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 3.7 | -0.4 | 3.7 | 0.3 | 0.3144 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.829 | 0.0015 | 0.0009 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5668 | 28 | -0.4
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 3.1 | -0.4 | 3.1 | 0 | 0.0086 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.001371 | 0.0009 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5351 | 28 | -0.4
2008-12-02 | 12:29 | 0.01 | -0.001 | -10 | 2.7 | -0.4 | 2.7 | 0 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0013 | 0.000871 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7111 | 11.7704 | 28 | -0.4
2008-12-02 | 12:29 | 0.006 | -0.001 | -10 | 2.5 | -0.5 | 2.5 | 0 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8712 | 0.0012 | 0.0008 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2883 | 11.5262 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.4 | -0.5 | 2.4 | 0 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0012 | 0.0008 | -0.000271 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7123 | 11.4939 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.4 | -0.5 | 2.4 | 0 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.7 | 28 | 5 | 195 | 59.9 | 0.0012 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6254 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.3 | -0.5 | 2.3 | -0.1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0012 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.3 | -0.5 | 2.3 | -0.1 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8699 | 0.0011 | 0.0008 | -0.00023 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.3 | -0.52931 | 2.3 | -0.1 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0011 | 0.0008 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6392 | 28 | -0.5
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 2.3 | -0.6 | 2.3 | -0.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8292 | 0.0011 | 0.0008 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5216 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.8 | -0.6 | 1.8 | -0.2 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.829 | 0.001071 | 0.0007 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7097 | 11.4177 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.5 | -0.6 | 1.5 | -0.2 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.001 | 0.0007 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4962 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.5 | -0.6 | 1.5 | -0.2 | 0 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.001 | 0.0007 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6493 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.5 | -0.6 | 1.5 | -0.2 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.000971 | 0.0007 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.2889 | 11.5938 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.4 | -0.6 | 1.4 | -0.2 | 0 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0009 | 0.000671 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.4 | -0.6 | 1.4 | -0.2 | 0 | 0.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8288 | 0.0009 | 0.0006 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7123 | 11.4901 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.3 | -0.6 | 1.3 | -0.2 | 0 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.000871 | 0.0006 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.412 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.3 | -0.6 | 1.3 | -0.3 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.0008 | 0.0006 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4667 | 28 | -0.6
2008-12-02 | 12:29 | 0 | -0.001 | -10 | 1.3 | -0.6 | 1.3 | -0.3 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0008 | 0.0006 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5906 | 28 | -0.6
2008-12-02 | 12:29 | 0.007 | -0.001 | -10 | 1.3 | -0.7 | 1.3 | -0.3 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8 | 0.0008 | 0.0006 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6762 | 28 | -0.7
2008-12-02 | 12:30 | 0.005 | -0.001 | -10 | 1.2 | -0.7 | 1.2 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0008 | 0.0006 | -0.000271 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5978 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.2 | -0.6 | 1.2 | -0.3 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8707 | 0.0007 | 0.000571 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5194 | 28 | -0.6
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.2 | -0.62909 | 1.2 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0007 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6182 | 28 | -0.6
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.1 | -0.7 | 1.1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.7 | 28 | 5 | 195 | 59.8 | 0.0007 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.7 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.1 | -0.7 | 1.1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0007 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7117 | 11.668 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.1 | -0.7 | 1.1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0007 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5111 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1.1 | -0.7 | 1.1 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0006 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4274 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.7 | 1 | -0.3 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0006 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4511 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.7 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0006 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4702 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.7 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0005 | 0.000371 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4082 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.7 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8291 | 0.0005 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4867 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.7 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.0005 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5646 | 28 | -0.7
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 1 | -0.8 | 1 | -0.3 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8293 | 0.0005 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5138 | 28 | -0.8
2008-12-02 | 12:30 | 0.003 | -0.001 | -10 | 1 | -0.8 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.829 | 0.0005 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.422 | 28 | -0.8
2008-12-02 | 12:30 | 0.01 | -0.001 | -10 | 1 | -0.8 | 1 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8288 | 0.0005 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4995 | 28 | -0.8
2008-12-02 | 12:30 | 0.01 | -0.001 | -10 | 0.9 | -0.8 | 0.9 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8712 | 0.0004 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4211 | 28 | -0.8
2008-12-02 | 12:30 | 0.01 | -0.001 | -10 | 0.9 | -0.8 | 0.9 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8713 | 0.0004 | 0.0003 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -0.8
2008-12-02 | 12:30 | 0.01 | -0.001 | -10 | 0.9 | -0.8 | 0.9 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8296 | 0.0004 | 0.0003 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5422 | 28 | -0.8
2008-12-02 | 12:30 | 0.005 | -0.001 | -10 | 0.8 | -0.9 | 0.8 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0004 | 0.0003 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.7574 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.5 | -0.9 | 0.5 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5219 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.5 | -0.9 | 0.5 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.871 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5492 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6275 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.687 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5301 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8712 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5634 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8288 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.6 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.56 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -0.9 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4027 | 28 | -0.9
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8706 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.5112 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8292 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.7083 | 11.6 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0003 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.6 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8709 | 0.0002 | 0.0002 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5306 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8711 | 0.0002 | 0.0001 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.5 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.3 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4738 | 28 | -1
2008-12-02 | 12:30 | 0 | -0.001 | -10 | 0.4 | -1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0003 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -1
Summary Information:
Post Processor DLL Version | 6.01
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-12-02
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 66
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 28261
Version | 6.907
CO Span(%) | 0.12
CO2 Span(%) | 11.9
C6H14 Span(ppm) | 2005
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 131
Version | 1.611
NO Span(ppm) | 1504
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 957.0
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | Backhoe J Deere MN 700212 310G
License Plate
Engine Displacement | 4.5
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 12:09
Test End Time | 12:30
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 991
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iAMBII_RPM
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
Faults:
0X5009 - 12/02/2008 11:06:44.976 - FID battery backed RAM corrupt
Warnings:
0X0000 - 12/02/2008 12:09:05.325 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 124
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-12-02 | 12:09 | [RELEASE_VER=2.018 BUILD=66 BDATE=10/21/2008 IP=10.10.1.55]
No Information Available

## THC Lin
Test Date:
Test Time:
12/2/2008
12:09:05
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## CO2 Lin
Test Date:
Test Time:
12/2/2008
12:09:05
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## CO Lin
Test Date:
Test Time:
12/2/2008
12:09:05
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## NO Lin
Test Date:
Test Time:
12/2/2008
12:09:05
NDUV Serial No:
131
SEMTECH Serial Number:
G05-SDS02

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
12/2/2008
12:09:05
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
12/2/2008
12:09:05
NDUV Serial No:
131
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
11.99 | Rm = | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555 | 0.98555
| C.F.= | 0.098698 | 0.19768 | 0.29695 | 0.39651 | 0.49636 | 0.5965 | 0.69693 | 0.79766 | 0.89868
| gas = | 1.1834 | 2.3702 | 3.5605 | 4.7542 | 5.9514 | 7.1521 | 8.3562 | 9.564 | 10.7752
Ro for CO2= | 0.8795