## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 321.8 | Scott gas | ALM054255 | ±1% | 2010-06-16 | 965.4
1200 ppm CO | 1205
12% CO2 | 12
1500 ppm NO | 1505
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.12 | Scott gas | ALM004744 | ±1% | 2010-06-04 | 93.36
200 ppm CO | 200.7
6% CO2 | 6.009
300 ppm NO | 297.2
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 37249 |  |  | CO - AMBII Serial Number: |  |  | 37249 |  | HC - AMBII Serial Number: |  |  | 37249 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-07-16 |  | 12 | X | 0.1 | 3 |  |  | 1205 |  | 50 | 3 |  | 321.8 |  | 4 | 3 |  | 1505 |  | 15 | 3 |  | 965.4 | 1000 | 5 | 2
Test Time: | 12:55
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | -0.002 | -0.002 |  |  | 0 | -7 | -7 |  |  | 0 | 1012.8 | 1012.8 |  |  | 0 | 0.6 | 0.6 |  |  | 0 | 0.1 | 0.1
90 | 105 | 1 | 1.2 | 1.184 | 1.16 | -0.024 | -2.03 |  | 120.5 | 128 | 8 | 6.64 |  | 32.18 | 1013.1 | 980.9 | 3048.17 | Fail | 150.5 | 150.1 | -0.4 | -0.27 |  | 96.54 | 96.7 | 0.2 | 0.21
150 | 165 | 2 | 2.4 | 2.372 | 2.36 | -0.012 | -0.51 |  | 241 | 246 | 5 | 2.07 |  | 64.36 | 1012.4 | 948 | 1472.96 | Fail | 301 | 297.5 | -3.5 | -1.16 |  | 193.08 | 192.2 | -0.9 | -0.47
210 | 225 | 3 | 3.6 | 3.563 | 3.61 | 0.047 | 1.32 |  | 361.5 | 367 | 6 | 1.66 |  | 96.54 | 1012.9 | 916.4 | 949.24 | Fail | 451.5 | 448.8 | -2.7 | -0.6 |  | 289.62 | 288 | -1.6 | -0.55
270 | 285 | 4 | 4.8 | 4.758 | 4.867 | 0.109 | 2.29 |  | 482 | 484 | 2 | 0.41 |  | 128.72 | 1012.9 | 884.2 | 686.92 | Fail | 602 | 600.7 | -1.3 | -0.22 |  | 386.16 | 383.9 | -2.3 | -0.6
330 | 345 | 5 | 6 | 5.956 | 6.118 | 0.162 | 2.72 |  | 602.5 | 600 | -3 | -0.5 |  | 160.9 | 1013 | 852.1 | 529.58 | Fail | 752.5 | 750.4 | -2.1 | -0.28 |  | 482.7 | 479.5 | -3.2 | -0.66
390 | 405 | 6 | 7.2 | 7.158 | 7.315 | 0.157 | 2.19 |  | 723 | 711 | -12 | -1.66 |  | 193.08 | 1012.9 | 819.8 | 424.59 | Fail | 903 | 902 | -1 | -0.11 |  | 579.24 | 576.1 | -3.1 | -0.54
450 | 465 | 7 | 8.4 | 8.363 | 8.538 | 0.175 | 2.09 |  | 843.5 | 835 | -9 | -1.07 |  | 225.26 | 1013.9 | 788.6 | 350.08 | Fail | 1053.5 | 1049.9 | -3.6 | -0.34 |  | 675.78 | 673.1 | -2.7 | -0.4
510 | 525 | 8 | 9.6 | 9.572 | 9.72 | 0.148 | 1.55 |  | 964 | 960 | -4 | -0.41 |  | 257.44 | 1013.3 | 755.9 | 293.62 | Fail | 1204 | 1201.7 | -2.3 | -0.19 |  | 772.32 | 770.3 | -2 | -0.26
570 | 585 | 9 | 10.8 | 10.784 | 10.872 | 0.088 | 0.82 |  | 1084.5 | 1074 | -11 | -1.01 |  | 289.62 | 1013.4 | 723.8 | 249.91 | Fail | 1354.5 | 1351 | -3.5 | -0.26 |  | 868.86 | 868.4 | -0.5 | -0.06
630 | 645 | 10 | 12 | 12 | 12.007 | 0.007 | 0.06 |  | 1205 | 1200 | -5 | -0.41 |  | 321.8 | 1012.4 | 690.6 | 214.61 | Fail | 1505 | 1507.5 | 2.5 | 0.17 |  | 965.4 | 967.7 | 2.3 | 0.24
690 | 705 | 9 | 10.8 | 10.784 | 10.861 | 0.077 | 0.71 |  | 1084.5 | 1070 | -15 | -1.38 |  | 289.62 | 1013 | 723.4 | 249.78 | Fail | 1354.5 | 1356.1 | 1.6 | 0.12 |  | 868.86 | 869.1 | 0.2 | 0.02
750 | 765 | 8 | 9.6 | 9.572 | 9.72 | 0.148 | 1.55 |  | 964 | 960 | -4 | -0.41 |  | 257.44 | 1012.5 | 755.1 | 293.31 | Fail | 1204 | 1201.5 | -2.5 | -0.21 |  | 772.32 | 770 | -2.3 | -0.3
810 | 825 | 7 | 8.4 | 8.363 | 8.532 | 0.169 | 2.02 |  | 843.5 | 840 | -4 | -0.47 |  | 225.26 | 1011.6 | 786.3 | 349.06 | Fail | 1053.5 | 1052.4 | -1.1 | -0.1 |  | 675.78 | 672.5 | -3.3 | -0.49
870 | 885 | 6 | 7.2 | 7.158 | 7.338 | 0.18 | 2.51 |  | 723 | 721 | -2 | -0.28 |  | 193.08 | 1012.9 | 819.8 | 424.59 | Fail | 903 | 901.9 | -1.1 | -0.12 |  | 579.24 | 576.3 | -2.9 | -0.5
930 | 945 | 5 | 6 | 5.956 | 6.11 | 0.154 | 2.59 |  | 602.5 | 600 | -3 | -0.5 |  | 160.9 | 1012.8 | 851.9 | 529.46 | Fail | 752.5 | 749.8 | -2.7 | -0.36 |  | 482.7 | 479.7 | -3 | -0.62
990 | 1005 | 4 | 4.8 | 4.758 | 4.88 | 0.122 | 2.56 |  | 482 | 480 | -2 | -0.41 |  | 128.72 | 1012.1 | 883.4 | 686.3 | Fail | 602 | 601.6 | -0.4 | -0.07 |  | 386.16 | 384.2 | -2 | -0.52
1050 | 1065 | 3 | 3.6 | 3.563 | 3.612 | 0.049 | 1.38 |  | 361.5 | 360 | -2 | -0.55 |  | 96.54 | 1012.4 | 915.9 | 948.73 | Fail | 451.5 | 449.9 | -1.6 | -0.35 |  | 289.62 | 288.5 | -1.1 | -0.38
1110 | 1125 | 2 | 2.4 | 2.372 | 2.365 | -0.007 | -0.3 |  | 241 | 240 | -1 | -0.41 |  | 64.36 | 1011.6 | 947.2 | 1471.72 | Fail | 301 | 300.5 | -0.5 | -0.17 |  | 193.08 | 192.2 | -0.9 | -0.47
1170 | 1185 | 1 | 1.2 | 1.184 | 1.16 | -0.024 | -2.03 |  | 120.5 | 120 | -1 | -0.83 |  | 32.18 | 1012.4 | 980.2 | 3045.99 | Fail | 150.5 | 151.5 | 1 | 0.66 |  | 96.54 | 96.2 | -0.3 | -0.31
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -10 | -10 |  |  | 0 | 1011.9 | 1011.9 |  |  | 0 | -0.8 | -0.8 |  |  | 0 | -0.1 | -0.1
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | 0.013 | 0.12 |  |  |  | Intercept | 1 | 12 |  |  | Intercept | 1012 | 3 | Fail |  | Intercept | -1.1 | 15.1 |  |  | Intercept | -1.1 | 9.7
|  |  | Slope | 1.01 | 0.99 - 1.01 |  |  |  | Slope | 0.99 | 0.99 - 1.01 |  |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0622 | 0.12 |  |  |  | SEE | 5.1846 | 12.05 |  |  | SEE | 0.5556 | 3.218 |  |  | SEE | 1.7187 | 15.05 |  |  | SEE | 1.511 | 9.677
|  |  | r2 | 0.9997 | 0.998 |  |  |  | r2 | 0.9998 | 0.998 |  |  | r2 | 0.1138 | 0.998 | Fail |  | r2 | 1 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-07-16 | 12:55 | 0.01 | -0.001 | -10 | -0.4 | 1.3 | 1.3 | 0.7 | 0.7349 | 0.1 | 12.9 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0016 | 0.0017 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.8 | 28.2 | 1.3
2008-07-16 | 12:55 | 0.007 | -0.001 | -10 | -0.3 | 1.2087 | 1.2 | 0.6 | 0.6 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0016 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7308 | 28.2 | 1.2
2008-07-16 | 12:55 | 0 | -0.001 | -10 | -0.3 | 1.2 | 1.2 | 0.7 | 0.7 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012.9125 | 30.6542 | 28.2 | 1.2
2008-07-16 | 12:55 | 0 | -0.001 | -10 | -0.2 | 1.1088 | 1.1 | 0.5 | 0.5261 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.001409 | 0.0015 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.645 | 28.2 | 1.1
2008-07-16 | 12:55 | 0 | -0.0008 | -8.1942 | -0.1 | 1.1 | 1.1 | 0.6 | 0.6 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0014 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.7985 | 28.2 | 1.1
2008-07-16 | 12:55 | 0 | 0 | -0.16854 | -0.1 | 1.0079 | 1 | 0.4 | 0.4 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.001308 | 0.0014 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012.0788 | 30.7241 | 28.2 | 1
2008-07-16 | 12:55 | 0 | 0 | 0 | 0 | 1 | 1 | 0.5 | 0.4951 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.0013 | 0.0013 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7525 | 28.2 | 1
2008-07-16 | 12:55 | 0 | 0 | 0 | 0 | 1 | 1 | 0.4 | 0.4483 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.001208 | 0.001208 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.8 | 28.2 | 1
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.1 | 1 | 1.1 | 0.4 | 0.4161 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.0012 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7945 | 28.2 | 1
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.2 | 0.9088 | 1.1 | 0.5 | 0.4501 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.001109 | 0.0011 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.7181 | 28.2 | 0.9
2008-07-16 | 12:55 | 0 | -0.0001 | -1.4821 | 0.2 | 0.9 | 1.1 | 0.4 | 0.351 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0011 | 0.0011 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6413 | 28.2 | 0.9
2008-07-16 | 12:55 | 0 | -0.001 | -9.8085 | 0.2 | 0.9 | 1.1 | 0.3 | 0.3324 | 0.1 | 12.9 | 950 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.001009 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.5645 | 28.2 | 0.9
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.2 | 0.8 | 1 | 0.3 | 0.3 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.5121 | 28.2121 | 0.8
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.8 | 1.1 | 0.4 | 0.4 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0.0001 | 9.1e-05 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.5889 | 28.2889 | 0.8
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.8 | 1.1 | 0.2 | 0.2475 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6656 | 28.2344 | 0.8
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.8 | 1.1 | 0.3 | 0.2758 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.001 | 0.0009 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7423 | 28.2 | 0.8
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.70825 | 1 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7811 | 28.2 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.7 | 1 | 0.4 | 0.3518 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7044 | 28.2 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.7 | 1 | 0.2 | 0.2494 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3083 | 0.0009 | 0.000808 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011.917 | 30.5556 | 28.2 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.3 | 0.7 | 1 | 0.3 | 0.2572 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.0009 | 0.0008 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6464 | 28.2 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | 0.2 | 0.2476 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.000709 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.8 | 28.2 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.0008 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.8 | 28.2021 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | 0.2 | 0.234 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.8 | 28.2787 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3075 | 0.0008 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.634 | 28.2447 | 0.7
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.60878 | 1 | 0.3 | 0.3 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.532 | 28.2 | 0.6
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.6 | 1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3088 | 0.0008 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.5914 | 28.2 | 0.6
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.4 | 0.6 | 1 | 0.2 | 0.1751 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0007 | 0.0007 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.5148 | 28.2 | 0.6
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.5 | 0.6 | 1.1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6855 | 28.2 | 0.6
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.5 | 0.6 | 1.1 | 0.2 | 0.1528 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.4 | 0.0007 | 0.0006 | 0.0001 | 9.2e-05 |  | n/a |  | 33 | 55 | 1048 | 1012.9157 | 30.723 | 28.2385 | 0.6
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.5 | 0.50862 | 1 | 0.1 | 0.1485 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0007 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.5848 | 28.3 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.2 | 0.1599 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.000609 | 0.000509 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.5082 | 28.3 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.2 | 0.1513 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0006 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.7054 | 28.2315 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0006 | 0.000409 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7097 | 28.2 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.2 | 0.2 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012.9125 | 30.6 | 28.2 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.1 | 0.1261 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6 | 28.2 | 0.5
2008-07-16 | 12:55 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.000509 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012.9122 | 30.7495 | 28.2 | 0.5
2008-07-16 | 12:55 | -0.005 | -0.001 | -10 | 0.6 | 0.40788 | 1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.7489 | 28.2 | 0.4
2008-07-16 | 12:55 | -0.01 | -0.001 | -10 | 0.6 | 0.4 | 1 | 0.2 | 0.1755 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013.9192 | 30.7277 | 28.2 | 0.4
2008-07-16 | 12:55 | -0.01 | -0.001 | -10 | 0.7 | 0.3 | 1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1014 | 30.7905 | 28.2 | 0.3
2008-07-16 | 12:55 | -0.01 | -0.0007 | -6.9641 | 0.6 | 0.3 | 0.9 | 0.1 | 0.1032 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1014 | 30.6374 | 28.2 | 0.3
2008-07-16 | 12:55 | -0.002 | 0 | 0 | 0.6 | 0.3 | 0.9 | 0.1 | 0.1479 | 0.1 | 12.9 | 951 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013.0882 | 30.7162 | 28.2581 | 0.3
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.3 | 0.9 | 0.1 | 0.06 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.6954 | 28.3 | 0.3
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.3 | 0.9 | 0.1 | 0.1479 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0005 | 0.0004 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012.0873 | 30.5115 | 28.2885 | 0.3
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.20877 | 0.8 | 0.1 | 0.0518 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0005 | 0.000309 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011.0876 | 30.5881 | 28.2119 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1323 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0005 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.6 | 28.2648 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0475 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.6 | 28.2586 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6 | 28.218 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6 | 28.2946 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.2 | 0.1519 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.7425 | 28.2287 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0133 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.7042 | 28.2479 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0255 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.0004 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.6 | 28.2755 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0475 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0004 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.6023 | 28.2011 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.0581 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0004 | 0.0003 | 9.1e-05 | 0 |  | n/a |  | 33 | 55 | 1048.4385 | 1013 | 30.7556 | 28.2778 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0257 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.0873 | 30.6911 | 28.3 | 0.2
2008-07-16 | 12:55 | 0 | 0 | 0 | 0.6 | 0.10893 | 0.7 | 0 | 0.0381 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.6311 | 28.3 | 0.1
2008-07-16 | 12:56 | 0 | 0 | 0 | 0.6 | 0.1 | 0.7 | 0.1 | 0.1474 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7 | 28.3 | 0.1
2008-07-16 | 12:56 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.0004 | 0.000208 | 8e-06 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.7 | 28.3 | 0
2008-07-16 | 12:56 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | 0.1 | 0.1 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28.8 | 5 | 195 | 59.3919 | 0.0004 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.9194 | 30.6391 | 28.3 | 0
2008-07-16 | 12:56 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | 0 | 0 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.4 | 0.0004 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5625 | 28.3 | 0
2008-07-16 | 12:56 | 0.561 | 0 | 0 | 0.6 | 0 | 0.6 | 0.1 | 0.0516 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.0003 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.2858 | 0
2008-07-16 | 12:56 | 0.995 | 0.0095 | 95.1555 | 28.4 | -0.1 | 28.4 | 5.5 | 5.4773 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3914 | 0.005864 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.2092 | -0.1
2008-07-16 | 12:56 | 1.13 | 0.0114 | 113.5492 | 114.1 | -0.008818 | 114.1 | 73.8 | 73.7799 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.022268 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6349 | 28.2 | 0
2008-07-16 | 12:56 | 1.13 | 0.0115 | 114.7692 | 144.5 | 0 | 144.5 | 91.7 | 91.6763 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.4 | 0.027617 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6121 | 28.2439 | 0
2008-07-16 | 12:56 | 1.13 | 0.012 | 120 | 146.7 | 0.09127 | 146.8 | 94.8 | 94.8115 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.3913 | 0.028183 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.9127 | 30.5 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.138 | 0.012 | 120 | 147.1 | 0.1 | 147.2 | 95.8 | 95.8 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0282 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.5 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 147.4 | 0.1 | 147.5 | 96 | 96 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.0283 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.5737 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 148.1 | 0.1 | 148.2 | 96.2 | 96.2297 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0284 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.6506 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 148.4 | 0.1 | 148.5 | 96.2 | 96.2 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3921 | 0.028492 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.6454 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 148.8 | 0.1 | 148.9 | 96.4 | 96.4 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3919 | 0.0285 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 0.1
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 148.8 | 0.19175 | 149 | 96.3 | 96.3482 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3918 | 0.0285 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 0.2
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 148.9 | 0.2 | 149.1 | 96.5 | 96.5 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0285 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6143 | 28.3 | 0.2
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.5 | 0.2 | 149.7 | 96.4 | 96.4162 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3914 | 0.0286 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.7 | 28.3 | 0.2
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.6 | 0.2 | 149.8 | 96.5 | 96.5002 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.6793 | 28.3 | 0.2
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.6 | 0.2 | 149.8 | 96.6 | 96.5933 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.9111 | 30.5261 | 28.3 | 0.2
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.6 | 0.29127 | 149.9 | 96.6 | 96.5609 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 0.3
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.7 | 0.3 | 150 | 96.5 | 96.5481 | 0.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.3913 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.2598 | 0.3
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.7 | 0.3 | 150 | 96.5 | 96.5 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.3088 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.2169 | 0.3
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.8 | 0.3 | 150.1 | 96.6 | 96.6 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.3923 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.2935 | 0.3
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.9 | 0.39212 | 150.3 | 96.5 | 96.548 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.4 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5702 | 28.3 | 0.4
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.9 | 0.4 | 150.3 | 96.7 | 96.6564 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.4 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5532 | 28.3 | 0.4
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.9 | 0.4 | 150.3 | 96.5 | 96.5482 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 0.4
2008-07-16 | 12:56 | 1.14 | 0.012 | 120 | 149.9 | 0.49157 | 150.4 | 96.7 | 96.6506 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 0.5
2008-07-16 | 12:56 | 1.148 | 0.012 | 120 | 150 | 1.1384 | 151.1 | 96.5 | 96.5177 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0287 | 0.0002 | 9.1e-05 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5 | 28.3 | 1.1
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 2.1129 | 152.1 | 96.5 | 96.5467 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0287 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5 | 28.3 | 2.1
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 5.473 | 155.5 | 96.5 | 96.5481 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0287 | 0.000291 | 0.000291 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 5.5
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 6.5213 | 156.5 | 96.6 | 96.6 | 0.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0003 | 0.000391 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.513 | 28.3 | 6.5
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 6.7912 | 156.8 | 96.7 | 96.65 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.0287 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.6663 | 28.3 | 6.8
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 6.8 | 156.8 | 96.5 | 96.5 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0287 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.6404 | 28.3 | 6.8
2008-07-16 | 12:56 | 1.15 | 0.012 | 120 | 150 | 6.8 | 156.8 | 96.7 | 96.7289 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.564 | 28.3 | 6.8
2008-07-16 | 12:56 | 1.156 | 0.012 | 120 | 150 | 6.7081 | 156.7 | 96.7 | 96.6501 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 6.7
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150 | 6.6 | 156.6 | 96.7 | 96.6589 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.0825 | 30.5 | 28.3 | 6.6
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150 | 6.4085 | 156.4 | 96.6 | 96.6494 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0004 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.9155 | 30.6321 | 28.3 | 6.4
2008-07-16 | 12:56 | 1.16 | 0.0125 | 124.5616 | 150 | 5.6689 | 155.7 | 96.7 | 96.7481 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3914 | 0.0288 | 0.0003 | 0.000309 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6145 | 28.3 | 5.7
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.6792 | 154.8 | 96.6 | 96.6 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4608 | 28.3 | 4.7
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.5087 | 154.6 | 96.7 | 96.7019 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3075 | 28.3 | 4.5
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.5 | 154.6 | 96.6 | 96.6491 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4456 | 28.3 | 4.5
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.5 | 154.6 | 96.6 | 96.6 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5494 | 28.3 | 4.5
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.4088 | 154.5 | 96.7 | 96.7473 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.3912 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5218 | 28.3 | 4.4
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.4 | 154.5 | 96.5 | 96.5487 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2973 | 28.3 | 4.4
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.4 | 154.5 | 96.7 | 96.6659 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2207 | 28.3 | 4.4
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.1 | 4.3082 | 154.4 | 96.6 | 96.6 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3918 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3119 | 28.3 | 4.3
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.2 | 4.3 | 154.5 | 96.7 | 96.7249 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3916 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4326 | 28.3 | 4.3
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.2 | 4.1172 | 154.3 | 96.7 | 96.6735 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.0288 | 0.0003 | 0.0003 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.9138 | 30.5 | 28.3 | 4.1
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.2 | 4.1 | 154.3 | 96.7 | 96.7002 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.000209 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 4.1
2008-07-16 | 12:56 | 1.16 | 0.013 | 130 | 150.2 | 4.1 | 154.3 | 96.6 | 96.5794 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.5624 | 28.3 | 4.1
2008-07-16 | 12:56 | 1.16 | 0.0127 | 127.0854 | 150.3 | 4 | 154.3 | 96.6 | 96.5788 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.0891 | 30.5609 | 28.3 | 4
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.3 | 4 | 154.3 | 96.7 | 96.7471 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3087 | 0.0288 | 0.0003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.9125 | 30.5 | 28.3 | 4
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.3 | 4 | 154.3 | 96.6 | 96.6418 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3912 | 0.0288 | 0.0003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.0876 | 30.5 | 28.3 | 4
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.3 | 3.9088 | 154.2 | 96.7 | 96.652 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0003 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 3.9
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.3 | 3.9 | 154.2 | 96.6 | 96.6 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3921 | 0.0288 | 0.0003 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5456 | 28.3 | 3.9
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.4 | 3.8 | 154.2 | 96.7 | 96.7308 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3919 | 0.0288 | 0.0003 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0806 | 30.6 | 28.3 | 3.8
2008-07-16 | 12:56 | 1.16 | 0.012 | 120 | 150.4 | 3.7082 | 154.1 | 96.6 | 96.6 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.028892 | 0.0003 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6 | 28.3 | 3.7
2008-07-16 | 12:56 | 1.168 | 0.012 | 120 | 150.5 | 2.9675 | 153.5 | 96.7 | 96.6674 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3084 | 0.0289 | 0.0003 | 0.000108 | -8e-06 |  | n/a |  | 33 | 55 | 1049 | 1012.0843 | 30.6 | 28.3 | 3
2008-07-16 | 12:56 | 1.17 | 0.012 | 120 | 150.5 | 2.9 | 153.4 | 96.7 | 96.7 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5482 | 28.3 | 2.9
2008-07-16 | 12:56 | 1.17 | 0.012 | 120 | 150.5 | 2.9 | 153.4 | 96.7 | 96.6535 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4435 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.17 | 0.012 | 120 | 150.5 | 2.9 | 153.4 | 96.6 | 96.6374 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.9127 | 30.3053 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.169 | 0.012 | 120 | 150.5 | 2.9 | 153.4 | 96.6 | 96.6012 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.0875 | 30.3817 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.16 | 0.012 | 120 | 150.6 | 2.9 | 153.5 | 96.6 | 96.6495 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3418 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.16 | 0.012 | 120 | 150.6 | 2.9 | 153.5 | 96.5 | 96.5423 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.335 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.16 | 0.012 | 120 | 150.6 | 2.9 | 153.5 | 96.6 | 96.6491 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.4 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3882 | 28.3 | 2.9
2008-07-16 | 12:57 | 1.627 | 0.0136 | 136.03 | 150.6 | 2.9 | 153.5 | 96.5 | 96.5248 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3081 | 0.0289 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3116 | 28.3 | 2.9
2008-07-16 | 12:57 | 2.215 | 0.0202 | 202.0521 | 176.9 | 2.9 | 179.8 | 107.7 | 107.6618 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3083 | 0.033511 | 0.0003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4301 | 28.3 | 2.9
2008-07-16 | 12:57 | 2.35 | 0.024 | 240 | 257.7 | 3.6325 | 261.3 | 164.9 | 164.9382 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3084 | 0.046916 | 0.0003 | 0.000192 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.9157 | 30.5 | 28.3 | 3.6
2008-07-16 | 12:57 | 2.35 | 0.024 | 240 | 287 | 3.8914 | 290.9 | 186.6 | 186.5887 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3086 | 0.051653 | 0.000391 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 3.9
2008-07-16 | 12:57 | 2.35 | 0.0241 | 241.1448 | 291.3 | 3.9912 | 295.2 | 190.7 | 190.6994 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.052374 | 0.0004 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 4
2008-07-16 | 12:57 | 2.35 | 0.025 | 249.8179 | 293.5 | 4 | 297.5 | 191.6 | 191.6036 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3913 | 0.052765 | 0.0004 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 4
2008-07-16 | 12:57 | 2.345 | 0.025 | 250 | 294.5 | 4.1913 | 298.7 | 191.9 | 191.9336 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3087 | 0.0529 | 0.0005 | 0.0002 | -9.1e-05 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 4.2
2008-07-16 | 12:57 | 2.34 | 0.025 | 250 | 294.6 | 4.2913 | 298.9 | 191.9 | 191.8529 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3913 | 0.0529 | 0.0005 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4505 | 28.3 | 4.3
2008-07-16 | 12:57 | 2.34 | 0.025 | 250 | 294.7 | 4.3 | 299 | 192.1 | 192.1 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3088 | 0.053 | 0.0005 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0876 | 30.3026 | 28.3 | 4.3
2008-07-16 | 12:57 | 2.34 | 0.025 | 250 | 294.7 | 4.3912 | 299.1 | 192.1 | 192.1425 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.053 | 0.0005 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.456 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.34 | 0.025 | 250 | 294.7 | 4.4 | 299.1 | 192.1 | 192.1 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.053 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5546 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.34 | 0.025 | 250 | 294.7 | 4.4 | 299.1 | 192.3 | 192.2514 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.053 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.345 | 0.025 | 250 | 294.7 | 4.4 | 299.1 | 192.3 | 192.3267 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3082 | 0.0531 | 0.0006 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.35 | 0.025 | 250 | 294.7 | 4.4 | 299.1 | 192.3 | 192.2988 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0531 | 0.0007 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.6 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.351 | 0.025 | 250 | 294.8 | 4.4 | 299.2 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3086 | 0.0532 | 0.0007 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4779 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 294.9 | 4.4 | 299.3 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3088 | 0.0532 | 0.0007 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4377 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 294.9 | 4.4 | 299.3 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.4 | 0.0532 | 0.000791 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.9147 | 30.4713 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 294.9 | 4.4 | 299.3 | 192 | 192.0146 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3089 | 0.0532 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3181 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 295.6 | 4.4 | 300 | 192.4 | 192.3507 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.4 | 0.053391 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4347 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.357 | 0.025 | 250 | 296.4 | 4.4 | 300.8 | 192.4 | 192.4 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.4 | 0.053491 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049.9066 | 1013 | 30.4126 | 28.3 | 4.4
2008-07-16 | 12:57 | 2.35 | 0.025 | 250 | 296.8 | 4.309 | 301.1 | 192.1 | 192.1229 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0535 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3608 | 28.3 | 4.3
2008-07-16 | 12:57 | 2.35 | 0.025 | 250 | 296.8 | 4.2 | 301 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3921 | 0.0536 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.5919 | 28.3 | 4.2
2008-07-16 | 12:57 | 2.358 | 0.025 | 250 | 296.9 | 4.1 | 301 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0536 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5262 | 28.3 | 4.1
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 296.9 | 4.0084 | 300.9 | 192.7 | 192.6517 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0536 | 0.0008 | 0.0002 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4495 | 28.3 | 4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297 | 4 | 301 | 192.2 | 192.2496 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3086 | 0.0536 | 0.0008 | 0.000109 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0863 | 30.3727 | 28.3 | 4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297 | 3.9 | 300.9 | 192.3 | 192.3265 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3913 | 0.0536 | 0.0008 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3039 | 28.3 | 3.9
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297 | 3.8087 | 300.8 | 192.3 | 192.3473 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3087 | 0.0536 | 0.0008 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3805 | 28.3 | 3.8
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.1 | 3.8 | 300.9 | 192.2 | 192.2485 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0536 | 0.0008 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4571 | 28.3 | 3.8
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.1 | 3.7 | 300.8 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3913 | 0.0536 | 0.0008 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 3.7
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.1 | 3.6088 | 300.7 | 192.3 | 192.3 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3088 | 0.053691 | 0.000891 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 3.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.2 | 3.5088 | 300.7 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3088 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0878 | 30.5 | 28.3 | 3.5
2008-07-16 | 12:57 | 2.36 | 0.0244 | 243.8103 | 297.2 | 3.5 | 300.7 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.2 | 29 | 5 | 195 | 59.3921 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3728 | 28.3 | 3.5
2008-07-16 | 12:57 | 2.36 | 0.0245 | 244.5293 | 297.3 | 3.4081 | 300.7 | 192.4 | 192.3752 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.4 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.2598 | 28.3 | 3.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.4 | 3.4 | 300.8 | 192.2 | 192.1622 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3082 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9175 | 30.2672 | 28.3 | 3.4
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.4 | 3.3 | 300.7 | 192.3 | 192.2517 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049.6557 | 1012 | 30.5739 | 28.3 | 3.3
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.4 | 3.2086 | 300.6 | 192.3 | 192.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3914 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3897 | 28.3 | 3.2
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.4 | 3.1088 | 300.5 | 192.3 | 192.2575 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3088 | 0.0537 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.912 | 30.3 | 28.3 | 3.1
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.4 | 3.1 | 300.5 | 192.2 | 192.1568 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.4 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0873 | 30.2767 | 28.3 | 3.1
2008-07-16 | 12:57 | 2.36 | 0.0245 | 245.4754 | 297.6 | 3 | 300.6 | 192.1 | 192.0507 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3087 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2001 | 28.3 | 3
2008-07-16 | 12:57 | 2.36 | 0.024 | 240 | 297.7 | 3 | 300.7 | 192.2 | 192.2121 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2766 | 28.3 | 3
2008-07-16 | 12:57 | 2.36 | 0.024 | 240 | 297.6 | 2.9088 | 300.5 | 192.2 | 192.2004 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4063 | 28.3 | 2.9
2008-07-16 | 12:57 | 2.36 | 0.024 | 240 | 297.7 | 2.9 | 300.6 | 192.1 | 192.1479 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3921 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4405 | 28.3 | 2.9
2008-07-16 | 12:57 | 2.36 | 0.024 | 240 | 297.7 | 2.9 | 300.6 | 192.1 | 192.0582 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.2 | 5 | 195 | 59.3919 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3 | 28.3 | 2.9
2008-07-16 | 12:57 | 2.36 | 0.024 | 240 | 297.6 | 2.8083 | 300.4 | 192.1 | 192.1014 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3917 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3 | 28.3 | 2.8
2008-07-16 | 12:57 | 2.36 | 0.0245 | 244.5431 | 297.6 | 2.8 | 300.4 | 192.3 | 192.3484 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3084 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4187 | 28.3 | 2.8
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.7 | 2.7086 | 300.4 | 192.4 | 192.3648 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.4 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4643 | 28.3 | 2.7
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.8 | 2.6085 | 300.4 | 192.2 | 192.2394 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.3915 | 0.0538 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3878 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.8 | 2.6 | 300.4 | 192.1 | 192.0895 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3087 | 0.053891 | 0.0009 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3108 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.3 | 192.2982 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.3087 | 0.0539 | 0.000991 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.2 | 192.1736 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.4 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.2574 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.3 | 192.3232 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.2193 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.3 | 192.2502 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3088 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011.9122 | 30.2959 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.3 | 192.2502 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.391 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.445 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 297.9 | 2.6 | 300.5 | 192.1 | 192.1432 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.4 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 2.6
2008-07-16 | 12:57 | 2.36 | 0.025 | 250 | 298 | 2.6 | 300.6 | 192.2 | 192.153 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3084 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4743 | 28.3 | 2.6
2008-07-16 | 12:58 | 2.36 | 0.025 | 250 | 298 | 2.6 | 300.6 | 192 | 192.0474 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.4 | 0.0539 | 0.001 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4023 | 28.3 | 2.6
2008-07-16 | 12:58 | 2.36 | 0.025 | 250 | 298 | 2.5 | 300.5 | 192.2 | 192.1655 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4789 | 28.3 | 2.5
2008-07-16 | 12:58 | 2.36 | 0.025 | 250 | 297.9 | 2.5 | 300.4 | 192.2 | 192.1502 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.3 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4445 | 28.3 | 2.5
2008-07-16 | 12:58 | 2.36 | 0.025 | 250 | 297.9 | 2.5 | 300.4 | 192.1 | 192.0636 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.4 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3679 | 28.3 | 2.5
2008-07-16 | 12:58 | 2.36 | 0.025 | 250 | 297.9 | 2.5 | 300.4 | 192 | 192.0496 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 29 | 5 | 195 | 59.4 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2913 | 28.3 | 2.5
2008-07-16 | 12:58 | 2.577 | 0.025 | 250 | 297.8 | 2.5 | 300.3 | 192 | 191.999 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.3914 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2146 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.06 | 0.0305 | 304.6076 | 297.9 | 2.5 | 300.4 | 192.2 | 192.2 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0539 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3241 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.0353 | 353.1419 | 359.8 | 2.5 | 362.3 | 229.7 | 229.734 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.063177 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0788 | 30.4 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 431.5 | 2.5 | 434 | 273.9 | 273.8857 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.8 | 5 | 195 | 59.3081 | 0.073439 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3848 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 442.7 | 2.5 | 445.2 | 284.1 | 284.1365 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 29 | 5 | 195 | 59.3 | 0.074942 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3081 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 443.3 | 2.5 | 445.8 | 286.6 | 286.6279 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 29 | 5 | 195 | 59.3916 | 0.075092 | 0.001 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2315 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 445.2 | 2.5 | 447.7 | 287.2 | 287.2061 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 29 | 5 | 195 | 59.3 | 0.075391 | 0.001091 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.2 | 2.5 | 448.7 | 287.5 | 287.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 29 | 5 | 195 | 59.3087 | 0.075491 | 0.0011 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2217 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.3 | 2.5 | 448.8 | 287.5 | 287.5258 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.3 | 0.0756 | 0.0011 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2983 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.4 | 2.5 | 448.9 | 287.7 | 287.6561 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0756 | 0.001191 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4498 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.3 | 2.5 | 448.8 | 287.3 | 287.2534 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.075691 | 0.0012 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.346 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.3 | 2.5 | 448.8 | 287.8 | 287.801 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.0757 | 0.0012 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2555 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.7 | 2.5 | 449.2 | 287.6 | 287.6156 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.075791 | 0.001291 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4086 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.7 | 2.5 | 449.2 | 287.6 | 287.5946 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.0758 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5626 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.7 | 2.5 | 449.2 | 288 | 288.0016 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.0758 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.542 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.6 | 2.5 | 449.1 | 287.9 | 287.9 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0758 | 0.0013 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4654 | 28.3 | 2.5
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.7 | 2.4 | 449.1 | 287.8 | 287.7988 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3084 | 0.0759 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 446.8 | 2.4 | 449.2 | 287.5 | 287.5014 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.0759 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 447.2 | 2.4 | 449.6 | 287.7 | 287.6614 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.075991 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2712 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 447.3 | 2.4 | 449.7 | 288 | 287.9509 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3085 | 0.076 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.241 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 447.4 | 2.4 | 449.8 | 287.7 | 287.6603 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3913 | 0.076 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3352 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 447.4 | 2.4 | 449.8 | 288.1 | 288.1482 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3 | 0.076 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4885 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.2 | 2.4 | 450.6 | 288.1 | 288.1006 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.076185 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.7 | 2.4 | 451.1 | 287.9 | 287.9178 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.4 | 0.0762 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0788 | 30.5 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.9 | 2.4 | 451.3 | 287.7 | 287.7271 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3081 | 0.0762 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.4519 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449.1 | 2.4 | 451.5 | 288.1 | 288.1 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.4 | 0.0762 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.3007 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449.3 | 2.4 | 451.7 | 288 | 287.9836 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3085 | 0.0762 | 0.0014 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.3773 | 28.3 | 2.4
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449.3 | 2.3 | 451.6 | 288 | 288.0254 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3 | 0.0763 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.3461 | 28.3 | 2.3
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449.2 | 2.2 | 451.4 | 287.8 | 287.8455 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3912 | 0.0763 | 0.0015 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.361 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449.1 | 2.2 | 451.3 | 288.1 | 288.0813 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0763 | 0.0015 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.4857 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449 | 2.2 | 451.2 | 287.5 | 287.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0763 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3325 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 449 | 2.2 | 451.2 | 288 | 287.9905 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3088 | 0.0763 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049.8894 | 1012 | 30.4207 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.7 | 2.2 | 450.9 | 288 | 288.0061 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3088 | 0.0763 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3892 | 28.3369 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.7 | 2.2 | 450.9 | 288.1 | 288.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3077 | 0.0763 | 0.0015 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.2405 | 28.3865 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.5 | 2.2 | 450.7 | 288 | 288 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3921 | 0.0763 | 0.001592 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.4693 | 28.3102 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.6 | 2.2 | 450.8 | 287.7 | 287.6996 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3081 | 0.0763 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.3677 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.5 | 2.2 | 450.7 | 288.1 | 288.1077 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0763 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.2571 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.7 | 2.2 | 450.9 | 288.2 | 288.179 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.8 | 5 | 195 | 59.4 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.2596 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.037 | 370 | 448.7 | 2.2 | 450.9 | 288 | 287.9983 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.3 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.4899 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.0367 | 367.4312 | 448.9 | 2.2 | 451.1 | 288 | 288.0233 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.0871 | 30.3536 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 448.9 | 2.2 | 451.1 | 288 | 287.9662 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3911 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.3996 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 448.9 | 2.2 | 451.1 | 288 | 288 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449 | 2.2 | 451.2 | 288 | 288 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0876 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.1 | 2.2 | 451.3 | 288 | 287.9974 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.391 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.2 | 2.2 | 451.4 | 288.2 | 288.1517 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.3 | 2.2 | 451.5 | 288.2 | 288.2132 | 0.1 | 12.9 | 955 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5328 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.3 | 2.2 | 451.5 | 287.7 | 287.7036 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.8 | 5 | 195 | 59.3918 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049.3006 | 1011 | 30.5906 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.3 | 2.2 | 451.5 | 288.3 | 288.3034 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.3 | 0.0764 | 0.0016 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049.2168 | 1011 | 30.514 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.2 | 2.2 | 451.4 | 287.9 | 287.8839 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0764 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.3122 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.2 | 2.2 | 451.4 | 288.2 | 288.2435 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0764 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.3178 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.2 | 2.2 | 451.4 | 288 | 288.0465 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0764 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.2 | 2.2 | 451.4 | 288.3 | 288.3 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0764 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.1 | 2.2 | 451.3 | 287.9 | 287.9451 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3691 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449.1 | 2.2 | 451.3 | 288.2 | 288.2336 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.5 | 28.3543 | 2.2
2008-07-16 | 12:58 | 3.61 | 0.036 | 360 | 449 | 2.2 | 451.2 | 288.3 | 288.3478 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.9212 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:59 | 3.61 | 0.036 | 360 | 448.9 | 2.2 | 451.1 | 288.1 | 288.0712 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:59 | 3.61 | 0.036 | 360 | 448.7 | 2.2 | 450.9 | 288 | 287.9995 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:59 | 3.61 | 0.036 | 360 | 448.6 | 2.2 | 450.8 | 287.9 | 287.901 | 0.1 | 12.9 | 955 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0765 | 0.0017 | 8e-06 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:59 | 3.61 | 0.036 | 360 | 448.6 | 2.2 | 450.8 | 287.8 | 287.8485 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.0765 | 0.0017 | 9.1e-05 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.4157 | 28.3 | 2.2
2008-07-16 | 12:59 | 3.61 | 0.036 | 360 | 448.6 | 2.2 | 450.8 | 288.1 | 288.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014.913 | 30.2136 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.039 | 0.0376 | 375.5997 | 448.6 | 2.2 | 450.8 | 288.3 | 288.2527 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0765 | 0.0017 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.4448 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.663 | 0.0448 | 447.6713 | 477.2 | 2.2 | 479.4 | 295.6 | 295.6033 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.080333 | 0.0017 | 9e-06 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.4417 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 560.5 | 2.2 | 562.7 | 349 | 348.9721 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.091135 | 0.0017 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.0875 | 30.4 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 592.5 | 2.2 | 594.7 | 376.7 | 376.6674 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.094906 | 0.001791 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.4 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 596.5 | 2.2 | 598.7 | 381.6 | 381.6242 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.095591 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.0878 | 30.4 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 597.4 | 2.2 | 599.6 | 382.7 | 382.7023 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.095692 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049.7819 | 1012 | 30.465 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 597.6 | 2.2 | 599.8 | 383.3 | 383.2975 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0957 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.4584 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 597.8 | 2.2 | 600 | 383.4 | 383.4102 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0957 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.4182 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 597.8 | 2.2 | 600 | 383.7 | 383.6517 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28.8 | 5 | 195 | 59.3084 | 0.095792 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.4948 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 598.4 | 2.2 | 600.6 | 383.9 | 383.8515 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 29 | 5 | 195 | 59.4 | 0.0958 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.5 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 598.9 | 2.2 | 601.1 | 383.7 | 383.6843 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 29 | 5 | 195 | 59.3913 | 0.0959 | 0.0018 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.452 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599.2 | 2.2 | 601.4 | 383.8 | 383.7533 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28.2 | 5 | 195 | 59.3 | 0.096 | 0.0019 | 9.1e-05 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.4 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599.2 | 2.2 | 601.4 | 383.5 | 383.5288 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.096 | 0.0019 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.3977 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599.1 | 2.2 | 601.3 | 383.7 | 383.6521 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.096 | 0.0019 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2444 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599 | 2.2 | 601.2 | 383.6 | 383.649 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0961 | 0.002 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3088 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 598.9 | 2.2 | 601.1 | 383.8 | 383.7601 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.309 | 0.0961 | 0.002 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.431 | 28.3 | 2.2
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599 | 2.1 | 601.1 | 383.9 | 383.949 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.8 | 5 | 195 | 59.3921 | 0.0961 | 0.002 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.4923 | 28.3077 | 2.1
2008-07-16 | 12:59 | 4.85 | 0.049 | 490 | 599.1 | 2.1 | 601.2 | 383.6 | 383.6497 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28.2 | 5 | 195 | 59.4 | 0.0962 | 0.002 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.4158 | 28.3842 | 2.1
2008-07-16 | 12:59 | 4.854 | 0.049 | 490 | 599.1 | 2.1 | 601.2 | 383.8 | 383.8 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0962 | 0.002092 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.2783 | 28.4 | 2.1
2008-07-16 | 12:59 | 4.86 | 0.049 | 490 | 599.2 | 2.1 | 601.3 | 383.6 | 383.5552 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0962 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3123 | 28.3626 | 2.1
2008-07-16 | 12:59 | 4.861 | 0.049 | 490 | 599.8 | 2.1 | 601.9 | 384.1 | 384.1051 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.096291 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.4579 | 28.3 | 2.1
2008-07-16 | 12:59 | 4.869 | 0.049 | 490 | 600.4 | 2.1 | 602.5 | 383.7 | 383.7167 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0963 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.228 | 28.3 | 2.1
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.5 | 2 | 602.5 | 383.7 | 383.7498 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.0963 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3668 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.5 | 2 | 602.5 | 383.9 | 383.8548 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.096391 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2867 | 28.4 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.6 | 2 | 602.6 | 383.8 | 383.765 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0964 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3595 | 28.4 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.6 | 2 | 602.6 | 383.9 | 383.8981 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3923 | 0.0964 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2067 | 28.4 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.6 | 2 | 602.6 | 383.7 | 383.6773 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3079 | 0.0964 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3265 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.6 | 2 | 602.6 | 383.7 | 383.7077 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3919 | 0.0964 | 0.0021 | 0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3005 | 28.3 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.6 | 2 | 602.6 | 384.1 | 384.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0964 | 0.0021 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.3462 | 28.3 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.7 | 2 | 602.7 | 384 | 384.0109 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2035 | 28.3 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.9 | 2 | 602.9 | 383.8 | 383.8 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2801 | 28.3 | 2
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 601 | 1.9 | 602.9 | 383.9 | 383.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.3913 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2433 | 28.3567 | 1.9
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 601 | 1.9 | 602.9 | 383.9 | 383.9297 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28.3667 | 1.9
2008-07-16 | 12:59 | 4.87 | 0.0487 | 487.418 | 600.9 | 1.9 | 602.8 | 384 | 383.9506 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.87 | 0.048 | 480 | 600.8 | 1.9 | 602.7 | 383.7 | 383.6963 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0965 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.87 | 0.048 | 480 | 600.8 | 1.9 | 602.7 | 384 | 383.9629 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0965 | 0.002291 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.1369 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.87 | 0.048 | 480 | 600.7 | 1.9 | 602.6 | 384 | 384.049 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.914 | 30.1794 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.867 | 0.048 | 480 | 600.7 | 1.9 | 602.6 | 384 | 384.0383 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.0806 | 30.3 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.86 | 0.048 | 480 | 600.7 | 1.9 | 602.6 | 384 | 384.0493 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3082 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3 | 28.3 | 1.9
2008-07-16 | 12:59 | 4.86 | 0.048 | 480 | 600.8 | 1.8 | 602.6 | 383.7 | 383.6976 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2305 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.86 | 0.048 | 480 | 600.7 | 1.8 | 602.5 | 383.8 | 383.7716 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3086 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.86 | 0.048 | 480 | 600.8 | 1.8 | 602.6 | 383.9 | 383.9489 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1545 | 28.3228 | 1.8
2008-07-16 | 12:59 | 4.864 | 0.048 | 480 | 600.8 | 1.8 | 602.6 | 384.1 | 384.0909 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0966 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.0012 | 28.3994 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.048 | 480 | 600.2 | 1.8 | 602 | 383.7 | 383.7373 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.0966 | 0.002391 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1519 | 28.4 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.0489 | 488.599 | 600.2 | 1.8 | 602 | 384 | 383.9738 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1474 | 28.4 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.05 | 500 | 600.2 | 1.8 | 602 | 384 | 383.9516 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.129 | 28.371 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.05 | 500 | 600.2 | 1.8 | 602 | 383.8 | 383.834 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.21 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.05 | 500 | 600.3 | 1.8 | 602.1 | 383.8 | 383.8256 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.3641 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.05 | 500 | 600.2 | 1.8 | 602 | 383.8 | 383.8443 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3081 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9194 | 30.2828 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.0494 | 494.0207 | 600.4 | 1.8 | 602.2 | 383.9 | 383.9491 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3083 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0828 | 30.2 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.3 | 1.8 | 602.1 | 383.6 | 383.5931 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3084 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.188 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.2 | 1.8 | 602 | 383.9 | 383.8858 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9138 | 30.1113 | 28.3 | 1.8
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.1 | 1.7 | 601.8 | 383.7 | 383.7488 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1 | 28.3654 | 1.7
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.2 | 1.7 | 601.9 | 384 | 384.0045 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.142 | 28.358 | 1.7
2008-07-16 | 12:59 | 4.87 | 0.049 | 490 | 600.2 | 1.7 | 601.9 | 383.6 | 383.5541 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2372 | 28.3186 | 1.7
2008-07-16 | 13:00 | 4.87 | 0.049 | 490 | 600.2 | 1.7 | 601.9 | 383.8 | 383.7517 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.3905 | 28.3952 | 1.7
2008-07-16 | 13:00 | 4.87 | 0.0486 | 485.6444 | 600.3 | 1.7 | 602 | 383.9 | 383.8516 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1845 | 28.4 | 1.7
2008-07-16 | 13:00 | 4.87 | 0.048 | 480 | 600.3 | 1.7 | 602 | 383.6 | 383.6364 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3921 | 0.0966 | 0.0024 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9212 | 30.1484 | 28.3516 | 1.7
2008-07-16 | 13:00 | 4.87 | 0.048 | 480 | 600.2 | 1.7 | 601.9 | 383.9 | 383.9399 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.002492 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.225 | 28.3 | 1.7
2008-07-16 | 13:00 | 4.87 | 0.048 | 480 | 600.2 | 1.7 | 601.9 | 383.9 | 383.927 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0966 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2968 | 28.3 | 1.7
2008-07-16 | 13:00 | 5.33 | 0.0532 | 531.7442 | 600.2 | 1.7 | 601.9 | 384 | 384.0495 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3084 | 0.0966 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1435 | 28.3 | 1.7
2008-07-16 | 13:00 | 5.972 | 0.058 | 579.7743 | 611.6 | 1.7 | 613.3 | 396 | 396.0392 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.098062 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1 | 28.3548 | 1.7
2008-07-16 | 13:00 | 6.1 | 0.06 | 600 | 701.7 | 1.7 | 703.4 | 451.9 | 451.8841 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.10856 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1314 | 28.3686 | 1.7
2008-07-16 | 13:00 | 6.1 | 0.06 | 600 | 738.6 | 1.7 | 740.3 | 472.8 | 472.7567 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3089 | 0.11268 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0889 | 30.192 | 28.3 | 1.7
2008-07-16 | 13:00 | 6.1 | 0.06 | 600 | 745.1 | 1.7 | 746.8 | 477.5 | 477.5087 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.11339 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1154 | 28.3 | 1.7
2008-07-16 | 13:00 | 6.106 | 0.06 | 600 | 745.1 | 1.6 | 746.7 | 478.5 | 478.4882 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1135 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1613 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.106 | 0.06 | 600 | 745.9 | 1.6 | 747.5 | 478.7 | 478.7 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.11359 | 0.0026 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9124 | 30.1621 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.1 | 0.06 | 600 | 746.1 | 1.6 | 747.7 | 478.7 | 478.6502 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.11369 | 0.0026 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1145 | 28.3145 | 1.6
2008-07-16 | 13:00 | 6.101 | 0.06 | 600 | 746.1 | 1.6 | 747.7 | 478.9 | 478.9284 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.1137 | 0.0027 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1911 | 28.3911 | 1.6
2008-07-16 | 13:00 | 6.109 | 0.06 | 600 | 746.1 | 1.6 | 747.7 | 479.1 | 479.126 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.4 | 0.1138 | 0.002792 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3323 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 746.3 | 1.6 | 747.9 | 478.9 | 478.9 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3918 | 0.1138 | 0.0028 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 746.5 | 1.6 | 748.1 | 479 | 478.9586 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1139 | 0.0028 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 746.6 | 1.6 | 748.2 | 479 | 479.0485 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1139 | 0.0028 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 746.7 | 1.6 | 748.3 | 479.5 | 479.4601 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1139 | 0.002891 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2737 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.111 | 0.06 | 600 | 746.6 | 1.6 | 748.2 | 479.4 | 479.4481 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.114 | 0.0029 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2495 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.119 | 0.06 | 600 | 746.8 | 1.6 | 748.4 | 479.5 | 479.5481 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.114 | 0.0029 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1727 | 28.3273 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 746.9 | 1.6 | 748.5 | 479.3 | 479.3102 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.114 | 0.0029 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1039 | 28.3961 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 747 | 1.6 | 748.6 | 479.4 | 479.4483 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.114 | 0.0029 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1805 | 28.3195 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 747 | 1.6 | 748.6 | 479.6 | 479.6 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.11409 | 0.002991 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 747.1 | 1.6 | 748.7 | 479.4 | 479.4 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1141 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1662 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.0604 | 604.1182 | 747.4 | 1.6 | 749 | 479.6 | 479.6039 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3081 | 0.1141 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0806 | 30.1103 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.061 | 610 | 747.4 | 1.6 | 749 | 479.3 | 479.3359 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3085 | 0.1141 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9155 | 30.187 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.113 | 0.061 | 610 | 747.5 | 1.6 | 749.1 | 479.8 | 479.8 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1141 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1364 | 28.3636 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.061 | 610 | 747.7 | 1.6 | 749.3 | 479.5 | 479.5491 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1142 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1803 | 28.4 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.061 | 610 | 748.6 | 1.6 | 750.2 | 479.6 | 479.5821 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1143 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9129 | 30.2832 | 28.3832 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.0604 | 604.0917 | 748.7 | 1.6 | 750.3 | 479.6 | 479.605 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1143 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2067 | 28.3067 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 749.9 | 1.6 | 751.5 | 479.3 | 479.2843 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.11449 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 750.1 | 1.6 | 751.7 | 479.4 | 479.4443 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1145 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2465 | 28.3465 | 1.6
2008-07-16 | 13:00 | 6.117 | 0.06 | 600 | 750.2 | 1.6 | 751.8 | 479.4 | 479.4051 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.1145 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2769 | 28.3769 | 1.6
2008-07-16 | 13:00 | 6.114 | 0.06 | 600 | 750.2 | 1.6 | 751.8 | 479.8 | 479.8478 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1145 | 0.003 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2003 | 28.3003 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 750.3 | 1.6 | 751.9 | 479.7 | 479.6634 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1145 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.11 | 0.06 | 600 | 750.3 | 1.6 | 751.9 | 479.6 | 479.5992 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1145 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.111 | 0.06 | 600 | 750.5 | 1.6 | 752.1 | 479.4 | 479.3554 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3916 | 0.1145 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0843 | 30.1704 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.119 | 0.06 | 600 | 750.4 | 1.6 | 752 | 479.4 | 479.4474 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1145 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9138 | 30.1 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.4 | 1.6 | 752 | 479.4 | 479.4 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1145 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.4 | 1.6 | 752 | 479.7 | 479.6854 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1146 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1 | 28.3591 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.3 | 1.6 | 751.9 | 479.2 | 479.2397 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1146 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.171 | 28.3645 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.3 | 1.6 | 751.9 | 479.4 | 479.3794 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1146 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2879 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.2 | 1.6 | 751.8 | 479.6 | 479.559 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.2 | 28 | 5 | 195 | 59.4 | 0.1146 | 0.0031 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.211 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.2 | 1.6 | 751.8 | 479.4 | 479.3975 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1146 | 0.003192 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1343 | 28.3657 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.1 | 1.6 | 751.7 | 479.5 | 479.4501 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3919 | 0.1146 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.9194 | 30.0576 | 28.3576 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.3 | 1.6 | 751.9 | 479.3 | 479.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1146 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.038 | 28.319 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.5 | 1.6 | 752.1 | 479.5 | 479.5495 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3916 | 0.1146 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.0843 | 30.1913 | 28.3956 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.8 | 1.6 | 752.4 | 479.6 | 479.5657 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1146 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.2 | 28.4 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751 | 1.6 | 752.6 | 479.5 | 479.4953 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1146 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1511 | 28.3511 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 750.9 | 1.6 | 752.5 | 479 | 479.0479 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.1 | 1.6 | 752.7 | 479.3 | 479.3152 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1021 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.1 | 1.6 | 752.7 | 479.3 | 479.3008 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1787 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.1 | 1.6 | 752.7 | 479.4 | 479.4066 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.2 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751 | 1.6 | 752.6 | 479.3 | 479.2826 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3923 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9232 | 30.1681 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.1 | 1.6 | 752.7 | 479.3 | 479.2501 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0788 | 30.0915 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.4 | 1.6 | 753 | 479.4 | 479.3628 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.4 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.0149 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.5 | 1.6 | 753.1 | 479 | 479.0496 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.3 | 1.6
2008-07-16 | 13:00 | 6.12 | 0.06 | 600 | 751.5 | 1.6 | 753.1 | 479.2 | 479.1761 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.3383 | 1.6
2008-07-16 | 13:00 | 6.124 | 0.06 | 600 | 751.8 | 1.6 | 753.4 | 479.1 | 479.0761 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.4 | 1.6
2008-07-16 | 13:01 | 6.13 | 0.06 | 600 | 752.4 | 1.6 | 754 | 479.3 | 479.2509 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1147 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.4 | 1.6
2008-07-16 | 13:01 | 6.13 | 0.06 | 600 | 752.6 | 1.6 | 754.2 | 479.2 | 479.2473 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30.0681 | 28.4 | 1.6
2008-07-16 | 13:01 | 6.13 | 0.06 | 600 | 752.5 | 1.6 | 754.1 | 479.5 | 479.4985 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30.1 | 28.3552 | 1.6
2008-07-16 | 13:01 | 6.122 | 0.06 | 600 | 752.3 | 1.6 | 753.9 | 479.4 | 479.4495 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1 | 28.3 | 1.6
2008-07-16 | 13:01 | 6.12 | 0.06 | 600 | 752.1 | 1.6 | 753.7 | 479.4 | 479.4006 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.1 | 28.3 | 1.6
2008-07-16 | 13:01 | 6.112 | 0.06 | 600 | 751.9 | 1.6 | 753.5 | 479.3 | 479.2597 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.0256 | 28.3 | 1.6
2008-07-16 | 13:01 | 6.359 | 0.0611 | 611.4187 | 751.7 | 1.6 | 753.3 | 479 | 479.0038 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3081 | 0.1148 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.9192 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.141 | 0.0653 | 653.3444 | 757.2 | 1.5082 | 758.7 | 480.7 | 480.738 | 0.1 | 12.9 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.11544 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30.0276 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.299 | 0.0703 | 702.5168 | 828.6 | 1.5 | 830.1 | 528.6 | 528.6417 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.12302 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.0957 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.32 | 0.071 | 710 | 886.7 | 1.5 | 888.2 | 566.7 | 566.6519 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3914 | 0.12885 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.019 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.32 | 0.071 | 710 | 893.8 | 1.5 | 895.3 | 573.7 | 573.7447 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.4 | 0.1296 | 0.0032 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.1152 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.32 | 0.0717 | 717.2394 | 896.9 | 1.5 | 898.4 | 575.5 | 575.5435 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.12988 | 0.0033 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0871 | 30.1316 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.32 | 0.072 | 720 | 901 | 1.5 | 902.5 | 575.4 | 575.3737 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.13037 | 0.0033 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.32 | 0.072 | 720 | 901.3 | 1.5913 | 902.9 | 575.4 | 575.3851 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1304 | 0.0033 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.072 | 720 | 901.4 | 1.6 | 903 | 575.5 | 575.5473 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.1304 | 0.0033 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9359 | 28.3641 | 1.6
2008-07-16 | 13:01 | 7.327 | 0.072 | 720 | 901.5 | 1.6 | 903.1 | 575.7 | 575.7222 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.309 | 0.1304 | 0.0033 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9815 | 28.3593 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.072 | 720 | 901.5 | 1.6 | 903.1 | 575.9 | 575.9352 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1304 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.0827 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.072 | 720 | 901.6 | 1.6 | 903.2 | 575.8 | 575.8 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1305 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.0061 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.072 | 720 | 901.5 | 1.6 | 903.1 | 576.3 | 576.2582 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.1305 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8589 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.0716 | 715.9582 | 901.5 | 1.6 | 903.1 | 576 | 576.0259 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.1305 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8943 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.0713 | 712.528 | 901.5 | 1.6 | 903.1 | 576 | 575.9971 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.1305 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.33 | 0.072 | 720 | 901.8 | 1.6 | 903.4 | 576.1 | 576.0704 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1305 | 0.0034 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.087 | 30 | 28.3004 | 1.6
2008-07-16 | 13:01 | 7.324 | 0.072 | 720 | 901.9 | 1.6 | 903.5 | 576.1 | 576.1039 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1305 | 0.003491 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30 | 28.377 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.072 | 720 | 901.8 | 1.6 | 903.4 | 576.1 | 576.0503 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.1306 | 0.0035 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.8928 | 28.4 | 1.6
2008-07-16 | 13:01 | 7.323 | 0.072 | 720 | 901.5 | 1.6 | 903.1 | 576.3 | 576.2919 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1306 | 0.0035 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8605 | 28.4 | 1.6
2008-07-16 | 13:01 | 7.329 | 0.072 | 720 | 901.2 | 1.6 | 902.8 | 576.3 | 576.2968 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1306 | 0.0035 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.9124 | 30 | 28.3932 | 1.6
2008-07-16 | 13:01 | 7.321 | 0.072 | 720 | 901.2 | 1.6 | 902.8 | 576.3 | 576.2502 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1306 | 0.003592 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 30 | 28.3165 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.072 | 720 | 901.1 | 1.6 | 902.7 | 576.1 | 576.1422 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3919 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.9398 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.309 | 0.072 | 720 | 901.1 | 1.6 | 902.7 | 576.1 | 576.1475 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.9 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.304 | 0.072 | 720 | 901.3 | 1.6 | 902.9 | 576.3 | 576.3448 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.0845 | 29.9131 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 901.8 | 1.6 | 903.4 | 576.1 | 576.0691 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9895 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 902 | 1.6 | 903.6 | 575.9 | 575.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0882 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.0715 | 715.0482 | 901.8 | 1.6 | 903.4 | 575.8 | 575.7878 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1307 | 0.0036 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9138 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.317 | 0.07 | 700 | 901.8 | 1.6 | 903.4 | 576.2 | 576.1508 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.8 | 5 | 195 | 59.3 | 0.13079 | 0.003691 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8197 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.07 | 700 | 901.7 | 1.6 | 903.3 | 576 | 575.9964 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.4 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8965 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.07 | 700 | 901.6 | 1.6 | 903.2 | 575.8 | 575.7856 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.9124 | 29.9731 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.07 | 700 | 901.6 | 1.6 | 903.2 | 576 | 576.0475 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.9005 | 28.3497 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.0704 | 704.2619 | 901.6 | 1.6 | 903.2 | 576.2 | 576.2 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.8264 | 28.3736 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.071 | 710 | 901.7 | 1.6 | 903.3 | 576 | 575.9919 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3919 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.903 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.32 | 0.071 | 710.4837 | 902 | 1.6 | 903.6 | 576 | 575.9518 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9797 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.314 | 0.0719 | 718.824 | 901.9 | 1.6 | 903.5 | 576.2 | 576.2 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.0563 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 901.9 | 1.6 | 903.5 | 576 | 575.9713 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 30.0011 | 28.333 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 901.9 | 1.6 | 903.5 | 576.1 | 576.1491 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8 | 28.3904 | 1.6
2008-07-16 | 13:01 | 7.311 | 0.0716 | 715.993 | 902 | 1.6 | 903.6 | 576.2 | 576.2 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8 | 28.3138 | 1.6
2008-07-16 | 13:01 | 7.319 | 0.071 | 710 | 901.9 | 1.6 | 903.5 | 576.1 | 576.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9256 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.313 | 0.071 | 710 | 902.1 | 1.6 | 903.7 | 576.1 | 576.0505 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1308 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9606 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.071 | 710 | 902.4 | 1.6 | 904 | 576.2 | 576.2185 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.13089 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0768 | 29.9 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.071 | 710 | 902.5 | 1.6 | 904.1 | 576.3 | 576.2505 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1309 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.0716 | 715.7028 | 902.8 | 1.6 | 904.4 | 576.1 | 576.1498 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1309 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 30.0386 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 903 | 1.6 | 904.6 | 576.1 | 576.1435 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3083 | 0.1309 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011.0827 | 30.0541 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 903.3 | 1.6 | 904.9 | 576.2 | 576.2 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3084 | 0.1309 | 0.0037 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 903.5 | 1.6 | 905.1 | 576.2 | 576.2103 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1309 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.072 | 720 | 903.6 | 1.6 | 905.2 | 576 | 576.0028 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1309 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30.0758 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.0727 | 727.1476 | 903.5 | 1.6 | 905.1 | 576.2 | 576.1568 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3931 | 0.1309 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30.1524 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 903.2 | 1.6 | 904.8 | 576 | 576.01 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1309 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 30.1133 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 903 | 1.6 | 904.6 | 576.2 | 576.2 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9103 | 28.3 | 1.6
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 902.7 | 1.5 | 904.2 | 576 | 576 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.9124 | 30.0644 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 902.8 | 1.5 | 904.3 | 576 | 576 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3077 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.865 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 902.9 | 1.5 | 904.4 | 576.1 | 576.1491 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3919 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7353 | 28.3 | 1.5
2008-07-16 | 13:01 | 7.31 | 0.073 | 730 | 903.1 | 1.5 | 904.6 | 576.1 | 576.147 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0825 | 29.824 | 28.3 | 1.5
2008-07-16 | 13:02 | 7.31 | 0.073 | 730 | 903.2 | 1.5 | 904.7 | 576.3 | 576.2863 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9776 | 28.3 | 1.5
2008-07-16 | 13:02 | 7.31 | 0.073 | 730 | 903.5 | 1.5 | 905 | 576 | 576 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011.0862 | 29.9346 | 28.3 | 1.5
2008-07-16 | 13:02 | 7.31 | 0.0728 | 727.8667 | 903.5 | 1.5 | 905 | 576.3 | 576.3006 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 29.8159 | 28.3421 | 1.5
2008-07-16 | 13:02 | 7.317 | 0.072 | 720 | 903.3 | 1.5 | 904.8 | 576.1 | 576.1351 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011.9129 | 29.7561 | 28.3813 | 1.5
2008-07-16 | 13:02 | 7.325 | 0.072 | 720 | 903.2 | 1.5 | 904.7 | 576.1 | 576.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3085 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9862 | 28.3046 | 1.5
2008-07-16 | 13:02 | 7.707 | 0.0734 | 734.4767 | 903.3 | 1.5 | 904.8 | 575.9 | 575.9 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.131 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.8559 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.337 | 0.0784 | 784.2052 | 921.9 | 1.5 | 923.4 | 581.6 | 581.5727 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.13282 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.8 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.0824 | 824.0103 | 1006 | 1.5 | 1007.5 | 640.6 | 640.6333 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.14053 | 0.0038 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.914 | 29.8 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1032.8 | 1.5 | 1034.3 | 665.3 | 665.3169 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.14326 | 0.003892 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8038 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1044.9 | 1.5 | 1046.4 | 670.4 | 670.3984 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3919 | 0.14436 | 0.0039 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9571 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1047 | 1.5 | 1048.5 | 671.6 | 671.6481 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.14459 | 0.0039 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8897 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1047.2 | 1.5 | 1048.7 | 672.2 | 672.2142 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1446 | 0.003992 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8318 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1047.2 | 1.5 | 1048.7 | 672.4 | 672.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1447 | 0.004 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1013 | 29.9 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1047.4 | 1.5 | 1048.9 | 672.4 | 672.3657 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1447 | 0.004 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1013.088 | 29.9 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.51 | 0.083 | 830 | 1047.8 | 1.5 | 1049.3 | 672.2 | 672.2 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1448 | 0.004091 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7765 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.517 | 0.083 | 830 | 1047.9 | 1.5 | 1049.4 | 672.4 | 672.3509 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1448 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1013 | 29.7384 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.52 | 0.083 | 830 | 1049.6 | 1.5 | 1051.1 | 672.6 | 672.5833 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.14499 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.815 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.52 | 0.083 | 830 | 1049.9 | 1.5 | 1051.4 | 672.7 | 672.7 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.145 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8918 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.52 | 0.083 | 830 | 1050.1 | 1.5 | 1051.6 | 672.5 | 672.5 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3077 | 0.145 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0768 | 29.9 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.52 | 0.083 | 830 | 1050.4 | 1.5 | 1051.9 | 672.4 | 672.4 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.145 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.8549 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.529 | 0.083 | 830 | 1050.3 | 1.5 | 1051.8 | 672.5 | 672.549 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.4 | 0.14509 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7783 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.3 | 1.5 | 1051.8 | 672.6 | 672.5754 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3917 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0827 | 29.7016 | 28.3 | 1.5
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.5 | 1.4 | 1051.9 | 672.8 | 672.8 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.9235 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.5 | 1.4 | 1051.9 | 672.7 | 672.697 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.5 | 1.4 | 1051.9 | 672.8 | 672.8 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3911 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.9159 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.5 | 1.4 | 1051.9 | 672.7 | 672.6709 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7045 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 672.9 | 672.8508 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1451 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7813 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 672.7 | 672.7 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1452 | 0.004291 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.9161 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.536 | 0.083 | 830 | 1050.2 | 1.4 | 1051.6 | 672.8 | 672.8026 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1452 | 0.0043 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1012.9122 | 30 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 672.8 | 672.8491 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3079 | 0.1452 | 0.0043 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0788 | 29.9886 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 672.8 | 672.7971 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3081 | 0.1452 | 0.0043 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.912 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 673 | 673.0206 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1452 | 0.0043 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.0825 | 29.7706 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1050.3 | 1.4 | 1051.7 | 672.9 | 672.8505 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3083 | 0.1452 | 0.0043 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.917 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.54 | 0.083 | 830.4344 | 1050.1 | 1.3 | 1051.4 | 673.2 | 673.2122 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.1452 | 0.004391 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.754 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.0839 | 838.7886 | 1050 | 1.3 | 1051.3 | 673.2 | 673.2 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32.7463 | 55 | 1050 | 1013 | 29.9839 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.084 | 840 | 1049.9 | 1.3 | 1051.2 | 673.4 | 673.3643 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9129 | 29.9287 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.084 | 840 | 1049.7 | 1.3 | 1051 | 673.3 | 673.2502 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.0836 | 836.0635 | 1049.7 | 1.3 | 1051 | 673.1 | 673.1 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.6 | 1.3 | 1050.9 | 673 | 672.9854 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3911 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8975 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.6 | 1.3 | 1050.9 | 672.9 | 672.9 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.1453 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7443 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.7 | 1.3 | 1051 | 672.9 | 672.9 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.4 | 0.1453 | 0.004492 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.9 | 1.3 | 1051.2 | 673 | 672.9794 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3082 | 0.14539 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7312 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.7 | 1.3 | 1051 | 673.1 | 673.1483 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.8 | 1.3 | 1051.1 | 673 | 673.0485 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3086 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830 | 1049.8 | 1.3 | 1051.1 | 672.9 | 672.9 | 0.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7388 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.54 | 0.083 | 830.4104 | 1049.8 | 1.3 | 1051.1 | 672.9 | 672.9 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7379 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.534 | 0.0839 | 838.7856 | 1049.8 | 1.3 | 1051.1 | 673.2 | 673.197 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8145 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.53 | 0.084 | 840 | 1050 | 1.3 | 1051.3 | 673 | 673.0347 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8912 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.53 | 0.084 | 840 | 1050.3 | 1.3 | 1051.6 | 673.1 | 673.0505 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8322 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.53 | 0.084 | 840 | 1050.6 | 1.3 | 1051.9 | 673.1 | 673.1 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3077 | 0.1454 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1013 | 29.8444 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.53 | 0.0838 | 837.5851 | 1050.6 | 1.3 | 1051.9 | 673.5 | 673.5 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.9209 | 28.3 | 1.3
2008-07-16 | 13:02 | 8.53 | 0.0831 | 830.6986 | 1050.8 | 1.4 | 1052.2 | 673.1 | 673.1499 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9973 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.0839 | 838.8221 | 1050.8 | 1.4 | 1052.2 | 673 | 673 | 0.1 | 13 | 954 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3082 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.9257 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.0833 | 832.7726 | 1050.8 | 1.4 | 1052.2 | 673.1 | 673.1 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3916 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1012 | 29.7982 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.8 | 1.4 | 1052.2 | 673.1 | 673.1322 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3086 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7549 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.5 | 1.4 | 1051.9 | 673.2 | 673.2 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8914 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.2 | 1.4 | 1051.6 | 673.1 | 673.1 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3088 | 0.1455 | 0.0046 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1013 | 29.7379 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050.1 | 1.4 | 1051.5 | 673.2 | 673.2016 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1455 | 0.0047 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1013 | 29.8731 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1050 | 1.4 | 1051.4 | 673.3 | 673.266 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.4 | 0.1455 | 0.0047 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9124 | 29.9313 | 28.3 | 1.4
2008-07-16 | 13:02 | 8.53 | 0.083 | 830 | 1049.9 | 1.4 | 1051.3 | 673.2 | 673.1525 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3077 | 0.14559 | 0.0047 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1014 | 29.811 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.532 | 0.083 | 830 | 1049.8 | 1.4 | 1051.2 | 673.2 | 673.2 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.1456 | 0.0047 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.8877 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.54 | 0.083 | 830 | 1050.1 | 1.4 | 1051.5 | 673.1 | 673.1468 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3081 | 0.1456 | 0.0047 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.0806 | 29.9 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.54 | 0.083 | 830 | 1050.1 | 1.4 | 1051.5 | 673 | 672.9503 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1456 | 0.004792 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1013 | 29.941 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.54 | 0.083 | 830 | 1049.7 | 1.4 | 1051.1 | 672.8 | 672.8 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3084 | 0.1456 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9824 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.54 | 0.083 | 830 | 1049.9 | 1.4 | 1051.3 | 673 | 673.0449 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1456 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9057 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.54 | 0.083 | 830 | 1050 | 1.4 | 1051.4 | 672.6 | 672.6497 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.14569 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.088 | 29.9709 | 28.3 | 1.4
2008-07-16 | 13:03 | 8.713 | 0.083 | 830 | 1050 | 1.4 | 1051.4 | 672.8 | 672.8434 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1457 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1013 | 29.9048 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.295 | 0.0879 | 879.0969 | 1052.1 | 1.4 | 1053.5 | 676 | 675.9836 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.14588 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1050 | 1013 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.64 | 0.0958 | 957.985 | 1098.4 | 1.4 | 1099.8 | 708.8 | 708.836 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.14991 | 0.0048 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1177.6 | 1.4 | 1179 | 754.7 | 754.7434 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.15657 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050.1025 | 1012 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1192.6 | 1.4 | 1194 | 766.4 | 766.3964 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.15777 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32.2533 | 55 | 1051 | 1013 | 29.9085 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1194.3 | 1.4 | 1195.7 | 768.9 | 768.878 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3921 | 0.15798 | 0.0049 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.9691 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1196.2 | 1.4 | 1197.6 | 769.4 | 769.4018 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1581 | 0.004992 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1051 | 1013 | 29.9 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1196.5 | 1.4 | 1197.9 | 769.6 | 769.597 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3918 | 0.1582 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1196.4 | 1.4 | 1197.8 | 769.7 | 769.7 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1582 | 0.005 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1051 | 1013 | 29.7782 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1196.7 | 1.4 | 1198.1 | 769.9 | 769.9 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1583 | 0.0051 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7375 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1197.1 | 1.4 | 1198.5 | 769.9 | 769.8535 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1583 | 0.0051 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1197.4 | 1.4 | 1198.8 | 770.1 | 770.0592 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1583 | 0.0051 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012.0871 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1198.4 | 1.4 | 1199.8 | 770 | 770 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1584 | 0.0051 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012 | 29.8 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.69 | 0.095 | 950 | 1199 | 1.4 | 1200.4 | 770.1 | 770.0992 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1585 | 0.005191 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012 | 29.7561 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.699 | 0.095 | 950 | 1199.4 | 1.4 | 1200.8 | 770.1 | 770.07 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.1585 | 0.0052 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.7 | 0.0951 | 950.6683 | 1199.2 | 1.4 | 1200.6 | 770.2 | 770.1503 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.15859 | 0.0052 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.7 | 0.0959 | 958.8183 | 1199 | 1.4 | 1200.4 | 770 | 770 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3921 | 0.1586 | 0.0052 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1012.9212 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.704 | 0.096 | 960 | 1199.1 | 1.4 | 1200.5 | 770.3 | 770.2532 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1586 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.71 | 0.096 | 960 | 1199.1 | 1.4 | 1200.5 | 770.3 | 770.2964 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1586 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.71 | 0.0956 | 956.1589 | 1199.2 | 1.4 | 1200.6 | 770.1 | 770.1366 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1586 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.719 | 0.095 | 950 | 1199.5 | 1.4 | 1200.9 | 770.3 | 770.3494 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1587 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32.7467 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.095 | 950.3294 | 1200.3 | 1.4 | 1201.7 | 770.1 | 770.0982 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1587 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.726 | 0.0959 | 958.7753 | 1201.3 | 1.4 | 1202.7 | 770.2 | 770.2148 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.15889 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7343 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.73 | 0.0953 | 952.8691 | 1201.5 | 1.4 | 1202.9 | 770.3 | 770.2504 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1589 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.8218 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.728 | 0.0955 | 955.4712 | 1201.5 | 1.4 | 1202.9 | 770.6 | 770.5989 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.1589 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.975 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.5 | 1.4 | 1202.9 | 770.5 | 770.4691 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1589 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050.099 | 1013 | 29.8073 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.6 | 1.4 | 1203 | 770.3 | 770.2998 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.1589 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.8 | 1.4 | 1203.2 | 770.4 | 770.4373 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1589 | 0.0053 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.8 | 1.4 | 1203.2 | 770.7 | 770.6518 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1589 | 0.005392 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.8 | 1.4 | 1203.2 | 770.4 | 770.3988 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1589 | 0.0054 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.7 | 1.4 | 1203.1 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1589 | 0.0054 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.7476 | 28.2524 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.8 | 1.4 | 1203.2 | 770.4 | 770.4 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1589 | 0.0054 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1051 | 1013 | 29.8243 | 28.2243 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.7 | 1.4 | 1203.1 | 770.1 | 770.1488 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.15899 | 0.0054 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8991 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.9 | 1.4 | 1203.3 | 770.3 | 770.2586 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1051 | 1013 | 29.8224 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.7 | 1.4 | 1203.1 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7457 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.7 | 1.4 | 1203.1 | 770.4 | 770.3505 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32.2537 | 55 | 1051 | 1013 | 29.7 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.4 | 1.4 | 1202.8 | 770.2 | 770.2352 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32.7467 | 55 | 1051 | 1013.9212 | 29.7152 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.4 | 1.4 | 1202.8 | 770.3 | 770.2519 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3081 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8686 | 28.3 | 1.4
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.5 | 1.4917 | 1203 | 770.3 | 770.3 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9 | 28.3 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.5 | 1.5 | 1203 | 770.4 | 770.4495 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3084 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9 | 28.2623 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.6 | 1.5 | 1203.1 | 770.3 | 770.3488 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8857 | 28.2143 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1201.7 | 1.5 | 1203.2 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013.0882 | 29.809 | 28.291 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1202.1 | 1.5 | 1203.6 | 770.3 | 770.3477 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.159 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8677 | 28.3 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1202.1 | 1.5 | 1203.6 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.2 | 28 | 5 | 195 | 59.4 | 0.159 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012.9127 | 29.9441 | 28.2559 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1202.1 | 1.5 | 1203.6 | 770.3 | 770.3 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9592 | 28.2 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1202.1 | 1.5 | 1203.6 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8051 | 28.2 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.096 | 960 | 1202.3 | 1.5 | 1203.8 | 770.2 | 770.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.9477 | 28.2 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.0953 | 953.0033 | 1202.3 | 1.5 | 1203.8 | 770.1 | 770.1 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 30 | 28.2507 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1202.1 | 1.5 | 1203.6 | 770.1 | 770.1 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9172 | 28.3 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1202.1 | 1.5 | 1203.6 | 770.1 | 770.1403 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3082 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7043 | 28.2957 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1202 | 1.5 | 1203.5 | 769.9 | 769.9 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3084 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.781 | 28.219 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1202 | 1.5 | 1203.5 | 770.3 | 770.3 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3914 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9154 | 28.2 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1201.8 | 1.5 | 1203.3 | 770.2 | 770.2179 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9313 | 28.2344 | 1.5
2008-07-16 | 13:03 | 9.72 | 0.095 | 950 | 1202 | 1.5 | 1203.5 | 770.2 | 770.2004 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.789 | 28.289 | 1.5
2008-07-16 | 13:04 | 9.72 | 0.0955 | 955.4146 | 1202 | 1.5 | 1203.5 | 770.1 | 770.1 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3088 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7123 | 28.2123 | 1.5
2008-07-16 | 13:04 | 9.72 | 0.0956 | 956.1448 | 1201.9 | 1.5 | 1203.4 | 770.2 | 770.1567 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7644 | 28.2 | 1.5
2008-07-16 | 13:04 | 9.72 | 0.095 | 950 | 1202.2 | 1.5 | 1203.7 | 770 | 770.0475 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7589 | 28.2 | 1.5
2008-07-16 | 13:04 | 9.72 | 0.095 | 950 | 1202.2 | 1.5 | 1203.7 | 770.2 | 770.1511 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.3 | 0.1591 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.2177 | 1.5
2008-07-16 | 13:04 | 9.72 | 0.095 | 950 | 1202 | 1.5 | 1203.5 | 770.1 | 770.0694 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3919 | 0.1591 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.2944 | 1.5
2008-07-16 | 13:04 | 10.032 | 0.095 | 950 | 1202 | 1.5 | 1203.5 | 770 | 769.9972 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3082 | 0.1591 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.2289 | 1.5
2008-07-16 | 13:04 | 10.606 | 0.1009 | 1009.383 | 1202.8 | 1.5 | 1204.3 | 770.3 | 770.2701 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.15919 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.1064 | 1063.8391 | 1283.3 | 1.5 | 1284.8 | 822.1 | 822.1336 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.16548 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050.0854 | 1013 | 29.7 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1341.4 | 1.5 | 1342.9 | 857.9 | 857.9429 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.16973 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7046 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1346.3 | 1.5 | 1347.8 | 865.7 | 865.6571 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.17028 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 30.0114 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1347.9 | 1.5 | 1349.4 | 867 | 867.0063 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3085 | 0.1704 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.991 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1348 | 1.5 | 1349.5 | 867.3 | 867.3071 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1704 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8688 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1348.2 | 1.5 | 1349.7 | 868 | 867.9572 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.17049 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1348.7 | 1.5 | 1350.2 | 868 | 868.0481 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1705 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1348.5 | 1.5 | 1350 | 868.5 | 868.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1705 | 0.005792 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1348.3 | 1.5 | 1349.8 | 868.6 | 868.5502 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3081 | 0.1705 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1349.8 | 1.5 | 1351.3 | 868.7 | 868.6502 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.4 | 0.1707 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8146 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1350 | 1.4 | 1351.4 | 868.5 | 868.465 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1707 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8913 | 28.2 | 1.4
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1350 | 1.4 | 1351.4 | 868.4 | 868.4474 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1707 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.832 | 28.2 | 1.4
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1349.8 | 1.4 | 1351.2 | 868.8 | 868.7509 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.17079 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8446 | 28.2 | 1.4
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1350.3 | 1.5 | 1351.8 | 868.4 | 868.4148 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1708 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8788 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.89 | 0.107 | 1070 | 1350.4 | 1.5 | 1351.9 | 868.5 | 868.4975 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1708 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8024 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.885 | 0.107 | 1070 | 1350.3 | 1.5 | 1351.8 | 868.7 | 868.6507 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1708 | 0.005991 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9493 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.3 | 1.5 | 1351.8 | 868.8 | 868.758 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1709 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8975 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.2 | 1.5 | 1351.7 | 868.7 | 868.7 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3923 | 0.1709 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.2 | 1.5 | 1351.7 | 868.9 | 868.8715 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1709 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014.0806 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.2 | 1.5 | 1351.7 | 869 | 869.047 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1709 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.2 | 1.5 | 1351.7 | 869 | 869.0033 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3916 | 0.171 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350 | 1.5 | 1351.5 | 869.1 | 869.0817 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.171 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014.9138 | 29.8 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.2 | 1.5 | 1351.7 | 869 | 868.9527 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.171 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014.087 | 29.8231 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.6 | 1.5 | 1352.1 | 869.1 | 869.0514 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.171 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9764 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.8 | 1.5 | 1352.3 | 868.8 | 868.8 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.171 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9351 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.871 | 0.107 | 1070 | 1351 | 1.5 | 1352.5 | 868.6 | 868.6 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.171 | 0.006191 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8584 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.8 | 1.5 | 1352.3 | 868.5 | 868.4686 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3911 | 0.171 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8365 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.8 | 1.5 | 1352.3 | 868.6 | 868.5978 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.171 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9899 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.8 | 1.5 | 1352.3 | 868.4 | 868.3956 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.171 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9284 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.872 | 0.107 | 1070 | 1350.6 | 1.5 | 1352.1 | 868.2 | 868.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013.9192 | 29.9483 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.107 | 1070 | 1350.6 | 1.5 | 1352.1 | 868.3 | 868.3292 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050.9156 | 1014 | 30 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.871 | 0.107 | 1070 | 1350.6 | 1.5 | 1352.1 | 868.5 | 868.5 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.0016 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.4 | 1.5 | 1351.9 | 868.4 | 868.3712 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.0783 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.7 | 1.5 | 1352.2 | 868.4 | 868.4 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.1 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.107 | 1070 | 1350.5 | 1.5 | 1352 | 868.5 | 868.5 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.0368 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.872 | 0.107 | 1070 | 1350.7 | 1.5 | 1352.2 | 868.3 | 868.3044 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3912 | 0.1711 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8834 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.88 | 0.1074 | 1073.7025 | 1350.9 | 1.5 | 1352.4 | 868.5 | 868.5 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1711 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.7301 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.871 | 0.108 | 1080 | 1351.1 | 1.5 | 1352.6 | 868.5 | 868.5475 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1711 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.0878 | 29.8233 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.1 | 1.5 | 1352.6 | 868.5 | 868.5 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3921 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1050 | 1013 | 29.8617 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.5 | 1.5 | 1353 | 868.4 | 868.4482 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.815 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.7 | 1.5 | 1353.2 | 868.4 | 868.4017 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3081 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.081 | 29.8918 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.7 | 1.5 | 1353.2 | 868.5 | 868.5339 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32.7463 | 55 | 1050 | 1012 | 29.8315 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.6 | 1.5 | 1353.1 | 868.4 | 868.4254 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9138 | 29.8452 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.108 | 1080 | 1351.4 | 1.5 | 1352.9 | 868.4 | 868.438 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3088 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8781 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.1079 | 1079.4399 | 1351.1 | 1.5 | 1352.6 | 868.2 | 868.2041 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3913 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9129 | 29.8014 | 28.2 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.1063 | 1062.9654 | 1350.9 | 1.5 | 1352.4 | 868.4 | 868.4492 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9505 | 28.1248 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1350.7 | 1.5 | 1352.2 | 868.3 | 868.2919 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8454 | 28.1515 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1350.9 | 1.5 | 1352.4 | 868.5 | 868.5 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.728 | 28.172 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1351.1 | 1.5 | 1352.6 | 868.6 | 868.6026 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8 | 28.105 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1351.1 | 1.5 | 1352.6 | 868.6 | 868.6 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3079 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8 | 28.1814 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1351.1 | 1.5 | 1352.6 | 868.2 | 868.1937 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8 | 28.1417 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1350.8 | 1.5 | 1352.3 | 868 | 868 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3916 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.8352 | 28.1 | 1.5
2008-07-16 | 13:04 | 10.87 | 0.106 | 1060 | 1350.8 | 1.5 | 1352.3 | 868 | 868.0357 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3086 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 29.9 | 28.1 | 1.5
2008-07-16 | 13:05 | 10.87 | 0.106 | 1060 | 1350.6 | 1.5 | 1352.1 | 867.8 | 867.8499 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1015 | 29.9 | 28.1 | 1.5
2008-07-16 | 13:05 | 10.87 | 0.1064 | 1063.7088 | 1350.2 | 1.5 | 1351.7 | 867.9 | 867.9 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1712 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1015 | 29.7689 | 28.1 | 1.5
2008-07-16 | 13:05 | 10.87 | 0.107 | 1070 | 1350.4 | 1.5 | 1351.9 | 867.7 | 867.7 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1712 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1015 | 29.8267 | 28.1 | 1.5
2008-07-16 | 13:05 | 10.87 | 0.107 | 1070 | 1350.6 | 1.5 | 1352.1 | 867.8 | 867.8 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1712 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1015 | 29.9434 | 28.1189 | 1.5
2008-07-16 | 13:05 | 10.993 | 0.107 | 1070 | 1351 | 1.5 | 1352.5 | 867.8 | 867.8 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1713 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014.0876 | 29.7133 | 28.1956 | 1.5
2008-07-16 | 13:05 | 11.409 | 0.1161 | 1160.5456 | 1351 | 1.5 | 1352.5 | 872.1 | 872.0929 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1713 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050.0872 | 1014 | 29.8446 | 28.1277 | 1.5
2008-07-16 | 13:05 | 11.942 | 0.1189 | 1189.1674 | 1408.2 | 1.6 | 1409.8 | 904.6 | 904.5561 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.17535 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013.0788 | 29.851 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.95 | 0.118 | 1180 | 1482.5 | 1.6 | 1484.1 | 951.5 | 951.4609 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1804 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.95 | 0.1182 | 1181.9231 | 1497.6 | 1.6 | 1499.2 | 963.6 | 963.6195 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.18148 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.95 | 0.119 | 1190 | 1501.1 | 1.6 | 1502.7 | 965.9 | 965.9413 | 0.1 | 12.9 | 958 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3916 | 0.1817 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.955 | 0.119 | 1190 | 1501.4 | 1.6 | 1503 | 966.5 | 966.4664 | 0.1 | 12.9 | 958 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.18179 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7442 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.96 | 0.119 | 1190 | 1501.7 | 1.6 | 1503.3 | 966.7 | 966.7041 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1818 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7325 | 28.1325 | 1.6
2008-07-16 | 13:05 | 11.966 | 0.119 | 1190 | 1501.8 | 1.6 | 1503.4 | 966.9 | 966.9 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1818 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1908 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502 | 1.6 | 1503.6 | 966.8 | 966.8063 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1818 | 0.006591 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1141 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502.4 | 1.6 | 1504 | 967.2 | 967.2014 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.18189 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7374 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502.4 | 1.6 | 1504 | 967 | 967.0087 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1819 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7785 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502.3 | 1.6 | 1503.9 | 966.9 | 966.9491 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1819 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.884 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502.4 | 1.6 | 1504 | 967 | 966.9501 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1819 | 0.006692 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8074 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.119 | 1190 | 1502.5 | 1.6 | 1504.1 | 967 | 966.956 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.18199 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.1193 | 1193.4305 | 1502.7 | 1.6 | 1504.3 | 967.2 | 967.2 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.182 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1502.8 | 1.6 | 1504.4 | 967.3 | 967.2583 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.182 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8228 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1503.1 | 1.6 | 1504.7 | 967.3 | 967.3351 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.1821 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8995 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1502.8 | 1.6 | 1504.4 | 967.6 | 967.5538 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1821 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8239 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1502.8 | 1.6 | 1504.4 | 967.7 | 967.7 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1821 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7471 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1503.3 | 1.6 | 1504.9 | 967.8 | 967.8 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1821 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7296 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1503.5 | 1.6 | 1505.1 | 967.8 | 967.7505 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1822 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.99 | 0.12 | 1200 | 1504 | 1.6 | 1505.6 | 967.7 | 967.6688 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1822 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.6
2008-07-16 | 13:05 | 11.997 | 0.12 | 1200 | 1504.1 | 1.6 | 1505.7 | 967.6 | 967.6 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.3912 | 0.1822 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7405 | 28.1 | 1.6
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1504.5 | 1.6919 | 1506.2 | 967.5 | 967.4544 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1822 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1504.3 | 1.7 | 1506 | 967.4 | 967.4108 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1822 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7248 | 28.1 | 1.7
2008-07-16 | 13:05 | 12.002 | 0.12 | 1200 | 1504.6 | 1.7 | 1506.3 | 967.6 | 967.6 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1822 | 0.006892 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8789 | 28.1 | 1.7
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1505 | 1.7 | 1506.7 | 967.7 | 967.6515 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.18229 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.834 | 28.1 | 1.7
2008-07-16 | 13:05 | 12.001 | 0.12 | 1200 | 1505.1 | 1.7 | 1506.8 | 967.6 | 967.5714 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1823 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1505.9 | 1.7 | 1507.6 | 967.5 | 967.5 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.18239 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1506.2 | 1.7 | 1507.9 | 967.8 | 967.7603 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1824 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1506.2 | 1.7 | 1507.9 | 967.7 | 967.7 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1824 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7267 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1506.3 | 1.7 | 1508 | 967.7 | 967.6945 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3914 | 0.1824 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.75 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1506.9 | 1.7 | 1508.6 | 967.4 | 967.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.18249 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12 | 0.12 | 1200 | 1507.2 | 1.7 | 1508.9 | 967.6 | 967.6497 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3921 | 0.1825 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12.007 | 0.12 | 1200 | 1507.3 | 1.7 | 1509 | 967.6 | 967.5743 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1825 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.7
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.7 | 1.8 | 1509.5 | 967.5 | 967.5168 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.1825 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012.0843 | 29.7431 | 28.1569 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.7 | 1.8 | 1509.5 | 967.5 | 967.5473 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1825 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7 | 28.1664 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1508 | 1.8 | 1509.8 | 967.7 | 967.6571 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1825 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7103 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1508.5 | 1.8 | 1510.3 | 967.9 | 967.9 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1825 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.787 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1508.6 | 1.8 | 1510.4 | 967.7 | 967.7469 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7364 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1508.4 | 1.8 | 1510.2 | 967.7 | 967.6572 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7 | 28.0596 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1508.2 | 1.8 | 1510 | 967.6 | 967.5516 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7342 | 28.0171 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.7 | 1.8 | 1509.5 | 967.8 | 967.7961 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3914 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8876 | 28.0938 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.5 | 1.8 | 1509.3 | 967.8 | 967.8055 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3079 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9704 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507 | 1.8 | 1508.8 | 967.8 | 967.8 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.4 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9057 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507 | 1.8 | 1508.8 | 967.7 | 967.7481 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8479 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1506.9 | 1.8 | 1508.7 | 967.9 | 967.9164 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3916 | 0.1826 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.9988 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.1 | 1.8 | 1508.9 | 967.9 | 967.8504 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1826 | 0.007091 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1014 | 29.8454 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.3 | 1.8 | 1509.1 | 967.7 | 967.7 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1826 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.5 | 1.8 | 1509.3 | 967.5 | 967.5021 | 0.1 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.1826 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7693 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.8 | 1.8 | 1509.6 | 967.4 | 967.4 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3085 | 0.1826 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7074 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.01 | 0.12 | 1200 | 1507.9 | 1.8 | 1509.7 | 967.9 | 967.8586 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3911 | 0.1826 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7841 | 28.1 | 1.8
2008-07-16 | 13:05 | 12.017 | 0.12 | 1200 | 1507.9 | 1.8 | 1509.7 | 967.8 | 967.8474 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.7392 | 28.0392 | 1.8
2008-07-16 | 13:05 | 12.02 | 0.12 | 1200 | 1507.9 | 1.8 | 1509.7 | 968 | 968 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3079 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.8125 | 28.0375 | 1.8
2008-07-16 | 13:05 | 12.02 | 0.12 | 1200 | 1507.9 | 1.8 | 1509.7 | 967.8 | 967.8029 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1013 | 29.9858 | 28.0858 | 1.8
2008-07-16 | 13:06 | 12.02 | 0.12 | 1200 | 1507.8 | 1.8 | 1509.6 | 967.8 | 967.8 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012.9175 | 29.9091 | 28.0091 | 1.8
2008-07-16 | 13:06 | 12.02 | 0.12 | 1200 | 1508 | 1.8 | 1509.8 | 967.8 | 967.8 | 0.1 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012.0843 | 29.8324 | 28 | 1.8
2008-07-16 | 13:06 | 12.02 | 0.12 | 1200 | 1508 | 1.8 | 1509.8 | 967.8 | 967.7732 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012.0862 | 29.7559 | 28 | 1.8
2008-07-16 | 13:06 | 12.02 | 0.12 | 1200 | 1508 | 1.8 | 1509.8 | 967.8 | 967.8 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7 | 28 | 1.8
2008-07-16 | 13:06 | 12.02 | 0.12 | 1200 | 1508 | 1.8 | 1509.8 | 967.8 | 967.8 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1051 | 1012 | 29.7 | 28 | 1.8
2008-07-16 | 13:06 | 11.721 | 0.12 | 1200 | 1507.8 | 1.8 | 1509.6 | 967.8 | 967.8 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.3 | 0.1827 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050.9215 | 1013 | 29.9223 | 28 | 1.8
2008-07-16 | 13:06 | 11.122 | 0.114 | 1139.8499 | 1492.6 | 1.8 | 1494.4 | 965.4 | 965.4182 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1817 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8982 | 28.0509 | 1.8
2008-07-16 | 13:06 | 10.97 | 0.108 | 1080 | 1407.2 | 1.8 | 1409 | 912.5 | 912.5177 | 0.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28.246 | 192 | 28 | 5 | 195 | 59.3912 | 0.17587 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7722 | 28.1 | 1.8
2008-07-16 | 13:06 | 10.961 | 0.108 | 1080 | 1372.8 | 1.8 | 1374.6 | 877.9 | 877.9053 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.754 | 192 | 28 | 5 | 195 | 59.3923 | 0.17344 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7129 | 28.1 | 1.8
2008-07-16 | 13:06 | 10.946 | 0.108 | 1080 | 1360.8 | 1.8 | 1362.6 | 870.8 | 870.8412 | 0.1 | 13 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.17251 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9435 | 28.1 | 1.8
2008-07-16 | 13:06 | 10.94 | 0.108 | 1080 | 1360.3 | 1.8 | 1362.1 | 869.3 | 869.3296 | 0.1 | 13 | 949 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1725 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0579 | 28.1 | 1.8
2008-07-16 | 13:06 | 10.923 | 0.108 | 1080 | 1360.3 | 1.8 | 1362.1 | 869.1 | 869.1493 | 0.1 | 13 | 949 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1725 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9962 | 28.1 | 1.8
2008-07-16 | 13:06 | 10.888 | 0.108 | 1080 | 1360.2 | 1.8 | 1362 | 868.9 | 868.9362 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.246 | 192 | 28 | 5 | 195 | 59.3 | 0.1725 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.0887 | 1.8
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1358.8 | 1.8 | 1360.6 | 868.9 | 868.8653 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.012 | 1.8
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1358.8 | 1.8 | 1360.6 | 868.9 | 868.8533 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0882 | 29.8 | 28.0648 | 1.8
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1358.6 | 1.8911 | 1360.5 | 868.9 | 868.9498 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.754 | 192 | 28 | 5 | 195 | 59.4 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.0585 | 1.9
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1358.4 | 1.8087 | 1360.2 | 868.8 | 868.8024 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.0183 | 1.8
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1358.3 | 1.9 | 1360.2 | 868.5 | 868.5 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.0949 | 1.9
2008-07-16 | 13:06 | 10.882 | 0.108 | 1080 | 1358.3 | 1.9 | 1360.2 | 868.7 | 868.7 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8 | 28.0284 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1358.2 | 1.9 | 1360.1 | 868.6 | 868.5509 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3921 | 0.1724 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.8483 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1357.1 | 1.9 | 1359 | 868.5 | 868.5468 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9251 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1357 | 1.9 | 1358.9 | 868.5 | 868.5 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1356.9 | 1.9 | 1358.8 | 868.7 | 868.7055 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3916 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1357.1 | 1.9 | 1359 | 868.7 | 868.7 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1357 | 1.9 | 1358.9 | 868.6 | 868.6 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1356.8 | 1.9 | 1358.7 | 868.6 | 868.6499 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9916 | 28 | 1.9
2008-07-16 | 13:06 | 10.89 | 0.108 | 1080 | 1356.8 | 1.9 | 1358.7 | 868.8 | 868.7515 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9147 | 28 | 1.9
2008-07-16 | 13:06 | 10.883 | 0.108 | 1080 | 1356.6 | 1.9 | 1358.5 | 868.9 | 868.9271 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9 | 28 | 1.9
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1356.7 | 1.9 | 1358.6 | 868.9 | 868.9043 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3925 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9386 | 28 | 1.9
2008-07-16 | 13:06 | 10.88 | 0.108 | 1080 | 1356.5 | 1.9 | 1358.4 | 869.3 | 869.2551 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.877 | 0.108 | 1080 | 1356.3 | 1.9 | 1358.2 | 869.2 | 869.1982 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.1077 | 1076.7 | 1356.3 | 1.9 | 1358.2 | 869.5 | 869.4561 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.246 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9194 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.6 | 1.9 | 1358.5 | 869.5 | 869.4898 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3 | 0.1723 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.8 | 1.9 | 1358.7 | 870 | 870.0024 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.754 | 192 | 28 | 5 | 195 | 59.3916 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0221 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.7 | 1.9 | 1358.6 | 869.9 | 869.9496 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0988 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.9 | 1.9 | 1358.8 | 870 | 869.9534 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.0245 | 28 | 1.9
2008-07-16 | 13:06 | 10.858 | 0.107 | 1070 | 1356.6 | 1.9 | 1358.5 | 869.8 | 869.8467 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.85 | 0.107 | 1070 | 1356.6 | 1.9 | 1358.5 | 869.5 | 869.4963 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9719 | 28 | 1.9
2008-07-16 | 13:06 | 10.859 | 0.107 | 1070 | 1356.3 | 1.9 | 1358.2 | 869.1 | 869.1355 | 0.1 | 13 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.9103 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.1 | 1.9 | 1358 | 868.8 | 868.8012 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0644 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.1 | 1.9 | 1358 | 868.9 | 868.9 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.1 | 1.9 | 1358 | 868.8 | 868.8 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1356.1 | 1.9 | 1358 | 868.9 | 868.9 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.246 | 192 | 28 | 5 | 195 | 59.3918 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1355.7 | 1.9 | 1357.6 | 868.6 | 868.6 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28.754 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1355.7 | 1.9 | 1357.6 | 868.9 | 868.8773 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1355.9 | 1.9 | 1357.8 | 869.1 | 869.1 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1424 | 28 | 1.9
2008-07-16 | 13:06 | 10.865 | 0.107 | 1070 | 1355.7 | 1.9 | 1357.6 | 869.1 | 869.0595 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.2462 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1809 | 28 | 1.9
2008-07-16 | 13:06 | 10.87 | 0.107 | 1070 | 1355.5 | 1.9 | 1357.4 | 868.8 | 868.8253 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1042 | 28 | 1.9
2008-07-16 | 13:06 | 10.87 | 0.107 | 1070 | 1355.4 | 1.9 | 1357.3 | 868.7 | 868.6992 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0275 | 28 | 1.9
2008-07-16 | 13:06 | 10.87 | 0.107 | 1070 | 1355.7 | 1.9 | 1357.6 | 868.9 | 868.935 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0983 | 28 | 1.9
2008-07-16 | 13:06 | 10.862 | 0.107 | 1070 | 1355.7 | 1.9 | 1357.6 | 868.9 | 868.8509 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9232 | 30.1483 | 28 | 1.9
2008-07-16 | 13:06 | 10.86 | 0.107 | 1070 | 1355.4 | 1.9 | 1357.3 | 869.2 | 869.1568 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.865 | 0.107 | 1070 | 1355.5 | 1.9 | 1357.4 | 868.8 | 868.8 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3919 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:06 | 10.87 | 0.1077 | 1076.6612 | 1355.7 | 1.9 | 1357.6 | 868.9 | 868.853 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.056 | 28 | 1.9
2008-07-16 | 13:06 | 10.865 | 0.108 | 1080 | 1355.7 | 1.9 | 1357.6 | 868.8 | 868.8029 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3084 | 0.1722 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.84 | 0.108 | 1080 | 1355.7 | 1.9 | 1357.6 | 868.6 | 868.5986 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.17211 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1094 | 28 | 1.9
2008-07-16 | 13:06 | 10.84 | 0.108 | 1080 | 1355.6 | 1.9 | 1357.5 | 868.5 | 868.4982 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.4 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.186 | 28 | 1.9
2008-07-16 | 13:06 | 10.847 | 0.108 | 1080 | 1355.4 | 1.9 | 1357.3 | 868.5 | 868.5 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1373 | 28 | 1.9
2008-07-16 | 13:06 | 10.85 | 0.108 | 1080 | 1355.3 | 1.9 | 1357.2 | 868.5 | 868.5 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.3087 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:06 | 10.85 | 0.108 | 1080 | 1355.1 | 1.9 | 1357 | 868.5 | 868.5 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191.2 | 28 | 5 | 195 | 59.4 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1161 | 28 | 1.9
2008-07-16 | 13:07 | 10.85 | 0.108 | 1080 | 1355 | 1.9 | 1356.9 | 868.5 | 868.4646 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1929 | 28 | 1.9
2008-07-16 | 13:07 | 10.85 | 0.108 | 1080 | 1354.6 | 1.9 | 1356.5 | 868.2 | 868.2477 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3079 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9212 | 30.2 | 28 | 1.9
2008-07-16 | 13:07 | 10.85 | 0.108 | 1080 | 1354.8 | 1.9 | 1356.7 | 868.7 | 868.738 | 0.1 | 13 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3919 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1538 | 28 | 1.9
2008-07-16 | 13:07 | 10.857 | 0.108 | 1080 | 1354.7 | 1.9 | 1356.6 | 868.5 | 868.5255 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3918 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1229 | 28 | 1.9
2008-07-16 | 13:07 | 10.86 | 0.108 | 1080 | 1354.6 | 1.9 | 1356.5 | 868.3 | 868.3 | 0.1 | 13 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3084 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1997 | 28 | 1.9
2008-07-16 | 13:07 | 10.603 | 0.108 | 1080 | 1354.6 | 1.9 | 1356.5 | 868.3 | 868.3 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.9
2008-07-16 | 13:07 | 10.033 | 0.1019 | 1019.1753 | 1354.7 | 1.9 | 1356.6 | 867.5 | 867.5232 | 0.1 | 13 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3 | 0.1721 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.9
2008-07-16 | 13:07 | 9.76 | 0.0963 | 963.4327 | 1273.4 | 1.9 | 1275.3 | 821.7 | 821.6795 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 28.754 | 191 | 28 | 5 | 195 | 59.4 | 0.16602 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1405 | 28 | 1.9
2008-07-16 | 13:07 | 9.76 | 0.097 | 970 | 1215.8 | 1.9 | 1217.7 | 781.7 | 781.6674 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.16166 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:07 | 9.76 | 0.097 | 970 | 1210.2 | 1.9 | 1212.1 | 773.2 | 773.159 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.16112 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30 | 28 | 1.9
2008-07-16 | 13:07 | 9.745 | 0.097 | 970 | 1209.4 | 1.9 | 1211.3 | 771.6 | 771.6 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.16101 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1193 | 28 | 1.9
2008-07-16 | 13:07 | 9.73 | 0.097 | 970 | 1208.9 | 1.9 | 1210.8 | 771.6 | 771.634 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1609 | 0.0071 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2361 | 28 | 1.9
2008-07-16 | 13:07 | 9.73 | 0.097 | 970 | 1207.8 | 1.9 | 1209.7 | 771.2 | 771.2464 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3081 | 0.16081 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2749 | 28 | 1.9
2008-07-16 | 13:07 | 9.73 | 0.097 | 970 | 1207.5 | 1.9 | 1209.4 | 771.4 | 771.3577 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.1608 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0825 | 30.1208 | 28 | 1.9
2008-07-16 | 13:07 | 9.738 | 0.097 | 970 | 1207.6 | 1.9 | 1209.5 | 771.1 | 771.0989 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.3 | 0.1607 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1661 | 28 | 1.9
2008-07-16 | 13:07 | 9.727 | 0.097 | 970 | 1207.5 | 1.9 | 1209.4 | 771 | 770.9515 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.1607 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.097 | 970 | 1206.8 | 1.9 | 1208.7 | 771 | 771.007 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.1607 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.912 | 30.1802 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.097 | 970 | 1206.5 | 1.9 | 1208.4 | 770.8 | 770.8478 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1607 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1034 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.097 | 970 | 1206.7 | 1.9 | 1208.6 | 770.5 | 770.4937 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1607 | 0.007 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1734 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.0967 | 966.797 | 1206.4 | 1.9 | 1208.3 | 770.4 | 770.3581 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1606 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0997 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1205.5 | 1.9 | 1207.4 | 770.2 | 770.2495 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1606 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.0268 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1204.8 | 1.9 | 1206.7 | 770.1 | 770.1487 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3077 | 0.1605 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1204.5 | 1.9 | 1206.4 | 770.1 | 770.1 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3079 | 0.1604 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1204.6 | 1.9 | 1206.5 | 770.1 | 770.1063 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1604 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.043 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1204.5 | 1.9 | 1206.4 | 770.4 | 770.3553 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1604 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.0673 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.7 | 1.8 | 1204.5 | 770.2 | 770.2417 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1603 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1793 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.7 | 1.8914 | 1204.6 | 770.2 | 770.1504 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.16021 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011.0862 | 30.0259 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.5 | 1.9 | 1204.4 | 770.1 | 770.1207 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.3088 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.0637 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.8 | 1.9 | 1204.7 | 770 | 770.0497 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.1 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.8 | 1.9 | 1204.7 | 770.1 | 770.1481 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.1172 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.7 | 1.8 | 1204.5 | 770 | 770 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.1939 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.5 | 1.8 | 1204.3 | 770.2 | 770.2473 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.5 | 1.8 | 1204.3 | 770 | 769.9502 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3077 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.1527 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202.4 | 1.8 | 1204.2 | 770.1 | 770.1297 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.076 | 28 | 1.8
2008-07-16 | 13:07 | 9.725 | 0.096 | 960 | 1202.4 | 1.8 | 1204.2 | 770 | 769.9768 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011.0806 | 30.0014 | 28 | 1.8
2008-07-16 | 13:07 | 9.73 | 0.096 | 960 | 1202.3 | 1.8 | 1204.1 | 769.9 | 769.8872 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3917 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.1548 | 28 | 1.8
2008-07-16 | 13:07 | 9.728 | 0.096 | 960 | 1202.2 | 1.8 | 1204 | 769.8 | 769.8483 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011.0843 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1202 | 1.8 | 1203.8 | 769.9 | 769.9008 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011.9137 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.6 | 1.8 | 1203.4 | 769.8 | 769.8179 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.6 | 1.8 | 1203.4 | 769.8 | 769.8019 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0871 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.4 | 1.8 | 1203.2 | 770 | 770 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2609 | 28 | 1.8
2008-07-16 | 13:07 | 9.717 | 0.096 | 960 | 1201.2 | 1.8 | 1203 | 769.9 | 769.9429 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2624 | 28 | 1.8
2008-07-16 | 13:07 | 9.71 | 0.096 | 960 | 1201.2 | 1.8 | 1203 | 769.9 | 769.9494 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.71 | 0.096 | 960 | 1201.1 | 1.8 | 1202.9 | 769.9 | 769.8527 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.718 | 0.096 | 960 | 1201.1 | 1.8 | 1202.9 | 770.1 | 770.0663 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.0806 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1200.9 | 1.8 | 1202.7 | 770.1 | 770.1 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.9 | 1203 | 770.2 | 770.1651 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.0843 | 30.2208 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.9 | 1203 | 770.1 | 770.1051 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2978 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.9 | 1203 | 770.2 | 770.2499 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3085 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2257 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.9 | 1203 | 770.2 | 770.2 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.9
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.8087 | 1202.9 | 770.3 | 770.2503 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3913 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.8 | 1202.9 | 770.4 | 770.4 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3087 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1957 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.8 | 1202.9 | 770.3 | 770.2701 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1186 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.1 | 1.8 | 1202.9 | 770.4 | 770.353 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.3 | 1.8 | 1203.1 | 770.2 | 770.243 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.309 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1349 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.096 | 960 | 1201.4 | 1.8 | 1203.2 | 770.4 | 770.3963 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3081 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.0958 | 958.2494 | 1201.4 | 1.8 | 1203.2 | 770.1 | 770.1482 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3084 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.095 | 950.2558 | 1201.8 | 1.8 | 1203.6 | 770.5 | 770.4633 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3914 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.095 | 950 | 1201.8 | 1.8 | 1203.6 | 770.5 | 770.5027 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:07 | 9.72 | 0.095 | 950 | 1201.8 | 1.8 | 1203.6 | 770.4 | 770.4 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 9.72 | 0.095 | 950 | 1201.6 | 1.8 | 1203.4 | 770.2 | 770.2 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 9.72 | 0.095 | 950 | 1201.3 | 1.8 | 1203.1 | 770.3 | 770.2785 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.4153 | 28 | 1.8
2008-07-16 | 13:08 | 9.72 | 0.095 | 950 | 1201.1 | 1.8 | 1202.9 | 770.3 | 770.2891 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3546 | 28 | 1.8
2008-07-16 | 13:08 | 9.72 | 0.095 | 950 | 1200.9 | 1.8 | 1202.7 | 770.2 | 770.175 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3077 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.9232 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 9.72 | 0.095 | 950 | 1200.8 | 1.8 | 1202.6 | 770.3 | 770.3487 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.1602 | 0.0069 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 9.723 | 0.095 | 950 | 1200.9 | 1.8 | 1202.7 | 770.2 | 770.2481 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1601 | 0.006808 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 9.588 | 0.0945 | 945.2961 | 1200.9 | 1.8 | 1202.7 | 770.3 | 770.2646 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.1601 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.1447 | 28 | 1.8
2008-07-16 | 13:08 | 8.741 | 0.0909 | 909.0678 | 1200.8 | 1.8 | 1202.6 | 762.2 | 762.2304 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3916 | 0.1601 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.1 | 28 | 1.8
2008-07-16 | 13:08 | 8.589 | 0.084 | 840 | 1127.9 | 1.8 | 1129.7 | 723.1 | 723.1062 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.15414 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1087 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1066.3 | 1.8 | 1068.1 | 682.4 | 682.4062 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.14917 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.1854 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1057 | 1.8 | 1058.8 | 675.2 | 675.2372 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3089 | 0.14793 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2622 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1054.9 | 1.8 | 1056.7 | 673.7 | 673.7343 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1478 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2611 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1054.7 | 1.8 | 1056.5 | 673.3 | 673.2784 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1478 | 0.0068 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1054 | 1.8 | 1055.8 | 673.3 | 673.3034 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1477 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 8.56 | 0.084 | 840 | 1053.8 | 1.8 | 1055.6 | 673.3 | 673.2656 | 0.1 | 12.9 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.1476 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2689 | 28 | 1.8
2008-07-16 | 13:08 | 8.553 | 0.084 | 840 | 1053.7 | 1.8 | 1055.5 | 673.4 | 673.3506 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1476 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.3 | 28 | 1.8
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1053.7 | 1.8 | 1055.5 | 673.1 | 673.053 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3082 | 0.1476 | 0.0067 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0825 | 30.3 | 28 | 1.8
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1053.2 | 1.8 | 1055 | 673 | 673 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1475 | 0.006608 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3 | 28 | 1.8
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1053.3 | 1.8 | 1055.1 | 672.8 | 672.8 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1475 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2242 | 28 | 1.8
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1053.4 | 1.8 | 1055.2 | 672.9 | 672.8632 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.1475 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.8
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1053 | 1.7 | 1054.7 | 672.9 | 672.8536 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.14741 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.6 | 1.7 | 1054.3 | 672.9 | 672.8507 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1474 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.5 | 1.7 | 1054.2 | 673.1 | 673.1077 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1474 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.4 | 1.7 | 1054.1 | 672.9 | 672.9244 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1473 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2588 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.4 | 1.7 | 1054.1 | 672.9 | 672.9487 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1473 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2645 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.2 | 1.7 | 1053.9 | 672.7 | 672.7 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1473 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2125 | 28 | 1.7
2008-07-16 | 13:08 | 8.55 | 0.084 | 840 | 1052.3 | 1.7 | 1054 | 672.8 | 672.7742 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1473 | 0.0066 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2891 | 28 | 1.7
2008-07-16 | 13:08 | 8.543 | 0.084 | 840 | 1052.4 | 1.7 | 1054.1 | 672.6 | 672.6481 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2341 | 28 | 1.7
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.6 | 1.7 | 1054.3 | 672.6 | 672.6 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.7
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.5 | 1.6 | 1054.1 | 672.4 | 672.4 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.5 | 1.6 | 1054.1 | 672.6 | 672.5596 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.4 | 1.6 | 1054 | 672.5 | 672.548 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1269 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.3 | 1.6 | 1053.9 | 672.6 | 672.6019 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1473 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1498 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.4 | 1.6 | 1054 | 672.6 | 672.5839 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.3087 | 0.14721 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.5 | 1.6 | 1054.1 | 672.6 | 672.6 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.3 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.0876 | 30.2032 | 28 | 1.6
2008-07-16 | 13:08 | 8.54 | 0.084 | 840 | 1052.5 | 1.6 | 1054.1 | 672.6 | 672.6462 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.3 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.28 | 28 | 1.6
2008-07-16 | 13:08 | 8.539 | 0.084 | 840 | 1052.6 | 1.6 | 1054.2 | 672.6 | 672.6085 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2433 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.8 | 1.6 | 1054.4 | 672.7 | 672.7 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.3001 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.8 | 1.6 | 1054.4 | 672.4 | 672.4 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3918 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.4699 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.8 | 1.6 | 1054.4 | 672.6 | 672.631 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2398 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.5 | 1.6 | 1054.1 | 672.3 | 672.3497 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1472 | 0.0065 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2635 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.6 | 1053.7 | 672.6 | 672.5724 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1472 | 0.006409 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2598 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.6 | 1053.7 | 672.3 | 672.252 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3911 | 0.1472 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1051.9 | 1.6 | 1053.5 | 672.4 | 672.4496 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1472 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052 | 1.6 | 1053.6 | 672.6 | 672.615 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1472 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052 | 1.6 | 1053.6 | 672.4 | 672.4 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.309 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.3 | 1.6 | 1053.9 | 672.5 | 672.5488 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.3 | 1.6 | 1053.9 | 672.3 | 672.3426 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3081 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.6
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.3 | 1.5082 | 1053.8 | 672.6 | 672.5518 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3918 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.2 | 1.5 | 1053.7 | 672.6 | 672.5505 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.5 | 1053.6 | 672.7 | 672.6654 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.5 | 1053.6 | 672.6 | 672.6 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.3 | 1.5 | 1053.8 | 672.6 | 672.6231 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049.5088 | 1011 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.5 | 1053.6 | 672.8 | 672.7521 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1052.1 | 1.5 | 1053.6 | 672.7 | 672.6755 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.1471 | 0.0064 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1051.8 | 1.5 | 1053.3 | 672.5 | 672.5121 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1471 | 0.006309 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1050.9 | 1.5 | 1052.4 | 672.4 | 672.4027 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.3921 | 0.147 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.53 | 0.084 | 840 | 1050.4 | 1.5 | 1051.9 | 672.4 | 672.4237 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.147 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9194 | 30.2 | 28 | 1.5
2008-07-16 | 13:08 | 8.535 | 0.084 | 840 | 1050.4 | 1.5 | 1051.9 | 672.5 | 672.458 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.14691 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2442 | 28 | 1.5
2008-07-16 | 13:08 | 8.54 | 0.0834 | 833.538 | 1050.4 | 1.4085 | 1051.8 | 672.1 | 672.1 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2793 | 28 | 1.4
2008-07-16 | 13:09 | 8.539 | 0.083 | 830 | 1050.4 | 1.4 | 1051.8 | 672.1 | 672.1252 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.203 | 28 | 1.4
2008-07-16 | 13:09 | 8.53 | 0.0833 | 833.107 | 1050.4 | 1.4 | 1051.8 | 672.2 | 672.2141 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 8.53 | 0.084 | 840 | 1050.5 | 1.4 | 1051.9 | 672.4 | 672.4 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3915 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049.5388 | 1012 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 8.53 | 0.084 | 840 | 1050.5 | 1.4 | 1051.9 | 672.3 | 672.3397 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049.9116 | 1012 | 30.2822 | 28 | 1.4
2008-07-16 | 13:09 | 8.53 | 0.084 | 840 | 1050.5 | 1.4 | 1051.9 | 672.3 | 672.3 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.4914 | 28 | 1.4
2008-07-16 | 13:09 | 8.399 | 0.084 | 840 | 1050.2 | 1.4 | 1051.6 | 672.3 | 672.3483 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3375 | 28 | 1.4
2008-07-16 | 13:09 | 7.997 | 0.078 | 779.7164 | 1050.2 | 1.4 | 1051.6 | 672.4 | 672.4212 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1469 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.358 | 28 | 1.4
2008-07-16 | 13:09 | 7.403 | 0.0713 | 712.975 | 991.1 | 1.4 | 992.5 | 644.9 | 644.9131 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3921 | 0.14168 | 0.0063 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9212 | 30.3653 | 28 | 1.4
2008-07-16 | 13:09 | 7.38 | 0.072 | 720 | 918.9 | 1.4 | 920.3 | 592.9 | 592.9095 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3081 | 0.13462 | 0.006208 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0806 | 30.3114 | 28 | 1.4
2008-07-16 | 13:09 | 7.38 | 0.072 | 720 | 908.8 | 1.4 | 910.2 | 580.4 | 580.413 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.13363 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3882 | 28 | 1.4
2008-07-16 | 13:09 | 7.38 | 0.072 | 720 | 904.7 | 1.4 | 906.1 | 577.8 | 577.7897 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.13321 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2702 | 27.9351 | 1.4
2008-07-16 | 13:09 | 7.377 | 0.072 | 720 | 903.5 | 1.4 | 904.9 | 576.9 | 576.8507 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1331 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 27.9416 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.9 | 1.4 | 904.3 | 576.7 | 576.6667 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.13301 | 0.0062 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2365 | 28 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.7 | 1.4 | 904.1 | 576.3 | 576.2927 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.3915 | 0.133 | 0.006109 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3899 | 28 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.6 | 1.4 | 904 | 576.4 | 576.3617 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.133 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2566 | 28 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.6 | 1.4 | 904 | 576.1 | 576.099 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1329 | 0.0061 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9124 | 30.3452 | 28 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.6 | 1.4 | 904 | 576 | 576 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1329 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.086 | 30.4248 | 28 | 1.4
2008-07-16 | 13:09 | 7.37 | 0.072 | 720 | 902.3 | 1.4 | 903.7 | 576.2 | 576.1797 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.309 | 0.1329 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9102 | 30.2053 | 28 | 1.4
2008-07-16 | 13:09 | 7.369 | 0.072 | 720 | 902.4 | 1.4 | 903.8 | 576 | 575.9519 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1329 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0806 | 30.4355 | 28 | 1.4
2008-07-16 | 13:09 | 7.36 | 0.072 | 720 | 902.1 | 1.4 | 903.5 | 576.1 | 576.147 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.13281 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3896 | 28 | 1.4
2008-07-16 | 13:09 | 7.343 | 0.072 | 720 | 902 | 1.4 | 903.4 | 576.3 | 576.3 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3916 | 0.1328 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2681 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.2 | 1.4 | 903.6 | 576.1 | 576.1488 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.1328 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0862 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.4 | 1.4 | 903.8 | 576.1 | 576.0509 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1328 | 0.006 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.3 | 1.4 | 903.7 | 575.8 | 575.8349 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1328 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.4 | 1.4 | 903.8 | 575.9 | 575.9 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.13271 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3164 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.4 | 1.4 | 903.8 | 575.9 | 575.9 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1327 | 0.0059 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.5 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.7 | 1.4 | 904.1 | 576.1 | 576.0506 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1327 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.5 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.5 | 1.4 | 903.9 | 576.2 | 576.2058 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3923 | 0.1327 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2932 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.3 | 1.4 | 903.7 | 576.1 | 576.1481 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1327 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.1544 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.2 | 1.4 | 903.6 | 576.4 | 576.4 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1326 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9192 | 30.1223 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.2 | 1.4 | 903.6 | 576.2 | 576.1693 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3082 | 0.1326 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.199 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 901.9 | 1.4 | 903.3 | 576.1 | 576.1 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1326 | 0.0058 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.1 | 1.4 | 903.5 | 576.2 | 576.1749 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1326 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013.0863 | 30.2 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.1 | 1.4 | 903.5 | 576.1 | 576.1253 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.1326 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2292 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.1 | 1.4 | 903.5 | 576.3 | 576.3031 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1325 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2941 | 28 | 1.4
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.1 | 1.5 | 903.6 | 576.3 | 576.3 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1325 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2175 | 28 | 1.5
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.1 | 1.5 | 903.6 | 576.1 | 576.1 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1325 | 0.0057 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2 | 28 | 1.5
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902.3 | 1.5 | 903.8 | 576.1 | 576.1 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2356 | 28 | 1.5
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 902 | 1.5 | 903.5 | 576.1 | 576.1 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.34 | 0.072 | 720 | 901.9 | 1.5 | 903.4 | 576.6 | 576.5834 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3081 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.345 | 0.072 | 720 | 902 | 1.5 | 903.5 | 576.4 | 576.4 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3918 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2341 | 28 | 1.5
2008-07-16 | 13:09 | 7.35 | 0.072 | 720 | 902.1 | 1.5 | 903.6 | 576.4 | 576.4012 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3085 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.0845 | 30.2427 | 28 | 1.5
2008-07-16 | 13:09 | 7.35 | 0.072 | 720 | 901.9 | 1.5 | 903.4 | 576.3 | 576.3427 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1325 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9138 | 30.3389 | 28 | 1.5
2008-07-16 | 13:09 | 7.35 | 0.072 | 720 | 901.9 | 1.5 | 903.4 | 576.6 | 576.5507 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.13241 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.912 | 30.4925 | 28 | 1.5
2008-07-16 | 13:09 | 7.35 | 0.072 | 720 | 901.9 | 1.5 | 903.4 | 576.6 | 576.6027 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1324 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9129 | 30.2812 | 28 | 1.5
2008-07-16 | 13:09 | 7.35 | 0.072 | 720 | 901.9 | 1.5 | 903.4 | 576.3 | 576.3308 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.1324 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9127 | 30.2993 | 28 | 1.5
2008-07-16 | 13:09 | 7.355 | 0.072 | 720 | 901.7 | 1.5 | 903.2 | 576.2 | 576.1766 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3927 | 0.1324 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.4264 | 28 | 1.5
2008-07-16 | 13:09 | 7.339 | 0.072 | 720 | 901.6 | 1.5 | 903.1 | 576.2 | 576.1584 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1324 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.4939 | 28 | 1.5
2008-07-16 | 13:09 | 7.29 | 0.0725 | 724.598 | 901.6 | 1.5 | 903.1 | 576.3 | 576.3498 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1324 | 0.0056 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3404 | 28 | 1.5
2008-07-16 | 13:09 | 7.29 | 0.0727 | 727.0613 | 901.7 | 1.5 | 903.2 | 576.5 | 576.5262 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1324 | 0.005508 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.29 | 0.072 | 720 | 901.5 | 1.5 | 903 | 576.6 | 576.6 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1324 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.29 | 0.072 | 720 | 901.3 | 1.5 | 902.8 | 576.4 | 576.4 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3917 | 0.1324 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.3198 | 28 | 1.5
2008-07-16 | 13:09 | 7.299 | 0.072 | 720 | 901.2 | 1.5 | 902.7 | 576.4 | 576.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1324 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.4733 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901.2 | 1.5 | 902.7 | 576 | 575.9646 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3086 | 0.1324 | 0.0055 | 0 | -9e-06 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.4367 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901 | 1.5 | 902.5 | 576.1 | 576.1036 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.1324 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.36 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901.1 | 1.5 | 902.6 | 576 | 576 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1324 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901.3 | 1.5 | 902.8 | 576.1 | 576.1462 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.0875 | 30.3 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901.1 | 1.5 | 902.6 | 576.1 | 576.0981 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.2299 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.072 | 720 | 901.3 | 1.5 | 902.8 | 576.2 | 576.1567 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.0878 | 30.2938 | 28 | 1.5
2008-07-16 | 13:09 | 7.31 | 0.0707 | 707.4548 | 901.6 | 1.5 | 903.1 | 576.1 | 576.1 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3764 | 28 | 1.5
2008-07-16 | 13:10 | 7.31 | 0.07 | 700 | 901.6 | 1.5 | 903.1 | 576 | 576.0041 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3003 | 28 | 1.5
2008-07-16 | 13:10 | 7.31 | 0.07 | 700 | 901.6 | 1.5 | 903.1 | 576.4 | 576.4147 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3918 | 0.1323 | 0.0055 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.377 | 28 | 1.5
2008-07-16 | 13:10 | 7.318 | 0.0701 | 701.2245 | 901.7 | 1.5 | 903.2 | 576.2 | 576.1988 | 0.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1323 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1011 | 30.3463 | 28 | 1.5
2008-07-16 | 13:10 | 7.32 | 0.071 | 709.7279 | 901.4 | 1.5 | 902.9 | 576.3 | 576.2515 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1323 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3609 | 28 | 1.5
2008-07-16 | 13:10 | 7.32 | 0.071 | 710 | 901.5 | 1.5 | 903 | 576.1 | 576.0843 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.1323 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.5 | 28 | 1.5
2008-07-16 | 13:10 | 7.004 | 0.0665 | 665.1591 | 901.2 | 1.5 | 902.7 | 576.2 | 576.2022 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.1323 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.5 | 28 | 1.5
2008-07-16 | 13:10 | 6.235 | 0.0616 | 616.2836 | 899.3 | 1.5 | 900.8 | 564.1 | 564.1469 | 0.1 | 13 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.13212 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.9127 | 30.5 | 28 | 1.5
2008-07-16 | 13:10 | 6.201 | 0.059 | 590 | 831 | 1.5 | 832.5 | 508.3 | 508.3446 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.2 | 28 | 5 | 195 | 59.3087 | 0.12521 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.5 | 28 | 1.5
2008-07-16 | 13:10 | 6.129 | 0.0592 | 591.5184 | 760.3 | 1.5 | 761.8 | 485.8 | 485.7757 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.11774 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.4721 | 28 | 1.5
2008-07-16 | 13:10 | 6.11 | 0.06 | 599.6816 | 753.4 | 1.4077 | 754.8 | 481.3 | 481.3267 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.11701 | 0.0055 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 30.3187 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 752 | 1.4 | 753.4 | 480.6 | 480.6089 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1168 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 751.9 | 1.4 | 753.3 | 480.2 | 480.1923 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.11671 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 752 | 1.4 | 753.4 | 480 | 479.9816 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1167 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2793 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 752 | 1.4 | 753.4 | 479.8 | 479.7746 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1167 | 0.0054 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2029 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 751.7 | 1.4 | 753.1 | 479.5 | 479.4964 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1167 | 0.005309 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 30.2741 | 28 | 1.4
2008-07-16 | 13:10 | 6.116 | 0.06 | 600 | 751.5 | 1.4 | 752.9 | 479.3 | 479.3433 | 0.1 | 13 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.11661 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9129 | 30.3508 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 751.5 | 1.4 | 752.9 | 479.7 | 479.7 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.1166 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3451 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 751.4 | 1.4 | 752.8 | 479.4 | 479.4 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1166 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9124 | 30.2 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 751.3 | 1.4 | 752.7 | 479.6 | 479.6315 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1166 | 0.0053 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 751 | 1.4 | 752.4 | 479.6 | 479.6018 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.1165 | 0.005209 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0878 | 30.3732 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.5 | 1.4 | 751.9 | 479.6 | 479.5944 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1164 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4656 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.3 | 1.4 | 751.7 | 479.6 | 479.5705 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.1164 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3778 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.2 | 1.4 | 751.6 | 479.9 | 479.9059 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.1164 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2244 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.3 | 1.4 | 751.7 | 479.9 | 479.9058 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.3915 | 0.1164 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2645 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.2 | 1.4 | 751.6 | 479.8 | 479.8037 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1164 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3825 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750.1 | 1.4 | 751.5 | 480.1 | 480.0562 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1163 | 0.0052 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.912 | 30.482 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 750 | 1.4 | 751.4 | 480 | 479.9891 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.1163 | 0.005109 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4053 | 28 | 1.4
2008-07-16 | 13:10 | 6.12 | 0.06 | 600 | 749.9 | 1.4 | 751.3 | 479.7 | 479.6519 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1163 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0873 | 30.2572 | 28 | 1.4
2008-07-16 | 13:10 | 6.111 | 0.06 | 600 | 749.9 | 1.4 | 751.3 | 479.6 | 479.5504 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.1163 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2481 | 28 | 1.4
2008-07-16 | 13:10 | 6.1 | 0.06 | 600 | 750 | 1.5 | 751.5 | 479.6 | 479.5643 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.1163 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.5
2008-07-16 | 13:10 | 6.1 | 0.06 | 600 | 750 | 1.5 | 751.5 | 479.4 | 479.3996 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.1163 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.5
2008-07-16 | 13:10 | 6.1 | 0.06 | 600 | 749.9 | 1.5 | 751.4 | 479.5 | 479.462 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3921 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.5
2008-07-16 | 13:10 | 6.108 | 0.06 | 600 | 749.9 | 1.5 | 751.4 | 479.7 | 479.652 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3919 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4098 | 28 | 1.5
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.9 | 1.5 | 751.4 | 479.5 | 479.5029 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3736 | 28 | 1.5
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.8 | 1.4084 | 751.2 | 479.3 | 479.3 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1332 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 479.4 | 479.397 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.3 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4401 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.7 | 1.4 | 751.1 | 479.6 | 479.6071 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1162 | 0.0051 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3765 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.8 | 1.4 | 751.2 | 479.5 | 479.5 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0871 | 30.3768 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.9 | 1.4 | 751.3 | 479.9 | 479.9075 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4849 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.8 | 1.4 | 751.2 | 480 | 480.0295 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4081 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.6 | 1.4 | 751 | 479.9 | 479.8995 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.391 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2629 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 480 | 479.9503 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.9212 | 30.3357 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 479.7 | 479.6615 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.1162 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.4342 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 479.9 | 479.8786 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.11611 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2041 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.6 | 1.4 | 751 | 479.6 | 479.5528 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1161 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0843 | 30.2 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.8 | 1.4 | 751.2 | 479.7 | 479.6931 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.1161 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3551 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750 | 1.4 | 751.4 | 479.5 | 479.5 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.1161 | 0.005 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750 | 1.4 | 751.4 | 479.8 | 479.751 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.1161 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4895 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750 | 1.4 | 751.4 | 479.5 | 479.4545 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1161 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3367 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.2 | 1.4 | 751.6 | 479.6 | 479.5992 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.1161 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4171 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.3 | 1.4 | 751.7 | 479.3 | 479.3024 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1161 | 0.0049 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.3938 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.5 | 1.4 | 751.9 | 479.3 | 479.3253 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.1161 | 0.0049 | -9.1e-05 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.5 | 1.4 | 751.9 | 479.3 | 479.3477 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.4 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.6 | 1.4 | 752 | 479.3 | 479.2567 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3919 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2657 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.6 | 1.4 | 752 | 479.3 | 479.3457 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.4 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750.3 | 1.4 | 751.7 | 479.3 | 479.35 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3086 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.9137 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 750 | 1.4 | 751.4 | 479.4 | 479.4 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.913 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.8 | 1.4 | 751.2 | 479.2 | 479.2 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.6 | 1.4 | 751 | 479.4 | 479.4491 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.9127 | 30.2508 | 28 | 1.4
2008-07-16 | 13:10 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 479.4 | 479.428 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3912 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2 | 28 | 1.4
2008-07-16 | 13:11 | 6.11 | 0.06 | 600 | 749.5 | 1.4 | 750.9 | 479.3 | 479.3 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2027 | 28 | 1.4
2008-07-16 | 13:11 | 6.11 | 0.06 | 600 | 749.7 | 1.4 | 751.1 | 479.2 | 479.2424 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3912 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2794 | 28 | 1.4
2008-07-16 | 13:11 | 6.11 | 0.06 | 600 | 749.8 | 1.4 | 751.2 | 479.6 | 479.5992 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3079 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2439 | 28 | 1.4
2008-07-16 | 13:11 | 6.11 | 0.06 | 600 | 750.1 | 1.4 | 751.5 | 479.4 | 479.4027 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2982 | 28 | 1.4
2008-07-16 | 13:11 | 6.11 | 0.06 | 600 | 750.2 | 1.4 | 751.6 | 479.1 | 479.1 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.116 | 0.0048 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.4
2008-07-16 | 13:11 | 5.815 | 0.0557 | 556.5852 | 750.3 | 1.4 | 751.7 | 479.3 | 479.3034 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.11591 | 0.0047 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.4
2008-07-16 | 13:11 | 5.09 | 0.0507 | 507.3537 | 750.3 | 1.4 | 751.7 | 476.5 | 476.4872 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.1159 | 0.0047 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3116 | 28 | 1.4
2008-07-16 | 13:11 | 4.906 | 0.048 | 480 | 685.4 | 1.4 | 686.8 | 420.4 | 420.4329 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3087 | 0.11225 | 0.0047 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2395 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 616.1 | 1.4 | 617.5 | 391.1 | 391.0864 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3913 | 0.10052 | 0.0047 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2837 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 607.3 | 1.4 | 608.7 | 386.1 | 386.1274 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.099417 | 0.0046 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2071 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 606 | 1.4 | 607.4 | 385.2 | 385.1718 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3087 | 0.099209 | 0.0046 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.0875 | 30.2696 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 604.8 | 1.4 | 606.2 | 384.9 | 384.8701 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3088 | 0.099109 | 0.004509 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3463 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 603.8 | 1.4 | 605.2 | 384.5 | 384.5326 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.099 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.377 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 602.6 | 1.4 | 604 | 384 | 384.049 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3081 | 0.098808 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3002 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 602.5 | 1.4 | 603.9 | 384.4 | 384.4496 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.098708 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.4
2008-07-16 | 13:11 | 4.9 | 0.048 | 480 | 602.5 | 1.4 | 603.9 | 383.9 | 383.9046 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3916 | 0.0987 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3532 | 28 | 1.4
2008-07-16 | 13:11 | 4.895 | 0.048 | 480 | 602.4 | 1.4 | 603.8 | 384.2 | 384.1577 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3086 | 0.0987 | 0.0045 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3701 | 28 | 1.4
2008-07-16 | 13:11 | 4.89 | 0.048 | 480 | 602.5 | 1.4 | 603.9 | 384 | 384 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0987 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.4
2008-07-16 | 13:11 | 4.89 | 0.048 | 480 | 602.5 | 1.4 | 603.9 | 384 | 384 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0987 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.4
2008-07-16 | 13:11 | 4.89 | 0.048 | 480 | 602.5 | 1.4 | 603.9 | 383.7 | 383.6984 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0987 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4199 | 28 | 1.4
2008-07-16 | 13:11 | 4.89 | 0.048 | 480 | 602.4 | 1.4 | 603.8 | 383.5 | 383.548 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.098609 | 0.0044 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.4
2008-07-16 | 13:11 | 4.89 | 0.048 | 480 | 602.4 | 1.4 | 603.8 | 384 | 384.0423 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3912 | 0.0986 | 0.004309 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9124 | 30.4619 | 28 | 1.4
2008-07-16 | 13:11 | 4.886 | 0.048 | 480 | 602.4 | 1.4 | 603.8 | 384 | 384.0064 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0986 | 0.0043 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2307 | 28 | 1.4
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.4 | 1.3 | 603.7 | 383.7 | 383.7134 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.4 | 0.098508 | 0.004208 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.4 | 1.3 | 603.7 | 383.7 | 383.7023 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0985 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.243 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.2 | 1.3 | 603.5 | 384.2 | 384.2008 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.2 | 28 | 5 | 195 | 59.3 | 0.0985 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2801 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.2 | 1.3 | 603.5 | 384 | 383.9816 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.8 | 28 | 5 | 195 | 59.3084 | 0.0985 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2034 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.1 | 1.3 | 603.4 | 383.8 | 383.8 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.0985 | 0.0042 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4208 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.1 | 1.3 | 603.4 | 384 | 383.9728 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0984 | 0.004109 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.9 | 1.3 | 603.2 | 383.9 | 383.936 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0984 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.446 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602 | 1.3 | 603.3 | 384 | 383.9521 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0984 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2963 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602 | 1.3 | 603.3 | 384 | 384.0478 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0984 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2196 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.9 | 1.3 | 603.2 | 384.1 | 384.1184 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.098309 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.3712 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.9 | 1.3 | 603.2 | 383.7 | 383.7004 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0983 | 0.004009 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0878 | 30.4663 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602 | 1.3 | 603.3 | 383.9 | 383.948 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.0983 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4104 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602 | 1.3 | 603.3 | 383.9 | 383.9 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0983 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4871 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.1 | 1.3 | 603.4 | 384.1 | 384.1459 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0983 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.3 | 1.3 | 603.6 | 384.1 | 384.1 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0983 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4594 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.4 | 1.3 | 603.7 | 383.7 | 383.7372 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.098209 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.4173 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.4 | 1.3 | 603.7 | 383.8 | 383.8055 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3085 | 0.0982 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.494 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.4 | 1.3 | 603.7 | 384.4 | 384.396 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0982 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2879 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.2 | 1.3 | 603.5 | 384.2 | 384.2 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0982 | 0.003909 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.2948 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 602.1 | 1.3 | 603.4 | 384.2 | 384.1988 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0982 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.4241 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.7 | 1.3 | 603 | 384.1 | 384.0819 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0982 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.2 | 1.3 | 602.5 | 384.3 | 384.275 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.098108 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 601.1 | 1.3 | 602.4 | 384.3 | 384.2994 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0981 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.9 | 1.3 | 602.2 | 384.5 | 384.4986 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0981 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4383 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.6 | 1.3 | 601.9 | 384.5 | 384.4583 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.0981 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.5 | 1.3 | 601.8 | 384.3 | 384.2519 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0981 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.4 | 1.3 | 601.7 | 384.5 | 384.4527 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.088 | 30.3609 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.4 | 1.3 | 601.7 | 384.3 | 384.3495 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4752 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.1 | 1.3 | 601.4 | 384.7 | 384.6876 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.5857 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.2 | 1.3 | 601.5 | 384.6 | 384.6 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.3 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.509 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.3 | 1.3 | 601.6 | 384.5 | 384.4515 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.4 | 1.3 | 601.7 | 384.5 | 384.5413 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.5 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.3 | 1.3 | 601.6 | 384.4 | 384.3954 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4374 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.3 | 1.3 | 601.6 | 384.8 | 384.8 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2083 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.048 | 480 | 600.3 | 1.3 | 601.6 | 384.4 | 384.4399 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3485 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.0488 | 487.732 | 600.2 | 1.3 | 601.5 | 384.6 | 384.6083 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4508 | 28 | 1.3
2008-07-16 | 13:11 | 4.88 | 0.049 | 490 | 600.1 | 1.3 | 601.4 | 384.5 | 384.5 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.5274 | 28 | 1.3
2008-07-16 | 13:12 | 4.88 | 0.049 | 490 | 600.2 | 1.3 | 601.5 | 384.8 | 384.7511 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.5911 | 28 | 1.3
2008-07-16 | 13:12 | 4.88 | 0.049 | 490 | 600.5 | 1.3 | 601.8 | 384.5 | 384.5492 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.098 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4376 | 28 | 1.3
2008-07-16 | 13:12 | 4.88 | 0.049 | 490 | 600.6 | 1.3 | 601.9 | 384.7 | 384.6761 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.097909 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2841 | 28 | 1.3
2008-07-16 | 13:12 | 4.88 | 0.049 | 490 | 600.7 | 1.3 | 602 | 384.3 | 384.325 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3911 | 0.0979 | 0.003809 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.3
2008-07-16 | 13:12 | 4.88 | 0.049 | 490 | 600.8 | 1.3 | 602.1 | 384.4 | 384.4481 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.0979 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.3
2008-07-16 | 13:12 | 4.558 | 0.044 | 440.4824 | 600.8 | 1.3 | 602.1 | 384.5 | 384.4793 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0979 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.3
2008-07-16 | 13:12 | 3.738 | 0.0388 | 388.0788 | 600.7 | 1.3 | 602 | 371.2 | 371.2335 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0979 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3942 | 28 | 1.3
2008-07-16 | 13:12 | 3.634 | 0.0363 | 362.629 | 562.4 | 1.3 | 563.7 | 319.7 | 319.7036 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3082 | 0.093271 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.9175 | 30.4172 | 28 | 1.3
2008-07-16 | 13:12 | 3.63 | 0.037 | 370 | 463.2 | 1.3 | 464.5 | 296.4 | 296.3895 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.080312 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0843 | 30.3362 | 28 | 1.3
2008-07-16 | 13:12 | 3.63 | 0.037 | 370 | 456.1 | 1.3 | 457.4 | 290.8 | 290.8269 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.079317 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4896 | 28 | 1.3
2008-07-16 | 13:12 | 3.63 | 0.037 | 370 | 453.5 | 1.3 | 454.8 | 289.7 | 289.6855 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.3087 | 0.078909 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3571 | 28 | 1.3
2008-07-16 | 13:12 | 3.63 | 0.0364 | 363.8833 | 452.8 | 1.3 | 454.1 | 289.3 | 289.3 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.078809 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2518 | 28 | 1.3
2008-07-16 | 13:12 | 3.629 | 0.036 | 360 | 452.4 | 1.3 | 453.7 | 289 | 288.9507 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0787 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2497 | 28 | 1.3
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 452.3 | 1.3 | 453.6 | 289.1 | 289.0879 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0787 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4031 | 28 | 1.3
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 452.1 | 1.3 | 453.4 | 288.6 | 288.6 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3912 | 0.078609 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9122 | 30.5564 | 28 | 1.3
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 452 | 1.3 | 453.3 | 288.7 | 288.6959 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0786 | 0.003508 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0788 | 30.3802 | 28 | 1.3
2008-07-16 | 13:12 | 3.624 | 0.036 | 360 | 451.3 | 1.2 | 452.5 | 288.5 | 288.5408 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3081 | 0.078508 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2633 | 28 | 1.2
2008-07-16 | 13:12 | 3.627 | 0.036 | 360 | 451.2 | 1.2 | 452.4 | 288.4 | 288.4119 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3917 | 0.078408 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 451.2 | 1.2 | 452.4 | 288.8 | 288.8259 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0784 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.7 | 1.2 | 451.9 | 288.7 | 288.7351 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3914 | 0.078309 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2767 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.6 | 1.2 | 451.8 | 288.8 | 288.8499 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.8 | 5 | 195 | 59.3 | 0.0783 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.315 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.4 | 1.2 | 451.6 | 288.5 | 288.515 | 0.1 | 13 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.2 | 5 | 195 | 59.3087 | 0.0783 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4549 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.4 | 1.2 | 451.6 | 288.6 | 288.5994 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.8 | 5 | 195 | 59.3087 | 0.0782 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2248 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.4 | 1.2 | 451.6 | 288.7 | 288.7 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 29 | 5 | 195 | 59.3088 | 0.0782 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4052 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.3 | 1.2 | 451.5 | 288.7 | 288.6847 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.2 | 5 | 195 | 59.4 | 0.0782 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9124 | 30.4098 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.4 | 1.2 | 451.6 | 288.8 | 288.7514 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3088 | 0.0782 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.2 | 1.2 | 451.4 | 288.8 | 288.7976 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.3 | 0.078108 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.6 | 288.5716 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.0781 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2248 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.4 | 288.3519 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0781 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3553 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.7 | 1.2 | 450.9 | 288.8 | 288.8 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.078 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0847 | 30.4154 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.6 | 288.5725 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.078 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2139 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.6 | 288.5765 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.078 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.445 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.4 | 288.4105 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3087 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3834 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.4 | 288.4073 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3913 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.615 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.3 | 288.3463 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0857 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.7 | 288.6842 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.6 | 288.6491 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2347 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.7 | 288.6975 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.078 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.4 | 288.4407 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2187 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 450 | 1.2 | 451.2 | 288.6 | 288.5506 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.003108 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2954 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.6 | 288.5992 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3914 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2279 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.5 | 288.5336 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3465 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.6 | 288.6019 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.4489 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.4 | 288.4423 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.5 | 288.4992 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.6 | 288.6 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.0779 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2445 | 28 | 1.2
2008-07-16 | 13:12 | 3.61 | 0.036 | 360 | 449.7 | 1.2 | 450.9 | 288.6 | 288.5751 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.003008 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2322 | 28 | 1.2
2008-07-16 | 13:12 | 3.613 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.4 | 288.448 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0779 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3089 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.2 | 288.249 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3856 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.8 | 1.2 | 451 | 288.1 | 288.1 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3377 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.3 | 288.2517 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0779 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.339 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.9 | 1.1 | 451 | 288.2 | 288.2273 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0862 | 30.3687 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450 | 1.1 | 451.1 | 288 | 287.9512 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2152 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450 | 1.1 | 451.1 | 288.5 | 288.5491 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2691 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450 | 1.1 | 451.1 | 288.2 | 288.2 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.2 | 288.1769 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 449.9 | 1.2 | 451.1 | 288.3 | 288.251 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.0778 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.3 | 28 | 1.2
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450 | 1.1079 | 451.1 | 288.4 | 288.3741 | 0.1 | 13 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3079 | 0.0778 | 0.002908 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1485 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.2 | 1.1 | 451.3 | 288.2 | 288.2 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3081 | 0.0778 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1525 | 28 | 1.1
2008-07-16 | 13:12 | 3.62 | 0.036 | 360 | 450.2 | 1.1 | 451.3 | 288.2 | 288.1986 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2872 | 28 | 1.1
2008-07-16 | 13:13 | 3.621 | 0.036 | 360 | 450.2 | 1.1 | 451.3 | 287.9 | 287.8587 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4835 | 28 | 1.1
2008-07-16 | 13:13 | 3.629 | 0.036 | 360 | 450.3 | 1.1 | 451.4 | 288.2 | 288.1759 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0778 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2543 | 28 | 1.1
2008-07-16 | 13:13 | 3.63 | 0.036 | 360 | 450.3 | 1.1 | 451.4 | 288.4 | 288.4 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0777 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 3.63 | 0.036 | 360 | 450.3 | 1.1 | 451.4 | 288.1 | 288.1319 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0777 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2355 | 28 | 1.1
2008-07-16 | 13:13 | 3.63 | 0.036 | 360 | 450.3 | 1.1 | 451.4 | 288.3 | 288.3 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.0777 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.1
2008-07-16 | 13:13 | 3.441 | 0.036 | 360 | 450.2 | 1.1 | 451.3 | 288 | 288.017 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0777 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3 | 28 | 1.1
2008-07-16 | 13:13 | 2.904 | 0.0302 | 301.6148 | 450.2 | 1.1 | 451.3 | 288.3 | 288.3483 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0777 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2341 | 28 | 1.1
2008-07-16 | 13:13 | 2.456 | 0.0233 | 232.7733 | 417 | 1.1 | 418.1 | 259.4 | 259.3842 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.073139 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2853 | 28 | 1.1
2008-07-16 | 13:13 | 2.422 | 0.024 | 240 | 318.8 | 1.1 | 319.9 | 206.5 | 206.4771 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.058864 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4193 | 28 | 1.1
2008-07-16 | 13:13 | 2.396 | 0.024 | 240 | 304.4 | 1.1 | 305.5 | 195.7 | 195.7042 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.05664 | 0.002808 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4959 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.0248 | 247.8212 | 302.6 | 1.1 | 303.7 | 193.7 | 193.6931 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.056308 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2821 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.025 | 250 | 301.5 | 1.1 | 302.6 | 193.2 | 193.1838 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0562 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1507 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.0246 | 245.6702 | 300.9 | 1.1 | 302 | 192.8 | 192.8465 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.056 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.178 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.8 | 192.8467 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.056 | 0.002709 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.3946 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.7 | 192.7 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.056 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2413 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.7 | 192.7481 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.056 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.144 | 28.056 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.5 | 192.5413 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.055909 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1327 | 28.0673 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.5 | 192.4538 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.2 | 28 | 5 | 195 | 59.3912 | 0.0559 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.8 | 1.1 | 301.9 | 192.5 | 192.5043 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0559 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.0243 | 242.7769 | 300.8 | 1.1 | 301.9 | 192.4 | 192.3547 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.055808 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.0806 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.0249 | 249.166 | 300.8 | 1.1 | 301.9 | 192.4 | 192.4 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.0558 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.9175 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.0241 | 240.8257 | 300.7 | 1.1 | 301.8 | 192.4 | 192.4 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0558 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2483 | 28 | 1.1
2008-07-16 | 13:13 | 2.38 | 0.024 | 240 | 300.3 | 1.1 | 301.4 | 192.5 | 192.5 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3914 | 0.0558 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.4784 | 28 | 1.1
2008-07-16 | 13:13 | 2.377 | 0.024 | 240 | 300.3 | 1.1 | 301.4 | 192.3 | 192.3253 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.361 | 28 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.3 | 1.1 | 301.4 | 192.3 | 192.3491 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2538 | 28 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 1.1 | 301.3 | 192.2 | 192.2355 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.9127 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 1.1 | 301.3 | 192.2 | 192.2 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1049 | 1013 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 1.1 | 301.3 | 192 | 192 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2 | 28 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.1 | 1.1 | 301.2 | 192 | 192.0491 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1049 | 1013 | 30.2 | 28.0528 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.1 | 1.1 | 301.2 | 191.9 | 191.9488 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0557 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28.0705 | 1.1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 1 | 301.2 | 192.1 | 192.1 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0556 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.3 | 1 | 301.3 | 192.1 | 192.0507 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.2 | 28 | 5 | 195 | 59.4 | 0.0556 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.3 | 1 | 301.3 | 192 | 192 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3917 | 0.0556 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.4 | 1 | 301.4 | 192.3 | 192.2551 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0556 | 0.002309 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.164 | 28 | 1
2008-07-16 | 13:13 | 2.366 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.1 | 192.1476 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.0556 | 0.0023 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.1129 | 28 | 1
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.4 | 1 | 301.4 | 192.3 | 192.2981 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.0555 | 0.0023 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.9127 | 30.1893 | 28 | 1
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.2 | 192.167 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3913 | 0.0555 | 0.0023 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1049 | 1011 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.2 | 192.1588 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0555 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192 | 192.0369 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3923 | 0.0555 | 0.0023 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011.9234 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.1 | 192.1253 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0555 | 0.002208 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.364 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.1 | 192.148 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3919 | 0.0555 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.273 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.3 | 192.3 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3917 | 0.0555 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2504 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.5 | 1 | 301.5 | 192.3 | 192.3467 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.0555 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28 | 1
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.5 | 0.90862 | 301.4 | 192.2 | 192.2 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.0555 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1969 | 28.0031 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.6 | 0.9 | 301.5 | 192.3 | 192.2553 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.1202 | 28.0798 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.6 | 0.9 | 301.5 | 192.1 | 192.0982 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.2 | 28 | 5 | 195 | 59.3913 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.1564 | 28.0436 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.6 | 0.9 | 301.5 | 192.2 | 192.2 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3085 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.367 | 0.024 | 240 | 300.5 | 0.9 | 301.4 | 192.1 | 192.1349 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.1903 | 28 | 0.9
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.4 | 0.9 | 301.3 | 192.3 | 192.301 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1136 | 28 | 0.9
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.3 | 0.9 | 301.2 | 192 | 191.9809 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3923 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1631 | 28 | 0.9
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.2 | 192.2093 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.0554 | 0.0022 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.9212 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.3 | 192.3 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0554 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1049 | 1013 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.36 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.2 | 192.2 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3918 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1049 | 1013 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.362 | 0.024 | 240 | 300 | 0.9 | 300.9 | 192.1 | 192.1 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1049 | 1013 | 30.2 | 28.0697 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.1 | 192.103 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.0536 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.2 | 192.1867 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.8 | 5 | 195 | 59.3913 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192 | 192 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.2 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.0855 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.2 | 192.2286 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.2 | 192.1686 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1049 | 1012 | 30.2 | 28 | 0.9
2008-07-16 | 13:13 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.3 | 192.2502 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011.9122 | 30.2 | 28 | 0.9
2008-07-16 | 13:14 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.2 | 192.2364 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3079 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1049 | 1011 | 30.1936 | 28.0064 | 0.9
2008-07-16 | 13:14 | 2.37 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.1 | 192.1032 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.9194 | 30.1169 | 28.0831 | 0.9
2008-07-16 | 13:14 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.1 | 192.1 | 0.1 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1595 | 28.0405 | 0.9
2008-07-16 | 13:14 | 2.37 | 0.024 | 240 | 300.2 | 0.9 | 301.1 | 192.1 | 192.1147 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.0361 | 0.9
2008-07-16 | 13:14 | 2.37 | 0.024 | 240 | 300.1 | 0.9 | 301 | 192.3 | 192.3 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.0553 | 0.0021 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.213 | 28.087 | 0.9
2008-07-16 | 13:14 | 2.126 | 0.021 | 210.0516 | 300.1 | 0.9 | 301 | 192.3 | 192.2535 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0553 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.0884 | 30.2894 | 28.0106 | 0.9
2008-07-16 | 13:14 | 1.365 | 0.016 | 159.9034 | 299.7 | 0.9 | 300.6 | 190.4 | 190.4306 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.055109 | 0.002 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1049 | 1012 | 30.3 | 28 | 0.9
2008-07-16 | 13:14 | 1.199 | 0.0122 | 122.4232 | 251.8 | 0.9 | 252.7 | 136.4 | 136.3673 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.047571 | 0.002 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.3 | 28 | 0.9
2008-07-16 | 13:14 | 1.18 | 0.0129 | 129.275 | 165 | 0.9 | 165.9 | 104.7 | 104.6736 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.033096 | 0.002 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2606 | 28 | 0.9
2008-07-16 | 13:14 | 1.174 | 0.0121 | 120.9417 | 155 | 0.9 | 155.9 | 99 | 98.9656 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.031353 | 0.002 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1049 | 1012 | 30.1069 | 28 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 153.1 | 0.9 | 154 | 97.6 | 97.6039 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3921 | 0.030916 | 0.001908 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2464 | 28 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152.8 | 0.9 | 153.7 | 97.1 | 97.0852 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.030808 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2502 | 28.0498 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152.7 | 0.9 | 153.6 | 96.8 | 96.8 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0308 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.1 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152.5 | 0.9 | 153.4 | 96.9 | 96.8756 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3916 | 0.030708 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.0968 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 151.9 | 0.9 | 152.8 | 96.6 | 96.6029 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.030609 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.2 | 28.0202 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152 | 0.9 | 152.9 | 96.7 | 96.6959 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0306 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.2 | 28 | 0.9
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 151.9 | 0.80882 | 152.7 | 96.5 | 96.5499 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0306 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.9118 | 30.2332 | 28 | 0.8
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.6 | 96.6 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.030509 | 0.001709 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.3 | 28 | 0.8
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.4 | 96.4493 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0305 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.3 | 28 | 0.8
2008-07-16 | 13:14 | 1.17 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.5 | 96.5055 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3088 | 0.0305 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.2368 | 28.0632 | 0.8
2008-07-16 | 13:14 | 1.163 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.4 | 96.435 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.0305 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.2399 | 28.1 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.6 | 96.5534 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.030408 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.2669 | 28.1 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 152 | 0.8 | 152.8 | 96.4 | 96.4364 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0304 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1048 | 1011 | 30.1135 | 28.1 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.8 | 152.7 | 96.4 | 96.3518 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3917 | 0.0304 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32.746 | 55 | 1048 | 1011.9173 | 30.1699 | 28.0301 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.8 | 152.7 | 96.3 | 96.3484 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3084 | 0.0304 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.8 | 152.7 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0304 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.8 | 152.7 | 96.4 | 96.3979 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0303 | 0.001509 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.8 | 152.7 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.0303 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.9129 | 30.3531 | 28 | 0.8
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.3 | 96.3 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0303 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2936 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.9 | 0.7 | 152.6 | 96.2 | 96.2233 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0303 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.0875 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.3 | 96.3 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3089 | 0.0303 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.2 | 96.2475 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.030209 | 0.001409 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.3 | 96.3018 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2599 | 28.0599 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.2 | 96.2348 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3081 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2634 | 28.0634 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.3 | 96.2507 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3082 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.7 | 0.7 | 152.4 | 96.2 | 96.2494 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3915 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.1 | 96.1067 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.266 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.2 | 96.2465 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.0302 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2572 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.7 | 0.7 | 152.4 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3087 | 0.030109 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.7 | 0.7 | 152.4 | 96.3 | 96.3 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3913 | 0.0301 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.0875 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.8 | 0.7 | 152.5 | 96.2 | 96.1527 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3086 | 0.0301 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 0.7
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.60878 | 152 | 96.3 | 96.2503 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.0878 | 30.2 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.2336 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2531 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.2261 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3919 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.3936 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.2342 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2403 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.1517 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.2485 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3914 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.1 | 96.1 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28.0099 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.2 | 96.2491 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3912 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048.7802 | 1012 | 30.2 | 28.0866 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.1 | 96.0984 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3089 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1368 | 28.0368 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.8 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.0875 | 30.1 | 28 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.1 | 96.1 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.03 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1 | 28.0165 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.3 | 96.2514 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3088 | 0.03 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1 | 28.0932 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.2 | 96.1691 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0302 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.1 | 96.1487 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0013 | -8e-06 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0465 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3916 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1231 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 96.2 | 96.2493 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.03 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1998 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.2 | 0.6 | 151.8 | 96.2 | 96.2492 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151.1 | 0.6 | 151.7 | 96.1 | 96.1154 | 0.1 | 12.9 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:14 | 1.16 | 0.012 | 120 | 151 | 0.6 | 151.6 | 96.3 | 96.3469 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.0702 | 0.6
2008-07-16 | 13:15 | 1.16 | 0.012 | 120 | 150.9 | 0.6 | 151.5 | 96.2 | 96.2493 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3087 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.0065 | 0.6
2008-07-16 | 13:15 | 1.16 | 0.012 | 120 | 150.9 | 0.6 | 151.5 | 96.3 | 96.3166 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3912 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.0831 | 0.6
2008-07-16 | 13:15 | 1.16 | 0.012 | 120 | 150.9 | 0.6 | 151.5 | 96.2 | 96.2 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3923 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 1.16 | 0.012 | 120 | 150.9 | 0.6 | 151.5 | 96.2 | 96.2498 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 1.16 | 0.012 | 120 | 150.8 | 0.6 | 151.4 | 96.2 | 96.2347 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 1.04 | 0.012 | 120 | 150.7 | 0.6 | 151.3 | 96.3 | 96.2972 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0827 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.654 | 0.0063 | 63.4639 | 150.6 | 0.6 | 151.2 | 96.3 | 96.3404 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1014 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.12 | -0.0015 | -14.802 | 142.6 | 0.6 | 143.2 | 73.5 | 73.5028 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.028438 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013.0862 | 30.2 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.028 | 0 | 0 | 44.9 | 0.6 | 45.5 | 16.8 | 16.8251 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3086 | 0.010177 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1803 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.02 | 0 | 0 | 4.4 | 0.6 | 5 | 3.3 | 3.2759 | 0.1 | 12.9 | 951 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.002117 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1036 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.02 | 0 | 0 | 0.8 | 0.6 | 1.4 | 1.1 | 1.1019 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.001417 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048.7348 | 1013 | 30.1731 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.02 | 0 | 0 | 0.2 | 0.6 | 0.8 | 0.3 | 0.299 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.8 | 28 | 5 | 195 | 59.3 | 0.001209 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28.0502 | 0.6
2008-07-16 | 13:15 | 0.02 | 0 | 0 | 0.2 | 0.6 | 0.8 | 0.2 | 0.2 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0012 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.2 | 28.0264 | 0.6
2008-07-16 | 13:15 | 0.02 | 0 | 0 | 0.2 | 0.6 | 0.8 | 0 | 0.0488 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0012 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.1939 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.013 | 0 | 0 | 0.1 | 0.6 | 0.7 | 0.1 | 0.0509 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.392 | 0.0012 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32.254 | 55 | 1048 | 1013 | 30.0405 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.01 | 0 | 0 | 0 | 0.6 | 0.6 | 0 | 0.0343 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3918 | 0.001108 | 0.001108 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.1128 | 28.1 | 0.6
2008-07-16 | 13:15 | 0.01 | -0.0007 | -7.4148 | 0 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.0011 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.2 | 28.1 | 0.5
2008-07-16 | 13:15 | 0.007 | -0.001 | -10 | 0 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0011 | 0.0011 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.1805 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | 0 | 0.5 | 0.5 | -0.1 | 0 | 0.1 | 12.9 | 952 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3913 | 0.0011 | 0.0011 | 0 | -9e-06 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.0273 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | 0 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3087 | 0.001009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1013 | 30.1261 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.1 | 0.5 | 0.5 | -0.1 | 0 | 0.1 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.1 | 0.5 | 0.5 | 0 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.1837 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.1 | 0.5 | 0.5 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.1071 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.1 | 0.5 | 0.5 | 0 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.001 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.1 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.2 | 0.5 | 0.5 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.000909 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.1927 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.2 | 0.5 | 0.5 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3919 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.277 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.2 | 0.5 | 0.5 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3082 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.2003 | 28.1 | 0.5
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.2 | 0.40843 | 0.4 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.8 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.289 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3085 | 0.0009 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.4686 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.000809 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.0854 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.0008 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.1192 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0008 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.2 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.000808 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0008 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0682 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.3 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3916 | 0.0008 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0428 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.6 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0007 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0804 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.7 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3912 | 0.0007 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0038 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.7 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3087 | 0.0007 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0729 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -0.8 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3913 | 0.0007 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1495 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1 | 0.4 | 0.4 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.000609 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1477 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.1 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3088 | 0.0006 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0057 | 28.1 | 0.4
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.30878 | 0.3 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0006 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1589 | 28.1 | 0.3
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.3 | 0.3 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.000508 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1439 | 28.1561 | 0.3
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.3 | 0.3 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 32.746 | 55 | 1049 | 1012 | 30.1328 | 28.1672 | 0.3
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.3 | 0.3 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3083 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1906 | 28.1 | 0.3
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.0843 | 30.1139 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.0013 | -12.5508 | -1.2 | 0.2 | 0.2 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.9138 | 30.1628 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.0019 | -19.3083 | -1.2 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.0011 | -10.975 | -1.2 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.9129 | 30.1678 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.2 | 0.2 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3071 | 0.0005 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0146 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3089 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1694 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.2 | 0.2 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3923 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0461 | 28.1539 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.2 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.0788 | 30.1227 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.3 | 0.2 | 0.2 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3919 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.1993 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.3 | 0.2 | 0.2 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.276 | 28.1 | 0.2
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.4 | 0.1 | 0.1 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.1947 | 28.1 | 0.1
2008-07-16 | 13:15 | 0 | -0.001 | -10 | -1.4 | 0.1 | 0.1 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.1293 | 28.1 | 0.1
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.4 | 0.1 | 0.1 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.194 | 28.1 | 0.1
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.5 | 0.1 | 0.1 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.9129 | 30.1174 | 28.1 | 0.1
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.5 | 0 | 0 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.2 | 5 | 195 | 59.3 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 0
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.8 | 0 | 0 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.0857 | 30.1 | 28.1357 | 0
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.8 | 0 | 0 | -0.2 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.2 | 0
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.8 | 0 | 0 | -0.1 | 0 | 0.1 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28.8 | 5 | 195 | 59.3 | 0.0003 | 0.0006 | 0 | 0 |  | n/a |  | 32.254 | 55 | 1049 | 1013 | 30.1 | 28.2 | 0
2008-07-16 | 13:16 | 0 | -0.001 | -10 | -1.8 | 0 | 0 | -0.2 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 29 | 5 | 195 | 59.3079 | 0.0003 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9683 | 28.1341 | 0
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
No Processing Errors Detected
Test Date | 2008-07-16
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.499
CO2 Span(%) | 6
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 300
NO2 Span(ppm) | 505
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 604.8
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | White Van
License Plate
Engine Displacement | 7.3
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 12:55
Test End Time | 13:16
Test Duration (s) | 1262
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 993
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.832
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 07/16/2008 12:47:45.026 - None Found
0X0000 - 07/16/2008 12:51:28.144 - None Found
0X0000 - 07/16/2008 12:54:35.631 - None Found
Warnings:
0X0000 - 07/16/2008 12:47:45.027 - None Found
0X0000 - 07/16/2008 12:51:28.144 - None Found
0X0000 - 07/16/2008 12:54:35.632 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 119
Brake Specific Dropout Limit Count | 1262
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-07-16
Time | 12:48
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
CO(ppm) | 0 | 0 | 0
CO2(%) | 0 | 0 | 0
NO(ppm) | 1.4438 | 0.63333 | 0.81042
NO2(ppm) | 0.66875 | -1 | 1.6687
THC(ppmC) | -0.56875 | 0.04 | -0.60875
Span
InfoVer | 1
Date | 2008-07-16
Time | 12:51
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 1200 | 12 | 20.9 | 1989 | x^0: 1505; x^1: 505; x^2: 0; x^3: 965.4
Gas | Previous | Current | Difference
CO(ppm) | 1168.75 | 1200 | -31.25
CO2(%) | 11.9569 | 12.0047 | -0.047792
NO(ppm) | 1479.0438 | 1509.0333 | -29.9896
THC(ppmC) | 980.1688 | 968.4267 | 11.7421
Test Information:
SEMTECH_DATA_FILE | 2008-07-16 | 12:47 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## THC Lin
Test Date:
Test Time:
7/16/2008
12:54:36
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
7/16/2008
12:54:36
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
7/16/2008
12:54:36
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
7/16/2008
12:54:36
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
7/16/2008
12:54:36
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
7/16/2008
12:54:36
NDUV Serial No:
208
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
12 | Rm = | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554
| C.F.= | 0.098697 | 0.19768 | 0.29695 | 0.39651 | 0.49636 | 0.5965 | 0.69693 | 0.79766 | 0.89868
| gas = | 1.1844 | 2.3722 | 3.5634 | 4.7581 | 5.9563 | 7.158 | 8.3632 | 9.5719 | 10.7842
Ro for CO2= | 0.8795