## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 26374
NDUV Serial Number: | 210
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 30.64 | Scott gas | ALM018306 | ±1% | 2008-06-27 | 91.92
200 ppm CO | 201.2
6% CO2 | 5.953
300 ppm NO | 302.6
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | NO2 NDUV Serial Number: |  |  | 210 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-05-07 |  | 505 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 14:59
Test software: | 10.09 SP1
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 1.3 | 1.3 |  |  | 0 | 0.1 | 0.1
90 | 105 | 1 | 50.5 | 48.9 | -1.6 | -3.17 |  | 2.09 | 2 | -0.09 | -4.31
150 | 165 | 2 | 101 | 102.5 | 1.5 | 1.49 |  | 4.18 | 4.1 | -0.08 | -1.91
210 | 225 | 3 | 151.5 | 152.4 | 0.9 | 0.59 |  | 6.27 | 6.1 | -0.17 | -2.71
270 | 285 | 4 | 202 | 202.9 | 0.9 | 0.45 |  | 8.36 | 8.2 | -0.16 | -1.91
330 | 345 | 5 | 252.5 | 254.1 | 1.6 | 0.63 |  | 10.45 | 10.3 | -0.15 | -1.44
390 | 405 | 6 | 303 | 306.7 | 3.7 | 1.22 |  | 12.54 | 12.4 | -0.14 | -1.12
450 | 465 | 7 | 353.5 | 356.1 | 2.6 | 0.74 |  | 14.63 | 14.5 | -0.13 | -0.89
510 | 525 | 8 | 404 | 406.3 | 2.3 | 0.57 |  | 16.72 | 16.6 | -0.12 | -0.72
570 | 585 | 9 | 454.5 | 458.5 | 4 | 0.88 |  | 18.81 | 18.7 | -0.11 | -0.58
630 | 645 | 10 | 505 | 507.1 | 2.1 | 0.42 |  | 20.9 | 20.8 | -0.1 | -0.48
690 | 705 | 9 | 454.5 | 459.2 | 4.7 | 1.03 |  | 18.81 | 18.8 | -0.01 | -0.05
750 | 765 | 8 | 404 | 407.5 | 3.5 | 0.87 |  | 16.72 | 16.7 | -0.02 | -0.12
810 | 825 | 7 | 353.5 | 358.6 | 5.1 | 1.44 |  | 14.63 | 14.58 | -0.05 | -0.34
870 | 885 | 6 | 303 | 307.6 | 4.6 | 1.52 |  | 12.54 | 12.45 | -0.09 | -0.72
930 | 945 | 5 | 252.5 | 255.4 | 2.9 | 1.15 |  | 10.45 | 10.3 | -0.15 | -1.44
990 | 1005 | 4 | 202 | 204.6 | 2.6 | 1.29 |  | 8.36 | 8.26 | -0.1 | -1.2
1050 | 1065 | 3 | 151.5 | 153.7 | 2.2 | 1.45 |  | 6.27 | 6.2 | -0.07 | -1.12
1110 | 1125 | 2 | 101 | 101.7 | 0.7 | 0.69 |  | 4.18 | 4.1 | -0.08 | -1.91
1170 | 1185 | 1 | 50.5 | 49.3 | -1.2 | -2.38 |  | 2.09 | 2 | -0.09 | -4.31
1230 | 1245 | 0 | 0 | -0.2 | -0.2 |  |  | 0 | 0 | 0
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | 0 | 5.1 |  | Intercept | -0.06 | 0.21
|  |  |  | Slope | 1.01 | 0.99 - 1.01 |  | Slope | 1 | 0.99 - 1.01
|  |  |  | SEE | 1.232 | 5.05 |  | SEE | 0.0635 | 0.209
|  |  |  | r2 | 0.9999 | 0.998 |  | r2 | 0.9999 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.7 | 1.9938 | 2 | 1 | 1 | 0.1 | 12.6 | 900 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0017 | 0.0021 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4 | 2
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.7 | 1.9 | 1.9 | 1 | 1 | 0.1 | 12.6 | 900 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0017 | 0.0021 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5691 | 26.4 | 1.9
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.6 | 1.9 | 1.9 | 1 | 1 | 0.1 | 12.6 | 901 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0021 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.481 | 26.4 | 1.9
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.6 | 1.9 | 1.9 | 1 | 1 | 0.1 | 12.6 | 901 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3244 | 26.4 | 1.9
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.5 | 1.8938 | 1.9 | 1 | 1 | 0.1 | 12.6 | 901 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4332 | 26.4 | 1.9
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.5 | 1.8 | 1.8 | 1 | 1 | 0.1 | 12.6 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.0019 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5454 | 26.4 | 1.8
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.4 | 1.8 | 1.8 | 1 | 1 | 0.1 | 12.6 | 901 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.001494 | 0.001894 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.576 | 26.4 | 1.8
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.4 | 1.8 | 1.8 | 1.1 | 1.0573 | 0.1 | 12.6 | 900 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0018 | 0.0001 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.5025 | 26.4 | 1.8
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.4 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.8 | 899 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0018 | 0.0001 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985 | 30.5812 | 26.4 | 1.8
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.4 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.9 | 898 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0018 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 1.8
2008-05-07 | 14:59 | 0 | 0 | 0 | -2.4 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.9 | 898 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0018 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5616 | 26.4 | 1.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.3 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.9 | 898 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001394 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.517 | 26.4 | 1.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.3 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.9 | 898 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5956 | 26.4 | 1.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.3 | 1.8 | 1.8 | 1.1 | 1.1 | 0.1 | 12.9 | 898 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5257 | 26.4 | 1.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.3 | 1.7 | 1.7 | 1 | 1.0379 | 0.1 | 12.9 | 898 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4471 | 26.4 | 1.7
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.2 | 1.7 | 1.7 | 1 | 1 | 0.1 | 12.9 | 898 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4315 | 26.4 | 1.7
2008-05-07 | 15:00 | 0 | -0.0001 | -0.84913 | -2.2 | 1.7 | 1.7 | 0.9 | 0.9456 | 0.1 | 12.9 | 898 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 986 | 30.5 | 26.4 | 1.7
2008-05-07 | 15:00 | 0 | -0.0009 | -9.0932 | -2.2 | 1.7 | 1.7 | 0.9 | 0.9 | 0.1 | 12.9 | 898 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0013 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985.938 | 30.5 | 26.4 | 1.7
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -2.1 | 1.7 | 1.7 | 0.9 | 0.9 | 0.1 | 12.9 | 898 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 985 | 30.5 | 26.4 | 1.7
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -2.1 | 1.7 | 1.7 | 0.9 | 0.9 | 0.1 | 12.8 | 899 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.546 | 26.4 | 1.7
2008-05-07 | 15:00 | 0.001 | -0.001 | -10 | -2.1 | 1.7 | 1.7 | 0.9 | 0.8541 | 0.1 | 12.9 | 899 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5755 | 26.4 | 1.7
2008-05-07 | 15:00 | 0.009 | -0.0007 | -7.2986 | -2.1 | 1.7 | 1.7 | 0.8 | 0.8 | 0.1 | 12.8 | 899 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5055 | 26.4 | 1.7
2008-05-07 | 15:00 | 0.01 | -0.0001 | -0.84704 | -2.1 | 1.6938 | 1.7 | 0.8 | 0.8 | 0.1 | 12.9 | 899 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.6636 | 26.4 | 1.7
2008-05-07 | 15:00 | 0.01 | -0.0009 | -9.0707 | -2.1 | 1.6 | 1.6 | 0.7 | 0.7 | 0.1 | 12.8 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.7605 | 26.4 | 1.6
2008-05-07 | 15:00 | 0.01 | -0.0003 | -2.6083 | -2.1 | 1.6 | 1.6 | 0.7 | 0.7 | 0.1 | 12.7 | 901 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.6825 | 26.4 | 1.6
2008-05-07 | 15:00 | 0.007 | 0 | 0 | -2.1 | 1.6 | 1.6 | 0.6 | 0.6412 | 0.1 | 12.8 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5177 | 26.4 | 1.6
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.1 | 1.594 | 1.6 | 0.7 | 0.6587 | 0.1 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.5962 | 26.4 | 1.6
2008-05-07 | 15:00 | 0 | 0 | 0 | -2.1 | 1.5 | 1.5 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0012 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4 | 1.5
2008-05-07 | 15:00 | -0.007 | 0 | 0 | -2.1 | 1.5 | 1.5 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0011 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4928 | 26.4 | 1.5
2008-05-07 | 15:00 | -0.01 | 0 | 0 | -2 | 1.5 | 1.5 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0011 | 0.001494 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4322 | 26.4 | 1.5
2008-05-07 | 15:00 | -0.01 | 0 | 0 | -1.9 | 1.5 | 1.5 | 0.6 | 0.6369 | 0.1 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0011 | 0.0014 | 0.0001 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 984.062 | 30.5106 | 26.4 | 1.5
2008-05-07 | 15:00 | -0.007 | -0.0006 | -5.9343 | -1.8 | 1.5 | 1.5 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0011 | 0.0014 | 0.0001 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.5895 | 26.4 | 1.5
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.8 | 1.4 | 1.4 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0011 | 0.0014 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 984.0676 | 30.5318 | 26.4 | 1.4
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.8 | 1.4 | 1.4 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.001094 | 0.0014 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5936 | 26.4 | 1.4
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.8 | 1.4 | 1.4 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 937 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 1.4
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.8 | 1.4 | 1.4 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 1.4
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.7 | 1.3 | 1.3 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 937 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 1.3
2008-05-07 | 15:00 | 0 | -0.001 | -10 | -1.7 | 1.3 | 1.3 | 0.7 | 0.7 | 0.1 | 12.6 | 901 | 936 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7612 | 26.4 | 1.3
2008-05-07 | 15:00 | 0 | 0.0003 | 3.239 | -1.7 | 1.2938 | 1.3 | 0.7 | 0.7 | 0.1 | 12.6 | 901 | 936 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7601 | 26.4 | 1.3
2008-05-07 | 15:00 | 0 | 0.0017 | 17.3165 | -1.7 | 1.2 | 1.2 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 9.4e-05 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.663 | 26.4 | 1.2
2008-05-07 | 15:00 | 0 | 0.001 | 10 | -1.7 | 1.1938 | 1.2 | 0.7 | 0.7 | 0.1 | 12.6 | 901 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5057 | 26.4 | 1.2
2008-05-07 | 15:00 | 0 | 0.001 | 10 | -1.7 | 1.1 | 1.1 | 0.7 | 0.7 | 0.1 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6516 | 26.4 | 1.1
2008-05-07 | 15:00 | 0 | 0.001 | 10 | -1.7 | 1 | 1 | 0.8 | 0.7665 | 0.1 | 12.7 | 901 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 1
2008-05-07 | 15:00 | 0 | 0.001 | 10 | -1.6 | 1 | 1 | 0.8 | 0.8 | 0.1 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6671 | 26.3671 | 1
2008-05-07 | 15:00 | 0 | 0.0006 | 5.6063 | -1.6 | 1 | 1 | 0.8 | 0.8 | 0.1 | 12.7 | 902 | 936 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6224 | 26.3112 | 1
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.6 | 0.9 | 0.9 | 0.8 | 0.8 | 0.1 | 12.6 | 902 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.7803 | 26.3901 | 0.9
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.6 | 0.89376 | 0.9 | 0.9 | 0.9 | 0.1 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.0624 | 30.7312 | 26.4 | 0.9
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.6 | 0.8 | 0.8 | 0.9 | 0.9 | 0.1 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.7474 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.6 | 0.8 | 0.8 | 0.9 | 0.9 | 0.1 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.8 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0.03 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0013 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.7907 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.001294 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.6334 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.538 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.5405 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.6191 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.6978 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.8528 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.845 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.8336 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985.0676 | 30.8755 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.7182 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 0.9 | 0.9 | 0 | 12.8 | 900 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1021.2429 | 985 | 30.7696 | 26.4 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | 1 | 0.9535 | 0 | 12.9 | 900 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.848 | 26.352 | 0.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.5 | 0.7 | 0.7 | 1 | 1 | 0 | 12.8 | 900 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.8471 | 26.3264 | 0.7
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.6 | 1.4006 | 1.4 | 1 | 1 | 0 | 12.9 | 900 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0009 | 0.0012 | 4.3e-05 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.7053 | 26.4 | 1.4
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.8 | 21.5168 | 21.5 | 1 | 1 | 0.6 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.000906 | 0.0012 | 0.001162 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.784 | 26.4 | 21.5
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.2 | 42.517 | 42.5 | 1.4 | 1.3777 | 1.42 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001206 | 0.001306 | 0.002412 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.8 | 26.4 | 42.5
2008-05-07 | 15:00 | 0 | 0 | 0 | -1 | 45.806 | 45.8 | 1.8 | 1.7553 | 1.83 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0015 | 0.0026 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985.0604 | 30.8411 | 26.4 | 45.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.1 | 46.7 | 46.7 | 2 | 1.9545 | 1.9 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0016 | 0.0026 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8608 | 26.4 | 46.7
2008-05-07 | 15:00 | 0 | 0 | 0 | -1.1 | 46.8062 | 46.8 | 2 | 2 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.001606 | 0.001706 | 0.0026 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7032 | 26.4 | 46.8
2008-05-07 | 15:00 | 0 | 0 | 0 | -0.4 | 46.9062 | 46.9 | 1.9 | 1.9 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001806 | 0.0018 | 0.0027 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 986 | 30.8541 | 26.4 | 46.9
2008-05-07 | 15:00 | 0 | 0 | 0 | -0.2 | 47.1 | 47.1 | 1.9 | 1.9 | 2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002 | 0.0019 | 0.0027 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.7888 | 26.4 | 47.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.2 | 47.2062 | 47.2 | 2 | 1.9669 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002106 | 0.002006 | 0.0027 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.7685 | 26.4 | 47.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.5 | 47.306 | 47.3 | 2 | 2 | 2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002218 | 0.002218 | 0.0027 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985 | 30.8611 | 26.4 | 47.3
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.6 | 48.2 | 48.2 | 2 | 2 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0026 | 0.0026 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6255 | 26.4 | 48.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.6 | 48.3062 | 48.3 | 1.9 | 1.9299 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.002606 | 0.002606 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6701 | 26.4 | 48.3
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.7 | 48.4 | 48.4 | 1.9 | 1.9 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0027 | 0.0027 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7968 | 26.4 | 48.4
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.7 | 48.5 | 48.5 | 1.9 | 1.9 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.002706 | 0.002706 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021.7593 | 986 | 30.8732 | 26.4 | 48.5
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.7 | 48.5 | 48.5 | 2 | 2 | 2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0028 | 0.0028 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.0601 | 30.794 | 26.4 | 48.5
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.8 | 48.5 | 48.5 | 2 | 2 | 2 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.002807 | 0.0029 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.9325 | 30.7154 | 26.4 | 48.5
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.8 | 48.6 | 48.6 | 2.1 | 2.1 | 2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0029 | 0.0029 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7632 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.8 | 48.5062 | 48.5 | 2.2 | 2.1573 | 2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.002906 | 0.002906 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.8419 | 26.4 | 48.5
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.9 | 48.6 | 48.6 | 2.2 | 2.2 | 2 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.003 | 0.003 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.9 | 48.6 | 48.6 | 2.2 | 2.2 | 2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.003 | 0.003006 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.9 | 48.6 | 48.6 | 2.3 | 2.3 | 2 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.003006 | 0.0031 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -1 | 48.6 | 48.6 | 2.2 | 2.2447 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0031 | 0.003106 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -1 | 48.6 | 48.6 | 2.2 | 2.2 | 2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0031 | 0.0032 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.6
2008-05-07 | 15:01 | 0 | 0 | 0 | -1 | 48.7 | 48.7 | 2.2 | 2.2 | 2 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.3 | 59.4 | 0.003106 | 0.0032 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.9 | 26.4 | 48.7
2008-05-07 | 15:01 | 0 | 0 | 0 | -1 | 48.8 | 48.8 | 2.1 | 2.1 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.6 | 59.4 | 0.0032 | 0.0032 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 48.8 | 48.8 | 2.1 | 2.1 | 2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 196 | 59.4 | 0.0032 | 0.003206 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9708 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 48.8 | 48.8 | 2 | 2 | 2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195.6 | 59.4 | 0.0032 | 0.0033 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9013 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 48.8 | 48.8 | 2 | 2 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 190 | 6 | 195.2 | 59.4 | 0.003206 | 0.0033 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.856 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 48.8 | 48.8 | 1.9 | 1.9382 | 2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0033 | 0.0034 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9934 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.8 | 48.8 | 1.9 | 1.9 | 2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0033 | 0.0034 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9148 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.8 | 48.8 | 1.8 | 1.8 | 2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0033 | 0.0034 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.8062 | 48.8 | 1.7 | 1.73 | 2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.003306 | 0.0034 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 48.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9 | 48.9 | 1.7 | 1.7 | 2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0034 | 0.0035 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9 | 48.9 | 1.7 | 1.7 | 2 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0034 | 0.0035 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9 | 48.9 | 1.6 | 1.6 | 2 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0034 | 0.0035 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9 | 48.9 | 1.6 | 1.6 | 2 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0034 | 0.0035 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9 | 48.9 | 1.5 | 1.5 | 2 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 190 | 6 | 195 | 59.4 | 0.003406 | 0.0035 | 0.0027 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.9353 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 48.9062 | 48.9 | 1.5 | 1.5 | 2 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -6e-06 |  | n/a |  | 32 | 59 | 1021 | 985.938 | 30.9859 | 26.4 | 48.9
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49 | 49 | 1.5 | 1.5 | 2 | 12.6 | 902 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.9072 | 26.4 | 49
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49 | 49 | 1.5 | 1.5 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 49
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49 | 49 | 1.5 | 1.5 | 2 | 12.6 | 902 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 49
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49 | 49 | 1.5 | 1.5 | 2 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9286 | 26.4 | 49
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49 | 49 | 1.6 | 1.571 | 2 | 12.7 | 902 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9928 | 26.4 | 49
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.1 | 49.1 | 1.6 | 1.5725 | 2 | 12.6 | 902 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9142 | 26.4 | 49.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.1 | 49.1 | 1.5 | 1.5 | 2 | 12.7 | 902 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0035 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9 | 26.4 | 49.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.1 | 49.1 | 1.5 | 1.5 | 2 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.003506 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.943 | 26.4 | 49.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.1 | 49.1 | 1.4 | 1.4477 | 2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 31 | 26.4 | 49.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.1062 | 49.1 | 1.4 | 1.4 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.9994 | 26.4 | 49.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.2 | 49.2 | 1.4 | 1.4 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0036 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.8421 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.2 | 49.2 | 49.2 | 1.3 | 1.3 | 2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.8575 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.3 | 1.3 | 2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.2 | 1.2424 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0027 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8852 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.2 | 1.2 | 2 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8067 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.2 | 1.2 | 2 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8716 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.1 | 1.1 | 2 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.9503 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.1 | 1.1 | 2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9708 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.2 | 49.2 | 1.1 | 1.1 | 2 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8922 | 26.4 | 49.2
2008-05-07 | 15:01 | 0 | 0 | 0 | -1.1 | 49.3 | 49.3 | 1.1 | 1.1 | 2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8135 | 26.4 | 49.3
2008-05-07 | 15:01 | 0 | 0 | 0 | -1 | 49.3062 | 49.3 | 1 | 1 | 2 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0036 | 0.0037 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8651 | 26.4 | 49.3
2008-05-07 | 15:01 | 0.001 | 0 | 0 | -1 | 50.33 | 50.3 | 1 | 1 | 2.05 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0037 | 0.0037 | 0.00285 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 50.3
2008-05-07 | 15:01 | 0.009 | 0 | 0 | -0.9 | 104.6656 | 104.7 | 1 | 1 | 3.47 | 12.9 | 901 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005581 | 0.005474 | 0.005868 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 104.7
2008-05-07 | 15:01 | 0.01 | 0 | 0 | -0.8 | 123.7722 | 123.8 | 1.4 | 1.3602 | 4.8 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.007206 | 0.007006 | 0.006975 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9009 | 26.4 | 123.8
2008-05-07 | 15:01 | 0.004 | 0 | 0 | -0.7 | 111.397 | 111.4 | 1.7 | 1.7 | 4.56 | 12.9 | 901 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0073 | 0.007106 | 0.006282 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9794 | 26.4 | 111.4
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.6 | 105.7886 | 105.8 | 1.7 | 1.7 | 4.28 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0073 | 0.0072 | 0.005994 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 31 | 26.4 | 105.8
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.6 | 104.1 | 104.1 | 1.6 | 1.6296 | 4.13 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0074 | 0.0072 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9633 | 26.4 | 104.1
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.6 | 103.994 | 104 | 1.6 | 1.6 | 4.1 | 12.9 | 901 | 933 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0074 | 0.0072 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 104
2008-05-07 | 15:01 | 0 | 0 | 0 | -0.7 | 103.7489 | 103.7 | 1.5 | 1.5 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0074 | 0.0073 | 0.005894 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 103.7
2008-05-07 | 15:01 | 0 | -0.0001 | -0.82797 | -0.7 | 103 | 103 | 1.5 | 1.5 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0074 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9722 | 26.4 | 103
2008-05-07 | 15:02 | 0.011 | -0.0009 | -8.8666 | -0.6 | 102.9 | 102.9 | 1.5 | 1.5 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0074 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 31 | 26.4 | 102.9
2008-05-07 | 15:02 | 0.016 | -0.001 | -10 | -0.7 | 102.9 | 102.9 | 1.5 | 1.5 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.007406 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9702 | 26.4 | 102.9
2008-05-07 | 15:02 | 0.01 | -0.001 | -10 | -0.7 | 102.9 | 102.9 | 1.5 | 1.5 | 4.1 | 12.7 | 902 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 102.9
2008-05-07 | 15:02 | 0.009 | -0.001 | -10 | -0.7 | 102.8938 | 102.9 | 1.5 | 1.5 | 4.1 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9 | 26.4 | 102.9
2008-05-07 | 15:02 | 0.001 | -0.001 | -10 | -0.7 | 102.8 | 102.8 | 1.5 | 1.5 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.007306 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9 | 26.4 | 102.8
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.8 | 102.8 | 1.5 | 1.5 | 4.1 | 12.7 | 902 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.9 | 26.4 | 102.8
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.7062 | 102.7 | 1.5 | 1.5 | 4.1 | 12.7 | 903 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.9 | 26.4 | 102.7
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.794 | 102.8 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9 | 26.4 | 102.8
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.7 | 102.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.9 | 26.4 | 102.7
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.7 | 102.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.9586 | 26.4 | 102.7
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.7 | 102.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0075 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9253 | 26.4 | 102.7
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.6933 | 102.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.007507 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8159 | 26.4 | 102.7
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.6 | 102.6 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8945 | 26.4 | 102.6
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.6 | 102.6 | 1.6 | 1.5568 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 102.6
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.6 | 102.6 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8483 | 26.4 | 102.6
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.6 | 102.6 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8607 | 26.4 | 102.6
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.6 | 102.6 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0074 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9911 | 26.4 | 102.6
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.6 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.007406 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9129 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8552 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.0005 | -4.5621 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.7531 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.6061 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8416 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.9 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.7 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.8622 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.8 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8164 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.0006 | -5.5831 | -0.8 | 102.5 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 933 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.8949 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.001 | -10 | -0.8 | 102.4938 | 102.5 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7528 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | -0.0008 | -7.9316 | -0.9 | 102.4936 | 102.5 | 1.5 | 1.5448 | 4.1 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.0075 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8046 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.9 | 102.494 | 102.5 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0076 | 0.007506 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.8382 | 26.4 | 102.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.4 | 102.4 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 935 | 926 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.007606 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7 | 26.4 | 102.4
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.3936 | 102.4 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0639 | 30.7 | 26.4 | 102.4
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.3 | 102.3 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8997 | 26.4 | 102.3
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.2938 | 102.3 | 1.5 | 1.5 | 4.1 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8186 | 26.4 | 102.3
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.2 | 102.2 | 1.4 | 1.4027 | 4.1 | 12.7 | 902 | 934 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.624 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 102.2 | 102.2 | 1.4 | 1.4489 | 4.1 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7075 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.9434 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8208 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.1 | 102.1 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7742 | 26.4 | 102.1
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.1 | 102.1 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.8 | 26.4 | 102.1
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.1 | 102.1 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7686 | 26.4 | 102.1
2008-05-07 | 15:02 | 0 | -0.0005 | -5.3211 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | -0.0006 | -6.2057 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.7 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5661 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5457 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0077 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6244 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.2 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0077 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7027 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.2 | 102.2 | 102.2 | 1.4 | 1.4 | 4.1 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0077 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7816 | 26.4 | 102.2
2008-05-07 | 15:02 | 0 | 0 | 0 | -1.1 | 102.8324 | 102.8 | 1.4 | 1.4 | 4.35 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0077 | 0.0077 | 0.005831 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7397 | 26.4 | 102.8
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.3 | 138.42 | 138.4 | 1.5 | 1.5126 | 5.67 | 12.9 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.008443 | 0.008143 | 0.007837 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6222 | 26.4 | 138.4
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.8 | 152.6062 | 152.6 | 1.8 | 1.7674 | 6.01 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.009706 | 0.0095 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.535 | 26.4 | 152.6
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.9 | 152.6 | 152.6 | 2 | 1.9528 | 6.1 | 12.9 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0099 | 0.009606 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.692 | 26.4 | 152.6
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.9 | 152.6 | 152.6 | 1.9 | 1.9443 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.009906 | 0.009706 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5507 | 26.4 | 152.6
2008-05-07 | 15:02 | 0 | 0 | 0 | -0.9 | 152.5 | 152.5 | 1.8 | 1.8003 | 6.1 | 12.7 | 902 | 933 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010006 | 0.009806 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5 | 26.4 | 152.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 152.5938 | 152.6 | 1.8 | 1.8 | 6.1 | 12.7 | 903 | 934 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0101 | 0.0099 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5 | 26.4319 | 152.6
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 152.4938 | 152.5 | 1.8 | 1.8 | 6.1 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010106 | 0.01 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5105 | 26.4895 | 152.5
2008-05-07 | 15:02 | 0 | 0 | 0 | -1 | 152.4 | 152.4 | 1.8 | 1.8 | 6.1 | 12.7 | 903 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010206 | 0.010006 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5891 | 26.4109 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.4 | 152.4 | 1.7 | 1.7 | 6.1 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010306 | 0.0101 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9378 | 30.6677 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.9 | 152.4 | 152.4 | 1.7 | 1.7 | 6.1 | 12.7 | 903 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | 0.0102 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5146 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.7 | 1.7 | 6.1 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0105 | 0.0103 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3744 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3062 | 152.3 | 1.7 | 1.7 | 6.1 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0105 | 0.0103 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.5969 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.4 | 152.4 | 1.7 | 1.7 | 6.1 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0105 | 0.0103 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5179 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.4 | 152.4 | 1.7 | 1.7 | 6.1 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.010506 | 0.010306 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4393 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0106 | 0.0104 | 0.0086 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4786 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0.0104 | 0.0086 | -6e-06 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5822 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0.0105 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5043 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.0105 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4251 | 26.4 | 152.4
2008-05-07 | 15:03 | -0.009 | 0 | 0 | -1.2 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.010506 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4537 | 26.4 | 152.4
2008-05-07 | 15:03 | -0.003 | 0 | 0 | -1.2 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.010706 | 0.0106 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0106 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.4 | 152.4 | 1.6 | 1.6 | 6.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.010606 | 0.0086 | -0.0001 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.4 | 152.4 | 1.6 | 1.5721 | 6.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0107 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5681 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.0004 | 3.7631 | -1.2 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.010806 | 0.0107 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5533 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0107 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5505 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.010706 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.6928 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.3 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.010906 | 0.0108 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5351 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.011 | 0.0108 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.5 | 1.5 | 6.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.011 | 0.010806 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5795 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.4 | 1.4387 | 6.1 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.011 | 0.0109 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.6816 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.2 | 152.4 | 152.4 | 1.4 | 1.4 | 6.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0111 | 0.0109 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.6031 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.3 | 152.4 | 152.4 | 1.4 | 1.4 | 6.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0111 | 0.0109 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1021.09 | 985 | 30.4489 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.001 | 10 | -1.3 | 152.4 | 152.4 | 1.4 | 1.4 | 6.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0111 | 0.010906 | 0.0086 | -0.0001 |  | n/a |  | 31.522 | 59 | 1022 | 985 | 30.4542 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0.0003 | 2.9284 | -1.2 | 152.3938 | 152.4 | 1.4 | 1.4 | 6.1 | 12.7 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0111 | 0.011 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.3938 | 152.4 | 1.4 | 1.4 | 6.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.011 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.4892 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.011 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4101 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.2 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.011 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4685 | 26.4685 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.011 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4529 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5258 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5956 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.517 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5616 | 26.4616 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5599 | 26.4599 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3938 | 152.4 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.5 | 26.4 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.3 | 1.3439 | 6.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0113 | 0.0111 | 0.0086 | -0.0001 |  | n/a |  | 31.478 | 59 | 1022 | 986 | 30.4241 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.509 | 26.4545 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.8 | 901 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 31.5215 | 59 | 1022 | 986 | 30.6 | 26.4671 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5887 | 26.4113 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5097 | 26.4903 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.5 | 26.4311 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2938 | 152.3 | 1.4 | 1.4 | 6.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.5 | 26.4 | 152.3
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2 | 152.2 | 1.4 | 1.4 | 6.1 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0112 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.5 | 26.4 | 152.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2 | 152.2 | 1.4 | 1.4 | 6.1 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.011206 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4953 | 26.4047 | 152.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2 | 152.2 | 1.4 | 1.4 | 6.1 | 12.8 | 902 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0114 | 0.0113 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.4167 | 26.4833 | 152.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2 | 152.2 | 1.4 | 1.4 | 6.1 | 12.7 | 903 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0115 | 0.0113 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.5 | 152.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.2 | 152.2 | 1.4 | 1.4 | 6.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0115 | 0.0113 | 0.0086 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4402 | 26.5 | 152.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -1.1 | 152.3922 | 152.4 | 1.4 | 1.4 | 6.1 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0115 | 0.0113 | 0.008612 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4814 | 26.4814 | 152.4
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.5 | 165.7798 | 165.8 | 1.4 | 1.4 | 6.51 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.011643 | 0.011325 | 0.00938 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4023 | 26.4023 | 165.8
2008-05-07 | 15:03 | 0 | 0 | 0 | 1 | 187.1641 | 188.2 | 1.5 | 1.4986 | 7.62 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.012319 | 0.011719 | 0.010625 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4 | 26.4 | 187.2
2008-05-07 | 15:03 | 0 | 0 | 0 | 0.3 | 199.4022 | 199.7 | 1.8 | 1.8136 | 8.06 | 12.6 | 904 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.012706 | 0.011306 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4549 | 26.4549 | 199.4
2008-05-07 | 15:03 | 0 | 0 | 0 | 0 | 201.1062 | 201.1 | 2 | 2 | 8.11 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.013606 | 0.013106 | 0.0114 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4332 | 26.5 | 201.1
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.1 | 201.9062 | 201.9 | 1.8 | 1.8411 | 8.2 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.013706 | 0.013306 | 0.0114 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3118 | 26.5 | 201.9
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.2 | 202.1062 | 202.1 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0139 | 0.013406 | 0.0114 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3906 | 26.5 | 202.1
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.3 | 202.2 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.014006 | 0.0136 | 0.0114 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3308 | 26.4308 | 202.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.3 | 202.2 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0141 | 0.013606 | 0.011406 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3955 | 26.4478 | 202.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.4 | 202.2062 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.014106 | 0.013706 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4737 | 26.4737 | 202.2
2008-05-07 | 15:03 | 0 | 0 | 0 | -0.5 | 202.3 | 202.3 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.014206 | 0.013806 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4099 | 26.405 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.5 | 202.3 | 202.3 | 1.8 | 1.8 | 8.2 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.014306 | 0.013906 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5673 | 26.4836 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.6 | 202.2938 | 202.3 | 1.8 | 1.8 | 8.2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0144 | 0.014006 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.4134 | 26.5 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.7 | 202.3 | 202.3 | 1.8 | 1.8 | 8.2 | 12.7 | 903 | 936 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0145 | 0.0141 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3812 | 26.4594 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.7 | 202.2938 | 202.3 | 1.7 | 1.7442 | 8.2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0145 | 0.014106 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.481 | 26.419 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.8 | 202.2 | 202.2 | 1.8 | 1.7776 | 8.2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0146 | 0.014206 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4021 | 26.4979 | 202.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.8 | 202.2 | 202.2 | 1.7 | 1.7422 | 8.2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.014606 | 0.0143 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.553 | 26.4235 | 202.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.9 | 202.2 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0147 | 0.0144 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4347 | 26.4 | 202.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.9 | 202.2062 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0148 | 0.0145 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3674 | 26.4 | 202.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 202.206 | 202.2 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.014806 | 0.0145 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4877 | 26.4123 | 202.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 202.3 | 202.3 | 1.8 | 1.8 | 8.2 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0149 | 0.0145 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4091 | 26.4909 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 202.3 | 202.3 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0149 | 0.0146 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.5 | 202.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.1 | 202.4 | 202.4 | 1.8 | 1.7507 | 8.2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.015 | 0.0147 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9399 | 30.3519 | 26.4519 | 202.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.1 | 202.4062 | 202.4 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.015 | 0.0147 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3534 | 26.4267 | 202.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.1 | 202.5062 | 202.5 | 1.8 | 1.8 | 8.2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.015 | 0.0147 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.5 | 26.5 | 202.5
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6 | 202.6 | 1.7 | 1.7453 | 8.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.015006 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.5 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6 | 202.6 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0151 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4376 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6 | 202.6 | 1.8 | 1.8 | 8.2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0151 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6062 | 202.6 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0151 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3806 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6 | 202.6 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0151 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3018 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.6 | 202.6 | 1.8 | 1.7661 | 8.2 | 12.7 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0151 | 0.0148 | 0.0115 | -0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5302 | 26.5 | 202.6
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.7 | 202.7 | 1.8 | 1.8 | 8.2 | 12.8 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.015106 | 0.0149 | 0.0115 | -9.4e-05 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4893 | 26.5 | 202.7
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.7 | 202.7 | 1.8 | 1.8 | 8.2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0152 | 0.0149 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.366 | 26.466 | 202.7
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 202.8938 | 202.9 | 1.8 | 1.8 | 8.2 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | 0.0149 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3251 | 26.4126 | 202.9
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 202.7938 | 202.8 | 1.8 | 1.8 | 8.2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | 0.0149 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4822 | 26.4911 | 202.8
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.9 | 202.9 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | 0.0149 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4303 | 26.5 | 202.9
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 202.9062 | 202.9 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.4519 | 202.9
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 203 | 203 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4265 | 26.4265 | 203
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.2 | 203 | 203 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.5 | 203
2008-05-07 | 15:04 | 0 | -0.0002 | -1.8565 | -1.2 | 203 | 203 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.5 | 203
2008-05-07 | 15:04 | 0 | -0.001 | -10 | -1.2 | 203 | 203 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4374 | 26.5 | 203
2008-05-07 | 15:04 | 0 | -0.001 | -10 | -1.3 | 203.0062 | 203 | 1.8 | 1.8 | 8.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.015 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3591 | 26.4591 | 203
2008-05-07 | 15:04 | 0 | -0.0003 | -3.1389 | -1.3 | 203.1938 | 203.2 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3192 | 26.4192 | 203.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 203.1062 | 203.1 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3984 | 26.4984 | 203.1
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 203.2062 | 203.2 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.169 | 26.5 | 203.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 203.3 | 203.3 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1556 | 26.5 | 203.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.3 | 203.3 | 203.3 | 1.8 | 1.8 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2342 | 26.5 | 203.3
2008-05-07 | 15:04 | 0.011 | 0 | 0 | -1.3 | 203.3 | 203.3 | 1.8 | 1.8 | 8.2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3255 | 26.5 | 203.3
2008-05-07 | 15:04 | 0.009 | 0 | 0 | -1.3 | 203.3 | 203.3 | 1.8 | 1.8 | 8.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | 0.0151 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4827 | 26.5 | 203.3
2008-05-07 | 15:04 | -0.01 | 0 | 0 | -1.2 | 203.4 | 203.4 | 1.8 | 1.8 | 8.2 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.015406 | 0.015106 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.2903 | 26.4301 | 203.4
2008-05-07 | 15:04 | -0.01 | 0 | 0 | -1.2 | 203.4 | 203.4 | 1.7 | 1.7443 | 8.2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.2969 | 26.4485 | 203.4
2008-05-07 | 15:04 | -0.001 | 0 | 0 | -1.1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 30.4 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1.1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0601 | 30.2115 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1413 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.3 | 203.3 | 1.8 | 1.7533 | 8.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2396 | 26.5 | 203.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.3062 | 203.3 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3972 | 26.5 | 203.3
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.4 | 203.4 | 1.7 | 1.7 | 8.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.4 | 203.4 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3443 | 26.4443 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -1 | 203.4 | 203.4 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.2658 | 26.4342 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.9 | 203.4 | 203.4 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0155 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2258 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.9 | 203.4 | 203.4 | 1.8 | 1.8 | 8.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.015506 | 0.0152 | 0.0115 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3832 | 26.5 | 203.4
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.9 | 204.1626 | 204.2 | 1.8 | 1.8 | 8.2 | 12.6 | 905 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0156 | 0.015206 | 0.011543 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3299 | 26.4299 | 204.2
2008-05-07 | 15:04 | 0 | 0 | 0 | -0.1 | 225.323 | 225.3 | 2 | 1.9829 | 8.85 | 12.6 | 904 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.01585 | 0.015343 | 0.012762 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2031 | 26.4484 | 225.3
2008-05-07 | 15:04 | 0 | 0 | 0 | 0.5 | 246.1116 | 246.6 | 2.2 | 2.2 | 9.68 | 12.6 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.017019 | 0.016325 | 0.014006 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1538 | 26.5 | 246.1
2008-05-07 | 15:04 | 0 | 0 | 0 | 0.5 | 250.9558 | 251.4 | 2.1 | 2.1 | 10.1 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.017512 | 0.016819 | 0.014206 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3 | 26.5 | 251
2008-05-07 | 15:04 | 0 | 0 | 0 | 0.2 | 251.9 | 252.1 | 2.2 | 2.2 | 10.2 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.017806 | 0.017206 | 0.0143 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0601 | 30.3 | 26.5 | 251.9
2008-05-07 | 15:04 | 0 | 0 | 0 | 0.1 | 251.9 | 252 | 2.2 | 2.2 | 10.2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0179 | 0.0173 | 0.0143 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.363 | 26.5 | 251.9
2008-05-07 | 15:04 | 0 | 0 | 0 | 0.1 | 252.8 | 252.9 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.018 | 0.017306 | 0.0143 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2752 | 26.5 | 252.8
2008-05-07 | 15:04 | 0 | 0 | 0 | 0 | 252.906 | 252.9 | 2.2 | 2.2 | 10.3 | 12.6 | 902 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018006 | 0.017406 | 0.014306 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1808 | 26.5 | 252.9
2008-05-07 | 15:05 | 0 | 0 | 0 | 0 | 253.1 | 253.1 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018106 | 0.017506 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 30.495 | 26.5 | 253.1
2008-05-07 | 15:05 | 0 | 0 | 0 | 0 | 253.1062 | 253.1 | 2.2 | 2.2 | 10.3 | 12.7 | 902 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018206 | 0.017606 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3453 | 26.5 | 253.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.1 | 253.2124 | 253.2 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018306 | 0.017706 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3558 | 26.5 | 253.2
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.4 | 253.4 | 2.2 | 2.2 | 10.3 | 12.7 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0184 | 0.0178 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3314 | 26.5 | 253.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.4 | 253.4 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0184 | 0.017806 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2262 | 26.5 | 253.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.4 | 253.4 | 2.2 | 2.2 | 10.3 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0185 | 0.0179 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.3833 | 26.5 | 253.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.018506 | 0.018 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3297 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0186 | 0.018007 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2512 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.2 | 2.2 | 10.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0187 | 0.0181 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.2 | 2.2265 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0187 | 0.0181 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.3 | 2.3 | 10.3 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0187 | 0.0182 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5 | 253.5 | 2.3 | 2.3 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0188 | 0.0182 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1368 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5062 | 253.5 | 2.3 | 2.3 | 10.3 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0188 | 0.018206 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1836 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.5062 | 253.5 | 2.3 | 2.3 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0189 | 0.0183 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.5 | 253.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.6062 | 253.6 | 2.3 | 2.3 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0189 | 0.0183 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.5 | 253.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 253.7 | 253.7 | 2.3 | 2.3 | 10.3 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0189 | 0.0183 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2225 | 26.5 | 253.7
2008-05-07 | 15:05 | 0 | 0.0002 | 1.8497 | -0.3 | 253.7 | 253.7 | 2.3 | 2.3 | 10.3 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.3062 | 0.0189 | 0.0184 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2559 | 26.5 | 253.7
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 253.7062 | 253.7 | 2.3 | 2.3 | 10.3 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.019 | 0.0184 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2311 | 26.5 | 253.7
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 253.8062 | 253.8 | 2.3 | 2.3 | 10.3 | 12.8 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.019 | 0.018406 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 253.8
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 254 | 254 | 2.2 | 2.2 | 10.3 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.019 | 0.0185 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 254
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 254 | 254 | 2.3 | 2.2551 | 10.3 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.019006 | 0.0185 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3818 | 26.5 | 254
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 254.1 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0191 | 0.0185 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4025 | 26.5 | 254.1
2008-05-07 | 15:05 | 0 | 0.001 | 10 | -0.3 | 254.1062 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0191 | 0.018506 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.5 | 254.1
2008-05-07 | 15:05 | 0 | 0.001 | 9.7083 | -0.3 | 254.1938 | 254.2 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0191 | 0.0186 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0624 | 30.2937 | 26.5 | 254.2
2008-05-07 | 15:05 | 0 | 0.0001 | 1.375 | -0.3 | 254.1 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.019106 | 0.0186 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2152 | 26.5 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.3 | 254.1 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0186 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2634 | 26.5634 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.3 | 254.0938 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0186 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022.5514 | 986 | 30.2159 | 26.6 | 254.1
2008-05-07 | 15:05 | 0 | -0.0003 | -3.3528 | -0.3 | 254 | 254 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0187 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1618 | 26.5794 | 254
2008-05-07 | 15:05 | 0 | -0.0008 | -8.262 | -0.3 | 254 | 254 | 2.1 | 2.1445 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0187 | 0.0144 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3976 | 26.5008 | 254
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.3 | 254.1 | 254.1 | 2.1 | 2.1468 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0187 | 0.014407 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1667 | 26.5778 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.3 | 254.1 | 254.1 | 2.2 | 2.2 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0192 | 0.0187 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.269 | 26.5437 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.4 | 254.1 | 254.1 | 2.1 | 2.1 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0193 | 0.0187 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3651 | 26.5 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.4 | 254.1 | 254.1 | 2.1 | 2.1 | 10.3 | 12.9 | 900 | 933 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0193 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022.4447 | 986 | 30.2592 | 26.5136 | 254.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.4 | 254.1 | 254.1 | 2.1 | 2.1 | 10.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0193 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0236 | 26.5921 | 254.1
2008-05-07 | 15:05 | 0 | -0.0002 | -1.7309 | -0.4 | 254.1 | 254.1 | 2.1 | 2.1 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.019306 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0708 | 26.5292 | 254.1
2008-05-07 | 15:05 | 0 | -0.001 | -10 | -0.5 | 254.1 | 254.1 | 2.1 | 2.1 | 10.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1492 | 26.5 | 254.1
2008-05-07 | 15:05 | 0 | -0.001 | -10 | -0.5 | 254.2062 | 254.2 | 2.1 | 2.1 | 10.3 | 12.7 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1723 | 26.5 | 254.2
2008-05-07 | 15:05 | 0 | -0.001 | -10 | -0.4 | 254.4062 | 254.4 | 2.1 | 2.1 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1065 | 26.5 | 254.4
2008-05-07 | 15:05 | 0 | -0.001 | -10 | -0.4 | 254.4 | 254.4 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1851 | 26.5 | 254.4
2008-05-07 | 15:05 | 0 | -0.0007 | -6.6526 | -0.3 | 254.4 | 254.4 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1365 | 26.5635 | 254.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.4 | 254.4 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0188 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9324 | 30.1421 | 26.5579 | 254.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.4938 | 254.5 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.018806 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1583 | 26.5208 | 254.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.4 | 254.4 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0189 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0009 | 26.5995 | 254.4
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.5 | 254.5 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0194 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 986 | 30.1562 | 26.5219 | 254.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.5938 | 254.6 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.019406 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31 | 59 | 1021 | 985.062 | 30.0873 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.5062 | 254.5 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31 | 59 | 1021 | 985.9381 | 30.0349 | 26.5 | 254.5
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.6 | 254.6 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31 | 59 | 1021 | 985.938 | 30.1 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.6 | 254.6 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31 | 59 | 1021 | 986 | 30.1 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 254.6 | 254.6 | 2.1 | 2.1 | 10.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31 | 59 | 1021 | 986 | 30.171 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.1 | 254.5938 | 254.6 | 2 | 2.0355 | 10.3 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 31.5215 | 59 | 1021.0831 | 986 | 30.2496 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.1 | 254.6 | 254.6 | 2 | 2 | 10.3 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0145 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.5 | 254.6
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.2 | 260.8 | 260.8 | 2.1 | 2.1 | 10.31 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0195 | 0.0189 | 0.0148 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2863 | 26.5 | 260.8
2008-05-07 | 15:05 | 0.009 | 0 | 0 | -0.1 | 274.3056 | 274.3 | 2 | 2.005 | 11.69 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.019712 | 0.019099 | 0.015636 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1292 | 26.5 | 274.3
2008-05-07 | 15:05 | 0.01 | 0 | 0 | -0.4 | 323.2 | 323.2 | 2.5 | 2.4704 | 13.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.022143 | 0.021443 | 0.0184 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0361 | 26.5 | 323.2
2008-05-07 | 15:05 | 0.01 | 0 | 0 | -0.8 | 316.7706 | 316.8 | 2.9 | 2.9 | 12.99 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0228 | 0.0222 | 0.017988 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0425 | 26.5 | 316.8
2008-05-07 | 15:05 | 0.007 | 0 | 0 | -0.8 | 309.9938 | 310 | 2.6 | 2.6406 | 12.59 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0229 | 0.0223 | 0.0176 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 310
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.8 | 309.1 | 309.1 | 2.4 | 2.4446 | 12.5 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0229 | 0.0223 | 0.0176 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 309.1
2008-05-07 | 15:05 | 0 | 0 | 0 | -0.7 | 308.2062 | 308.2 | 2.4 | 2.3552 | 12.4 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0229 | 0.0223 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 308.2
2008-05-07 | 15:05 | 0.009 | 0 | 0 | -0.7 | 307.3 | 307.3 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.022906 | 0.0223 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0433 | 26.5 | 307.3
2008-05-07 | 15:05 | 0.01 | 0 | 0 | -0.7 | 307.288 | 307.3 | 2.3 | 2.3213 | 12.4 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.494 | 0.023 | 0.0224 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0601 | 30.0706 | 26.5 | 307.3
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.8 | 307.1 | 307.1 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.023006 | 0.022406 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1859 | 26.5 | 307.1
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.8 | 307.0062 | 307 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0231 | 0.0225 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1073 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.8 | 307.1 | 307.1 | 2.4 | 2.4 | 12.4 | 12.7 | 903 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0231 | 0.0225 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307.1
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.8 | 307.094 | 307.1 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.023106 | 0.0226 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307.1
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307 | 307 | 2.5 | 2.4558 | 12.4 | 12.7 | 902 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0232 | 0.0226 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307.0062 | 307 | 2.4 | 2.4432 | 12.4 | 12.6 | 903 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0232 | 0.0226 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 306.994 | 307 | 2.5 | 2.4639 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0232 | 0.0226 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307 | 307 | 2.5 | 2.5 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0232 | 0.0226 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5638 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307.0938 | 307.1 | 2.5 | 2.5 | 12.4 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.023206 | 0.0227 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0575 | 26.5575 | 307.1
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307.1 | 307.1 | 2.5 | 2.5 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0233 | 0.0227 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0214 | 26.5 | 307.1
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307 | 307 | 2.5 | 2.5 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0233 | 0.0227 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -0.9 | 307 | 307 | 2.5 | 2.5 | 12.4 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0233 | 0.0227 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.01 | 0 | 0 | -1 | 306.994 | 307 | 2.5 | 2.5 | 12.4 | 12.7 | 902 | 935 | 922 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.5 | 0.0233 | 0.0228 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 307
2008-05-07 | 15:06 | 0.007 | 0 | 0 | -1 | 306.8938 | 306.9 | 2.5 | 2.5 | 12.4 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.023306 | 0.0228 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1 | 306.9 | 306.9 | 2.5 | 2.5015 | 12.4 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.0228 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1144 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1 | 306.8938 | 306.9 | 2.5 | 2.5449 | 12.4 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.0228 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1929 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1 | 306.8938 | 306.9 | 2.5 | 2.5 | 12.4 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.0228 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1285 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1 | 306.8 | 306.8 | 2.5 | 2.5 | 12.4 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.022806 | 0.0175 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 986 | 30.1 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.7937 | 306.8 | 2.5 | 2.5 | 12.4 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.0229 | 0.0175 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 986 | 30.1287 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.9 | 306.9 | 2.5 | 2.5 | 12.4 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0234 | 0.0229 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1927 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.8 | 306.8 | 2.5 | 2.5 | 12.4 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.0229 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1141 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.8 | 306.8 | 2.5 | 2.5 | 12.4 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.0229 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.894 | 306.9 | 2.5 | 2.5 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.0229 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0601 | 30.0571 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.1 | 306.8 | 306.8 | 2.5 | 2.5 | 12.4 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0217 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.7 | 306.7 | 2.4 | 2.444 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.5938 | 306.6 | 2.4 | 2.4 | 12.4 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.5 | 306.5 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1575 | 26.5 | 306.5
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6 | 306.6 | 2.4 | 2.4473 | 12.4 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1279 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.5062 | 306.5 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0292 | 26.5 | 306.5
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6938 | 306.7 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1865 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6062 | 306.6 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0235 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.0571 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6062 | 306.6 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.023 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.0502 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6931 | 306.7 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.023007 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022.4643 | 985.0694 | 30.1581 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6 | 306.6 | 2.4 | 2.4 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.0231 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2847 | 26.4924 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.2 | 306.6 | 306.6 | 2.4 | 2.4 | 12.4 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.0231 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1275 | 26.4138 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.6 | 306.6 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.0231 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2296 | 26.4648 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.6062 | 306.6 | 2.4 | 2.4 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.0231 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 30.2566 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.7 | 306.7 | 2.4 | 2.4 | 12.4 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0236 | 0.0231 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.178 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.706 | 306.7 | 2.3 | 2.3451 | 12.4 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0231 | 0.0175 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 985 | 30.1 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.9 | 306.9 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.5 | 306.9
2008-05-07 | 15:06 | 0 | -0.0002 | -1.5872 | -1.3 | 306.8062 | 306.8 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | -0.001 | -9.7504 | -1.3 | 306.8062 | 306.8 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | -0.0001 | -1.4309 | -1.3 | 306.8 | 306.8 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.13 | 26.5 | 306.8
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.7 | 306.7 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2871 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.3 | 306.6938 | 306.7 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.062 | 30.2281 | 26.5 | 306.7
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.6 | 306.6 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.2 | 26.4494 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.6062 | 306.6 | 2.3 | 2.3 | 12.4 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.938 | 30.1706 | 26.4294 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.5062 | 306.5 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 986 | 30.1159 | 26.5 | 306.5
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.6 | 306.6 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2731 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.5938 | 306.6 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1046 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.5938 | 306.6 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0437 | 26.5 | 306.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -1.4 | 306.5 | 306.5 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.0175 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 306.5
2008-05-07 | 15:06 | -0.009 | 0 | 0 | -1.4 | 307.3353 | 307.3 | 2.3 | 2.3 | 12.4 | 12.7 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0237 | 0.0232 | 0.017543 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1019 | 26.5 | 307.3
2008-05-07 | 15:06 | -0.01 | 0 | 0 | -1.8 | 330.5548 | 330.6 | 2.7 | 2.6643 | 13.08 | 12.6 | 904 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.023825 | 0.023325 | 0.018831 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2584 | 26.5 | 330.6
2008-05-07 | 15:06 | -0.005 | 0 | 0 | -2.3 | 348.5863 | 348.6 | 2.8 | 2.8449 | 13.91 | 12.7 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025206 | 0.0248 | 0.019812 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1841 | 26.5 | 348.6
2008-05-07 | 15:06 | 0 | 0 | 0 | -2.1 | 353.4176 | 353.4 | 2.7 | 2.7 | 14.3 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.025419 | 0.024925 | 0.020106 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 353.4
2008-05-07 | 15:06 | 0 | 0 | 0 | -2.4 | 355.3 | 355.3 | 2.7 | 2.7 | 14.4 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0258 | 0.025306 | 0.0202 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0847 | 26.5 | 355.3
2008-05-07 | 15:06 | 0 | 0 | 0 | -2.5 | 355.3 | 355.3 | 2.6 | 2.6433 | 14.4 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0259 | 0.0255 | 0.0202 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0061 | 26.5 | 355.3
2008-05-07 | 15:06 | 0 | 0 | 0 | -2.6 | 355.4 | 355.4 | 2.6 | 2.6 | 14.4 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025906 | 0.025506 | 0.0202 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0724 | 26.4276 | 355.4
2008-05-07 | 15:06 | 0 | 0 | 0 | -2.7 | 355.4 | 355.4 | 2.6 | 2.6 | 14.4 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0261 | 0.0257 | 0.0202 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 355.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.7 | 355.4124 | 355.4 | 2.7 | 2.6533 | 14.4 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.026206 | 0.0258 | 0.0202 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0703 | 26.4297 | 355.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.4 | 355.4938 | 355.5 | 2.7 | 2.7 | 14.4 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.026306 | 0.0259 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0083 | 26.4917 | 355.5
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.5 | 355.5 | 355.5 | 2.7 | 2.7 | 14.4 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0264 | 0.026 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0869 | 26.4131 | 355.5
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.5 | 355.6 | 355.6 | 2.7 | 2.7 | 14.4 | 12.7 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.5 | 0.0265 | 0.026 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4656 | 355.6
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.6 | 355.7062 | 355.7 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0265 | 0.0261 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4559 | 355.7
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.6 | 355.8 | 355.8 | 2.7 | 2.7 | 14.5 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0266 | 0.0262 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 355.8
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.6 | 355.7124 | 355.7 | 2.7 | 2.7 | 14.5 | 12.8 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0267 | 0.0263 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1013 | 26.4013 | 355.7
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.7 | 355.9124 | 355.9 | 2.7 | 2.7 | 14.5 | 12.9 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0268 | 0.0264 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1795 | 26.4795 | 355.9
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.7 | 356 | 356 | 2.7 | 2.7 | 14.5 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0268 | 0.026406 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2583 | 26.4417 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 356 | 356 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0269 | 0.0265 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4371 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 355.9 | 355.9 | 2.7 | 2.7 | 14.5 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0269 | 0.0265 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4843 | 355.9
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 356.0062 | 356 | 2.7 | 2.7 | 14.5 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.026906 | 0.026506 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 30.3 | 26.4057 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 356 | 356 | 2.7 | 2.7 | 14.5 | 12.8 | 902 | 935 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.027 | 0.0266 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1541 | 26.4 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.9 | 355.994 | 356 | 2.7 | 2.7 | 14.5 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.027 | 0.0267 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4516 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.9 | 356 | 356 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0271 | 0.0267 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4698 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 356.1 | 356.1 | 2.7 | 2.7 | 14.5 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0271 | 0.0267 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4088 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.8 | 356.0938 | 356.1 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0271 | 0.026706 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1 | 26.4874 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.9 | 356.1 | 356.1 | 2.7 | 2.7 | 14.5 | 12.8 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.027106 | 0.0268 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.034 | 26.434 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.9 | 356.1124 | 356.1 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0272 | 0.0268 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0445 | 26.4445 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -2.9 | 356.1124 | 356.1 | 2.7 | 2.7 | 14.5 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0272 | 0.0269 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1462 | 26.5 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356 | 356 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0273 | 0.0269 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022.4486 | 986 | 30.2951 | 26.4984 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.1 | 356.1 | 2.7 | 2.7 | 14.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0273 | 0.0269 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0598 | 26.4199 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.0938 | 356.1 | 2.7 | 2.7 | 14.5 | 12.7 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0273 | 0.027 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0588 | 26.4588 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.0062 | 356 | 2.7 | 2.7 | 14.5 | 12.7 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0273 | 0.027 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1752 | 26.5 | 356
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.0938 | 356.1 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0274 | 0.027 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0622 | 30.2676 | 26.5 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.2 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0274 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1104 | 26.5 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.206 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0274 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2467 | 26.4266 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.2 | 356.2 | 2.7 | 2.7 | 14.5 | 12.7 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0274 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.452 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.1938 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0274 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.2388 | 26.4694 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.1 | 356.1 | 2.7 | 2.7 | 14.5 | 12.7 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.027406 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1183 | 26.4 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3 | 356.1062 | 356.1 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.0271 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.2743 | 26.4 | 356.1
2008-05-07 | 15:07 | 0 | -0.0007 | -6.5981 | -3.1 | 356.2 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.027106 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | -0.001 | -10 | -3.1 | 356.2 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.345 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | -0.001 | -10 | -3.1 | 356.2 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.329 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | -0.0009 | -8.5833 | -3.1 | 356.2062 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1046 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | 0 | -0.25 | -3.1 | 356.3 | 356.3 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0275 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2617 | 26.4 | 356.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.2938 | 356.3 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0276 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.2406 | 26.4 | 356.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.1062 | 356.1 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0276 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.162 | 26.4381 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.1938 | 356.2 | 2.7 | 2.7 | 14.5 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0276 | 0.0272 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1166 | 26.4834 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.1062 | 356.1 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0276 | 0.027206 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.195 | 26.405 | 356.1
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.3 | 356.3 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1262 | 26.4 | 356.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.1 | 356.3 | 356.3 | 2.7 | 2.7 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1524 | 26.4 | 356.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.4 | 356.4 | 2.8 | 2.7937 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.406 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9399 | 30.1689 | 26.4 | 356.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.3938 | 356.4 | 2.8 | 2.7533 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.4 | 59.4 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 30.0903 | 26.4 | 356.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.3 | 356.3 | 2.8 | 2.8 | 14.5 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.8 | 59.4 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0123 | 26.4 | 356.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.2 | 356.2 | 2.8 | 2.8 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.9 | 59.5 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2001 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.2 | 356.2 | 2.8 | 2.8 | 14.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.5 | 59.4 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4 | 356.2
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.4938 | 356.5 | 2.8 | 2.8 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.2 | 59.4 | 0.0276 | 0.0273 | 0.0203 | 0 |  | n/a |  | 31.4785 | 59 | 1022 | 985 | 30.2517 | 26.4 | 356.5
2008-05-07 | 15:07 | 0 | 0 | 0 | -3.2 | 356.4062 | 356.4 | 2.8 | 2.8 | 14.5 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0277 | 0.0273 | 0.0203 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4 | 356.4
2008-05-07 | 15:07 | 0 | -0.0007 | -6.6611 | -3.2 | 356.5 | 356.5 | 2.9 | 2.857 | 14.5 | 12.8 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0277 | 0.0273 | 0.0203 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4 | 356.5
2008-05-07 | 15:07 | 0 | -0.0005 | -5.0735 | -3.2 | 357.0076 | 357 | 2.8 | 2.8 | 14.5 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.027706 | 0.027306 | 0.020337 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1599 | 26.4 | 357
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.2 | 377.4502 | 377.5 | 2.8 | 2.8 | 14.87 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.028125 | 0.027419 | 0.021543 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.062 | 30.2385 | 26.4 | 377.5
2008-05-07 | 15:07 | 0 | 0 | 0 | -0.6 | 389.7734 | 389.8 | 2.9 | 2.9141 | 16.14 | 12.9 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.02855 | 0.02775 | 0.022256 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3 | 26.4171 | 389.8
2008-05-07 | 15:07 | 0 | 0 | 0 | -1 | 404.3938 | 404.4 | 3.6 | 3.5819 | 16.49 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.029606 | 0.028806 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3 | 26.4955 | 404.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.1 | 404.3062 | 404.3 | 3.4 | 3.4434 | 16.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.029706 | 0.029006 | 0.0231 | 0 |  | n/a |  | 31.522 | 59 | 1023 | 986 | 30.3 | 26.4256 | 404.3
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.3 | 404.4 | 404.4 | 3.3 | 3.266 | 16.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030106 | 0.0294 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2469 | 26.4 | 404.4
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.4 | 404.6062 | 404.6 | 3.2 | 3.2 | 16.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030206 | 0.0295 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1683 | 26.4 | 404.6
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.5 | 404.7 | 404.7 | 3.2 | 3.2 | 16.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0303 | 0.0296 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1307 | 26.4 | 404.7
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.5 | 404.7 | 404.7 | 3.2 | 3.2 | 16.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0304 | 0.0297 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3665 | 26.4 | 404.7
2008-05-07 | 15:07 | 0 | 0 | 0 | -1.6 | 405.0433 | 405 | 3.2 | 3.2 | 16.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0305 | 0.0298 | 0.0231 | 0 |  | n/a |  | 31.478 | 59 | 1023 | 986 | 30.1976 | 26.4 | 405
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.7 | 405.7 | 405.7 | 3.2 | 3.2 | 16.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0299 | 0.0231 | 0 |  | n/a |  | 31.522 | 59 | 1023 | 986 | 30.1461 | 26.4 | 405.7
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.7 | 405.794 | 405.8 | 3.2 | 3.2 | 16.5 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030606 | 0.03 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1755 | 26.4 | 405.8
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.7 | 405.8 | 405.8 | 3.2 | 3.2 | 16.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.0307 | 0.0301 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 405.8
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 405.9 | 405.9 | 3.3 | 3.2541 | 16.6 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0308 | 0.0301 | 0.0231 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 405.9
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 405.9 | 405.9 | 3.3 | 3.3 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0308 | 0.030106 | 0.0231 | 0 |  | n/a |  | 31.4785 | 59 | 1022 | 986 | 30.1 | 26.4 | 405.9
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 405.9062 | 405.9 | 3.3 | 3.3 | 16.6 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0309 | 0.0302 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.1783 | 26.4 | 405.9
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 405.9938 | 406 | 3.3 | 3.3 | 16.6 | 12.7 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0309 | 0.030206 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.2469 | 26.4177 | 406
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 406.1 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.031 | 0.0303 | 0.0231 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 986 | 30.0118 | 26.4961 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 406.1062 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.031 | 0.030306 | 0.0231 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 985.9381 | 30.0749 | 26.4251 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.2 | 406.2 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.031006 | 0.0304 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0465 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.3 | 406.3 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.5 | 0.0311 | 0.030406 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0321 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | 0 | 0 | -2 | 406.2 | 406.2 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0311 | 0.0305 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1106 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | 0 | 0 | -2 | 406.1 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.031106 | 0.0305 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.1894 | 26.4 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.1 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0312 | 0.030506 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.0639 | 26.4 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 406.0938 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0312 | 0.0306 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.0466 | 26.4 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.8 | 406.1 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0312 | 0.0306 | 0.0231 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.0749 | 26.4251 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.1124 | 406.1 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.031206 | 0.0306 | 0.023106 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.938 | 30.0032 | 26.4968 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.3 | 406.3 | 3.3 | 3.3 | 16.6 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0313 | 0.0306 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0823 | 26.4177 | 406.3
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.4 | 406.4 | 3.4 | 3.4 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0313 | 0.0307 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.461 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.3938 | 406.4 | 3.4 | 3.4 | 16.6 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0313 | 0.0307 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4604 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.3062 | 406.3 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.031306 | 0.0307 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.0003 | -3.1015 | -1.9 | 406.3 | 406.3 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0314 | 0.0307 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -1.9 | 406.2938 | 406.3 | 3.4 | 3.4 | 16.6 | 12.6 | 902 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0314 | 0.0307 | 0.0232 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985.062 | 30.0245 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -1.9 | 406.2938 | 406.3 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0314 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1083 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.0002 | -1.9377 | -1.9 | 406.4062 | 406.4 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0314 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1673 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.4 | 406.4 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0314 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.3062 | 406.3 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0315 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.4 | 406.4 | 3.3 | 3.3445 | 16.6 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0315 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 406.5876 | 406.6 | 3.4 | 3.4 | 16.6 | 12.7 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0315 | 0.0308 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.4 | 406.6
2008-05-07 | 15:08 | -0.007 | 0 | 0 | -2 | 406.4876 | 406.5 | 3.4 | 3.4 | 16.6 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0315 | 0.0309 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.062 | 30.0742 | 26.4 | 406.5
2008-05-07 | 15:08 | -0.005 | 0 | 0 | -2 | 406.2 | 406.2 | 3.4 | 3.4 | 16.6 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.031506 | 0.0309 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0088 | 26.4 | 406.2
2008-05-07 | 15:08 | 0.003 | -0.0006 | -6.3644 | -2 | 406.2 | 406.2 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.030906 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.166 | 26.4 | 406.2
2008-05-07 | 15:08 | 0.009 | -0.001 | -10 | -2 | 406.2 | 406.2 | 3.4 | 3.3693 | 16.6 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1384 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2 | 406.194 | 406.2 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0599 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2 | 406.1062 | 406.1 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0185 | 26.4 | 406.1
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.2124 | 406.2 | 3.4 | 3.4 | 16.6 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0975 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.406 | 406.4 | 3.4 | 3.4 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 31.4785 | 59 | 1022 | 985.9399 | 30.1 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.6 | 406.6 | 3.4 | 3.4 | 16.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.0316 | 0.031 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.4 | 406.6
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.4938 | 406.5 | 3.4 | 3.4 | 16.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0317 | 0.031 | 0.0232 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.0619 | 30.1332 | 26.4 | 406.5
2008-05-07 | 15:08 | 0 | -0.0007 | -6.8124 | -2.1 | 406.4938 | 406.5 | 3.3 | 3.3451 | 16.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985.062 | 30.2 | 26.4 | 406.5
2008-05-07 | 15:08 | 0 | 0 | 0 | -2.1 | 406.4062 | 406.4 | 3.4 | 3.3628 | 16.6 | 12.7 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.2 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -2.1 | 406.3938 | 406.4 | 3.4 | 3.4 | 16.6 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1309 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -2.1 | 406.4 | 406.4 | 3.4 | 3.4 | 16.6 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1022.4216 | 986 | 30.1477 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -2.1 | 406.3 | 406.3 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1473 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.0005 | -4.6747 | -2.1 | 406.4 | 406.4 | 3.4 | 3.4 | 16.6 | 12.8 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0317 | 0.0311 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.005 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.4 | 406.4 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.031706 | 0.031106 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0836 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.1 | 406.3938 | 406.4 | 3.4 | 3.4 | 16.6 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0378 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | -0.001 | -10 | -2.2 | 406.3062 | 406.3 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 406.3
2008-05-07 | 15:08 | 0 | -0.0002 | -1.8276 | -2.2 | 406.4814 | 406.5 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 406.5
2008-05-07 | 15:08 | 0 | 0 | 0 | -2.1 | 406.206 | 406.2 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 406.2
2008-05-07 | 15:08 | 0 | 0 | 0 | -2 | 406.1124 | 406.1 | 3.4 | 3.4 | 16.6 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.0232 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 406.1
2008-05-07 | 15:08 | 0 | 0 | 0 | -2 | 406.4116 | 406.4 | 3.5 | 3.4618 | 16.6 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0318 | 0.0312 | 0.023206 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0552 | 26.4 | 406.4
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.9 | 409.8796 | 409.9 | 3.5 | 3.5 | 16.94 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.031842 | 0.031224 | 0.023402 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0662 | 26.4 | 409.9
2008-05-07 | 15:08 | 0 | 0 | 0 | -0.3 | 457.9548 | 458 | 3.6 | 3.6091 | 18.87 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.033874 | 0.032874 | 0.02615 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0125 | 26.4 | 458
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.1 | 475.5404 | 475.5 | 4.4 | 4.3841 | 19.5 | 12.8 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0354 | 0.034406 | 0.027081 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0911 | 26.4 | 475.5
2008-05-07 | 15:08 | 0 | 0 | 0 | -1 | 464.6016 | 464.6 | 4.3 | 4.3272 | 19.1 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0345 | 0.026488 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 464.6
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.1 | 460.6566 | 460.7 | 3.9 | 3.937 | 18.86 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0355 | 0.0346 | 0.026294 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 460.7
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.1 | 459.2938 | 459.3 | 3.8 | 3.8437 | 18.8 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0355 | 0.0346 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0733 | 26.4 | 459.3
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.1 | 459.2 | 459.2 | 3.8 | 3.8 | 18.7 | 12.6 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0355 | 0.0346 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0055 | 26.4 | 459.2
2008-05-07 | 15:08 | 0 | 0 | 0 | -1.2 | 459.3062 | 459.3 | 3.8 | 3.8 | 18.7 | 12.6 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0355 | 0.0346 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0841 | 26.4 | 459.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.2 | 459.3062 | 459.3 | 3.9 | 3.8551 | 18.7 | 12.6 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.035506 | 0.0347 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4 | 459.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.2 | 459.2 | 459.2 | 3.9 | 3.9 | 18.7 | 12.6 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.0356 | 0.0347 | 0.0262 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 985 | 30.1 | 26.4 | 459.2
2008-05-07 | 15:09 | 0 | -0.0005 | -4.6419 | -1.3 | 459.2938 | 459.3 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0356 | 0.034706 | 0.0262 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 985 | 30.1 | 26.4 | 459.3
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -1.4 | 459.1875 | 459.2 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0356 | 0.0348 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0622 | 30.1 | 26.4 | 459.2
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -1.4 | 458.9938 | 459 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.035606 | 0.0348 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0228 | 26.4 | 459
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -1.4 | 458.8 | 458.8 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.406 | 0.0357 | 0.0348 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1114 | 26.4 | 458.8
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -1.5 | 458.8 | 458.8 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0357 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1314 | 26.4 | 458.8
2008-05-07 | 15:09 | 0 | -0.0004 | -3.8199 | -1.5 | 458.8 | 458.8 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0357 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.013 | 26.4 | 458.8
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.5 | 458.612 | 458.6 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0357 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0916 | 26.4 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.5 | 458.6067 | 458.6 | 3.9 | 3.9 | 18.7 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0357 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.5 | 458.7 | 458.7 | 3.9 | 3.9 | 18.7 | 12.6 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.035706 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0513 | 26.4 | 458.7
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.5062 | 458.5 | 3.9 | 3.9 | 18.7 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.4 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.5 | 458.6 | 458.6 | 3.9 | 3.9424 | 18.7 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.0349 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0121 | 26.4061 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.5876 | 458.6 | 4 | 3.9563 | 18.7 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.034906 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1693 | 26.4847 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.5 | 458.5 | 4 | 4 | 18.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2 | 26.5 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.5 | 458.5 | 4 | 4 | 18.7 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2 | 26.4581 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.5124 | 458.5 | 4 | 4 | 18.7 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.159 | 26.4 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.6 | 458.7 | 458.7 | 4.1 | 4.0564 | 18.7 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0019 | 26.4 | 458.7
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.8 | 458.8 | 4.1 | 4.1 | 18.7 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.0262 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8447 | 26.4 | 458.8
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.7876 | 458.8 | 4.2 | 4.2 | 18.7 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0358 | 0.035 | 0.026194 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9118 | 26.4 | 458.8
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.7 | 458.7 | 4.2 | 4.2 | 18.7 | 12.7 | 903 | 935 | 926 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.4 | 458.7
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.7 | 458.7 | 4.2 | 4.2 | 18.7 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0135 | 26.4135 | 458.7
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.7 | 458.7 | 4.1 | 4.1 | 18.7 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0921 | 26.4921 | 458.7
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.6 | 458.6 | 4.1 | 4.144 | 18.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0293 | 26.4293 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.7 | 458.588 | 458.6 | 4.2 | 4.1567 | 18.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4062 | 458.4 | 4.2 | 4.2 | 18.7 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.4
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4938 | 458.5 | 4.1 | 4.1 | 18.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0351 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4938 | 458.5 | 4.1 | 4.1297 | 18.7 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4 | 458.4 | 4.1 | 4.1 | 18.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0359 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.4
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4 | 458.4 | 4.1 | 4.1 | 18.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0843 | 26.4422 | 458.4
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.8 | 458.4 | 458.4 | 4.1 | 4.1 | 18.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1582 | 26.5 | 458.4
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.9 | 458.3 | 458.3 | 4.1 | 4.1 | 18.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0359 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0009 | 26.5 | 458.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -1.9 | 458.2876 | 458.3 | 4.1 | 4.1 | 18.7 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.036 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0781 | 26.4219 | 458.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -2 | 458.1 | 458.1 | 4.1 | 4.1 | 18.7 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.036 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1564 | 26.4 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2 | 458.1 | 458.1 | 4.1 | 4.1 | 18.7 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.036 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.165 | 26.4 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2 | 458.3 | 458.3 | 4.2 | 4.1676 | 18.7 | 12.9 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.036 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.086 | 26.4 | 458.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -2 | 458.2938 | 458.3 | 4.2 | 4.2 | 18.7 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.036 | 0.0352 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0075 | 26.4 | 458.3
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.1 | 458.1062 | 458.1 | 4.2 | 4.2 | 18.7 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.036 | 0.035206 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.1 | 458.2124 | 458.2 | 4.3 | 4.2689 | 18.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.036 | 0.0353 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.4 | 458.2
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.1 | 458.1938 | 458.2 | 4.3 | 4.3 | 18.7 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.036 | 0.0353 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0283 | 26.4283 | 458.2
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.2 | 458.1 | 458.1 | 4.3 | 4.3 | 18.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.494 | 0.036 | 0.0353 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4931 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.2 | 458.1 | 458.1 | 4.4 | 4.3577 | 18.7 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0361 | 0.035306 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4145 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.2 | 458.1 | 458.1 | 4.4 | 4.4 | 18.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0359 | 26.4 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.2 | 458.1 | 458.1 | 4.4 | 4.4 | 18.7 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0427 | 26.4427 | 458.1
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.2 | 458 | 458 | 4.4 | 4.4 | 18.7 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 458
2008-05-07 | 15:09 | 0 | -0.0003 | -3.0137 | -2.2 | 458 | 458 | 4.4 | 4.4 | 18.7 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 458
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.2 | 458 | 458 | 4.4 | 4.4 | 18.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 458
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.2 | 458.0062 | 458 | 4.4 | 4.4 | 18.7 | 12.7 | 902 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 458
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.2 | 458.1 | 458.1 | 4.4 | 4.4 | 18.7 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 977 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4645 | 458.1
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.2 | 458.1 | 458.1 | 4.4 | 4.4 | 18.7 | 12.7 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 458.1
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.2 | 458.3062 | 458.3 | 4.4 | 4.4 | 18.7 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 458.3
2008-05-07 | 15:09 | 0 | -0.0007 | -6.9583 | -2.3 | 458.4062 | 458.4 | 4.4 | 4.4 | 18.7 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4062 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.1 | 26.4 | 458.4
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.3 | 458.5 | 458.5 | 4.5 | 4.4531 | 18.7 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4938 | 0.0361 | 0.0354 | 0.0261 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.15 | 26.45 | 458.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.5 | 474.8827 | 474.9 | 5 | 4.9732 | 19.18 | 12.6 | 904 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0362 | 0.0355 | 0.027041 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1714 | 26.4714 | 474.9
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.9 | 498.6495 | 498.6 | 5.1 | 5.1233 | 20.08 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.037025 | 0.036319 | 0.0284 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.4073 | 498.6
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.5 | 502.4938 | 502.5 | 5 | 5 | 20.49 | 12.7 | 904 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0378 | 0.037006 | 0.0286 | 0 |  | n/a |  | 31.478 | 59 | 1023 | 985.062 | 30.1 | 26.4859 | 502.5
2008-05-07 | 15:09 | 0 | 0 | 0 | -2.5 | 504.394 | 504.4 | 5 | 5 | 20.7 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.037906 | 0.037106 | 0.0287 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.0601 | 30.228 | 26.436 | 504.4
2008-05-07 | 15:09 | 0 | -0.0003 | -3.0369 | -2.7 | 505.1062 | 505.1 | 5 | 5 | 20.7 | 12.6 | 904 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.038006 | 0.037306 | 0.0288 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2147 | 26.4426 | 505.1
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.7 | 505.2062 | 505.2 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.038106 | 0.0374 | 0.0288 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.5 | 505.2
2008-05-07 | 15:09 | 0 | -0.001 | -10 | -2.9 | 505.312 | 505.3 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0382 | 0.0375 | 0.0288 | 0 |  | n/a |  | 31 | 59 | 1022.1728 | 985.9399 | 30.1 | 26.4998 | 505.3
2008-05-07 | 15:10 | 0 | -0.001 | -10 | -2.9 | 505.3124 | 505.3 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0383 | 0.0376 | 0.0288 | 0 |  | n/a |  | 31.522 | 59 | 1022 | 986 | 30.1 | 26.4212 | 505.3
2008-05-07 | 15:10 | 0 | -0.0004 | -3.7928 | -3 | 505.4124 | 505.4 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0384 | 0.0377 | 0.0288 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1574 | 26.4574 | 505.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -3 | 505.5876 | 505.6 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0385 | 0.0378 | 0.0288 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.164 | 26.464 | 505.6
2008-05-07 | 15:10 | 0 | 0 | 0 | -3 | 505.5496 | 505.5 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.038506 | 0.037806 | 0.028806 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1292 | 26.4 | 505.5
2008-05-07 | 15:10 | 0 | 0 | 0 | -3 | 506.3068 | 506.3 | 5 | 5 | 20.8 | 12.6 | 904 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0386 | 0.0379 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2865 | 26.4 | 506.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.9 | 506.2062 | 506.2 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.038606 | 0.037906 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4717 | 506.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.9 | 506.3938 | 506.4 | 5 | 5 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0387 | 0.038 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2497 | 26.5 | 506.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -3 | 506.4938 | 506.5 | 5 | 5 | 20.8 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0388 | 0.0381 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2289 | 26.4711 | 506.5
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.1 | 506.4938 | 506.5 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.038806 | 0.0382 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4075 | 506.5
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.1 | 506.4062 | 506.4 | 5 | 5 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.0389 | 0.038206 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4861 | 506.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.4938 | 506.5 | 5 | 5 | 20.8 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.038906 | 0.0383 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4356 | 506.5
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.4124 | 506.4 | 4.9 | 4.9298 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.039 | 0.0384 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2139 | 26.4431 | 506.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.3 | 506.412 | 506.4 | 4.9 | 4.9 | 20.8 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0391 | 0.038406 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1437 | 26.4781 | 506.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.3 | 506.8 | 506.8 | 4.9 | 4.9 | 20.8 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0392 | 0.0385 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2989 | 26.4 | 506.8
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.7 | 506.7 | 4.9 | 4.9 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0392 | 0.0386 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1418 | 26.4 | 506.7
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.694 | 506.7 | 4.8 | 4.8 | 20.8 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0386 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2154 | 26.4 | 506.7
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.5938 | 506.6 | 4.8 | 4.8 | 20.8 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0386 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2637 | 26.4 | 506.6
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.6 | 506.6 | 4.8 | 4.8 | 20.8 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.0393 | 0.0387 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1851 | 26.4149 | 506.6
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.2 | 506.7876 | 506.8 | 4.7 | 4.7 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0394 | 0.0387 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1065 | 26.4935 | 506.8
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.1 | 506.6 | 506.6 | 4.7 | 4.7 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0387 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1716 | 26.5 | 506.6
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.1 | 506.8 | 506.8 | 4.6 | 4.6455 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.039406 | 0.038706 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2503 | 26.5 | 506.8
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.1 | 506.9124 | 506.9 | 4.6 | 4.6 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0388 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.2123 | 26.5 | 506.9
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.1 | 507.1 | 507.1 | 4.6 | 4.6 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0388 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0234 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.1 | 507.0938 | 507.1 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0388 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.062 | 30.2594 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.1 | 507.2 | 507.2 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0388 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1701 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.1 | 507.2 | 507.2 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0397 | 0.0388 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.2 | 507.2 | 507.2 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0397 | 0.0389 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.2 | 507.1062 | 507.1 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.039706 | 0.0389 | 0.0289 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.1008 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.3 | 507.206 | 507.2 | 4.4 | 4.4 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0398 | 0.038906 | 0.0289 | 6e-06 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1793 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.3 | 507.3 | 507.3 | 4.5 | 4.5 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0398 | 0.039 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0842 | 26.5 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.4 | 507.2062 | 507.2 | 4.4 | 4.4403 | 20.8 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0398 | 0.039006 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.0366 | 26.4634 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.1879 | 507.2 | 4.4 | 4.4 | 20.8 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.039806 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.4 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.2062 | 507.2 | 4.4 | 4.4 | 20.8 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.938 | 30.1 | 26.4 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.2814 | 507.3 | 4.4 | 4.4 | 20.8 | 12.8 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.062 | 30.1 | 26.472 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507 | 507 | 4.4 | 4.4 | 20.8 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022.6915 | 986 | 30.0493 | 26.4492 | 507
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.1876 | 507.2 | 4.4 | 4.4 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0296 | 26.4296 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.0186 | 507 | 4.4 | 4.4 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1082 | 26.5 | 507
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.3 | 507.3 | 4.4 | 4.4 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.0399 | 0.0391 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1867 | 26.5 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.2062 | 507.2 | 4.4 | 4.4 | 20.8 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0399 | 0.039106 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1347 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.5 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0399 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4561 | 507.5
2008-05-07 | 15:10 | 0.008 | 0 | 0 | -2.6 | 507.5 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0399 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0775 | 26.4 | 507.5
2008-05-07 | 15:10 | 0.01 | 0 | 0 | -2.6 | 507.5124 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0399 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0011 | 26.4011 | 507.5
2008-05-07 | 15:10 | 0.01 | 0 | 0 | -2.6 | 507.4938 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0793 | 26.4793 | 507.5
2008-05-07 | 15:10 | 0.007 | 0 | 0 | -2.6 | 507.4938 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 507.5
2008-05-07 | 15:10 | 0.001 | 0 | 0 | -2.6 | 507.494 | 507.5 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0631 | 26.5 | 507.5
2008-05-07 | 15:10 | 0.01 | 0 | 0 | -2.6 | 507.3876 | 507.4 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 30 | 26.5 | 507.4
2008-05-07 | 15:10 | 0.002 | 0 | 0 | -2.6 | 507.2 | 507.2 | 4.3 | 4.3 | 20.9 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.194 | 507.2 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.4274 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 507.1933 | 507.2 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1024 | 26.4512 | 507.2
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.0938 | 507.1 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1701 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.1 | 507.1 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.1124 | 507.1 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 507.1
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.3 | 507.3 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.04 | 0.0392 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4344 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.3 | 507.3 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4938 | 0.04 | 0.039206 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4443 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.3124 | 507.3 | 4.3 | 4.3 | 20.9 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.04 | 0.0393 | 0.0289 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1 | 26.5 | 507.3
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.6 | 507.0156 | 507 | 4.3 | 4.3 | 20.9 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.04 | 0.0393 | 0.028881 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0985 | 26.5 | 507
2008-05-07 | 15:10 | 0 | 0 | 0 | -2.5 | 493.0026 | 493 | 4.3 | 4.3 | 20.6 | 12.7 | 901 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.039994 | 0.039294 | 0.028138 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0203 | 26.5 | 493
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.1 | 469.3652 | 469.4 | 4.2 | 4.2 | 19.68 | 12.6 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.039788 | 0.039194 | 0.026781 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0585 | 26.4415 | 469.4
2008-05-07 | 15:10 | 0 | 0 | 0 | -4 | 463.976 | 464 | 4.2 | 4.2 | 19.14 | 12.7 | 901 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.039094 | 0.026494 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0628 | 26.4 | 464
2008-05-07 | 15:10 | 0 | 0 | 0 | -3.7 | 460.1938 | 460.2 | 4.2 | 4.2 | 18.92 | 12.6 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.039194 | 0.038694 | 0.026206 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0158 | 26.4 | 460.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.6 | 460.0938 | 460.1 | 4.1 | 4.1 | 18.9 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.039 | 0.0385 | 0.0263 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0944 | 26.4 | 460.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.5 | 460.0062 | 460 | 4 | 4 | 18.9 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0389 | 0.0384 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.473 | 460
2008-05-07 | 15:11 | 0 | -0.0004 | -4.3693 | -3.5 | 460 | 460 | 4 | 4 | 18.9 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.038794 | 0.038294 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1516 | 26.5 | 460
2008-05-07 | 15:11 | 0 | -0.001 | -10 | -3.4 | 459.988 | 460 | 4 | 4 | 18.8 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0387 | 0.038194 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0601 | 30.1698 | 26.5 | 460
2008-05-07 | 15:11 | 0 | -0.001 | -10 | -3.4 | 459.8 | 459.8 | 4 | 4 | 18.8 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.038694 | 0.0381 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.8
2008-05-07 | 15:11 | 0 | -0.001 | -10 | -3.3 | 459.8876 | 459.9 | 4 | 4 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0385 | 0.038 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1 | 26.5 | 459.9
2008-05-07 | 15:11 | 0 | -0.001 | -10 | -3.3 | 459.7 | 459.7 | 4 | 4 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0384 | 0.037894 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0341 | 26.4341 | 459.7
2008-05-07 | 15:11 | 0 | -0.0004 | -3.9474 | -3.3 | 459.6936 | 459.7 | 4 | 4 | 18.8 | 12.8 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0383 | 0.037794 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0892 | 26.4446 | 459.7
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.3 | 459.6 | 459.6 | 4 | 4 | 18.8 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.038294 | 0.0377 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1768 | 26.5 | 459.6
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.6 | 459.6 | 4 | 4 | 18.8 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0382 | 0.037694 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.6
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.7 | 459.7 | 4 | 4 | 18.8 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.038194 | 0.0376 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 459.7
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.6 | 459.6 | 4 | 4 | 18.8 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0381 | 0.037594 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.6
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.5 | 459.5 | 4 | 4 | 18.8 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0381 | 0.0375 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.5
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.4062 | 459.4 | 4 | 4 | 18.8 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.038 | 0.0375 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1161 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.3876 | 459.4 | 4 | 4 | 18.8 | 12.8 | 900 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4938 | 0.038 | 0.0374 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022.2854 | 986 | 30.1947 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.2 | 459.3875 | 459.4 | 4 | 4 | 18.8 | 12.8 | 900 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.037994 | 0.0374 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1267 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.2 | 459.2 | 4 | 4 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0379 | 0.0374 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0481 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.3876 | 459.4 | 4 | 4 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0379 | 0.0373 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.061 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.2 | 459.2 | 4 | 4 | 18.8 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.037894 | 0.0373 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1909 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.006 | 459 | 4 | 4 | 18.8 | 12.9 | 900 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0378 | 0.037294 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1127 | 26.5 | 459
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.1124 | 459.1 | 4 | 4 | 18.8 | 12.9 | 899 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0378 | 0.0372 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.2062 | 459.2 | 4 | 4 | 18.8 | 12.9 | 899 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0378 | 0.0372 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1897 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.294 | 459.3 | 4 | 4 | 18.8 | 12.9 | 899 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.037794 | 0.0372 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2531 | 26.5 | 459.3
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.3 | 459.3 | 3.9 | 3.9417 | 18.8 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0377 | 0.0372 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.3
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.1 | 459.5 | 459.5 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0377 | 0.0371 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.5 | 459.5
2008-05-07 | 15:11 | 0 | 0 | 0 | -3 | 459.4 | 459.4 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0377 | 0.0371 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -3 | 459.3938 | 459.4 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0376 | 0.037094 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.062 | 30.1 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.1062 | 459.1 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0376 | 0.037 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0376 | 0.037 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0376 | 0.037 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.1 | 459.1 | 3.9 | 3.9 | 18.8 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0376 | 0.037 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.1938 | 459.2 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.037594 | 0.036994 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.9 | 3.9 | 18.8 | 12.8 | 900 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.9 | 3.9 | 18.8 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.5 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2062 | 459.2 | 3.8 | 3.8433 | 18.8 | 12.8 | 901 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0335 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.218 | 459.2 | 3.8 | 3.8455 | 18.8 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9399 | 30.0452 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.3124 | 459.3 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4762 | 459.3
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.5 | 459.5 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0375 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0976 | 26.4024 | 459.5
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.494 | 459.5 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.406 | 0.037494 | 0.0369 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.019 | 26.481 | 459.5
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.4 | 459.4 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0595 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.4 | 459.4 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0619 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.3938 | 459.4 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0167 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.3938 | 459.4 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0949 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0262 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.8 | 3.8 | 18.8 | 12.6 | 902 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.1062 | 459.1 | 3.8 | 3.8 | 18.8 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0311 | 26.5 | 459.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2 | 459.2 | 3.8 | 3.8 | 18.8 | 12.7 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0903 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2124 | 459.2 | 3.8 | 3.8 | 18.8 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4938 | 0.0374 | 0.0368 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0124 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.9 | 459.2125 | 459.2 | 3.8 | 3.8 | 18.8 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0374 | 0.036794 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.5 | 459.2
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.8 | 459.4 | 459.4 | 3.8 | 3.8 | 18.8 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.037394 | 0.0367 | 0.0262 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.5 | 459.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -2.8 | 459.0822 | 459.1 | 3.9 | 3.8692 | 18.8 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4938 | 0.0373 | 0.0367 | 0.026194 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0722 | 26.5 | 459.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -3 | 431.4697 | 431.5 | 3.9 | 3.8522 | 18.54 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.037016 | 0.036528 | 0.024634 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2947 | 26.5 | 431.5
2008-05-07 | 15:11 | 0 | 0 | 0 | -4.2 | 397 | 397 | 3.8 | 3.8 | 16.65 | 12.6 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0349 | 0.0346 | 0.0227 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1375 | 26.5 | 397
2008-05-07 | 15:11 | 0 | 0 | 0 | -3.9 | 387.3658 | 387.4 | 3.5 | 3.4834 | 16 | 12.6 | 901 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.033794 | 0.0335 | 0.022119 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0402 | 26.5 | 387.4
2008-05-07 | 15:11 | 0 | 0 | 0 | -4.1 | 403.7 | 403.7 | 3.5 | 3.4546 | 16.33 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.033694 | 0.0334 | 0.023 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0384 | 26.5 | 403.7
2008-05-07 | 15:11 | 0 | 0 | 0 | -4 | 403.8984 | 403.9 | 3.5 | 3.5 | 16.58 | 12.7 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0336 | 0.0333 | 0.023012 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1 | 26.5 | 403.9
2008-05-07 | 15:11 | 0 | 0 | 0 | -4 | 407 | 407 | 3.5 | 3.5 | 16.63 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0336 | 0.0333 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 407
2008-05-07 | 15:11 | 0 | 0 | 0 | -4 | 407.0062 | 407 | 3.5 | 3.5 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.033594 | 0.033294 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 407
2008-05-07 | 15:11 | 0 | 0 | 0 | -4 | 407.1 | 407.1 | 3.5 | 3.5 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0335 | 0.0332 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9415 | 26.5 | 407.1
2008-05-07 | 15:11 | 0 | 0 | 0 | -4 | 407.1 | 407.1 | 3.5 | 3.5 | 16.7 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.494 | 0.0335 | 0.0332 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 407.1
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 407.2 | 407.2 | 3.4 | 3.4423 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0335 | 0.0332 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8201 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 407.2 | 407.2 | 3.4 | 3.4318 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0335 | 0.0332 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9774 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -4 | 407.1 | 407.1 | 3.4 | 3.4448 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0335 | 0.0332 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0673 | 26.5673 | 407.1
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 407.1062 | 407.1 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0541 | 26.5082 | 407.1
2008-05-07 | 15:12 | 0 | -0.0001 | -0.85477 | -3.9 | 407.2 | 407.2 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.5 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9756 | 26.4244 | 407.2
2008-05-07 | 15:12 | 0 | -0.0009 | -9.1535 | -3.9 | 407.2 | 407.2 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9055 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | -0.001 | -10 | -3.9 | 407.2 | 407.2 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0634 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | -0.001 | -10 | -3.9 | 407.306 | 407.3 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | -0.0006 | -5.5649 | -3.9 | 407.4 | 407.4 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.0331 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1389 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 407.3938 | 407.4 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.5 | 0.0334 | 0.033094 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.165 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | -0.0001 | -0.84577 | -3.8 | 407.3 | 407.3 | 3.4 | 3.4 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.033 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0082 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | -0.0009 | -9.1376 | -3.8 | 407.2876 | 407.3 | 3.4 | 3.4 | 16.7 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0334 | 0.033 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 29.9254 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | -0.0003 | -2.5396 | -3.8 | 407.1 | 407.1 | 3.4 | 3.4 | 16.7 | 12.7 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.033394 | 0.033 | 0.0232 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0066 | 26.5 | 407.1
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 406.9938 | 407 | 3.4 | 3.4 | 16.7 | 12.8 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0333 | 0.033 | 0.023206 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0681 | 26.5 | 407
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.9 | 406.8062 | 406.8 | 3.4 | 3.4 | 16.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0333 | 0.033 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986.938 | 30.0211 | 26.5 | 406.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.8 | 407.0062 | 407 | 3.3 | 3.3462 | 16.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0333 | 0.033 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1783 | 26.5 | 407
2008-05-07 | 15:12 | 0 | -0.0001 | -1.0308 | -3.8 | 407.2 | 407.2 | 3.4 | 3.4 | 16.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0333 | 0.033 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9967 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | -0.0009 | -9.1477 | -3.8 | 407.3062 | 407.3 | 3.4 | 3.4 | 16.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0333 | 0.033 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9463 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | -0.001 | -10 | -3.8 | 407.3 | 407.3 | 3.4 | 3.4 | 16.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0333 | 0.032994 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9752 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | -0.001 | -10 | -3.8 | 407.4 | 407.4 | 3.3 | 3.3 | 16.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0333 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.906 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | -0.0006 | -5.6924 | -3.8 | 407.4 | 407.4 | 3.4 | 3.4 | 16.7 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0641 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.3876 | 407.4 | 3.4 | 3.4 | 16.7 | 12.9 | 899 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.2 | 407.2 | 3.3 | 3.3237 | 16.7 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0607 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.2062 | 407.2 | 3.4 | 3.3675 | 16.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9821 | 26.5 | 407.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.3 | 407.3 | 3.3 | 3.3 | 16.7 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9034 | 26.5 | 407.3
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.4938 | 407.5 | 3.3 | 3.3 | 16.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.0332 | 0.0329 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9752 | 26.5 | 407.5
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.406 | 407.4 | 3.2 | 3.2 | 16.7 | 12.9 | 901 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.032894 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9462 | 26.5 | 407.4
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.6062 | 407.6 | 3.3 | 3.3 | 16.7 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.9381 | 29.9323 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 407.7 | 407.7 | 3.2 | 3.2094 | 16.7 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.5 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.6 | 407.7 | 407.7 | 3.3 | 3.3 | 16.7 | 12.8 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.494 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.6 | 407.8 | 407.8 | 3.3 | 3.3 | 16.7 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8637 | 26.5 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.6 | 407.8938 | 407.9 | 3.3 | 3.3 | 16.7 | 12.8 | 901 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 407.9
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.7 | 407.7 | 3.3 | 3.3 | 16.7 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.8 | 407.8 | 3.2 | 3.2432 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4062 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.8 | 407.8 | 3.3 | 3.2688 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0332 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.8 | 407.8 | 3.2 | 3.2 | 16.7 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.033194 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8613 | 26.5 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.7062 | 407.7 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4938 | 0.0331 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.9798 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.6062 | 407.6 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0328 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0631 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.6 | 407.6 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9058 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 407.6 | 407.6 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4938 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.4 | 407.6 | 407.6 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9543 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.4 | 407.6062 | 407.6 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9342 | 26.5 | 407.6
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.4 | 407.7 | 407.7 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8115 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.4 | 407.7062 | 407.7 | 3.2 | 3.2 | 16.7 | 12.7 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9376 | 29.8901 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.3 | 407.8 | 407.8 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9688 | 26.5688 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.3 | 407.7 | 407.7 | 3.2 | 3.2 | 16.7 | 12.7 | 902 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5526 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.3 | 407.7 | 407.7 | 3.2 | 3.2 | 16.7 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0259 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.2 | 407.7062 | 407.7 | 3.2 | 3.2 | 16.7 | 12.7 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0954 | 26.5 | 407.7
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.2 | 407.8 | 407.8 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0327 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0169 | 26.5 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.2 | 407.8938 | 407.9 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.032694 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.5 | 407.9
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.2 | 407.8 | 407.8 | 3.2 | 3.2 | 16.7 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0331 | 0.0326 | 0.0233 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9599 | 26.5401 | 407.8
2008-05-07 | 15:12 | 0 | 0 | 0 | -3 | 392.2304 | 392.2 | 3.2 | 3.2 | 16.48 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.032994 | 0.032594 | 0.022363 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.5629 | 392.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -4.6 | 363.1 | 363.1 | 3.2 | 3.2 | 15.23 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0316 | 0.0315 | 0.0207 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.4049 | 363.1
2008-05-07 | 15:12 | 0 | 0 | 0 | -4.6 | 361.0884 | 361.1 | 3 | 3.0384 | 14.73 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.031481 | 0.031381 | 0.020594 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9761 | 26.4761 | 361.1
2008-05-07 | 15:12 | 0 | 0 | 0 | -4 | 359.3938 | 359.4 | 3 | 3 | 14.6 | 12.6 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.030881 | 0.030681 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9452 | 26.5 | 359.4
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.7 | 359.3 | 359.3 | 3 | 3.0401 | 14.6 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.030594 | 0.0303 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.5 | 359.3
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.6 | 359.3 | 359.3 | 3 | 3.0451 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0305 | 0.0302 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.912 | 26.5 | 359.3
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.6 | 359.2 | 359.2 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0304 | 0.030194 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9906 | 26.5 | 359.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 359.2 | 359.2 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0303 | 0.0301 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8615 | 26.5 | 359.2
2008-05-07 | 15:12 | 0 | 0 | 0 | -3.5 | 359.1062 | 359.1 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.030294 | 0.03 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8478 | 26.5 | 359.1
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.4 | 359 | 359 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0299 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9264 | 26.5 | 359
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.4 | 358.9 | 358.9 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.406 | 0.0301 | 0.0298 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.995 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.3 | 358.8938 | 358.9 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4938 | 0.030094 | 0.029794 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9164 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.3 | 358.8 | 358.8 | 3 | 3 | 14.6 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.03 | 0.0297 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 358.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.3 | 358.8 | 358.8 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.03 | 0.0297 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9407 | 26.5 | 358.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.3 | 358.9 | 358.9 | 3 | 3 | 14.6 | 12.8 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.029994 | 0.029694 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.3 | 358.8938 | 358.9 | 3 | 3 | 14.6 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0299 | 0.0296 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | -0.0006 | -5.8246 | -3.2 | 358.9 | 358.9 | 3 | 3 | 14.6 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4938 | 0.029894 | 0.0295 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9233 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | -0.001 | -10 | -3.2 | 358.9 | 358.9 | 3 | 3 | 14.6 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0298 | 0.0295 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9553 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | -0.001 | -10 | -3.1 | 358.9 | 358.9 | 3 | 3 | 14.6 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0298 | 0.0294 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.5 | 358.9
2008-05-07 | 15:13 | 0 | -0.0009 | -9.1592 | -3.1 | 358.8 | 358.8 | 3 | 3 | 14.58 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0297 | 0.0294 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.5 | 358.8
2008-05-07 | 15:13 | 0 | -0.0001 | -0.99592 | -3.1 | 358.7062 | 358.7 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0297 | 0.0294 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.5 | 358.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.1 | 358.6938 | 358.7 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0297 | 0.0293 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.5 | 358.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.5 | 358.5 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0296 | 0.0293 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.494 | 358.5 | 3 | 3 | 14.5 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0296 | 0.0293 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9472 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.5062 | 358.5 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.5 | 0.0296 | 0.0292 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.811 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.6 | 358.6 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0296 | 0.0292 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9681 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.7 | 358.7 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0295 | 0.0292 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0627 | 26.5 | 358.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.8 | 358.8 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0295 | 0.0291 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0176 | 26.5 | 358.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.8062 | 358.8 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0295 | 0.0291 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.5197 | 358.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.8 | 358.8 | 3 | 3 | 14.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.0291 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.5985 | 358.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.7062 | 358.7 | 3 | 3 | 14.5 | 12.8 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.0291 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.5229 | 358.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.6 | 358.6 | 3 | 3 | 14.6 | 12.7 | 902 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9557 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.6938 | 358.7 | 3 | 3 | 14.6 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.9316 | 26.5342 | 358.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.6062 | 358.6 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.8 | 26.6 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.6 | 358.6 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0294 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.6 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5876 | 358.6 | 3 | 3 | 14.6 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0294 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9404 | 26.5298 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.4938 | 358.5 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.029394 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5484 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.5 | 358.5 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5732 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.4124 | 358.4 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.0293 | 0.029 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.9881 | 26.5 | 358.4
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.5 | 358.5 | 3 | 3 | 14.6 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.028994 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8308 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5938 | 358.6 | 3 | 3 | 14.6 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0619 | 29.8 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5062 | 358.5 | 3 | 3 | 14.6 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.8 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.6 | 358.6 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.6 | 358.6 | 3 | 3 | 14.6 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0293 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5876 | 358.6 | 3 | 3 | 14.6 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.0293 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.8 | 26.5 | 358.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.4062 | 358.4 | 2.9 | 2.9452 | 14.6 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.029294 | 0.0289 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.8 | 26.5 | 358.4
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5 | 358.5 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8694 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.8 | 358.5 | 358.5 | 3 | 3 | 14.6 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9594 | 26.5 | 358.5
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.4 | 358.4 | 2.9 | 2.9 | 14.6 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.7236 | 26.5 | 358.4
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.3938 | 358.4 | 3 | 2.9679 | 14.6 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.7708 | 26.5 | 358.4
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.3938 | 358.4 | 2.9 | 2.9 | 14.6 | 12.6 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.8 | 26.5 | 358.4
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.2 | 358.2 | 2.9 | 2.9 | 14.58 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0292 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.8 | 26.5 | 358.2
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358.2938 | 358.3 | 2.9 | 2.9 | 14.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0288 | 0.0205 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5066 | 358.3
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 358.1 | 358.1 | 2.9 | 2.9 | 14.5 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0288 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5852 | 358.1
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 357.9938 | 358 | 2.9 | 2.9 | 14.5 | 12.6 | 902 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0288 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5362 | 358
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.9 | 357.9 | 2.9 | 2.9 | 14.5 | 12.7 | 902 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0288 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8425 | 26.5 | 357.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.9 | 357.9 | 2.9 | 2.9 | 14.5 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.028794 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9211 | 26.5 | 357.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.9 | 357.9 | 2.9 | 2.9 | 14.5 | 12.7 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0287 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9998 | 26.5 | 357.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.806 | 357.8 | 2.8 | 2.8436 | 14.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0287 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9216 | 26.5 | 357.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.9 | 357.9 | 2.9 | 2.9 | 14.5 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0287 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 357.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 358 | 358 | 2.9 | 2.9 | 14.5 | 12.8 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0287 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9353 | 26.5 | 358
2008-05-07 | 15:13 | 0 | 0 | 0 | -3 | 357.994 | 358 | 2.9 | 2.8544 | 14.5 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0291 | 0.0287 | 0.0204 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9717 | 26.5 | 358
2008-05-07 | 15:13 | 0 | 0 | 0 | -2.9 | 349.6552 | 349.7 | 2.9 | 2.9 | 14.11 | 12.8 | 900 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.028994 | 0.0287 | 0.019963 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8145 | 26.5 | 349.7
2008-05-07 | 15:13 | 0 | 0 | 0 | -4.1 | 318.8474 | 318.8 | 2.7 | 2.7452 | 13.23 | 12.8 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.028263 | 0.028157 | 0.018175 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9425 | 26.5 | 318.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.5 | 311.8896 | 311.9 | 2.6 | 2.6474 | 12.73 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.027293 | 0.027193 | 0.017787 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9501 | 26.5 | 311.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.5 | 308.8938 | 308.9 | 2.6 | 2.6 | 12.5 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.027194 | 0.026994 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.5 | 308.9
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.2 | 308.7876 | 308.8 | 2.6 | 2.6 | 12.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.026794 | 0.026594 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9072 | 26.5 | 308.8
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.1 | 308.5876 | 308.6 | 2.5 | 2.5429 | 12.5 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.026694 | 0.0264 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 29.9858 | 26.5 | 308.6
2008-05-07 | 15:13 | 0 | 0 | 0 | -3.1 | 308.4 | 308.4 | 2.5 | 2.5 | 12.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0266 | 0.026394 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8711 | 26.5 | 308.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -3 | 308.1938 | 308.2 | 2.5 | 2.5 | 12.5 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0265 | 0.026294 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 308.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -3 | 308.1938 | 308.2 | 2.5 | 2.5 | 12.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.026494 | 0.0262 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 308.2
2008-05-07 | 15:14 | 0 | -0.0008 | -8.0034 | -3 | 308.1 | 308.1 | 2.5 | 2.5 | 12.5 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.026394 | 0.0261 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8003 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.002 | -20 | -3 | 308.1 | 308.1 | 2.5 | 2.5 | 12.5 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0263 | 0.026094 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8789 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.0019 | -19.1585 | -3 | 308.1 | 308.1 | 2.5 | 2.5 | 12.5 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0262 | 0.026 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8425 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.0011 | -10.9886 | -3 | 308.106 | 308.1 | 2.5 | 2.5 | 12.5 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.406 | 0.0262 | 0.026 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8362 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -3 | 308.2865 | 308.3 | 2.5 | 2.5 | 12.5 | 12.7 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0262 | 0.0259 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0675 | 29.9148 | 26.5 | 308.3
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -3 | 308.2 | 308.2 | 2.4 | 2.4 | 12.5 | 12.6 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0261 | 0.0259 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9934 | 26.5 | 308.2
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.9 | 308.1 | 308.1 | 2.4 | 2.4 | 12.5 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.026094 | 0.0258 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9284 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.9 | 308.0938 | 308.1 | 2.4 | 2.4 | 12.4 | 12.7 | 902 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.026 | 0.025794 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8496 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.9 | 308.0062 | 308 | 2.4 | 2.4 | 12.4 | 12.7 | 902 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.026 | 0.0257 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 308
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.9 | 308.1 | 308.1 | 2.4 | 2.4 | 12.4 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.025994 | 0.0257 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.9 | 308.0938 | 308.1 | 2.4 | 2.4 | 12.4 | 12.7 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0259 | 0.025694 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.8 | 308.0938 | 308.1 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025894 | 0.0256 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8651 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.001 | -10 | -2.8 | 308.0938 | 308.1 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0258 | 0.0255 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8563 | 26.5 | 308.1
2008-05-07 | 15:14 | 0 | -0.0008 | -7.7128 | -2.7 | 308 | 308 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025794 | 0.0255 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 308
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 308 | 308 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0254 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 308
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.9 | 307.9 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0254 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 307.9
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.9 | 307.9 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0254 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 307.9
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.7938 | 307.8 | 2.4 | 2.4 | 12.43 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.025694 | 0.0254 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8738 | 26.5 | 307.8
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.7 | 307.7 | 2.4 | 2.4 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0253 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 29.9845 | 26.5 | 307.7
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.606 | 307.6 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.406 | 0.0256 | 0.0253 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9399 | 29.9059 | 26.5 | 307.6
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.6938 | 307.7 | 2.3 | 2.3 | 12.5 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0253 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9722 | 26.5 | 307.7
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.7 | 307.7 | 2.4 | 2.4 | 12.5 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.025594 | 0.025294 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8981 | 26.4491 | 307.7
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.612 | 307.6 | 2.3 | 2.3316 | 12.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.494 | 0.0255 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.4299 | 307.6
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.6938 | 307.7 | 2.3 | 2.3 | 12.5 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 307.7
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.6938 | 307.7 | 2.3 | 2.3 | 12.5 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8 | 26.5 | 307.7
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.6 | 307.6 | 2.3 | 2.3 | 12.4 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9317 | 26.5 | 307.6
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.4932 | 307.5 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0676 | 29.9112 | 26.5 | 307.5
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.4 | 307.4 | 2.3 | 2.3 | 12.4 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.025494 | 0.0252 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.823 | 26.5 | 307.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.4 | 307.4 | 2.3 | 2.3 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 307.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.3 | 307.3 | 2.3 | 2.2536 | 12.4 | 12.7 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 307.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.2 | 307.2 | 2.2 | 2.2409 | 12.4 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.2 | 307.2 | 2.2 | 2.2 | 12.4 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31.478 | 59 | 1022.7034 | 986 | 29.9 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.2 | 307.2 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.1938 | 307.2 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4062 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307 | 307 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8266 | 26.5 | 307
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.6 | 307.094 | 307.1 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.494 | 0.025394 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8 | 26.5 | 307.1
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307 | 307 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.0251 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8307 | 26.5 | 307
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.0932 | 307.1 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.025093 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8907 | 26.5 | 307.1
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.106 | 307.1 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8122 | 26.5 | 307.1
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.2 | 307.2 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.933 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.3 | 307.3 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0253 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9097 | 26.4549 | 307.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.3062 | 307.3 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.938 | 29.8 | 26.4 | 307.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.5 | 307.5 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0253 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8047 | 26.4 | 307.5
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.4062 | 307.4 | 2.1 | 2.1 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.025294 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 31.522 | 59 | 1023 | 985 | 29.9621 | 26.4 | 307.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.4 | 307.4 | 2.2 | 2.1667 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0252 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9403 | 26.4597 | 307.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.4 | 307.4 | 2.2 | 2.2 | 12.4 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0252 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9383 | 26.5 | 307.4
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.8 | 307.2062 | 307.2 | 2.1 | 2.1 | 12.4 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0252 | 0.025 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5 | 307.2
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.3 | 307.3 | 2.1 | 2.1 | 12.4 | 12.7 | 902 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0252 | 0.0249 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.5 | 307.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -2.7 | 307.3064 | 307.3 | 2.1 | 2.1 | 12.4 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0252 | 0.0249 | 0.0176 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9361 | 30 | 26.4261 | 307.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -3.1 | 297.4788 | 297.5 | 2.2 | 2.1937 | 12.26 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.025013 | 0.024844 | 0.016988 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9473 | 26.4 | 297.5
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.9 | 248.3056 | 248.3 | 2.1 | 2.1451 | 10.62 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.02235 | 0.02275 | 0.014257 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9313 | 26.4 | 248.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.6 | 237.7579 | 237.8 | 1.8 | 1.7929 | 9.63 | 12.6 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021582 | 0.021982 | 0.0136 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30 | 26.4099 | 237.8
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.7 | 245.2517 | 245.3 | 1.7 | 1.7 | 9.92 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021294 | 0.0216 | 0.014025 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30 | 26.4885 | 245.3
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.7 | 252.9624 | 253 | 1.8 | 1.8 | 10.17 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0216 | 0.0145 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30 | 26.4329 | 253
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.6 | 253.906 | 253.9 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.406 | 0.0212 | 0.0215 | 0.0145 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30 | 26.4 | 253.9
2008-05-07 | 15:14 | 0 | 0 | 0 | -5.1 | 254.8 | 254.8 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.021206 | 0.0215 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30 | 26.4243 | 254.8
2008-05-07 | 15:14 | 0 | 0 | 0 | -4.7 | 254.9 | 254.9 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0213 | 0.021494 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9972 | 26.4972 | 254.9
2008-05-07 | 15:14 | 0 | 0 | 0 | -4.7 | 254.9 | 254.9 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0213 | 0.0214 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9185 | 26.4185 | 254.9
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.6 | 254.9 | 254.9 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0213 | 0.0214 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9 | 26.4602 | 254.9
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.6 | 255 | 255 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0214 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9 | 26.5 | 255
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.5 | 255.0062 | 255 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0214 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.938 | 29.9 | 26.4826 | 255
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.4 | 255.1 | 255.1 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0213 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.938 | 29.9 | 26.4045 | 255.1
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.4 | 255.2 | 255.2 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0213 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.8255 | 26.4745 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.3 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0213 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9067 | 26.4467 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.3 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0213 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.968 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.2 | 255.3062 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 902 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.021194 | 0.021294 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.2 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0211 | 0.0212 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -4.1 | 255.2938 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 902 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0211 | 0.0212 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.0357 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -4 | 255.2 | 255.2 | 1.8 | 1.8 | 10.3 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0211 | 0.0212 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0535 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -4 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.7 | 902 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4938 | 0.0211 | 0.0212 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -4 | 255.2 | 255.2 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0212 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9932 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.9 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.7 | 902 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 29.8355 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.0004 | -3.7031 | -3.8 | 255.3062 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.8 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.8 | 255.4062 | 255.4 | 1.8 | 1.8 | 10.3 | 12.7 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.8396 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.7 | 255.5 | 255.5 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 29.9365 | 26.4182 | 255.5
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.7 | 255.4938 | 255.5 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0937 | 26.4969 | 255.5
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.7 | 255.4 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4245 | 255.4
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.7 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0211 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.7 | 255.3938 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021094 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0673 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | -0.0008 | -7.7877 | -3.6 | 255.4938 | 255.5 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0112 | 26.4 | 255.5
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.6 | 255.4 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0899 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.6 | 255.4 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | -0.0007 | -7.1002 | -3.6 | 255.4062 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0529 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | -0.0004 | -4.4502 | -3.6 | 255.4 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4938 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.6 | 255.3938 | 255.4 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.6 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.7 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.6 | 255.394 | 255.4 | 1.8 | 1.8 | 10.3 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.494 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.0604 | 30.0616 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.5 | 255.3062 | 255.3 | 1.8 | 1.8 | 10.3 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 30.0599 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.5 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0185 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.4 | 255.2938 | 255.3 | 1.8 | 1.8 | 10.3 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0974 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.4 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.176 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.4 | 255.3938 | 255.4 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.021 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0907 | 26.4 | 255.4
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.3 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0333 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -3.3 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.0007 | -7.0887 | -3.3 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.2 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.2 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0522 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.2 | 255.2938 | 255.3 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0264 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.1 | 255.2 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.1 | 255.1938 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3.1 | 255.1938 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0209 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.4623 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -3 | 255.1 | 255.1 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.020894 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0594 | 26.4594 | 255.1
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255 | 255 | 1.8 | 1.8 | 10.3 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.019 | 26.4 | 255
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.0062 | 255 | 1.8 | 1.8 | 10.3 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.020994 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0981 | 26.4 | 255
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.0232 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.8 | 901 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.2 | 255.2 | 1.8 | 1.8 | 10.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0678 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.2062 | 255.2 | 1.8 | 1.8 | 10.3 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1749 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.9 | 255.3 | 255.3 | 1.8 | 1.8 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0179 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.8 | 255.2938 | 255.3 | 1.8 | 1.843 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.0698 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.8 | 255.2062 | 255.2 | 1.9 | 1.8701 | 10.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.1 | 26.4 | 255.2
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.8 | 255.3 | 255.3 | 1.9 | 1.9 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.0146 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4 | 255.3
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.8 | 254.928 | 254.9 | 1.9 | 1.9 | 10.3 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0208 | 0.014581 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1057 | 26.4 | 254.9
2008-05-07 | 15:15 | 0 | -0.001 | -10 | -2.8 | 232.3606 | 232.4 | 1.9 | 1.9 | 9.84 | 12.6 | 902 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.020794 | 0.020694 | 0.013269 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.1843 | 26.4 | 232.4
2008-05-07 | 15:15 | 0 | -0.0006 | -6.2282 | -2.9 | 211.9886 | 212 | 1.7 | 1.7453 | 8.95 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.019881 | 0.019881 | 0.012188 | 0.0001 |  | n/a |  | 32 | 59 | 1021.8349 | 985 | 30.137 | 26.4 | 212
2008-05-07 | 15:15 | 0 | 0 | 0 | -2.1 | 207.0566 | 207.1 | 1.6 | 1.6187 | 8.53 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.019 | 0.018894 | 0.011894 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4 | 207.1
2008-05-07 | 15:15 | 0 | 0 | 0 | -2.1 | 206.2938 | 206.3 | 1.6 | 1.6 | 8.39 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.018894 | 0.018794 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1404 | 26.4 | 206.3
2008-05-07 | 15:15 | 0 | 0 | 0 | -2 | 206.2 | 206.2 | 1.6 | 1.6 | 8.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0188 | 0.018694 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2976 | 26.4 | 206.2
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.9 | 206.0579 | 206.1 | 1.6 | 1.6 | 8.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.018794 | 0.018594 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.4 | 206.1
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.8 | 205.2938 | 205.3 | 1.5 | 1.5434 | 8.3 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.018694 | 0.018494 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2439 | 26.4 | 205.3
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.8 | 205.1938 | 205.2 | 1.5 | 1.5 | 8.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018594 | 0.018394 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1306 | 26.4 | 205.2
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.8 | 205.1 | 205.1 | 1.5 | 1.5 | 8.3 | 12.7 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018494 | 0.0183 | 0.0118 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0133 | 26.4 | 205.1
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.7 | 205.1 | 205.1 | 1.5 | 1.5 | 8.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0184 | 0.0182 | 0.011794 | 0.0001 |  | n/a |  | 32 | 59 | 1021.1521 | 985 | 30.0919 | 26.4 | 205.1
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.7 | 205.0938 | 205.1 | 1.5 | 1.5 | 8.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0184 | 0.0182 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0294 | 26.4 | 205.1
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.6 | 205 | 205 | 1.6 | 1.6 | 8.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0183 | 0.0181 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.4 | 205
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.6 | 205 | 205 | 1.6 | 1.6 | 8.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0182 | 0.018 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0275 | 26.4 | 205
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.9 | 204.9 | 1.6 | 1.6 | 8.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018194 | 0.0179 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1064 | 26.4 | 204.9
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.8938 | 204.9 | 1.6 | 1.5554 | 8.3 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0181 | 0.017894 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1851 | 26.4 | 204.9
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.8 | 204.8 | 1.5 | 1.548 | 8.3 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.018094 | 0.0178 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.8 | 204.8 | 1.5 | 1.5395 | 8.3 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.018 | 0.0178 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1578 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.7062 | 204.7 | 1.6 | 1.5558 | 8.3 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.018 | 0.017794 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 30.1209 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.8 | 204.8 | 1.5 | 1.5364 | 8.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4938 | 0.017994 | 0.0177 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1995 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.8 | 204.8 | 1.6 | 1.5519 | 8.3 | 12.6 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0179 | 0.0177 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1219 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.8062 | 204.8 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0179 | 0.0176 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.9 | 204.9 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0179 | 0.0176 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.4 | 204.9
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.9 | 204.9 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.017894 | 0.0176 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1281 | 26.4 | 204.9
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.8 | 204.8 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0178 | 0.0176 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2854 | 26.4 | 204.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.494 | 0.0178 | 0.0175 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2287 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4062 | 0.017794 | 0.0175 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1501 | 26.4 | 204.7
2008-05-07 | 15:16 | -0.009 | 0 | 0 | -1.5 | 204.706 | 204.7 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0177 | 0.0175 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1571 | 26.4 | 204.7
2008-05-07 | 15:16 | -0.003 | 0 | 0 | -1.4 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0177 | 0.017494 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3072 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.7 | 204.7 | 1.5 | 1.5444 | 8.3 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0177 | 0.0174 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3858 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.7 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.017694 | 0.0174 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2721 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0176 | 0.0174 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2853 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.7 | 204.7 | 1.5 | 1.5 | 8.3 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0176 | 0.0174 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3783 | 26.4 | 204.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.6 | 204.6 | 1.6 | 1.6 | 8.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0176 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2991 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.8 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0176 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0632 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0176 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1726 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.5 | 204.4 | 204.4 | 1.6 | 1.6 | 8.2 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 192 | 6 | 195 | 59.4 | 0.017594 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2638 | 26.4 | 204.4
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.6 | 204.4 | 204.4 | 1.6 | 1.6 | 8.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.0175 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2148 | 26.4 | 204.4
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.6 | 204.4 | 204.4 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.0175 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2934 | 26.4 | 204.4
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.6 | 204.4 | 204.4 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 192 | 6 | 195 | 59.4 | 0.0175 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1565 | 26.4 | 204.4
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.7 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 192 | 6 | 195 | 59.4 | 0.0175 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2513 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.7 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 192 | 6 | 195 | 59.4 | 0.0175 | 0.0173 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3415 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.7 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.017294 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2159 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10.4254 | -1.8 | 204.5 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3731 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.002 | -20 | -1.8 | 204.5062 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2697 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.002 | -20 | -1.8 | 204.4938 | 204.5 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1562 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.002 | -20 | -1.8 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4932 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 984.9324 | 30.1224 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.002 | -20 | -1.8 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 984 | 30.201 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.002 | -20 | -1.8 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 984 | 30.2795 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.0015 | -14.7568 | -1.8 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0174 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 984 | 30.3 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0172 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2631 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.6 | 204.6 | 1.6 | 1.6 | 8.2 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.231 | 26.4 | 204.6
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.5 | 204.5 | 1.7 | 1.6549 | 8.2 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3884 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.4938 | 204.5 | 1.7 | 1.7 | 8.27 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1106 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.5 | 204.5 | 1.7 | 1.7 | 8.3 | 12.8 | 901 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0511 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.5 | 204.5 | 1.7 | 1.7 | 8.3 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.7 | 204.5 | 204.5 | 1.7 | 1.7 | 8.3 | 12.7 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0173 | 0.0171 | 0.0117 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0914 | 26.4 | 204.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.6 | 196.8277 | 196.8 | 1.7 | 1.7 | 8.19 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.017276 | 0.017076 | 0.011234 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.0129 | 26.4 | 196.8
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.2 | 167.386 | 167.4 | 1.7 | 1.7 | 7.27 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.016476 | 0.016282 | 0.009576 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1972 | 26.4 | 167.4
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.4 | 157.4884 | 157.5 | 1.6 | 1.6469 | 6.48 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.015781 | 0.015581 | 0.008994 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2556 | 26.4 | 157.5
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.2 | 155.6442 | 155.6 | 1.5 | 1.5 | 6.3 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.015375 | 0.015181 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1769 | 26.4 | 155.6
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.5 | 154.8 | 154.8 | 1.5 | 1.5 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0149 | 0.0148 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.1035 | 26.4 | 154.8
2008-05-07 | 15:16 | 0 | -0.001 | -10 | -1.5 | 154.8 | 154.8 | 1.5 | 1.4554 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.014794 | 0.014694 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985.9375 | 30.2603 | 26.4 | 154.8
2008-05-07 | 15:16 | 0 | -0.0008 | -8.1572 | -1.4 | 154.8 | 154.8 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.014694 | 0.014594 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3589 | 26.4 | 154.8
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 154.706 | 154.7 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0146 | 0.014494 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3624 | 26.4 | 154.7
2008-05-07 | 15:16 | 0 | 0 | 0 | -1.4 | 154.7 | 154.7 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0145 | 0.0144 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3162 | 26.4 | 154.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.4 | 154.6 | 154.6 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.014494 | 0.0143 | 0.0089 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3948 | 26.4 | 154.6
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.5 | 154.5938 | 154.6 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0144 | 0.014294 | 0.008894 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1797 | 26.4 | 154.6
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.5 | 154.4938 | 154.5 | 1.4 | 1.4 | 6.2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0143 | 0.0142 | 0.0088 | 0.0001 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2041 | 26.4 | 154.5
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.4 | 1.4 | 6.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.014294 | 0.014194 | 0.0088 | 0.0001 |  | n/a |  | 32 | 59 | 1021.5464 | 985 | 30.3 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.4 | 1.4276 | 6.2 | 12.7 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0142 | 0.0141 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.5 | 1.5 | 6.2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0141 | 0.014 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.5 | 1.5 | 6.2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0141 | 0.014 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.5 | 1.5 | 6.2 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0141 | 0.013994 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3452 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.4 | 154.4 | 1.5 | 1.5 | 6.2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.014 | 0.0139 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3762 | 26.4 | 154.4
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.6 | 154.3 | 154.3 | 1.5 | 1.5308 | 6.2 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.014 | 0.0139 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.3
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 154.3 | 154.3 | 1.6 | 1.6 | 6.2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0139 | 0.0138 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.3
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 154.3 | 154.3 | 1.6 | 1.6 | 6.2 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0139 | 0.0138 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.4 | 154.3
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 154.2938 | 154.3 | 1.7 | 1.6539 | 6.2 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4938 | 0.013894 | 0.0138 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2617 | 26.4 | 154.3
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 154.2062 | 154.2 | 1.7 | 1.7 | 6.2 | 12.7 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0138 | 0.0137 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021.4077 | 985 | 30.2505 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | -0.0003 | -3.4638 | -1.7 | 154.2062 | 154.2 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0138 | 0.0137 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4852 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | -0.0008 | -8.178 | -1.7 | 154.2 | 154.2 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0137 | 0.0137 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3521 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154.2 | 154.2 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0137 | 0.013694 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154.2 | 154.2 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0137 | 0.0136 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3314 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.9 | 154.2 | 154.2 | 1.8 | 1.7537 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0137 | 0.0136 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3901 | 26.4 | 154.2
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.9 | 154.1 | 154.1 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.013694 | 0.0136 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3122 | 26.4 | 154.1
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.9 | 154.1 | 154.1 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0136 | 0.0136 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.4 | 154.1
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154.0938 | 154.1 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0136 | 0.0135 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3917 | 26.4 | 154.1
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154.0938 | 154.1 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.013594 | 0.0135 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4512 | 26.4 | 154.1
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154 | 154 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0135 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3058 | 26.4 | 154
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154 | 154 | 1.8 | 1.8 | 6.2 | 12.6 | 904 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0135 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4633 | 26.4 | 154
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154 | 154 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0134 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3794 | 26.4 | 154
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.8 | 154 | 154 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0134 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1021.0928 | 985 | 30.3389 | 26.4 | 154
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 153.9938 | 154 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0134 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3825 | 26.4 | 154
2008-05-07 | 15:17 | 0 | 0 | -0.29167 | -1.7 | 153.9 | 153.9 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0134 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3044 | 26.4 | 153.9
2008-05-07 | 15:17 | 0 | -0.0009 | -8.625 | -1.7 | 153.8 | 153.8 | 1.7 | 1.7 | 6.2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0135 | 0.0134 | 0.0088 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4494 | 26.4 | 153.8
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.7 | 152.9 | 152.9 | 1.7 | 1.7 | 6.2 | 12.8 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4062 | 0.013494 | 0.0134 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4465 | 26.4 | 152.9
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.7 | 152.8 | 152.8 | 1.7 | 1.7 | 6.2 | 12.8 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.013394 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4321 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | -0.0006 | -6.4897 | -1.7 | 152.8 | 152.8 | 1.7 | 1.7 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | -0.0002 | -2.005 | -1.7 | 152.7938 | 152.8 | 1.7 | 1.7 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.5 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3142 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2001 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.0005 | -4.7635 | -1.6 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4931 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.0004 | -3.5997 | -1.7 | 152.6938 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0134 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3353 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.0008 | -8.2308 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.013394 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.6062 | 152.6 | 1.6 | 1.6 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3792 | 26.4 | 152.6
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0133 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4818 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.7 | 152.7 | 1.6 | 1.6 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4031 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.7 | 152.7 | 1.7 | 1.7 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.9381 | 30.4754 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.7 | 152.7 | 1.7 | 1.7 | 6.2 | 12.9 | 901 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.392 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.8 | 152.8 | 1.7 | 1.7 | 6.2 | 12.8 | 902 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3327 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.8 | 152.8 | 1.7 | 1.7 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3886 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.8 | 152.8 | 1.7 | 1.7 | 6.2 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3101 | 26.4 | 152.8
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.7 | 152.7062 | 152.7 | 1.7 | 1.7 | 6.2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0133 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3686 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.0005 | -5.1977 | -1.8 | 152.7 | 152.7 | 1.7 | 1.7 | 6.2 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.013294 | 0.0132 | 0.0087 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4 | 26.4 | 152.7
2008-05-07 | 15:17 | 0 | -0.001 | -10 | -1.8 | 152.204 | 152.2 | 1.8 | 1.8 | 6.2 | 12.7 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0132 | 0.0132 | 0.008675 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.4 | 26.4 | 152.2
2008-05-07 | 15:17 | 0 | -0.0008 | -8.2464 | -1.5 | 132.774 | 132.8 | 1.8 | 1.8 | 5.78 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.013157 | 0.013057 | 0.007619 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.4044 | 26.4 | 132.8
2008-05-07 | 15:17 | 0 | 0 | 0 | -1.5 | 104.5 | 104.5 | 1.7 | 1.7242 | 4.58 | 12.6 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.012082 | 0.012076 | 0.006 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.0601 | 30.483 | 26.4 | 104.5
2008-05-07 | 15:17 | 0 | 0 | 0 | -1 | 103.5938 | 103.6 | 1.6 | 1.6 | 4.24 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.011494 | 0.011394 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3767 | 26.4 | 103.6
2008-05-07 | 15:17 | 0 | 0 | 0 | -1 | 103.4938 | 103.5 | 1.5 | 1.5205 | 4.2 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.011394 | 0.011294 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.3401 | 26.4 | 103.5
2008-05-07 | 15:17 | 0.002 | 0 | 0 | -1 | 103.4 | 103.4 | 1.5 | 1.5 | 4.2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0112 | 0.011194 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.4 | 103.4
2008-05-07 | 15:17 | 0.01 | 0 | 0 | -1 | 103.4 | 103.4 | 1.5 | 1.5 | 4.19 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0111 | 0.011 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.4 | 103.4
2008-05-07 | 15:17 | 0.01 | 0 | 0 | -0.9 | 103.4 | 103.4 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.010994 | 0.010894 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4762 | 26.4 | 103.4
2008-05-07 | 15:17 | 0.01 | 0 | 0 | -0.9 | 103.2942 | 103.3 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.010894 | 0.010794 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.0582 | 30.3905 | 26.4 | 103.3
2008-05-07 | 15:17 | 0.01 | 0 | 0 | -1.1 | 103.2 | 103.2 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0107 | 0.0059 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3661 | 26.4 | 103.2
2008-05-07 | 15:18 | 0.01 | 0 | 0 | -1.4 | 102.3938 | 102.4 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010694 | 0.0106 | 0.005894 | 0 |  | n/a |  | 32 | 59 | 1021.919 | 985 | 30.4887 | 26.4 | 102.4
2008-05-07 | 15:18 | 0.01 | 0 | 0 | -1.4 | 102.2 | 102.2 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010594 | 0.010594 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4095 | 26.4 | 102.2
2008-05-07 | 15:18 | 0.01 | -0.0003 | -3.3444 | -1.4 | 102.1938 | 102.2 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010494 | 0.0104 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.0622 | 30.4692 | 26.4 | 102.2
2008-05-07 | 15:18 | 0.001 | -0.001 | -10 | -1.4 | 102.0938 | 102.1 | 1.6 | 1.6 | 4.1 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | 0.0104 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.5 | 26.4 | 102.1
2008-05-07 | 15:18 | 0 | -0.001 | -9.7164 | -1.3 | 102 | 102 | 1.5 | 1.532 | 4.1 | 12.6 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010394 | 0.0103 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4473 | 26.4 | 102
2008-05-07 | 15:18 | 0 | -0.0002 | -1.6126 | -1.3 | 102 | 102 | 1.5 | 1.5 | 4.1 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0103 | 0.0103 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.3151 | 26.4 | 102
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.4 | 101.9938 | 102 | 1.6 | 1.5714 | 4.1 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0103 | 0.0103 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 986 | 30.5511 | 26.4 | 102
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.9 | 101.9 | 1.6 | 1.5557 | 4.1 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0102 | 0.0102 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4 | 101.9
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8932 | 101.9 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0101 | 0.0102 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4 | 101.9
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8938 | 101.9 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0101 | 0.010194 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5809 | 26.4 | 101.9
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8 | 101.8 | 1.6 | 1.6 | 4.1 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.01 | 0.0101 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5019 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8 | 101.8 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.01 | 0.0101 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6535 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8 | 101.8 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.01 | 0.01 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6447 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8 | 101.8 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0099 | 0.01 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.566 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.8 | 101.8 | 1.6 | 1.6 | 4.1 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0099 | 0.009994 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5251 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.7938 | 101.8 | 1.5 | 1.5307 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.009894 | 0.0099 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6822 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.7 | 101.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0098 | 0.0099 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0098 | 0.0098 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.5 | 1.5 | 4.1 | 12.6 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0098 | 0.0098 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.8 | 101.8 | 1.4 | 1.4 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0097 | 0.0098 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6943 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.8 | 101.8 | 1.4 | 1.4 | 4.1 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.3938 | 0.0097 | 0.0098 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6157 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7062 | 101.7 | 1.3 | 1.3469 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0097 | 0.0097 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.8 | 101.8 | 1.3 | 1.3 | 4.1 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0097 | 0.0097 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.8 | 101.8 | 1.3 | 1.3 | 4.1 | 12.8 | 902 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0096 | 0.0097 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.2 | 1.2 | 4.1 | 12.8 | 902 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0096 | 0.0097 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.2 | 1.2 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0096 | 0.009694 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.6773 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.2 | 1.2 | 4.1 | 12.7 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0096 | 0.0096 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.6443 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.7 | 101.7 | 1.1 | 1.144 | 4.1 | 12.8 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.009594 | 0.0096 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.6343 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.6736 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5164 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.7062 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4511 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.6 | 101.8 | 101.8 | 1.1 | 1.0543 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4826 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.794 | 101.8 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0095 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7 | 26.4 | 101.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.009494 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.574 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1.1 | 1.1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0095 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.462 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1 | 1.0442 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.009494 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6979 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 101.7 | 101.7 | 1 | 1 | 4.1 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.7 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | 0 | -0.29118 | -1.7 | 101.7 | 101.7 | 1 | 1 | 4.1 | 12.9 | 901 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6439 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | -0.0009 | -8.6106 | -1.7 | 101.7 | 101.7 | 1 | 1.0475 | 4.1 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6347 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | -0.001 | -10 | -1.7 | 101.7 | 101.7 | 1 | 1 | 4.1 | 12.9 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6863 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | -0.001 | -10 | -1.7 | 101.6938 | 101.7 | 1 | 1 | 4.1 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0094 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6077 | 26.4 | 101.7
2008-05-07 | 15:18 | 0 | -0.001 | -10 | -1.7 | 101.6 | 101.6 | 1 | 1 | 4.1 | 12.9 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.009394 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.6
2008-05-07 | 15:18 | 0 | -0.001 | -10 | -1.7 | 101.6 | 101.6 | 1 | 1 | 4.1 | 12.7 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.6 | 26.4 | 101.6
2008-05-07 | 15:18 | 0 | -0.001 | -9.7164 | -1.8 | 101.6 | 101.6 | 1 | 1 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5719 | 26.4 | 101.6
2008-05-07 | 15:18 | 0 | -0.0001 | -1.3761 | -1.8 | 101.6 | 101.6 | 1 | 1 | 4.1 | 12.6 | 904 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 984 | 30.5068 | 26.4 | 101.6
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.5 | 101.5 | 0.9 | 0.9296 | 4.1 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5853 | 26.4 | 101.5
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.5 | 101.5 | 0.9 | 0.9467 | 4.1 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0094 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1022 | 984.0601 | 30.6 | 26.4 | 101.5
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.8 | 101.5 | 101.5 | 0.9 | 0.9 | 4.1 | 12.6 | 905 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0093 | 0.0058 | 0 |  | n/a |  | 32 | 59 | 1021.919 | 985 | 30.6 | 26.4 | 101.5
2008-05-07 | 15:18 | 0 | 0 | 0 | -1.7 | 99.383 | 99.4 | 0.9 | 0.9 | 4.1 | 12.6 | 904 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0093 | 0.0093 | 0.005688 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6211 | 26.4 | 99.4
2008-05-07 | 15:18 | 0 | 0 | 0 | -3.4 | 55.7894 | 55.8 | 0.9 | 0.9 | 2.74 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.007532 | 0.008026 | 0.003238 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6997 | 26.4 | 55.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.2 | 27 | 27 | 0.6 | 0.5976 | 1.45 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005894 | 0.006194 | 0.001506 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7 | 26.4 | 27
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.3 | 34.396 | 34.4 | 0.3 | 0.3 | 1.61 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0058 | 0.0061 | 0.002019 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5866 | 26.4 | 34.4
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.4 | 45.7922 | 45.8 | 0.5 | 0.4563 | 1.9 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3938 | 0.0057 | 0.006 | 0.002612 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6058 | 26.4 | 45.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.4 | 48.7 | 48.7 | 0.5 | 0.5 | 2 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0057 | 0.006 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7859 | 26.4 | 48.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.4 | 48.7062 | 48.7 | 0.5 | 0.5 | 2 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0057 | 0.005994 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7072 | 26.4 | 48.7
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.4 | 48.8062 | 48.8 | 0.5 | 0.5 | 2 | 12.6 | 903 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0056 | 0.0059 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7713 | 26.4 | 48.8
2008-05-07 | 15:18 | 0 | 0 | 0 | -2.3 | 48.9 | 48.9 | 0.5 | 0.5 | 2 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0056 | 0.0059 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7501 | 26.4 | 48.9
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.3 | 48.9062 | 48.9 | 0.4 | 0.4477 | 2 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0056 | 0.0059 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6429 | 26.4 | 48.9
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.3 | 49.0062 | 49 | 0.4 | 0.4 | 2 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0056 | 0.0059 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5143 | 26.4 | 49
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.3 | 49.1 | 49.1 | 0.4 | 0.4 | 2 | 12.6 | 903 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0056 | 0.0058 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6714 | 26.4 | 49.1
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.3 | 49.1 | 49.1 | 0.4 | 0.4 | 2 | 12.6 | 903 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0055 | 0.0058 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6362 | 26.4 | 49.1
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.2 | 49.1 | 49.1 | 0.5 | 0.5 | 2 | 12.7 | 903 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0055 | 0.0058 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7273 | 26.4 | 49.1
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.2 | 49.2 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0055 | 0.0057 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.8571 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.2 | 49.2 | 49.2 | 0.4 | 0.4447 | 2 | 12.7 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0055 | 0.0057 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7001 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.2 | 49.2 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005494 | 0.0057 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.7787 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.2 | 49.2 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0057 | 0.0028 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.7427 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0057 | 0.0028 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.7 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2 | 49.2 | 0.5 | 0.4688 | 2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.005694 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6855 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2 | 49.2 | 0.4 | 0.449 | 2 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.6069 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2062 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5285 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2938 | 49.3 | 0.4 | 0.4448 | 2 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.4498 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0054 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.4577 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2.1 | 49.2062 | 49.2 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005394 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5851 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.4 | 0.4 | 2 | 12.7 | 903 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0056 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.428 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.4 | 0.4 | 2 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4643 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.4 | 0.4 | 2 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.457 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.4 | 0.4 | 2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4218 | 26.4218 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.4 | 0.4 | 2 | 12.8 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4996 | 26.4996 | 49.3
2008-05-07 | 15:19 | 0 | 0 | 0 | -2 | 49.3 | 49.3 | 0.3 | 0.3324 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.0601 | 30.421 | 26.421 | 49.3
2008-05-07 | 15:19 | 0 | 0 | -0.042517 | -2 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.4576 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.0009 | -8.5459 | -1.9 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.5 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3938 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 31.5215 | 59 | 1022 | 986 | 30.5 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0053 | 0.0055 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.2 | 49.2 | 0.3 | 0.3 | 2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.005494 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4 | 49.2
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 900 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31.478 | 59 | 1022 | 986 | 30.4708 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.3922 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.3 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.938 | 30.3137 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.3062 | 49.3 | 0.3 | 0.3 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3 | 26.4 | 49.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.4 | 49.4 | 0.3 | 0.3446 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3868 | 26.4 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.4 | 49.4 | 0.4 | 0.3787 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5 | 26.4 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5 | 26.4006 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.7 | 902 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5 | 26.4792 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.7 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4422 | 26.4422 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.9 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.7 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0054 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.4 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.005394 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.4151 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.8 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0052 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.4936 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.7 | 49.4 | 49.4 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.3281 | 26.4281 | 49.4
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.7 | 49.5 | 49.5 | 0.4 | 0.4 | 2 | 12.6 | 904 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.4011 | 26.4506 | 49.5
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.7 | 49.5 | 49.5 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5 | 26.5 | 49.5
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.7 | 49.5 | 49.5 | 0.4 | 0.4 | 2 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5 | 26.5 | 49.5
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.7 | 49.5 | 49.5 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.5 | 26.5 | 49.5
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 905 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.4349 | 26.5 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0053 | 0.0028 | 0 |  | n/a |  | 31 | 59 | 1021 | 985 | 30.4 | 26.5 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 905 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.005294 | 0.0028 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.3777 | 26.4777 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0052 | 0.0028 | 0 |  | n/a |  | 31.478 | 59 | 1021 | 985 | 30.3019 | 26.4 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 905 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0052 | 0.0028 | 0 |  | n/a |  | 31.522 | 59 | 1021 | 985 | 30.4584 | 26.4 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.6062 | 49.6 | 0.4 | 0.4 | 2 | 12.6 | 904 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0052 | 0.002806 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4421 | 26.4579 | 49.6
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.7062 | 49.7 | 0.4 | 0.4 | 2 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0052 | 0.0029 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.5 | 49.7
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.6 | 49.7938 | 49.8 | 0.4 | 0.4 | 2 | 12.6 | 904 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | 0.0052 | 0.0029 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.4847 | 49.8
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.5 | 33.9 | 33.9 | 0.4 | 0.4 | 1.76 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | 0.0052 | 0.002 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.4061 | 33.9
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -1.3 | 23.2976 | 23.3 | 0.2 | 0.229 | 0.82 | 12.6 | 903 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195.2 | 59.4 | 0.004857 | 0.005057 | 0.001344 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5447 | 26.4724 | 23.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -0.9 | 4.876 | 4.9 | 0 | 0 | 0.36 | 12.6 | 903 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195.6 | 59.4 | 0.003881 | 0.003981 | 0.000294 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.4472 | 26.5 | 4.9
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -0.7 | 2.8519 | 2.9 | -0.2 | 0 | 0.11 | 12.6 | 903 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195.9 | 59.4 | 0.003494 | 0.003594 | 0.000194 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3887 | 26.4704 | 2.9
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -0.9 | 1.2938 | 1.3 | -0.2 | 0 | 0.1 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195.7 | 59.4 | 0.003294 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4082 | 1.3
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -0.9 | 1.1 | 1.1 | -0.3 | 0 | 0.1 | 12.6 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195.4 | 59.4 | 0.003194 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.6 | 26.4868 | 1.1
2008-05-07 | 15:19 | 0 | -0.001 | -10 | -0.9 | 1.1 | 1.1 | -0.3 | 0 | 0.1 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.003094 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4692 | 26.5 | 1.1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.9 | 1.1 | 1.1 | -0.3 | 0 | 0.1 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.003 | 0.003194 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.356 | 26.5 | 1.1
2008-05-07 | 15:20 | 0 | -0.0005 | -5.0735 | -0.8 | 1.0938 | 1.1 | -0.3 | 0 | 0.1 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0029 | 0.003 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.5 | 1.1
2008-05-07 | 15:20 | 0 | -0.0003 | -3.1348 | -0.7 | 1 | 1 | -0.4 | 0 | 0.1 | 12.6 | 903 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002794 | 0.002894 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3012 | 26.5 | 1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.6 | 0.9 | 0.9 | -0.3 | 0 | 0 | 12.7 | 903 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002694 | 0.002794 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3794 | 26.5 | 0.9
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.6 | 0.8938 | 0.9 | -0.4 | 0 | 0 | 12.6 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.002594 | 0.002694 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1021 | 985.062 | 30.4 | 26.5 | 0.9
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.5 | 0.8 | 0.8 | -0.4 | 0 | 0 | 12.7 | 903 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0025 | 0.0026 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3631 | 26.5 | 0.8
2008-05-07 | 15:20 | 0 | -0.0003 | -3.4107 | -0.5 | 0.7504 | 0.8 | -0.4 | 0 | 0 | 12.6 | 903 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.002494 | 0.002594 | 9.4e-05 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3155 | 26.5 | 0.8
2008-05-07 | 15:20 | 0 | 0 | 0 | -0.4 | -0.1 | 0 | -0.4 | 0 | 0 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0024 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3941 | 26.5 | -0.1
2008-05-07 | 15:20 | 0 | 0 | 0 | -0.4 | -0.1 | 0 | -0.5 | 0 | 0 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0023 | 0.0024 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3278 | 26.5 | -0.1
2008-05-07 | 15:20 | 0 | -0.0002 | -1.7308 | -0.3 | -0.1 | 0 | -0.5 | 0 | 0 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0022 | 0.0023 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4019 | 26.5 | -0.1
2008-05-07 | 15:20 | 0 | -0.001 | -9.7436 | -0.3 | -0.1 | 0 | -0.5 | 0 | 0 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0022 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4701 | 26.5 | -0.1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.1 | 0 | -0.5 | 0 | 0 | 12.8 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.002194 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3915 | 26.4915 | -0.1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.1 | 0 | -0.5 | 0 | 0 | 12.9 | 901 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0021 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.313 | 26.413 | -0.1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -0.5 | 0 | 0 | 12.8 | 902 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0021 | 0.002194 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3656 | 26.4656 | -0.1
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.002094 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.3115 | 26.4558 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.002 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2228 | 26.4228 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 901 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.002 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.5 | -0.2 | 0 | -0.6 | 0 | 0 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0019 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.062 | 30.3 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 902 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0019 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.938 | 30.3586 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0019 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4372 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 901 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001894 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.5 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.8 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0018 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.5 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 902 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3062 | 0.0018 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.427 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.0002 | -1.6464 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.9 | 901 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3938 | 0.0018 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | 0 | 0 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.7 | 903 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3062 | 0.0018 | 0.001894 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3699 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | 0 | 0 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.7 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001794 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2826 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | 0 | 0 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1259 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.0001 | -1.4126 | -0.4 | -0.2 | 0 | -0.6 | 0 | 0 | 12.7 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2316 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0017 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.7 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3462 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4948 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.7 | 903 | 936 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2607 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3938 | 0.0016 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1414 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 903 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0016 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1748 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.001594 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2679 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1108 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 903 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2554 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 903 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3938 | 0.0015 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3696 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0015 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.291 | 26.5 | -0.2
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.3938 | 0.0015 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2125 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 905 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2662 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3448 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3533 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.3 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2056 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.4408 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 905 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3821 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.3 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.3 | 0 | -0.7 | 0 | 0 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.2673 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.4 | -0.3 | 0 | -0.8 | 0 | 0 | 12.6 | 904 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0014 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.1102 | 26.5 | -0.3
2008-05-07 | 15:20 | 0 | -0.001 | -10 | -0.5 | -0.3 | 0 | -0.8 | 0 | 0 | 12.6 | 904 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.001394 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1021 | 985 | 30.541 | 26.5 | -0.3
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-05-07
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 26374
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 2004
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 210
Version | 1.611
NO Span(ppm) | 298
NO2 Span(ppm) | 505
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.12
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | CAT TH83
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 4
Torque (ecm or calc) | ecm
Use Frictional Torque | NO
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 14:59
Test End Time | 15:20
Test Duration (s) | 1262
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 976.199
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
iVEH_SPEED | iSCB_EAI1_XF
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/07/2008 14:57:34.427 - None Found
0X0000 - 05/07/2008 14:59:19.755 - None Found
Warnings:
0X0000 - 05/07/2008 14:57:34.428 - None Found
0X0000 - 05/07/2008 14:59:19.756 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1262
Brake Specific Dropout Limit Count | 1262
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 | iVEH_SPEED | Vehicle Speed (Dyno) | MPH | 1 | x^0: 0; x^1: 6
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Span
InfoVer | 1
Date | 2008-05-07
Time | 14:58
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 200 | 6 | 20.9 | 2004 | x^0: 298; x^1: 505; x^2: 0; x^3: 93.12
Gas | Previous | Current | Difference
NO2(ppm) | 506.9937 | 505.1467 | 1.8471
Test Information:
SEMTECH_DATA_FILE | 2008-05-07 | 14:56 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
No Information Available

## NO2 Linearity
Test Date:
Test Time:
5/7/2008
14:59:20
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
210

## NDUV
Test Date:
Test Time:
5/7/2008
14:59:20
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
210

## O2 Linearity
Test Date:
Test Time:
5/7/2008
14:59:20
SEMTECH Serial Number:
G05-SDS02

## O2
Test Date:
Test Time:
5/7/2008
14:59:20
SEMTECH Serial Number:
G05-SDS02