## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 321.8 | Scott gas | ALM054255 | ±1% | 2010-06-16 | 965.4
1200 ppm CO | 1205
12% CO2 | 12
1500 ppm NO | 1505
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.12 | Scott gas | ALM004744 | ±1% | 2010-06-04 | 93.36
200 ppm CO | 200.7
6% CO2 | 6.009
300 ppm NO | 297.2
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | NO2 NDUV Serial Number: |  |  | 208 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-07-16 |  | 505 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 13:33
Test software: | 10.09 SP1
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0.4 | 0.4 |  |  | 0 | 0.1 | 0.1
90 | 105 | 1 | 50.5 | 50.4 | -0.1 | -0.2 |  | 2.09 | 2.1 | 0.01 | 0.48
150 | 165 | 2 | 101 | 99.9 | -1.1 | -1.09 |  | 4.18 | 4.2 | 0.02 | 0.48
210 | 225 | 3 | 151.5 | 151.3 | -0.2 | -0.13 |  | 6.27 | 6.2 | -0.07 | -1.12
270 | 285 | 4 | 202 | 203 | 1 | 0.5 |  | 8.36 | 8.3 | -0.06 | -0.72
330 | 345 | 5 | 252.5 | 254.6 | 2.1 | 0.83 |  | 10.45 | 10.4 | -0.05 | -0.48
390 | 405 | 6 | 303 | 306.3 | 3.3 | 1.09 |  | 12.54 | 12.45 | -0.09 | -0.72
450 | 465 | 7 | 353.5 | 357.1 | 3.6 | 1.02 |  | 14.63 | 14.6 | -0.03 | -0.21
510 | 525 | 8 | 404 | 407.1 | 3.1 | 0.77 |  | 16.72 | 16.7 | -0.02 | -0.12
570 | 585 | 9 | 454.5 | 458 | 3.5 | 0.77 |  | 18.81 | 18.7 | -0.11 | -0.58
630 | 645 | 10 | 505 | 509.5 | 4.5 | 0.89 |  | 20.9 | 20.8 | -0.1 | -0.48
690 | 705 | 9 | 454.5 | 461.6 | 7.1 | 1.56 |  | 18.81 | 18.7 | -0.11 | -0.58
750 | 765 | 8 | 404 | 410.2 | 6.2 | 1.53 |  | 16.72 | 16.6 | -0.12 | -0.72
810 | 825 | 7 | 353.5 | 359.6 | 6.1 | 1.73 |  | 14.63 | 14.5 | -0.13 | -0.89
870 | 885 | 6 | 303 | 309.6 | 6.6 | 2.18 |  | 12.54 | 12.4 | -0.14 | -1.12
930 | 945 | 5 | 252.5 | 258.5 | 6 | 2.38 |  | 10.45 | 10.4 | -0.05 | -0.48
990 | 1005 | 4 | 202 | 206.3 | 4.3 | 2.13 |  | 8.36 | 8.3 | -0.06 | -0.72
1050 | 1065 | 3 | 151.5 | 155.2 | 3.7 | 2.44 |  | 6.27 | 6.3 | 0.03 | 0.48
1110 | 1125 | 2 | 101 | 103.8 | 2.8 | 2.77 |  | 4.18 | 4.2 | 0.02 | 0.48
1170 | 1185 | 1 | 50.5 | 52.5 | 2 | 3.96 |  | 2.09 | 2.2 | 0.11 | 5.26
1230 | 1245 | 0 | 0 | 0.8 | 0.8 |  |  | 0 | 0.1 | 0.1
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | 0.5 | 5.1 |  | Intercept | 0.06 | 0.21
|  |  |  | Slope | 1.01 | 0.99 - 1.01 |  | Slope | 0.99 | 0.99 - 1.01
|  |  |  | SEE | 1.7582 | 5.05 |  | SEE | 0.0454 | 0.209
|  |  |  | r2 | 0.9999 | 0.998 |  | r2 | 1 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1.4 | 1.4 | 0 | 0 | 0.1 | 12.9 | 950 | 961 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0023 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30 | 28 | 1.4
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1.3 | 1.3 | 0 | 0 | 0.1 | 12.9 | 950 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.002238 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.088 | 28 | 1.3
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1.2382 | 1.2 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3382 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1.2 | 1.2 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.2 | 1.2 | 0 | 0.0195 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3377 | 0.002138 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.126 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.2 | 1.2 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0021 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0493 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.2 | 1.2 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.002038 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0274 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.2 | 1.2 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.002 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 1.2
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.1377 | 1.1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3377 | 0.001938 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 30.1 | 28 | 1.1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.1 | 1.1 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.377 | 30.1 | 28 | 1.1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1.1 | 1.1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0019 | 0.001938 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 1.1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0019 | 0.0019 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.001838 | 0.0019 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.2 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0018 | 0.0019 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.382 | 30.0359 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3 | 0.0018 | 0.001838 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.6218 | 30 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0018 | 0.0018 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0168 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 1 | 1 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.001738 | 0.0018 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0938 | 28 | 1
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.1 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3 | 0.0017 | 0.001738 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0297 | 28 | 0.9
2008-07-16 | 13:33 | -0.004 | 0 | 0 | -0.1 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0017 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0471 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.1 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0017 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.076 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.1 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0017 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.1 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3621 | 0.0017 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.379 | 30 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.2 | 0.9 | 0.9 | 0 | 0 | 0.1 | 12.9 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 4 | 195 | 59.3 | 0.0016 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0543 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.2 | 0.9 | 0.9 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3377 | 0.0016 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1 | 28 | 0.9
2008-07-16 | 13:33 | -0.01 | 0 | 0 | -0.2 | 0.8 | 0.8 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3383 | 0.0016 | 0.0017 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 30.0924 | 28 | 0.8
2008-07-16 | 13:33 | -0.007 | 0 | 0 | -0.2 | 0.8 | 0.8 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3621 | 0.0016 | 0.001638 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 30.0157 | 28 | 0.8
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.73792 | 0.7 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0016 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010.6208 | 30.1219 | 28 | 0.7
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.7 | 0.7 | -0.1 | 0 | 0.1 | 12.9 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.362 | 0.0016 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010.3805 | 30.1624 | 28 | 0.7
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.6 | 0.6 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3381 | 0.001538 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1143 | 28 | 0.6
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.53758 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.191 | 28 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.5 | 0.5 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3377 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1323 | 28 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0556 | 28.0444 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.5 | 0.5 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30 | 28.079 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0015 | 0.0016 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30 | 28.0023 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0014 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0744 | 28 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0489 | 28 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.621 | 30 | 28 | 0.5
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.4 | 0.4 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0014 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0044 | 28 | 0.4
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.4 | 0.4 | 0.4 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3618 | 0.0014 | 0.0015 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0811 | 28 | 0.4
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.5 | 0.4 | 0.4 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 0.4
2008-07-16 | 13:33 | 0 | 0 | 0 | -0.6 | 0.4 | 0.4 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3376 | 0.001338 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 0.4
2008-07-16 | 13:33 | 0 | 0 | 0 | -1.1 | 0.3377 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3 | 0.0013 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.111 | 28 | 0.3
2008-07-16 | 13:33 | 0 | 0 | 0 | -1.1 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.4 | 0.0013 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.1877 | 28 | 0.3
2008-07-16 | 13:33 | 0 | 0 | 0 | -1.1 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1355 | 28 | 0.3
2008-07-16 | 13:33 | 0 | 0 | 0 | -1.1 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 0.3
2008-07-16 | 13:33 | 0 | 0.0003 | 2.5363 | -1.1 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1177 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0.001 | 10 | -1.1 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.001438 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.194 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0.001 | 10 | -1.2 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.2 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0.001 | 10 | -1.2 | 0.3 | 0.3 | -0.1 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1523 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0.0004 | 4.2237 | -1.2 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.2 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0992 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.2 | 0.3 | 0.3 | 0 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.377 | 30.0223 | 28 | 0.3
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.2 | 0.3 | 0.3 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0012 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30 | 28.0544 | 0.3
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.2 | 0.23808 | 0.2 | 0 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.001138 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9689 | 28.1 | 0.2
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.2 | 0.2 | -0.1 | 0 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3381 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9157 | 28.0922 | 0.2
2008-07-16 | 13:34 | 0 | -0.0006 | -5.7374 | -1.3 | 0.2 | 0.2 | 0 | 0 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.618 | 30.0691 | 28.0154 | 0.2
2008-07-16 | 13:34 | 0 | -0.0006 | -5.9173 | -1.3 | 0.2 | 0.2 | 0 | 0 | 0.1 | 13 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9775 | 28.0613 | 0.2
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.2 | 0.2 | 0 | 0 | 0.1 | 13 | 950 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9759 | 28.062 | 0.2
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.2 | 0.2 | 0 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0854 | 28 | 0.2
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.13805 | 0.1 | 0 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 4 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0087 | 28 | 0.1
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.1 | 0.1 | 0 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.6188 | 30.0679 | 28.0679 | 0.1
2008-07-16 | 13:34 | 0 | 0 | 0 | -1.3 | 0.1 | 0.1 | -0.1 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1446 | 28.1 | 0.1
2008-07-16 | 13:34 | 0 | -0.0004 | -4.0447 | -1.2 | 6.1429 | 6.1 | 0.2 | 0.2163 | 0.13 | 13 | 950 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3623 | 0.0011 | 0.0014 | 0.000374 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1788 | 28.1 | 6.1
2008-07-16 | 13:34 | 0 | -0.001 | -10 | 1.9 | 26.8256 | 28.7 | 1 | 1.0123 | 0.97 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.001687 | 0.0014 | 0.001635 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6217 | 30.1021 | 28.1 | 26.8
2008-07-16 | 13:34 | 0 | -0.001 | -10 | 3.2 | 43.8476 | 47 | 0.4 | 0.356 | 1.72 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.001962 | 0.001462 | 0.002624 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1746 | 28.0254 | 43.8
2008-07-16 | 13:34 | 0 | -0.001 | -10 | 2.9 | 48.2 | 51.1 | 0.1 | 0.1235 | 1.97 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.002362 | 0.001962 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2513 | 28 | 48.2
2008-07-16 | 13:34 | 0.004 | -0.001 | -10 | 2.6 | 48.3618 | 51 | 0.5 | 0.4616 | 2 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.002762 | 0.0024 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.2442 | 28 | 48.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2.4 | 48.5 | 50.9 | 0.5 | 0.4766 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.002862 | 0.002562 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.3797 | 30.1 | 28 | 48.5
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2.3 | 48.562 | 50.9 | 0.4 | 0.4206 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.003 | 0.002662 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3797 | 30.1 | 28 | 48.6
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2.3 | 48.6 | 50.9 | 0.4 | 0.412 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.003 | 0.0027 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.042 | 28 | 48.6
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2.1 | 48.7 | 50.8 | 0.3 | 0.3103 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.003162 | 0.0029 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30 | 28.0347 | 48.7
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2 | 49.2994 | 51.3 | 0.4 | 0.3756 | 2.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.003262 | 0.003062 | 0.0029 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.0113 | 28.1 | 49.3
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 2 | 50.4 | 52.4 | 0.3 | 0.3199 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3378 | 0.0033 | 0.0031 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.088 | 28.1 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 1.9 | 50.4 | 52.3 | 0.5 | 0.4522 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3379 | 0.0034 | 0.0032 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0354 | 28.0354 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 1.8 | 50.4 | 52.2 | 0.3 | 0.3 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0035 | 0.0033 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0827 | 28.0413 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 1.8 | 50.4 | 52.2 | 0.5 | 0.477 | 2.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0035 | 0.0033 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.164 | 28.082 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 1.7 | 50.4 | 52.1 | 0.3 | 0.34 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0035 | 0.003362 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0107 | 28.0054 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.001 | -10 | 1.6 | 50.4 | 52 | 0.4 | 0.4226 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0036 | 0.003462 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0713 | 28 | 50.4
2008-07-16 | 13:34 | 0.01 | -0.0002 | -2.4767 | 1.6 | 50.4 | 52 | 0.4 | 0.3772 | 2.1 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.003662 | 0.0035 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0523 | 28 | 50.4
2008-07-16 | 13:34 | 0.01 | 0 | 0 | 1.5 | 50.438 | 51.9 | 0.4 | 0.4218 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0037 | 0.003562 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0243 | 28.0243 | 50.4
2008-07-16 | 13:34 | 0.01 | 0 | 0 | 1.4 | 50.438 | 51.9 | 0.3 | 0.3104 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.0037 | 0.0036 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0986 | 28.1 | 50.4
2008-07-16 | 13:34 | 0.01 | 0 | 0 | 1.4 | 50.4 | 51.8 | 0.2 | 0.2228 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0038 | 0.0036 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0222 | 28.1 | 50.4
2008-07-16 | 13:34 | 0.01 | 0 | 0 | 1.4 | 50.5 | 51.9 | 0.3 | 0.2698 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0038 | 0.0037 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1091 | 28.0454 | 50.5
2008-07-16 | 13:34 | 0.004 | 0 | 0 | 1.4 | 50.5 | 51.9 | 0.1 | 0.1239 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0038 | 0.0037 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1373 | 28.0314 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.5 | 50.462 | 51.9 | 0.3 | 0.2785 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.008 | 28.092 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.5 | 50.5 | 52 | 0 | 0.0232 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0039 | 0.003762 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0847 | 28.0153 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.5 | 50.5 | 52 | 0.2 | 0.1824 | 2.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0039 | 0.0038 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.382 | 30.1 | 28 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.5 | 50.5 | 52 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0039 | 0.0038 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.138 | 28 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.4 | 50.5 | 51.9 | 0.1 | 0.1 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0039 | 0.0038 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.4 | 50.5 | 51.9 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.004 | 0.003862 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.4 | 50.5 | 51.9 | 0 | 0.0221 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.004 | 0.0039 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6195 | 30.0641 | 28 | 50.5
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.3 | 50.4 | 51.7 | 0.1 | 0.0572 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3619 | 0.004 | 0.0039 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0446 | 28 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.2 | 50.4 | 51.6 | 0 | 0.0221 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3618 | 0.004 | 0.004 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0787 | 28.0213 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.2 | 50.4 | 51.6 | 0.1 | 0.0751 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3377 | 0.0041 | 0.004 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0019 | 28.0981 | 50.4
2008-07-16 | 13:34 | -0.008 | 0 | 0 | 1.2 | 50.4 | 51.6 | 0 | 0 | 2.1 | 13 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.004 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0747 | 28.0253 | 50.4
2008-07-16 | 13:34 | -0.01 | 0 | 0 | 1.2 | 50.4 | 51.6 | 0.1 | 0.0763 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.004 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.3795 | 30.1 | 28 | 50.4
2008-07-16 | 13:34 | -0.006 | 0 | 0 | 1.1 | 50.4 | 51.5 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.004 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6192 | 30.1 | 28.028 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.1 | 50.4 | 51.5 | 0.1 | 0.1 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.004062 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1 | 28.1 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.1 | 50.4 | 51.5 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.0041 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1.1 | 50.4 | 51.5 | 0.1 | 0.1312 | 2.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0041 | 0.0041 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.158 | 28.1 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.004162 | 0.0041 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1307 | 28.1 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0.0208 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0042 | 0.0041 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.0887 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0.1 | 0.0534 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30 | 28.012 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0.0205 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30 | 28.0646 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0.0262 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3622 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6222 | 30 | 28.1 | 50.4
2008-07-16 | 13:34 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.018 | 28.1 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0.1 | 0.0773 | 2.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.0947 | 28.1 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4 | 51.3 | 0 | 0 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3381 | 0.0042 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1714 | 28.0286 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4 | 51.3 | 0.1 | 0.1 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3623 | 0.004262 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1042 | 28 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 1 | 50.4 | 51.4 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0042 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6223 | 30.0243 | 28 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4 | 51.3 | 0.1 | 0.0675 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.004262 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0993 | 28.0007 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4 | 51.3 | 0.1 | 0.0887 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 4.6 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0223 | 28.0777 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4 | 51.3 | 0 | 0.0211 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 4.4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0546 | 28.1 | 50.4
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.4624 | 51.4 | 0.1 | 0.1088 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 50.5
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.9 | 50.5 | 51.4 | 0 | 0.02 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.084 | 28.1 | 50.5
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.8 | 50.5 | 51.3 | 0.1 | 0.0788 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9307 | 28.1 | 50.5
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.8 | 50.5 | 51.3 | 0 | 0 | 2.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.362 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0227 | 28.1 | 50.5
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.8 | 50.5 | 51.3 | 0.1 | 0.0782 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3381 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 50.5
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.8 | 50.5 | 51.3 | 0 | 0 | 2.1 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1147 | 28.1 | 50.5
2008-07-16 | 13:35 | 0.006 | 0 | 0 | 0.8 | 50.5 | 51.3 | 0 | 0 | 2.1 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0043 | 0.0043 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1914 | 28.1 | 50.5
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.7 | 50.5 | 51.2 | 0 | 0 | 2.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3621 | 0.0043 | 0.004362 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 30.2 | 28.1 | 50.5
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 50.5 | 51.4 | 0 | 0.0291 | 2.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3619 | 0.004362 | 0.0044 | 0.003 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0212 | 28.1 | 50.5
2008-07-16 | 13:35 | 0.01 | -0.0004 | -3.8032 | 2 | 71.943 | 74 | 0.1 | 0.1 | 2.47 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.362 | 0.00481 | 0.004648 | 0.004272 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8642 | 28.1 | 71.9
2008-07-16 | 13:35 | 0.01 | -0.001 | -10 | 2.4 | 92.5 | 94.9 | 0.1 | 0.1411 | 3.41 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0057 | 0.0056 | 0.0055 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0942 | 28.1 | 92.5
2008-07-16 | 13:35 | 0.01 | -0.001 | -10 | 1.8 | 96.4776 | 98.3 | 0.3 | 0.255 | 3.92 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3618 | 0.006285 | 0.006247 | 0.005724 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 96.5
2008-07-16 | 13:35 | 0.01 | -0.001 | -10 | 1.6 | 98.3607 | 100 | 0.1 | 0.1221 | 4.07 | 12.9 | 955 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 4 | 195 | 59.3377 | 0.0065 | 0.0065 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1 | 28.1 | 98.4
2008-07-16 | 13:35 | 0.01 | -0.0003 | -2.5062 | 1.5 | 98.638 | 100.1 | 0.3 | 0.2791 | 4.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.006662 | 0.0067 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.128 | 28.1 | 98.6
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 1 | 98.7 | 99.7 | 0.1 | 0.1235 | 4.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3382 | 0.007062 | 0.007162 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1907 | 28.1 | 98.7
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 98.7619 | 99.7 | 0.3 | 0.3 | 4.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3619 | 0.007162 | 0.007262 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.0374 | 28.1 | 98.8
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 98.8 | 99.7 | 0.1 | 0.1 | 4.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0072 | 0.0073 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1 | 98.8
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 98.9621 | 99.9 | 0.2 | 0.2208 | 4.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3621 | 0.0073 | 0.0074 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1 | 99
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.8 | 99.162 | 100 | 0.2 | 0.179 | 4.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.007362 | 0.007562 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0227 | 28.1 | 99.2
2008-07-16 | 13:35 | 0.006 | 0 | 0 | 0.8 | 99.2 | 100 | 0.2 | 0.1668 | 4.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3621 | 0.0074 | 0.0076 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1761 | 28.1 | 99.2
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.7 | 99.2 | 99.9 | 0.2 | 0.1794 | 4.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0075 | 0.0077 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.2 | 28.1 | 99.2
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.7 | 99.2 | 99.9 | 0.1 | 0.1159 | 4.1 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.007562 | 0.007762 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1172 | 28.1 | 99.2
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.7 | 99.3 | 100 | 0.2 | 0.1795 | 4.17 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0076 | 0.0078 | 0.0058 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.018 | 28.1 | 99.3
2008-07-16 | 13:35 | 0.007 | 0 | 0 | 0.6 | 99.3622 | 100 | 0.2 | 0.1615 | 4.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.007662 | 0.007862 | 0.005862 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0946 | 28.1 | 99.4
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.6 | 99.4 | 100 | 0.3 | 0.2759 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.362 | 0.0077 | 0.0079 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0292 | 28.1 | 99.4
2008-07-16 | 13:35 | 0.01 | -0.0007 | -7.4663 | 0.8 | 99.4 | 100.2 | 0.1 | 0.1 | 4.2 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3619 | 0.007762 | 0.007962 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1 | 99.4
2008-07-16 | 13:35 | 0.01 | -0.0004 | -4.39 | 0.9 | 99.5 | 100.4 | 0.2 | 0.2 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 4 | 195 | 59.4 | 0.007862 | 0.008 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3823 | 30 | 28.1 | 99.5
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 99.6 | 100.5 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.0079 | 0.008062 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6218 | 30.0008 | 28.1 | 99.6
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 99.6 | 100.5 | 0.2 | 0.2 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0079 | 0.0081 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0778 | 28.1 | 99.6
2008-07-16 | 13:35 | 0.01 | 0 | 0 | 0.9 | 99.6 | 100.5 | 0.2 | 0.1785 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.008 | 0.0081 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0454 | 28.1 | 99.6
2008-07-16 | 13:35 | 0.004 | 0 | 0 | 0.8 | 99.7 | 100.5 | 0.1 | 0.134 | 4.2 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.008 | 0.0082 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.9686 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 0.8 | 99.7 | 100.5 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3624 | 0.008062 | 0.0082 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.9161 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | -0.0006 | -5.6055 | 0.8 | 99.7 | 100.5 | 0.1 | 0.1221 | 4.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3623 | 0.0081 | 0.0083 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 30.0693 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | -0.0006 | -6.2302 | 1 | 99.7378 | 100.7 | 0.2 | 0.1563 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.008162 | 0.0083 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6217 | 29.9773 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 99.7 | 100.8 | 0 | 0.0227 | 4.2 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0082 | 0.0083 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0141 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 99.7 | 100.8 | 0.2 | 0.1749 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0082 | 0.008362 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1853 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 99.7618 | 100.9 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.008262 | 0.0084 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.382 | 30.1087 | 28.1 | 99.8
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7377 | 100.9 | 0.2 | 0.177 | 4.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3623 | 0.0083 | 0.0084 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.032 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7 | 100.9 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0083 | 0.0084 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7 | 100.9 | 0.1 | 0.1227 | 4.2 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0083 | 0.0084 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7 | 100.9 | 0.1 | 0.0869 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0083 | 0.0084 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7 | 100.9 | 0.1 | 0.0765 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.008362 | 0.008462 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.0746 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7 | 100.9 | 0.1 | 0.0813 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0084 | 0.0085 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1513 | 28.1 | 99.7
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.7623 | 101 | 0 | 0.0213 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3623 | 0.0084 | 0.0085 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.1441 | 28.1 | 99.8
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 99.9 | 101.1 | 0.2 | 0.1547 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0085 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0092 | 28.1 | 99.9
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 100 | 101.2 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0084 | 0.0085 | 0.0059 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1011.3808 | 30.1626 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 100 | 101.1 | 0.2 | 0.1732 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0085 | 0.0086 | 0.0059 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 30.1421 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 100 | 101.1 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0085 | 0.0086 | 0.0059 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1012 | 30.0308 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 100 | 101.2 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.0085 | 0.0086 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9225 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 100 | 101.2 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.338 | 0.0085 | 0.0086 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0758 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 100 | 101.2 | 0.1 | 0.119 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0085 | 0.0086 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.1 | 28.1 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.2 | 100 | 101.2 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3619 | 0.0085 | 0.0087 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0175 | 28.1413 | 100
2008-07-16 | 13:35 | 0 | 0 | 0 | 1.1 | 100 | 101.1 | 0 | 0.0328 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0086 | 0.0087 | 0.0059 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1013 | 29.9179 | 28.1821 | 100
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 100 | 101.1 | 0.1 | 0.0759 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3377 | 0.0086 | 0.0087 | 0.0059 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.9944 | 28.1056 | 100
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100 | 101 | 0 | 0.0234 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.0086 | 0.008762 | 0.0059 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1013 | 30.0708 | 28.1 | 100
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100 | 101 | 0.1 | 0.0531 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.338 | 0.008662 | 0.0088 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 30.0048 | 28.1 | 100
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 100.0618 | 101 | 0 | 0.0233 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0087 | 0.0088 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.3815 | 29.9978 | 28.1 | 100.1
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 100.1 | 101 | 0.2 | 0.1765 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3377 | 0.0087 | 0.0088 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.2975 | 28.1 | 100.1
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 100.1 | 101 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3378 | 0.0087 | 0.0088 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0669 | 28.1 | 100.1
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 100.162 | 101.1 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0087 | 0.008862 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0545 | 28.1 | 100.2
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 100.2 | 101.1 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0087 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0064 | 28.1 | 100.2
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.3 | 101.3 | 0 | 0.0225 | 4.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0087 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8 | 28.1 | 100.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.3 | 101.3 | 0.1 | 0.0782 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3381 | 0.0087 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8 | 28.1 | 100.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.3 | 101.3 | 0 | 0.0329 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0087 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9225 | 28.1 | 100.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.4 | 101.4 | 0.1 | 0.0889 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0087 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9621 | 28.1379 | 100.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.4 | 101.4 | 0 | 0.0221 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.008762 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.9437 | 28.1854 | 100.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.4 | 101.4 | 0.1 | 0.1 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.362 | 0.0088 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3795 | 30.1736 | 28.1088 | 100.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 100.4 | 101.4 | 0 | 0 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0088 | 0.0089 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.0642 | 28.1679 | 100.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 109.9671 | 111.1 | 0.2 | 0.2 | 4.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0088 | 0.008962 | 0.00646 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9107 | 28.2 | 110
2008-07-16 | 13:36 | 0 | 0 | 0 | 2.3 | 134.7972 | 137.1 | 0.2 | 0.1764 | 4.84 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.009335 | 0.009311 | 0.007922 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.6218 | 29.8852 | 28.1787 | 134.8
2008-07-16 | 13:36 | 0 | 0 | 0 | 2.6 | 145.7 | 148.3 | 0.2 | 0.2228 | 5.74 | 12.9 | 954 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0099 | 0.0099 | 0.0086 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.1917 | 28.1021 | 145.7
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.8 | 147.5785 | 149.3 | 0.1 | 0.1 | 6.01 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.010634 | 0.010796 | 0.008662 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.3797 | 30.0509 | 28.1746 | 147.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 148.4 | 149.3 | 0.2 | 0.2 | 6.1 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.011448 | 0.01181 | 0.0087 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.6195 | 30.0512 | 28.1488 | 148.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.6 | 149.0242 | 149.6 | 0.2 | 0.2 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3624 | 0.0117 | 0.012062 | 0.008762 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.3758 | 30.1 | 28.1 | 149
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.5 | 149.5 | 150 | 0.2 | 0.1655 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.0122 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1045 | 28.1 | 149.5
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.8 | 149.5622 | 150.4 | 0.2 | 0.1761 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3378 | 0.011962 | 0.0123 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3783 | 30.1811 | 28.1 | 149.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 149.5 | 150.5 | 0.1 | 0.122 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0121 | 0.0124 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.2 | 28.1 | 149.5
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 149.5 | 150.5 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0121 | 0.0124 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.131 | 28.1 | 149.5
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 149.6 | 150.5 | 0.1 | 0.1232 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0122 | 0.0125 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30 | 28.1 | 149.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.8 | 149.5622 | 150.4 | 0.3 | 0.2746 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.012262 | 0.0126 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.6218 | 30 | 28.1 | 149.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.7 | 149.6622 | 150.4 | 0.2 | 0.1806 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.012362 | 0.0127 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0645 | 28.1645 | 149.7
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.7 | 149.7 | 150.4 | 0.2 | 0.2173 | 6.2 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.012462 | 0.0128 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.6203 | 30.0589 | 28.1589 | 149.7
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.8 | 149.762 | 150.6 | 0.1 | 0.1 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0125 | 0.0128 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30 | 28.1177 | 149.8
2008-07-16 | 13:36 | 0 | 0 | 0 | 0.9 | 149.8626 | 150.8 | 0.1 | 0.1444 | 6.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.012563 | 0.012863 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.3737 | 30 | 28.1938 | 149.9
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 149.9 | 150.9 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0126 | 0.0129 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.9293 | 28.1293 | 149.9
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 150 | 151 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3377 | 0.0126 | 0.012962 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0421 | 28.1 | 150
2008-07-16 | 13:36 | 0 | -0.0006 | -5.551 | 1 | 150.1 | 151.1 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0127 | 0.013 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.176 | 28.124 | 150.1
2008-07-16 | 13:36 | 0 | -0.0006 | -6.2173 | 1 | 150.262 | 151.3 | 0.1 | 0.1198 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0127 | 0.013 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1 | 28.2 | 150.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 150.3619 | 151.4 | 0.2 | 0.2 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3 | 0.0127 | 0.013 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3808 | 30.1 | 28.2 | 150.4
2008-07-16 | 13:36 | 0 | -0.0001 | -0.66667 | 1 | 150.4 | 151.4 | 0.2 | 0.1777 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3623 | 0.012762 | 0.013062 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30.0456 | 28.1456 | 150.4
2008-07-16 | 13:36 | 0 | -0.0009 | -8.6988 | 1 | 150.4621 | 151.5 | 0.3 | 0.2508 | 6.2 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0128 | 0.0131 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 30 | 28.131 | 150.5
2008-07-16 | 13:36 | -0.004 | -0.0003 | -2.8253 | 1 | 150.5619 | 151.6 | 0.1 | 0.1115 | 6.2 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3381 | 0.0128 | 0.0131 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.619 | 30.0077 | 28.1923 | 150.6
2008-07-16 | 13:36 | -0.008 | 0 | 0 | 1 | 150.6 | 151.6 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.338 | 0.0128 | 0.0131 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0844 | 28.1156 | 150.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 150.6 | 151.6 | 0.1 | 0.1 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0128 | 0.0132 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.039 | 28.161 | 150.6
2008-07-16 | 13:36 | -0.009 | 0 | 0 | 1 | 150.6 | 151.6 | 0.1 | 0.1183 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0129 | 0.0132 | 0.0088 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30 | 28.2 | 150.6
2008-07-16 | 13:36 | -0.01 | 0 | 0 | 1 | 150.5 | 151.5 | 0.2 | 0.1779 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3377 | 0.0129 | 0.0132 | 0.008862 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0143 | 28.1857 | 150.5
2008-07-16 | 13:36 | -0.005 | 0 | 0 | 1 | 150.6 | 151.6 | 0.1 | 0.134 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3378 | 0.0129 | 0.0132 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.091 | 28.109 | 150.6
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 150.7 | 151.7 | 0.2 | 0.1785 | 6.2 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0129 | 0.0132 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0324 | 28.1676 | 150.7
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 150.7 | 151.7 | 0.1 | 0.1235 | 6.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3382 | 0.0129 | 0.0132 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30 | 28.2 | 150.7
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 150.8 | 151.9 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.012962 | 0.0133 | 0.0089 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1013 | 30 | 28.2 | 150.8
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 150.9623 | 152.1 | 0.1 | 0.1 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.013 | 0.0133 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.2 | 151
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 151 | 152.1 | 0.2 | 0.227 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3379 | 0.013 | 0.0133 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9258 | 28.2 | 151
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 151.162 | 152.3 | 0.1 | 0.1 | 6.2 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.338 | 0.013 | 0.0133 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3797 | 29.9 | 28.2 | 151.2
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 151.2621 | 152.4 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3621 | 0.013 | 0.0133 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.9826 | 28.2 | 151.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.3381 | 152.3 | 0.1 | 0.1 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3381 | 0.013062 | 0.013362 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6188 | 30.1873 | 28.2 | 151.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.3 | 152.3 | 0.1 | 0.1186 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3377 | 0.0131 | 0.0134 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9573 | 28.2 | 151.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.3622 | 152.4 | 0.1 | 0.1389 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3378 | 0.0131 | 0.0134 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9576 | 28.2 | 151.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.4 | 152.4 | 0 | 0.0338 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0131 | 0.0134 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6205 | 29.9657 | 28.2 | 151.4
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.4619 | 152.5 | 0.1 | 0.1 | 6.2 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0131 | 0.0134 | 0.0089 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.9109 | 28.2 | 151.5
2008-07-16 | 13:36 | 0 | 0 | 0 | 1 | 151.9358 | 152.9 | 0.1 | 0.1 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3377 | 0.0131 | 0.013462 | 0.008962 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.9876 | 28.2 | 151.9
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 152.3 | 153.4 | 0.1 | 0.1 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.4 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 30.0642 | 28.2 | 152.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.1 | 152.3 | 153.4 | 0 | 0 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1013 | 30.0592 | 28.2 | 152.3
2008-07-16 | 13:36 | 0 | 0 | 0 | 1.2 | 152.362 | 153.5 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0348 | 28.2 | 152.4
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.2 | 152.4 | 153.6 | 0.1 | 0.1 | 6.2 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.1876 | 28.2 | 152.4
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.2 | 152.4 | 153.6 | 0.2 | 0.2223 | 6.2 | 12.9 | 954 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.338 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0596 | 28.1298 | 152.4
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.2 | 152.4 | 153.6 | 0.1 | 0.1 | 6.2 | 12.9 | 954 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0132 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.953 | 28.147 | 152.4
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.3 | 152.5 | 153.8 | 0.1 | 0.1231 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3622 | 0.013262 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9239 | 28.1761 | 152.5
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.4 | 152.6 | 154 | 0.2 | 0.1786 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3622 | 0.0133 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30 | 28.1007 | 152.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.4 | 152.662 | 154.1 | 0.1 | 0.1332 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0133 | 0.0135 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3795 | 30 | 28.1774 | 152.7
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.4 | 152.7 | 154.1 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3381 | 0.0133 | 0.013562 | 0.009 | 0 |  | n/a |  | 32.8931 | 55 | 1049 | 1011 | 30 | 28.1459 | 152.7
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.4 | 152.7 | 154.1 | 0.2 | 0.2 | 6.2 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3382 | 0.0133 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.382 | 30.0307 | 28.1 | 152.7
2008-07-16 | 13:37 | 0 | -0.0007 | -7.1737 | 1.4 | 152.7 | 154.1 | 0.3 | 0.2563 | 6.2 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3622 | 0.0133 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6218 | 30.1074 | 28.1074 | 152.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.7 | 154.1 | 0.1 | 0.1379 | 6.2 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.013362 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.184 | 28.184 | 152.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.6381 | 154 | 0.3 | 0.3 | 6.2 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3619 | 0.0134 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0786 | 28.2 | 152.6
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.6 | 154 | 0.2 | 0.2 | 6.2 | 13 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0134 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.2 | 152.6
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.6619 | 154.1 | 0.2 | 0.2221 | 6.2 | 13 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3381 | 0.0134 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.2 | 152.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.6 | 154 | 0.3 | 0.2651 | 6.2 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3383 | 0.0134 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6168 | 30 | 28.2 | 152.6
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.5377 | 153.9 | 0.2 | 0.2221 | 6.2 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3623 | 0.0134 | 0.0136 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8655 | 28.2 | 152.5
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.4 | 152.6 | 154 | 0.3 | 0.2797 | 6.2 | 13 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0134 | 0.0137 | 0.009 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.9317 | 28.2 | 152.6
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.7 | 160.1464 | 161.9 | 0.3 | 0.2688 | 6.22 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.338 | 0.013462 | 0.0137 | 0.00941 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0794 | 28.2 | 160.1
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 3 | 189.1255 | 192.1 | 0.6 | 0.553 | 7.12 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 4 | 195 | 59.3619 | 0.014633 | 0.014595 | 0.01111 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0028 | 28.2 | 189.1
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.7 | 195.7617 | 198.5 | 0.3 | 0.3219 | 7.91 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.015087 | 0.015149 | 0.011549 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0739 | 28.2 | 195.8
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.3 | 200.7244 | 203 | 0.4 | 0.4 | 8.16 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.015862 | 0.0161 | 0.0118 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1012 | 29.9989 | 28.2 | 200.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.2 | 201.3582 | 203.6 | 0.4 | 0.3808 | 8.2 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.016262 | 0.016562 | 0.011862 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9272 | 28.2 | 201.4
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.3 | 201.7 | 204 | 0.4 | 0.4194 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.016462 | 0.016662 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0039 | 28.2 | 201.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.6 | 201.6 | 204.2 | 0.4 | 0.3835 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0166 | 0.016762 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0805 | 28.2 | 201.6
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.5 | 201.7 | 204.2 | 0.4 | 0.4453 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.016662 | 0.016862 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 30.0428 | 28.2 | 201.7
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.4 | 201.7623 | 204.2 | 0.5 | 0.4604 | 8.3 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0167 | 0.0169 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.377 | 30.0676 | 28.2 | 201.8
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 2.4 | 201.8 | 204.2 | 0.3 | 0.3478 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0168 | 0.017 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.1896 | 28.2 | 201.8
2008-07-16 | 13:37 | 0 | -0.0003 | -2.8745 | 2.3 | 201.938 | 204.2 | 0.5 | 0.4858 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.016862 | 0.0171 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.113 | 28.2 | 201.9
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.3 | 202 | 204.3 | 0.4 | 0.4415 | 8.3 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3382 | 0.0169 | 0.017162 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.6185 | 29.9735 | 28.2 | 202
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.2 | 202.1617 | 204.4 | 0.5 | 0.4767 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.017 | 0.017262 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9 | 28.2 | 202.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.2 | 202.2 | 204.4 | 0.3 | 0.3204 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3621 | 0.017062 | 0.0173 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8833 | 28.2 | 202.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.2 | 202.1621 | 204.4 | 0.5 | 0.5 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3621 | 0.0171 | 0.017362 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8069 | 28.2 | 202.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.2 | 202.2619 | 204.5 | 0.4 | 0.3768 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.017162 | 0.0174 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8701 | 28.2 | 202.3
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.2 | 202.462 | 204.7 | 0.4 | 0.426 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0172 | 0.017462 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30.0408 | 28.2 | 202.5
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.1 | 202.5619 | 204.7 | 0.4 | 0.3772 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.0172 | 0.0175 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3812 | 30.1527 | 28.2 | 202.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 2.1 | 202.5376 | 204.6 | 0.4 | 0.4 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0173 | 0.0175 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 30 | 28.2005 | 202.5
2008-07-16 | 13:37 | 0 | 0 | 0 | 2 | 202.5622 | 204.6 | 0.4 | 0.4 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0173 | 0.0176 | 0.0119 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 30 | 28.2771 | 202.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 2 | 202.6 | 204.6 | 0.3 | 0.3234 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0173 | 0.0176 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0537 | 28.2463 | 202.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 2 | 202.6 | 204.6 | 0.4 | 0.3795 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0174 | 0.017662 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0697 | 28.2 | 202.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.9 | 202.6 | 204.5 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0174 | 0.0177 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.2 | 202.6
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.9 | 202.7 | 204.6 | 0.4 | 0.3783 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3378 | 0.017462 | 0.017762 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.2 | 202.7
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.8 | 202.7621 | 204.6 | 0.3 | 0.2766 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0178 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 30 | 28.2 | 202.8
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.8 | 202.8 | 204.6 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0175 | 0.017862 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6203 | 30 | 28.2 | 202.8
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.7 | 202.8 | 204.5 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.017562 | 0.0179 | 0.0119 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0273 | 28.2 | 202.8
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.7 | 202.7626 | 204.5 | 0.3 | 0.327 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0179 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.1805 | 28.2 | 202.8
2008-07-16 | 13:37 | 0 | -0.0007 | -7.087 | 1.7 | 202.8 | 204.5 | 0.3 | 0.2796 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3376 | 0.0176 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30.0662 | 28.2 | 202.8
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.7 | 202.8622 | 204.6 | 0.2 | 0.2205 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9564 | 28.2 | 202.9
2008-07-16 | 13:37 | 0 | -0.0007 | -6.5221 | 1.7 | 202.9 | 204.6 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.017662 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6217 | 29.9202 | 28.2 | 202.9
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.6 | 202.962 | 204.6 | 0.2 | 0.222 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.0177 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3795 | 29.9969 | 28.2 | 203
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.6 | 203.0619 | 204.7 | 0.4 | 0.3765 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3619 | 0.0177 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.9265 | 28.2 | 203.1
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.6 | 203 | 204.6 | 0.2 | 0.2316 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0177 | 0.018 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.8498 | 28.2 | 203
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.6 | 203.1 | 204.7 | 0.4 | 0.3781 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0177 | 0.018062 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.9073 | 28.2268 | 203.1
2008-07-16 | 13:37 | 0 | -0.0005 | -5.2988 | 1.6 | 203.038 | 204.6 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0177 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6203 | 30.1827 | 28.2965 | 203
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.6 | 203.062 | 204.7 | 0.3 | 0.3431 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0177 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7993 | 28.2199 | 203.1
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.6 | 203.0374 | 204.6 | 0.3 | 0.3113 | 8.3 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3374 | 0.0177 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7568 | 28.2 | 203
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.6 | 203.0624 | 204.7 | 0.4 | 0.4 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3624 | 0.0177 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8 | 28.2 | 203.1
2008-07-16 | 13:37 | 0 | -0.001 | -10 | 1.6 | 203.1 | 204.7 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.017762 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7898 | 28.2 | 203.1
2008-07-16 | 13:37 | 0 | -0.0003 | -3.2193 | 1.7 | 203.1622 | 204.8 | 0.3 | 0.312 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0178 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7132 | 28.2 | 203.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.7 | 203.2 | 204.9 | 0.4 | 0.4 | 8.3 | 13 | 954 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.0178 | 0.0181 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8895 | 28.2 | 203.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.7 | 203.2 | 204.9 | 0.3 | 0.3119 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0178 | 0.0182 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9202 | 28.2 | 203.2
2008-07-16 | 13:37 | 0 | 0 | 0 | 1.7 | 203.3 | 205 | 0.4 | 0.3777 | 8.3 | 13 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3618 | 0.0179 | 0.0182 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.8 | 28.2 | 203.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.3 | 204.9 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.0179 | 0.0182 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.8 | 28.2 | 203.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.3 | 204.9 | 0.5 | 0.4776 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0179 | 0.0183 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.87 | 28.2 | 203.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.238 | 204.8 | 0.3 | 0.2768 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0179 | 0.0183 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9 | 28.2 | 203.2
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.362 | 205 | 0.3 | 0.345 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.018 | 0.0183 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9235 | 28.2235 | 203.4
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.3381 | 204.9 | 0.3 | 0.3 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9997 | 28.2998 | 203.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.3 | 204.9 | 0.3 | 0.3242 | 8.3 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8464 | 28.2232 | 203.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.3623 | 205 | 0.4 | 0.3764 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7465 | 28.2 | 203.4
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.4 | 205 | 0.2 | 0.2333 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7603 | 28.2 | 203.4
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.462 | 205.1 | 0.3 | 0.3 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.3795 | 29.8932 | 28.2 | 203.5
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.5624 | 205.2 | 0.2 | 0.2218 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8165 | 28.2 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.6 | 205.2 | 0.4 | 0.3762 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3618 | 0.018 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9203 | 28.2602 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.6 | 205.2 | 0.2 | 0.2 | 8.3 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3622 | 0.018062 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8896 | 28.2632 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.6 | 205.2 | 0.4 | 0.3766 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3378 | 0.0181 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7135 | 28.2 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.6 | 205.2 | 0.3 | 0.2782 | 8.3 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.338 | 0.0181 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7901 | 28.2 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 203.638 | 205.2 | 0.2 | 0.2441 | 8.3 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0181 | 0.0184 | 0.012 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8 | 28.2 | 203.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.7 | 213.9908 | 215.7 | 0.2 | 0.2218 | 8.36 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3619 | 0.0181 | 0.018462 | 0.012657 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7566 | 28.2434 | 214
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.4 | 240.5994 | 243 | 0.4 | 0.4249 | 9.15 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.018949 | 0.019249 | 0.014224 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6242 | 29.7201 | 28.2799 | 240.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.4 | 250.246 | 252.6 | 0.4 | 0.4 | 9.91 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.019762 | 0.020162 | 0.014762 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.377 | 29.7968 | 28.2032 | 250.2
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.4 | 251.5596 | 254 | 0.4 | 0.418 | 10.2 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.019862 | 0.020262 | 0.014862 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8734 | 28.2 | 251.6
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.1 | 252.4948 | 254.6 | 0.5 | 0.461 | 10.3 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.020186 | 0.020609 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8499 | 28.2501 | 252.5
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.9 | 252.8618 | 254.8 | 0.3 | 0.3227 | 10.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0204 | 0.020862 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8536 | 28.2732 | 252.9
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 252.9 | 254.5 | 0.5 | 0.4775 | 10.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.0208 | 0.021362 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.621 | 29.9966 | 28.2 | 252.9
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.5 | 253.1863 | 254.7 | 0.4 | 0.4 | 10.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.020962 | 0.021462 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 29.9199 | 28.2 | 253.2
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.4 | 253.1 | 254.5 | 0.5 | 0.459 | 10.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.021062 | 0.0216 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7865 | 28.2 | 253.1
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.4 | 253.1 | 254.5 | 0.4 | 0.4 | 10.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.338 | 0.0211 | 0.021662 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.2334 | 253.1
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.4 | 253.1619 | 254.6 | 0.4 | 0.4439 | 10.35 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0212 | 0.021762 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7201 | 28.29 | 253.2
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.4 | 253.1 | 254.5 | 0.5 | 0.4887 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0213 | 0.021862 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8723 | 28.2139 | 253.1
2008-07-16 | 13:38 | 0 | 0 | 0 | 1.6 | 253.2 | 254.8 | 0.4 | 0.4296 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0214 | 0.021962 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.837 | 28.2 | 253.2
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.3 | 253.1378 | 255.5 | 0.5 | 0.4779 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0216 | 0.022 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8397 | 28.2397 | 253.1
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.4 | 253.262 | 255.7 | 0.4 | 0.4232 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0217 | 0.0221 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8675 | 28.3 | 253.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.4 | 253.3 | 255.7 | 0.5 | 0.5 | 10.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3381 | 0.0217 | 0.0221 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7134 | 28.3 | 253.3
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.7 | 253.4623 | 256.2 | 0.4 | 0.421 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3377 | 0.021862 | 0.0222 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.63 | 28.3 | 253.5
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.7 | 253.6622 | 256.4 | 0.4 | 0.4413 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3378 | 0.0219 | 0.0222 | 0.0149 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3782 | 29.6 | 28.3 | 253.7
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.7 | 253.762 | 256.5 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.338 | 0.0219 | 0.022262 | 0.014962 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.6234 | 28.3 | 253.8
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.7 | 253.862 | 256.6 | 0.6 | 0.6 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.362 | 0.022 | 0.0223 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7001 | 28.3 | 253.9
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.6 | 253.8 | 256.4 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3626 | 0.022 | 0.0223 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7768 | 28.3 | 253.8
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.6 | 253.8 | 256.4 | 0.4 | 0.4428 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.4 | 0.022 | 0.022362 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8534 | 28.3 | 253.8
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.6 | 253.9 | 256.5 | 0.4 | 0.3764 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0221 | 0.0224 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.84 | 28.27 | 253.9
2008-07-16 | 13:38 | 0 | 0 | 0 | 2.6 | 253.8378 | 256.4 | 0.4 | 0.3509 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0221 | 0.0224 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.2067 | 253.8
2008-07-16 | 13:38 | 0 | -0.0005 | -5.2606 | 2.6 | 253.9 | 256.5 | 0.4 | 0.3785 | 10.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0221 | 0.0224 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.2834 | 253.9
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.8 | 254.1238 | 256.9 | 0.3 | 0.3364 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3381 | 0.022162 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.3808 | 29.8201 | 28.24 | 254.1
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.3 | 257.2 | 0.5 | 0.4603 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0222 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.618 | 29.8266 | 28.2 | 254.3
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.3 | 257.2 | 0.3 | 0.3 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0222 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.2133 | 254.3
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.3 | 257.2 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0222 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.29 | 254.3
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.362 | 257.3 | 0.3 | 0.3 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0222 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8334 | 28.3 | 254.4
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.438 | 257.3 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0222 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8566 | 28.3 | 254.4
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.9 | 254.4626 | 257.4 | 0.3 | 0.3 | 10.4 | 12.9 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.022263 | 0.0225 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7799 | 28.3 | 254.5
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.8 | 254.5 | 257.3 | 0.3 | 0.3456 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0223 | 0.022562 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7033 | 28.3 | 254.5
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.8 | 254.5623 | 257.4 | 0.3 | 0.3397 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3377 | 0.0223 | 0.0226 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7734 | 28.2266 | 254.6
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.8 | 254.6 | 257.4 | 0.3 | 0.3294 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0223 | 0.0226 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.75 | 28.25 | 254.6
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.8 | 254.6 | 257.4 | 0.4 | 0.3763 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0223 | 0.0226 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7534 | 28.2733 | 254.6
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.7 | 254.5618 | 257.3 | 0.3 | 0.3401 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0223 | 0.0226 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8966 | 28.2034 | 254.6
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.5 | 254.6 | 257.1 | 0.5 | 0.4573 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3621 | 0.0223 | 0.0227 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.82 | 28.28 | 254.6
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.5 | 254.7621 | 257.3 | 0.3 | 0.3194 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0224 | 0.0227 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7433 | 28.3 | 254.8
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.4 | 254.738 | 257.2 | 0.5 | 0.4808 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0224 | 0.022762 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 254.7
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.4 | 254.7 | 257.1 | 0.3 | 0.3 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0224 | 0.0228 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6195 | 29.7 | 28.3 | 254.7
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.5 | 254.8 | 257.3 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.0224 | 0.0228 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 254.8
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.5 | 254.9624 | 257.5 | 0.4 | 0.4 | 10.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0225 | 0.0228 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 255
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.5 | 254.9377 | 257.4 | 0.4 | 0.4428 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0225 | 0.022862 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 254.9
2008-07-16 | 13:38 | 0 | -0.001 | -10 | 2.4 | 255 | 257.4 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0225 | 0.0229 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7163 | 28.3 | 255
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 255 | 257.4 | 0.4 | 0.4264 | 10.4 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0225 | 0.0229 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.793 | 28.3 | 255
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 255.1 | 257.5 | 0.5 | 0.4573 | 10.4 | 12.9 | 954 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0225 | 0.0229 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7301 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 255.1 | 257.5 | 0.3 | 0.3442 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.022562 | 0.0229 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.6534 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 255.1 | 257.5 | 0.5 | 0.5 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0226 | 0.022962 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.6466 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 255.1 | 257.5 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3379 | 0.0226 | 0.023 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7998 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 255.038 | 257.3 | 0.5 | 0.4885 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0226 | 0.023 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7234 | 28.3 | 255
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 255.062 | 257.4 | 0.4 | 0.357 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0226 | 0.023 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8598 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 255.0381 | 257.3 | 0.4 | 0.4237 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0226 | 0.023 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 30 | 28.3 | 255
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 254.9383 | 257.2 | 0.4 | 0.3783 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0226 | 0.023 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6168 | 29.9739 | 28.3 | 254.9
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 255 | 257.3 | 0.4 | 0.435 | 10.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.022662 | 0.023062 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.6674 | 28.3 | 255
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 255.0622 | 257.4 | 0.4 | 0.3795 | 10.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.3783 | 29.6598 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.2 | 255.1 | 257.3 | 0.3 | 0.3416 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.362 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.2 | 255.1 | 257.3 | 0.5 | 0.4765 | 10.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 255.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.2 | 255.1619 | 257.4 | 0.4 | 0.4232 | 10.4 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3619 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 255.2
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.2 | 255.2 | 257.4 | 0.6 | 0.562 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3378 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 255.2
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.2 | 255.262 | 257.5 | 0.4 | 0.4 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0227 | 0.0231 | 0.015 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7431 | 28.3 | 255.3
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.3 | 260.1141 | 262.4 | 0.5 | 0.5 | 10.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0227 | 0.0232 | 0.01531 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7802 | 28.3 | 260.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3.5 | 288.005 | 291.5 | 0.5 | 0.5 | 11.01 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.02321 | 0.023548 | 0.016972 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7036 | 28.3 | 288
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3.8 | 301.08 | 304.9 | 0.6 | 0.6214 | 11.9 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0239 | 0.0241 | 0.0176 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7731 | 28.3 | 301.1
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3.4 | 304.2623 | 307.7 | 0.6 | 0.5796 | 12.25 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.024687 | 0.025049 | 0.0179 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.8 | 28.3 | 304.3
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3.2 | 304.4 | 307.6 | 0.6 | 0.5879 | 12.3 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.024962 | 0.025362 | 0.0179 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7736 | 28.3 | 304.4
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3.1 | 304.362 | 307.5 | 0.6 | 0.5881 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0251 | 0.0255 | 0.0179 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7 | 28.3 | 304.4
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 3 | 304.3381 | 307.4 | 0.5 | 0.5273 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3381 | 0.025262 | 0.025662 | 0.0179 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6192 | 29.7 | 28.3 | 304.3
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.9 | 304.9555 | 307.9 | 0.6 | 0.5768 | 12.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.025362 | 0.025862 | 0.017962 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 305
2008-07-16 | 13:39 | 0 | -0.001 | -9.6806 | 2.9 | 305.3621 | 308.3 | 0.5 | 0.5226 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.025462 | 0.0259 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7 | 28.3 | 305.4
2008-07-16 | 13:39 | 0 | -0.0001 | -1.4906 | 2.8 | 305.4621 | 308.3 | 0.7 | 0.6633 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.025562 | 0.026062 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 29.7 | 28.3 | 305.5
2008-07-16 | 13:39 | 0 | -0.0007 | -6.7381 | 2.7 | 305.438 | 308.2 | 0.5 | 0.5 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.025662 | 0.026162 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6203 | 29.7 | 28.3 | 305.4
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.7 | 305.462 | 308.2 | 0.7 | 0.6852 | 12.4 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.025762 | 0.026262 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.8254 | 28.3 | 305.5
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.6 | 305.5 | 308.1 | 0.5 | 0.5 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.025862 | 0.026362 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.9 | 28.2607 | 305.5
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.5 | 305.5 | 308 | 0.6 | 0.5897 | 12.4 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0259 | 0.026462 | 0.018 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1012 | 29.8679 | 28.2161 | 305.5
2008-07-16 | 13:39 | 0 | -0.001 | -10 | 2.4 | 305.6 | 308 | 0.5 | 0.5 | 12.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026 | 0.0265 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7141 | 28.2929 | 305.6
2008-07-16 | 13:39 | 0 | -0.0002 | -1.5878 | 2.4 | 305.5378 | 307.9 | 0.6 | 0.5927 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026 | 0.0266 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7696 | 28.3 | 305.5
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.3 | 305.7241 | 308 | 0.7 | 0.6785 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0261 | 0.0267 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3795 | 29.8 | 28.3 | 305.7
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.2 | 305.6382 | 307.9 | 0.5 | 0.5257 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0262 | 0.026762 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7541 | 28.3 | 305.6
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.2 | 305.5623 | 307.8 | 0.7 | 0.6803 | 12.4 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0262 | 0.0268 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.6008 | 28.3 | 305.6
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.1 | 305.4378 | 307.5 | 0.5 | 0.5216 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.0263 | 0.0269 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6763 | 28.3 | 305.4
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.1 | 305.462 | 307.6 | 0.7 | 0.7 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0263 | 0.0269 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 305.5
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.1 | 305.6 | 307.7 | 0.5 | 0.5228 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0263 | 0.0269 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7296 | 28.3 | 305.6
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.1 | 305.7619 | 307.9 | 0.7 | 0.7 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026362 | 0.027 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7752 | 28.3 | 305.8
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 305.8624 | 308.3 | 0.6 | 0.5683 | 12.4 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026462 | 0.027 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.4684 | 28.3 | 305.9
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 306 | 308.7 | 0.7 | 0.6889 | 12.4 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3623 | 0.026562 | 0.027062 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.638 | 28.3 | 306
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 305.9 | 308.6 | 0.6 | 0.5778 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3378 | 0.0266 | 0.0271 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8 | 28.3 | 305.9
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 305.838 | 308.5 | 0.6 | 0.5925 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3 | 0.0266 | 0.0271 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8128 | 28.3 | 305.8
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 305.8 | 308.5 | 0.7 | 0.6636 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0266 | 0.0271 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8894 | 28.3 | 305.8
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 305.8618 | 308.6 | 0.5 | 0.5229 | 12.4 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3618 | 0.0266 | 0.0271 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.9 | 28.3 | 305.9
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.7 | 305.9618 | 308.7 | 0.6 | 0.5808 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026662 | 0.027162 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8145 | 28.3 | 306
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.6 | 306.1621 | 308.8 | 0.5 | 0.5229 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.4 | 0.0267 | 0.0272 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7194 | 28.3 | 306.2
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.6 | 306.262 | 308.9 | 0.7 | 0.7 | 12.4 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.026762 | 0.0273 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.3797 | 29.7961 | 28.3 | 306.3
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 306.4 | 308.9 | 0.5 | 0.5 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0268 | 0.0273 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.7272 | 28.3 | 306.4
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 306.5 | 309 | 0.6 | 0.6 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0268 | 0.0273 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.7 | 28.3 | 306.5
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 306.4 | 308.9 | 0.6 | 0.5646 | 12.4 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.0274 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.7 | 28.3 | 306.4
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 306.4 | 308.9 | 0.6 | 0.6241 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.026862 | 0.0274 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.7054 | 28.3 | 306.4
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.4 | 306.562 | 309 | 0.6 | 0.6 | 12.5 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0269 | 0.0274 | 0.018 | 0 |  | n/a |  | 32.8931 | 55 | 1049 | 1011 | 29.8587 | 28.3 | 306.6
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.4 | 306.5381 | 308.9 | 0.7 | 0.689 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3619 | 0.0269 | 0.0274 | 0.018 | 0 |  | n/a |  | 32.1061 | 55 | 1049 | 1011 | 29.7886 | 28.3 | 306.5
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.4 | 306.3381 | 308.7 | 0.7 | 0.6571 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0269 | 0.0274 | 0.018 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29.7 | 28.3 | 306.3
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.4 | 306.3 | 308.7 | 0.5 | 0.5 | 12.5 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0269 | 0.0275 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7092 | 28.3 | 306.3
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.3 | 306.238 | 308.5 | 0.7 | 0.6788 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.027 | 0.0275 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.6198 | 29.7856 | 28.3 | 306.2
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.4 | 306.1 | 308.5 | 0.5 | 0.5222 | 12.5 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.027 | 0.0275 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8623 | 28.3 | 306.1
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.5 | 306.2 | 308.7 | 0.7 | 0.7 | 12.5 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.027 | 0.0275 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.8218 | 28.3 | 306.2
2008-07-16 | 13:39 | 0 | 0 | 0 | 2.6 | 306.2619 | 308.9 | 0.5 | 0.5 | 12.5 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.027 | 0.0275 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7 | 28.3 | 306.3
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.7 | 306.3 | 309 | 0.7 | 0.7 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.027 | 0.0275 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7 | 28.3 | 306.3
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.6 | 306.2 | 308.8 | 0.5 | 0.5411 | 12.5 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0271 | 0.0276 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.4923 | 28.3 | 306.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.6 | 306.3 | 308.9 | 0.6 | 0.6221 | 12.5 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0271 | 0.0276 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.5377 | 28.3 | 306.3
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.6 | 306.1761 | 308.8 | 0.6 | 0.5785 | 12.5 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0271 | 0.0276 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.6775 | 28.3 | 306.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.6 | 306.0619 | 308.7 | 0.6 | 0.5883 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0271 | 0.0276 | 0.018 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.3808 | 29.6008 | 28.3 | 306.1
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.5 | 306.1 | 308.6 | 0.6 | 0.5795 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0271 | 0.027662 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.6758 | 28.3 | 306.1
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.5 | 306.2 | 308.7 | 0.5 | 0.5236 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3378 | 0.027162 | 0.0277 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7 | 28.3 | 306.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.5 | 306.3242 | 308.8 | 0.6 | 0.6229 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3621 | 0.0272 | 0.0277 | 0.0181 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1013 | 29.7 | 28.3 | 306.3
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.5 | 306.462 | 309 | 0.5 | 0.5 | 12.5 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0272 | 0.0277 | 0.0181 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29.7 | 28.3 | 306.5
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.4 | 306.6238 | 309.1 | 0.7 | 0.6805 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3619 | 0.0272 | 0.0277 | 0.0181 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1011.3812 | 29.7 | 28.3 | 306.6
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.4 | 306.6381 | 309 | 0.6 | 0.5525 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0272 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1011.6188 | 29.7 | 28.3 | 306.6
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.4 | 306.8248 | 309.2 | 0.6 | 0.6186 | 12.5 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3624 | 0.0272 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7 | 28.3 | 306.8
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.4 | 306.9 | 309.3 | 0.5 | 0.5 | 12.5 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0272 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.7123 | 28.3 | 306.9
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.4 | 306.8 | 309.2 | 0.5 | 0.5234 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3622 | 0.0272 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.789 | 28.3 | 306.8
2008-07-16 | 13:40 | 0 | -0.0007 | -6.6232 | 2.4 | 306.838 | 309.2 | 0.5 | 0.5 | 12.5 | 12.9 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.338 | 0.0273 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.6688 | 28.3 | 306.8
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 2.4 | 306.7382 | 309.1 | 0.5 | 0.5418 | 12.5 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3382 | 0.0273 | 0.0278 | 0.0181 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.6185 | 29.6845 | 28.3 | 306.7
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 2.6 | 314.1577 | 316.7 | 0.7 | 0.6768 | 12.5 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3623 | 0.0273 | 0.0278 | 0.018549 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7811 | 28.3 | 314.2
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.7 | 341.4112 | 346.1 | 0.7 | 0.7236 | 13.28 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3619 | 0.027948 | 0.028148 | 0.02011 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.7044 | 28.3 | 341.4
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.3 | 350.6 | 354.9 | 0.9 | 0.8659 | 14.09 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0291 | 0.0207 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.7 | 28.3 | 350.6
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.2 | 354.838 | 359 | 0.7 | 0.7 | 14.4 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3 | 0.0296 | 0.029962 | 0.0209 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1012.6195 | 29.7 | 28.3 | 354.8
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.3 | 355.024 | 359.3 | 0.9 | 0.8786 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.338 | 0.0298 | 0.0301 | 0.0209 | 0 |  | n/a |  | 32.893 | 55 | 1049 | 1012.38 | 29.6748 | 28.3 | 355
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.3 | 355.8624 | 360.2 | 0.8 | 0.7554 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.029962 | 0.0303 | 0.021 | 0 |  | n/a |  | 32.107 | 55 | 1049 | 1012.3758 | 29.6013 | 28.3 | 355.9
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.3 | 355.9377 | 360.2 | 0.8 | 0.8 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.4 | 0.030062 | 0.030462 | 0.021 | 6.2e-05 |  | n/a |  | 33 | 55 | 1049 | 1012.623 | 29.6783 | 28.3 | 355.9
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.2 | 355.9 | 360.1 | 0.7 | 0.7 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.4 | 0.030262 | 0.030562 | 0.021 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 355.9
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.2 | 356 | 360.2 | 0.8 | 0.7853 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0303 | 0.0307 | 0.021 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 356
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4.1 | 356.1238 | 360.2 | 0.8 | 0.7782 | 14.5 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.4 | 0.030362 | 0.030762 | 0.021 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.3808 | 29.6915 | 28.3 | 356.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 4 | 356.2 | 360.2 | 0.7 | 0.7 | 14.5 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0305 | 0.030862 | 0.021 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.6225 | 29.6147 | 28.3 | 356.2
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.9 | 356.1379 | 360 | 0.8 | 0.7791 | 14.5 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.030662 | 0.031 | 0.021 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.662 | 28.3 | 356.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.8 | 356.2 | 360 | 0.7 | 0.7212 | 14.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0307 | 0.0311 | 0.021062 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 356.2
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.8 | 356.138 | 359.9 | 0.9 | 0.8594 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.338 | 0.030762 | 0.0312 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 356.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.8 | 356.1 | 359.9 | 0.7 | 0.7 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0308 | 0.0312 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.7 | 28.3 | 356.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.8 | 356.1 | 359.9 | 0.8 | 0.8 | 14.6 | 13 | 954 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.030862 | 0.0313 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.6175 | 29.6315 | 28.3685 | 356.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.8 | 356.2624 | 360.1 | 0.7 | 0.7 | 14.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0309 | 0.0313 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6451 | 28.3549 | 356.3
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 356.4622 | 360.2 | 0.8 | 0.8 | 14.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3622 | 0.0309 | 0.0314 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6782 | 28.3 | 356.5
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 356.6241 | 360.3 | 0.8 | 0.7779 | 14.6 | 13 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.031 | 0.0314 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.3795 | 29.6015 | 28.3 | 356.6
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 356.7619 | 360.5 | 0.7 | 0.6902 | 14.6 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.031 | 0.0315 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.8253 | 28.3 | 356.8
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.6 | 356.8 | 360.4 | 0.8 | 0.7797 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0311 | 0.031562 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.7966 | 28.3 | 356.8
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.6 | 356.9378 | 360.6 | 0.7 | 0.7 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0311 | 0.0316 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.6149 | 28.3284 | 356.9
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.6 | 356.9 | 360.5 | 0.8 | 0.7766 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3621 | 0.0312 | 0.0316 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.4152 | 28.3949 | 356.9
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.6 | 357 | 360.6 | 0.7 | 0.7 | 14.6 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0312 | 0.0317 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.6451 | 28.3183 | 357
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.6 | 357.138 | 360.7 | 0.8 | 0.8 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0312 | 0.0317 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011.6195 | 29.5251 | 28.3 | 357.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 357.0381 | 360.7 | 0.7 | 0.6777 | 14.6 | 13 | 955 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0313 | 0.0317 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049.5293 | 1012 | 29.4 | 28.3 | 357
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 357 | 360.7 | 0.8 | 0.8 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0313 | 0.0318 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4116 | 28.3 | 357
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 357.1 | 360.8 | 0.7 | 0.7 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0313 | 0.0318 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4882 | 28.3 | 357.1
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.7 | 357.1 | 360.8 | 0.7 | 0.7 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.031362 | 0.0318 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049.0373 | 1012 | 29.5648 | 28.3 | 357.1
2008-07-16 | 13:40 | 0 | -0.0005 | -5.0703 | 3.6 | 356.9374 | 360.6 | 0.7 | 0.6773 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0314 | 0.0319 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.6257 | 29.7244 | 28.3 | 356.9
2008-07-16 | 13:40 | 0 | 0 | 0 | 3.6 | 357 | 360.6 | 0.7 | 0.6899 | 14.6 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.031462 | 0.0319 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.8638 | 28.3181 | 357
2008-07-16 | 13:40 | 0 | 0 | 0 | 3.5 | 357.1 | 360.6 | 0.7 | 0.6883 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0315 | 0.032 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.7106 | 28.3947 | 357.1
2008-07-16 | 13:40 | 0 | 0 | 0 | 3.5 | 357.1 | 360.6 | 0.6 | 0.6214 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0315 | 0.032 | 0.0211 | 0.0001 |  | n/a |  | 32.893 | 55 | 1050 | 1012 | 29.7 | 28.4 | 357.1
2008-07-16 | 13:40 | 0 | -0.0008 | -8.431 | 3.4 | 357.2 | 360.6 | 0.7 | 0.708 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.031562 | 0.032062 | 0.0211 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7 | 28.3523 | 357.2
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.3 | 357.262 | 360.6 | 0.6 | 0.621 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0316 | 0.0321 | 0.0211 | 0.0001 |  | n/a |  | 32.107 | 55 | 1050 | 1012 | 29.6758 | 28.3 | 357.3
2008-07-16 | 13:40 | 0 | -0.001 | -10 | 3.2 | 357.3 | 360.5 | 0.8 | 0.777 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3374 | 0.0316 | 0.032163 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6 | 28.3 | 357.3
2008-07-16 | 13:40 | 0 | -0.0007 | -6.7176 | 3.2 | 357.1376 | 360.3 | 0.7 | 0.6549 | 14.6 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3376 | 0.0316 | 0.0322 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6242 | 29.6 | 28.3 | 357.1
2008-07-16 | 13:40 | 0 | 0 | 0 | 3.1 | 357.1623 | 360.3 | 0.7 | 0.7221 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0316 | 0.0322 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.377 | 29.6 | 28.3 | 357.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 3.1 | 357.2 | 360.3 | 0.7 | 0.7 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0316 | 0.0322 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6 | 28.3312 | 357.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.9 | 357.2 | 360.1 | 0.7 | 0.7202 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0317 | 0.032262 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6205 | 29.5922 | 28.3922 | 357.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.9 | 357 | 359.9 | 0.7 | 0.7 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0317 | 0.0323 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.5156 | 28.3156 | 357
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.8 | 357.2 | 360 | 0.7 | 0.7414 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3622 | 0.0317 | 0.0323 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.5611 | 28.3 | 357.2
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.8 | 357.4 | 360.2 | 0.8 | 0.777 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0317 | 0.0323 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6377 | 28.3 | 357.4
2008-07-16 | 13:40 | 0 | 0 | 0 | 2.8 | 357.2379 | 360 | 0.7 | 0.7224 | 14.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3621 | 0.0317 | 0.0323 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.621 | 29.7143 | 28.3 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.8 | 357.2 | 360 | 0.8 | 0.754 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0317 | 0.0323 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.791 | 28.3 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.2 | 359.9 | 0.7 | 0.7 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0317 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7324 | 28.3676 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.2 | 359.9 | 0.8 | 0.7611 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0317 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7442 | 28.3558 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.2 | 359.9 | 0.7 | 0.7 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3376 | 0.0318 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7791 | 28.3 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.1623 | 359.9 | 0.7 | 0.7216 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.0318 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.7025 | 28.3 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.3242 | 360 | 0.6 | 0.6 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0318 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.379 | 29.6258 | 28.3 | 357.3
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.4 | 360.1 | 0.7 | 0.7202 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0318 | 0.0324 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6508 | 28.3 | 357.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.5 | 360.2 | 0.7 | 0.6818 | 14.6 | 13 | 954 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.0318 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.7 | 28.3 | 357.5
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.4146 | 360.1 | 0.6 | 0.6357 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3618 | 0.0318 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6878 | 28.3 | 357.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.4242 | 360.1 | 0.7 | 0.6774 | 14.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3379 | 0.0318 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4579 | 28.3 | 357.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.5 | 360.2 | 0.6 | 0.6199 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0319 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4 | 28.3 | 357.5
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 357.238 | 359.9 | 0.7 | 0.6627 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0319 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 32.893 | 55 | 1050 | 1012.6203 | 29.4679 | 28.3339 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.6 | 357.1 | 359.7 | 0.5 | 0.5199 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0319 | 0.0325 | 0.0211 | 0.0001 |  | n/a |  | 32.107 | 55 | 1050 | 1013 | 29.6106 | 28.3894 | 357.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.6 | 357.2 | 359.8 | 0.7 | 0.6782 | 14.6 | 13 | 954 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3626 | 0.0319 | 0.0326 | 0.0211 | 0.0001 |  | n/a |  | 32.893 | 55 | 1050 | 1013 | 29.6872 | 28.3128 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.6 | 357.1624 | 359.8 | 0.6 | 0.6 | 14.6 | 13 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0319 | 0.0326 | 0.0211 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.3635 | 357.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 2.7 | 364.7311 | 367.4 | 0.7 | 0.7223 | 14.6 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0319 | 0.0326 | 0.021549 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.6599 | 28.3599 | 364.7
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 390.7091 | 394.7 | 0.8 | 0.8 | 15.27 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.032711 | 0.033249 | 0.023111 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.617 | 28.3 | 390.7
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.5 | 402.6578 | 407.2 | 0.9 | 0.8886 | 16.19 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.033648 | 0.034048 | 0.023824 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.6934 | 28.3 | 402.7
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.2 | 404.9 | 409.1 | 0.9 | 0.8773 | 16.49 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.4 | 0.0341 | 0.0346 | 0.0239 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.7 | 28.3 | 404.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 405.1 | 409.1 | 0.7 | 0.7338 | 16.6 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034386 | 0.034948 | 0.0239 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.7 | 28.3468 | 405.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.7 | 405.2621 | 409 | 0.8 | 0.7792 | 16.6 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.034562 | 0.035224 | 0.023962 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.6291 | 28.4 | 405.3
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.6 | 405.362 | 409 | 0.7 | 0.721 | 16.6 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.034762 | 0.035362 | 0.024 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.4006 | 28.3997 | 405.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.5 | 405.9956 | 409.5 | 0.8 | 0.8 | 16.6 | 13 | 954 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.034862 | 0.0355 | 0.024 | 0.0001 |  | n/a |  | 32.107 | 55 | 1050 | 1012 | 29.5539 | 28.3231 | 406
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.5 | 406.4381 | 409.9 | 0.7 | 0.7 | 16.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0349 | 0.0356 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4929 | 28.3 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.4 | 406.4 | 409.8 | 0.8 | 0.8 | 16.6 | 13 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.035 | 0.0357 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4302 | 28.3302 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.3 | 406.4243 | 409.8 | 0.7 | 0.6777 | 16.65 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3378 | 0.035062 | 0.035762 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4932 | 28.4 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.2 | 406.4378 | 409.7 | 0.7 | 0.7239 | 16.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.035162 | 0.0359 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6217 | 29.4166 | 28.4 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.2 | 406.438 | 409.6 | 0.6 | 0.6 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.6 | 5 | 195 | 59.4 | 0.035262 | 0.035962 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6205 | 29.4 | 28.4 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.8 | 406.3619 | 410.2 | 0.7 | 0.729 | 16.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28.4 | 5 | 195 | 59.4 | 0.0354 | 0.036 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.5103 | 28.3632 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.8 | 406.4 | 410.2 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.035462 | 0.0361 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.6598 | 28.3134 | 406.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.8 | 406.6 | 410.4 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.4 | 0.0355 | 0.0361 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4299 | 28.39 | 406.6
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 406.5 | 410.4 | 0.8 | 0.7609 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0355 | 0.0361 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.3333 | 406.5
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 406.638 | 410.5 | 0.6 | 0.6215 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.035562 | 0.036162 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4866 | 28.3433 | 406.6
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 406.7238 | 410.6 | 0.8 | 0.7775 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0356 | 0.0362 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.6199 | 28.3801 | 406.7
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 406.8 | 410.7 | 0.6 | 0.6 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0357 | 0.036263 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.6966 | 28.3034 | 406.8
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 406.9246 | 410.8 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0357 | 0.0363 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.377 | 29.7 | 28.3731 | 406.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 407 | 410.9 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0357 | 0.0363 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6503 | 28.3503 | 407
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 407.1 | 411 | 0.7 | 0.6866 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.338 | 0.0358 | 0.036362 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6 | 28.3 | 407.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 407.3618 | 411.3 | 0.7 | 0.6782 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0358 | 0.0364 | 0.024 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011.3815 | 29.6 | 28.303 | 407.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 407.2377 | 411.1 | 0.6 | 0.626 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3377 | 0.035862 | 0.0364 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6 | 28.3796 | 407.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.2 | 407.2 | 411.4 | 0.7 | 0.677 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0359 | 0.0365 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.6563 | 28.3437 | 407.2
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.3 | 407.1379 | 411.4 | 0.5 | 0.5434 | 16.7 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3621 | 0.036 | 0.0365 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.6345 | 28.3 | 407.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.2 | 407 | 411.2 | 0.7 | 0.7 | 16.7 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.036 | 0.0365 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.5 | 28.3091 | 407
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.2 | 407 | 411.2 | 0.5 | 0.5225 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.036 | 0.0365 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.5 | 28.3861 | 407
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.2 | 406.8 | 411 | 0.7 | 0.6789 | 16.7 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.036062 | 0.0366 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.5626 | 28.4 | 406.8
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.1 | 406.9248 | 411.1 | 0.6 | 0.5658 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3376 | 0.0361 | 0.0366 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6 | 28.3608 | 406.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.1 | 406.9377 | 411 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0361 | 0.0367 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6 | 28.3 | 406.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 406.9 | 410.9 | 0.6 | 0.6 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4379 | 0.036162 | 0.0367 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6 | 28.3 | 406.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 406.9 | 410.9 | 0.6 | 0.5902 | 16.7 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0362 | 0.0367 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.6 | 28.3693 | 406.9
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 406.9619 | 411 | 0.6 | 0.6 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0362 | 0.036762 | 0.0241 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1013 | 29.508 | 28.354 | 407
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.0622 | 411.1 | 0.7 | 0.6733 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0362 | 0.0368 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.4227 | 28.3 | 407.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.1 | 411.1 | 0.7 | 0.6808 | 16.7 | 12.9 | 954 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0362 | 0.0368 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.4992 | 28.3 | 407.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.138 | 411.1 | 0.5 | 0.5337 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.6 | 5 | 195 | 59.338 | 0.036262 | 0.0368 | 0.0241 | 0.0001 |  | n/a |  | 32.107 | 55 | 1049 | 1013 | 29.5 | 28.3759 | 407.1
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.2859 | 411.3 | 0.8 | 0.8 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28.4 | 5 | 195 | 59.4 | 0.0363 | 0.0368 | 0.0241 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1012.3805 | 29.4475 | 28.4 | 407.3
2008-07-16 | 13:41 | 0 | 0 | 0 | 3.9 | 407.4 | 411.3 | 0.7 | 0.7114 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0363 | 0.0368 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.3708 | 28.3708 | 407.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.3 | 411.3 | 0.8 | 0.8 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3624 | 0.0363 | 0.036862 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.3233 | 28.3 | 407.3
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.4 | 411.4 | 0.7 | 0.6673 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0363 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.6298 | 28.3 | 407.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.3758 | 411.4 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0363 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.621 | 29.5228 | 28.3591 | 407.4
2008-07-16 | 13:41 | 0 | 0 | 0 | 4 | 407.5621 | 411.6 | 0.7 | 0.6798 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0363 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.4 | 28.4 | 407.6
2008-07-16 | 13:41 | 0 | 0 | 0 | 4.1 | 407.5381 | 411.6 | 0.7 | 0.6908 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0363 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.4247 | 28.4 | 407.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.2 | 407.6 | 411.8 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.5779 | 28.4 | 407.6
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.2 | 407.6 | 411.8 | 0.7 | 0.656 | 16.7 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32.107 | 55 | 1049 | 1012 | 29.4688 | 28.4 | 407.6
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.8242 | 412.1 | 0.8 | 0.7807 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1011.379 | 29.4 | 28.4 | 407.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.938 | 412.2 | 0.7 | 0.7232 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32.107 | 55 | 1049 | 1011 | 29.4566 | 28.4 | 407.9
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.762 | 412.1 | 0.9 | 0.8601 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.6864 | 28.4 | 407.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.675 | 412 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3375 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29.4113 | 28.4 | 407.7
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.6 | 411.9 | 0.8 | 0.8 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 32.1069 | 55 | 1049 | 1012 | 29.3 | 28.4 | 407.6
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 407.3754 | 411.7 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.623 | 29.3254 | 28.4 | 407.4
2008-07-16 | 13:42 | 0 | 0.0007 | 6.5657 | 4.3 | 407.3 | 411.6 | 0.8 | 0.8 | 16.7 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.4 | 28.4 | 407.3
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.3 | 407.4 | 411.7 | 0.7 | 0.7386 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0364 | 0.0369 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 29.4 | 28.4 | 407.4
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.2 | 407.5 | 411.7 | 0.8 | 0.7943 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0365 | 0.037 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.4 | 28.4 | 407.5
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.2 | 407.7619 | 412 | 0.7 | 0.7 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4.4 | 195 | 59.3381 | 0.0365 | 0.037 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.4315 | 28.3685 | 407.8
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.1 | 407.9 | 412 | 0.7 | 0.7273 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4.6 | 195 | 59.4 | 0.0365 | 0.037 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.4833 | 28.3 | 407.9
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.1 | 407.9621 | 412.1 | 0.8 | 0.7806 | 16.7 | 12.9 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0365 | 0.0371 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.3298 | 28.3 | 408
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4 | 407.938 | 411.9 | 0.6 | 0.6173 | 16.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0365 | 0.0371 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.3 | 28.3618 | 407.9
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4 | 407.9 | 411.9 | 0.8 | 0.7616 | 16.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.036562 | 0.0371 | 0.0241 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.4152 | 28.4 | 407.9
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 4.5 | 424.5091 | 429 | 0.6 | 0.6308 | 16.79 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.036725 | 0.0372 | 0.025126 | 0.0001 |  | n/a |  | 33 | 55 | 1049.9021 | 1011.6263 | 29.5549 | 28.4 | 424.5
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 5.8 | 447.717 | 453.6 | 0.9 | 0.8778 | 17.78 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4.4 | 195 | 59.3376 | 0.038336 | 0.038699 | 0.026475 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3251 | 28.4 | 447.7
2008-07-16 | 13:42 | 0 | 0.001 | 10 | 5.3 | 454.7191 | 460 | 0.8 | 0.766 | 18.38 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.039387 | 0.039849 | 0.026862 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3 | 28.4 | 454.7
2008-07-16 | 13:42 | 0 | 0.0002 | 2.0383 | 5.1 | 455.462 | 460.6 | 0.8 | 0.8237 | 18.55 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.039562 | 0.040062 | 0.026962 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.39 | 28.355 | 455.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 5 | 455.7238 | 460.8 | 0.8 | 0.777 | 18.6 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4 | 195 | 59.4 | 0.039662 | 0.040162 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4784 | 28.3216 | 455.7
2008-07-16 | 13:42 | 0 | 0 | 0 | 5 | 455.8 | 460.8 | 0.8 | 0.7857 | 18.6 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4.6 | 195 | 59.4 | 0.0397 | 0.040262 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4018 | 28.3982 | 455.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 5 | 455.8244 | 460.8 | 0.9 | 0.8791 | 18.69 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0398 | 0.0403 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3782 | 29.2503 | 28.4 | 455.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.9 | 456 | 460.9 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0399 | 0.0404 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2515 | 28.4 | 456
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.8 | 456.1 | 460.9 | 0.9 | 0.8799 | 18.7 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0399 | 0.0405 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3281 | 28.3719 | 456.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.8 | 456 | 460.8 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.04 | 0.040563 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.3 | 456
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.7 | 456.7618 | 461.5 | 0.9 | 0.8669 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0401 | 0.0407 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.3 | 456.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.7 | 457.1623 | 461.9 | 0.7 | 0.7 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.040162 | 0.040762 | 0.027 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.516 | 28.358 | 457.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.6 | 457.1622 | 461.8 | 0.9 | 0.8879 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3378 | 0.0402 | 0.0408 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4961 | 28.3654 | 457.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.1378 | 461.6 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3378 | 0.0403 | 0.0409 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6217 | 29.2887 | 28.3113 | 457.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.1 | 461.6 | 0.9 | 0.9 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.041 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2121 | 28.3879 | 457.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 457.1619 | 461.6 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3381 | 0.040362 | 0.041 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3808 | 29.3291 | 28.4 | 457.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.1 | 461.4 | 0.8 | 0.8447 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0404 | 0.0411 | 0.027062 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.4 | 457.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 456.9378 | 461.2 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.040462 | 0.0411 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.4 | 456.9
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.1 | 461.4 | 0.7 | 0.7231 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0405 | 0.0412 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.4 | 28.4 | 457.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.1239 | 461.4 | 0.9 | 0.877 | 18.7 | 12.9 | 956 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0405 | 0.0412 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2581 | 28.4 | 457.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.3 | 461.6 | 0.8 | 0.8206 | 18.7 | 12.9 | 956 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0405 | 0.0412 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2476 | 28.4 | 457.3
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.2376 | 461.5 | 1 | 0.9676 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.040562 | 0.0412 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3 | 28.4 | 457.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 457.3246 | 461.6 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0406 | 0.041262 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3007 | 28.4 | 457.3
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 457.6255 | 462 | 0.9 | 0.9 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0406 | 0.0413 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3723 | 29.3776 | 28.4 | 457.6
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 457.862 | 462.3 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0406 | 0.0413 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2915 | 28.4 | 457.9
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 458.0238 | 462.5 | 0.9 | 0.8602 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0407 | 0.0413 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1691 | 28.4 | 458
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0407 | 0.0413 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 29.115 | 28.4 | 458.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.8 | 0.8117 | 18.7 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0407 | 0.0413 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1011 | 29.2684 | 28.4 | 458.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 458.1 | 462.6 | 0.9 | 0.8766 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3378 | 0.0407 | 0.041362 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1783 | 28.4 | 458.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.9382 | 462.4 | 0.8 | 0.8283 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0407 | 0.0414 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2123 | 28.4 | 457.9
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.8 | 462.3 | 1 | 0.9611 | 18.7 | 12.9 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0408 | 0.0414 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.386 | 28.386 | 457.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.6 | 462.1 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0408 | 0.0414 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3093 | 28.3093 | 457.6
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.5 | 462 | 1 | 0.9655 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0408 | 0.0414 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3673 | 28.3673 | 457.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.5 | 457.4377 | 461.9 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0408 | 0.0414 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.623 | 29.2683 | 28.4 | 457.4
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 457.5241 | 461.9 | 0.9 | 0.8615 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0408 | 0.0415 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.3795 | 29.1616 | 28.4 | 457.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 457.8 | 462.2 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.0408 | 0.0415 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3915 | 28.4 | 457.8
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 458.0618 | 462.4 | 0.8 | 0.8178 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 4 | 195 | 59.4 | 0.0409 | 0.0415 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.382 | 29.1786 | 28.4 | 458.1
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.2244 | 462.6 | 0.8 | 0.8 | 18.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3622 | 0.0409 | 0.0415 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2008 | 28.4 | 458.2
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.3 | 462.7 | 0.8 | 0.7883 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3621 | 0.0409 | 0.0415 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.327 | 28.4 | 458.3
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.5 | 462.9 | 0.8 | 0.779 | 18.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.338 | 0.0409 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4 | 28.3963 | 458.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.4 | 462.8 | 0.7 | 0.7273 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.438 | 0.0409 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.4 | 28.3196 | 458.4
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.5253 | 462.9 | 0.8 | 0.7772 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0409 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.286 | 28.357 | 458.5
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.3 | 458.3 | 462.6 | 0.7 | 0.7111 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3617 | 0.0409 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2336 | 28.4 | 458.3
2008-07-16 | 13:42 | 0 | 0 | 0 | 4.4 | 458.4623 | 462.9 | 0.8 | 0.8 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3103 | 28.4 | 458.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 458.4 | 462.8 | 0.7 | 0.7 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3379 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3868 | 28.4 | 458.4
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 458.5238 | 462.9 | 0.8 | 0.8161 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3368 | 28.3368 | 458.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.6 | 463.1 | 0.8 | 0.767 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3 | 28.3396 | 458.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 458.6234 | 463.2 | 0.8 | 0.8227 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3617 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3161 | 28.4 | 458.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 458.6 | 463.2 | 0.8 | 0.7781 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.338 | 0.041 | 0.0416 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.3931 | 28.4 | 458.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 458.5 | 463.1 | 0.7 | 0.7478 | 18.7 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.041062 | 0.0417 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3301 | 28.4 | 458.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 458.438 | 463 | 0.8 | 0.7768 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0411 | 0.0417 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2534 | 28.4 | 458.4
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 458.338 | 462.9 | 0.7 | 0.722 | 18.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0411 | 0.0417 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1768 | 28.4 | 458.3
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.1 | 462.6 | 0.9 | 0.889 | 18.7 | 12.9 | 956 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.0411 | 0.0417 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.1002 | 28.4 | 458.1
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.1 | 462.6 | 0.8 | 0.8 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.4 | 0.0411 | 0.041762 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.0235 | 28.4 | 458.1
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.0377 | 462.5 | 0.9 | 0.9 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.0411 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.2127 | 28.4 | 458
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.1243 | 462.6 | 0.8 | 0.7783 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.041162 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.3404 | 28.4 | 458.1
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.8 | 0.822 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0412 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.2 | 28.3936 | 458.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.8 | 0.7664 | 18.7 | 12.9 | 955 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0412 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.2 | 28.3169 | 458.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.7 | 0.7097 | 18.7 | 12.9 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.4 | 0.0412 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.1403 | 28.3597 | 458.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 458.2 | 462.7 | 0.8 | 0.7787 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 4 | 195 | 59.4 | 0.0412 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1014 | 29.1 | 28.4 | 458.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 458.5863 | 463 | 0.7 | 0.744 | 18.7 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0412 | 0.0418 | 0.0271 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.113 | 28.387 | 458.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 5 | 467.7561 | 472.7 | 0.8 | 0.8 | 18.7 | 12.9 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.4 | 0.041262 | 0.041862 | 0.027658 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1896 | 28.3104 | 467.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 6.6 | 489.3879 | 495.9 | 0.8 | 0.8279 | 19.43 | 12.9 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3626 | 0.041851 | 0.042151 | 0.028976 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3737 | 29.2 | 28.3662 | 489.4
2008-07-16 | 13:43 | 0 | 0 | 0 | 6.7 | 499.817 | 506.5 | 1 | 0.981 | 20.28 | 12.9 | 957 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.042499 | 0.042799 | 0.029575 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2 | 28.4 | 499.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.9 | 505.946 | 511.9 | 0.8 | 0.8 | 20.6 | 12.9 | 958 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.043474 | 0.043936 | 0.029925 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1805 | 28.4 | 505.9
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.5 | 507.321 | 512.9 | 1 | 1 | 20.64 | 12.9 | 958 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.043762 | 0.044362 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1038 | 28.4 | 507.3
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.4 | 507.6 | 513 | 0.9 | 0.8763 | 20.7 | 12.9 | 958 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0439 | 0.0445 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2456 | 28.4 | 507.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.3 | 507.6375 | 513 | 1 | 1.0225 | 20.7 | 12.9 | 958 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.044 | 0.044662 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.2011 | 28.4 | 507.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.2 | 507.5 | 512.7 | 1 | 0.9685 | 20.7 | 12.9 | 958 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.044162 | 0.0448 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1261 | 28.3739 | 507.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.2 | 507.5622 | 512.8 | 1 | 1.0213 | 20.7 | 12.9 | 958 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0442 | 0.044862 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2054 | 28.3027 | 507.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.1 | 507.8242 | 513 | 1 | 0.9806 | 20.7 | 13 | 957 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0443 | 0.045 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.379 | 29.3586 | 28.3793 | 507.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 5.1 | 507.762 | 512.9 | 1 | 0.9881 | 20.7 | 13 | 957 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0444 | 0.0451 | 0.03 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2322 | 28.4 | 507.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 5 | 508.538 | 513.6 | 1.1 | 1.0775 | 20.7 | 13 | 957 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.362 | 0.044462 | 0.045162 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1649 | 28.3675 | 508.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.9 | 508.8 | 513.7 | 0.9 | 0.9278 | 20.7 | 13 | 957 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.044562 | 0.045262 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2911 | 28.3089 | 508.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.9 | 508.8 | 513.7 | 1.1 | 1.0888 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 4.4 | 195 | 59.4 | 0.0446 | 0.045362 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2146 | 28.3854 | 508.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.8 | 508.8623 | 513.7 | 1 | 0.9776 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.044662 | 0.0454 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2622 | 28.3378 | 508.9
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.8 | 508.7758 | 513.6 | 1.1 | 1.0778 | 20.8 | 13 | 956 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.044762 | 0.0455 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.339 | 28.3 | 508.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.7 | 508.7237 | 513.4 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0448 | 0.045562 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.3371 | 28.3157 | 508.7
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 508.6762 | 513.3 | 1 | 1.0212 | 20.8 | 13 | 957 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.044862 | 0.045662 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.0304 | 28.3924 | 508.7
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 508.8 | 513.4 | 1 | 1 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0449 | 0.0457 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1381 | 28.4 | 508.8
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 508.9 | 513.5 | 1 | 1.0234 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3621 | 0.045 | 0.0457 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1543 | 28.4 | 508.9
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.6 | 509.0861 | 513.7 | 1 | 0.979 | 20.8 | 13 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045 | 0.045762 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3797 | 29.1 | 28.3777 | 509.1
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 509.2 | 513.7 | 1 | 0.9894 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045 | 0.0458 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1 | 28.3011 | 509.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 509.2 | 513.7 | 1.1 | 1.0772 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045062 | 0.045862 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1756 | 28.3 | 509.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.5 | 509.3 | 513.8 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0451 | 0.0459 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1478 | 28.3522 | 509.3
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 509.2 | 513.6 | 1 | 0.9912 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0451 | 0.0459 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1576 | 28.3712 | 509.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 509.2 | 513.6 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045162 | 0.046 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2891 | 28.3 | 509.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.4 | 509.2 | 513.6 | 1 | 1.0386 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0452 | 0.046 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1359 | 28.3 | 509.2
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.3 | 513.6 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.0452 | 0.046 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1587 | 28.3 | 509.3
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.4622 | 513.8 | 1 | 1 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0452 | 0.0461 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1647 | 28.3 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.4244 | 513.7 | 1 | 0.9644 | 20.8 | 13 | 956 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045262 | 0.0461 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3782 | 29.1 | 28.312 | 509.4
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.5 | 513.8 | 0.9 | 0.9229 | 20.8 | 13 | 956 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0453 | 0.0461 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1 | 28.3886 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.638 | 513.9 | 0.9 | 0.9398 | 20.8 | 13 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0453 | 0.046162 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1652 | 28.4 | 509.6
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.4763 | 513.8 | 0.9 | 0.9439 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0453 | 0.0462 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.6183 | 29.2 | 28.4 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.3376 | 513.6 | 1 | 0.9782 | 20.8 | 13 | 956 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0453 | 0.0462 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1815 | 28.4 | 509.3
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.3 | 509.5 | 513.8 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0453 | 0.0462 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1049 | 28.4 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.4 | 513.6 | 1 | 1 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.045362 | 0.0462 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.1717 | 28.4 | 509.4
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.138 | 513.3 | 0.9 | 0.9 | 20.8 | 13 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0454 | 0.0462 | 0.0301 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013.6205 | 29.2 | 28.3519 | 509.1
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.5238 | 513.7 | 1 | 1.0288 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0454 | 0.0462 | 0.030162 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2 | 28.3244 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.6618 | 513.9 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0454 | 0.046262 | 0.0302 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2 | 28.3992 | 509.7
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.5 | 513.7 | 1 | 1.0195 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0454 | 0.0463 | 0.0302 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29.2 | 28.3221 | 509.5
2008-07-16 | 13:43 | 0 | 0 | 0 | 4.2 | 509.8241 | 514 | 1 | 0.9662 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0454 | 0.0463 | 0.0302 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3797 | 29.1454 | 28.3546 | 509.8
2008-07-16 | 13:43 | 0 | -0.0003 | -2.7727 | 4.1 | 509.9 | 514 | 0.9 | 0.9095 | 20.8 | 13 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0454 | 0.0463 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1 | 28.4 | 509.9
2008-07-16 | 13:43 | 0 | -0.0009 | -8.9 | 4.1 | 509.9764 | 514.1 | 1 | 0.977 | 20.8 | 13 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0454 | 0.0463 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1 | 28.4 | 510
2008-07-16 | 13:44 | 0 | -0.0001 | -0.56667 | 4.1 | 509.8381 | 513.9 | 0.9 | 0.9453 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.0454 | 0.0463 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012.6188 | 29.1 | 28.4 | 509.8
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.9248 | 514 | 1.1 | 1.0778 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3624 | 0.045462 | 0.0463 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012.3758 | 29.0385 | 28.3385 | 509.9
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.0377 | 514.1 | 0.9 | 0.9 | 20.8 | 13 | 956 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29 | 28.3382 | 510
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510 | 514.1 | 1 | 0.9778 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29 | 28.3852 | 510
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.9 | 514 | 0.9 | 0.9 | 20.8 | 13 | 955 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 33 | 55 | 1050 | 1013 | 29 | 28.3086 | 509.9
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.7762 | 513.9 | 1 | 1.0339 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 32.893 | 55 | 1050 | 1013 | 29.068 | 28.3 | 509.8
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.8236 | 513.9 | 0.9 | 0.9 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.1 | 28.3 | 509.8
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.9 | 514 | 0.9 | 0.9236 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0455 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.1 | 28.3 | 509.9
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 509.8379 | 513.9 | 1 | 0.9657 | 20.8 | 13 | 955 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3379 | 0.0456 | 0.0464 | 0.0302 | 0.0002 |  | n/a |  | 32 | 55 | 1050 | 1013 | 29.1 | 28.3 | 509.8
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.1 | 514.2 | 0.9 | 0.9192 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0456 | 0.0464 | 0.0302 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.2491 | 28.3 | 510.1
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.038 | 514.1 | 1 | 0.9775 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012.6195 | 29.1976 | 28.3512 | 510
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.2253 | 514.3 | 0.9 | 0.9421 | 20.8 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.0443 | 28.3722 | 510.2
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.3 | 514.4 | 1.1 | 1.0778 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.909 | 28.3045 | 510.3
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.3 | 514.4 | 0.9 | 0.9 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 32 | 55 | 1050 | 1012 | 29.0623 | 28.3811 | 510.3
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.2 | 510.238 | 514.4 | 1 | 1 | 20.8 | 13 | 955 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 32.107 | 55 | 1050 | 1012 | 29.1 | 28.3422 | 510.2
2008-07-16 | 13:44 | 0 | 0 | 0 | 4.1 | 510.2 | 514.3 | 0.9 | 0.9 | 20.8 | 13 | 955 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0456 | 0.0465 | 0.0302 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012 | 29.1 | 28.3344 | 510.2
2008-07-16 | 13:44 | 0 | 0 | 0 | 3.7 | 492.2008 | 495.9 | 1 | 1 | 20.64 | 13 | 956 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0456 | 0.0465 | 0.029129 | 0.0001 |  | n/a |  | 33 | 55 | 1050 | 1012.3815 | 29.1 | 28.389 | 492.2
2008-07-16 | 13:44 | 0 | 0 | 0 | 1.9 | 472.579 | 474.5 | 1.1 | 1.1 | 19.59 | 13 | 954 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3383 | 0.04503 | 0.046192 | 0.027992 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.1 | 28.3123 | 472.6
2008-07-16 | 13:44 | 0 | 0 | 0 | 1.7 | 465.7621 | 467.4 | 1 | 1.0212 | 19.04 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.044738 | 0.045938 | 0.0276 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.1 | 28.3 | 465.8
2008-07-16 | 13:44 | 0 | 0 | 0 | 1.7 | 464.6215 | 466.3 | 1 | 0.9723 | 18.87 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.044638 | 0.045838 | 0.027538 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.1 | 28.341 | 464.6
2008-07-16 | 13:44 | 0 | 0 | 0 | 1.8 | 463.2044 | 465 | 1 | 0.9604 | 18.8 | 13 | 949 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.044538 | 0.045676 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0649 | 28.4 | 463.2
2008-07-16 | 13:44 | 0 | -0.0005 | -4.5687 | 1.9 | 462.9 | 464.8 | 1 | 0.9762 | 18.76 | 13 | 949 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.044338 | 0.045538 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.9125 | 28.4 | 462.9
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.8 | 462.8752 | 464.7 | 0.8 | 0.8 | 18.7 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0443 | 0.0454 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0414 | 28.3293 | 462.9
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.6 | 462.7378 | 464.3 | 1 | 0.9795 | 18.7 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0441 | 0.0453 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0053 | 28.3 | 462.7
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.5 | 462.6 | 464.1 | 0.9 | 0.8731 | 18.7 | 13 | 950 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.044 | 0.0452 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.9239 | 28.3 | 462.6
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.6 | 462.6 | 464.2 | 0.9 | 0.9 | 18.7 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.4 | 0.043938 | 0.045138 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0007 | 28.3 | 462.6
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.7 | 462.5763 | 464.3 | 0.8 | 0.8 | 18.7 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0438 | 0.045 | 0.0274 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29.0775 | 28.3 | 462.6
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.8 | 462.5383 | 464.3 | 0.8 | 0.8349 | 18.7 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0437 | 0.0449 | 0.0274 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3542 | 462.5
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.9 | 462.4 | 464.3 | 0.8 | 0.8 | 18.7 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3622 | 0.043638 | 0.0448 | 0.0274 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0384 | 28.3692 | 462.4
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.9 | 462.4379 | 464.3 | 0.8 | 0.8234 | 18.7 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3621 | 0.0436 | 0.044738 | 0.0274 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9074 | 28.3 | 462.4
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.9 | 462.338 | 464.2 | 0.9 | 0.8645 | 18.7 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.043538 | 0.0447 | 0.0274 | 0.0001 |  | n/a |  | 32.107 | 55 | 1049 | 1012 | 28.9841 | 28.3 | 462.3
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2 | 462.238 | 464.2 | 0.7 | 0.7221 | 18.7 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0435 | 0.0447 | 0.0274 | 0.0001 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29 | 28.3 | 462.2
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2 | 462.1381 | 464.1 | 0.8 | 0.7772 | 18.7 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0435 | 0.0446 | 0.0274 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9627 | 28.3 | 462.1
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2 | 461.9622 | 464 | 0.7 | 0.7 | 18.7 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0434 | 0.044538 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.3778 | 28.928 | 28.3 | 462
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 1.9 | 461.8758 | 463.8 | 0.8 | 0.7781 | 18.7 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.4 | 0.043338 | 0.0445 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.621 | 29.0813 | 28.3 | 461.9
2008-07-16 | 13:44 | -0.008 | -0.001 | -10 | 1.8 | 461.7759 | 463.6 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0433 | 0.0445 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.6205 | 29.1673 | 28.3 | 461.8
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.8 | 461.8237 | 463.6 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.4 | 0.043238 | 0.0444 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.3815 | 29.1561 | 28.3 | 461.8
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.9 | 461.9 | 463.8 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0432 | 0.0444 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3 | 461.9
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.9 | 462.1244 | 464 | 0.7 | 0.7394 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.4 | 0.0432 | 0.044338 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.3782 | 29.1 | 28.3 | 462.1
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.9 | 462.3 | 464.2 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 4.4 | 195 | 59.4 | 0.043138 | 0.0443 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.8782 | 28.3739 | 462.3
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.9 | 462.3618 | 464.3 | 0.7 | 0.7232 | 18.7 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 4 | 195 | 59.4 | 0.0431 | 0.0443 | 0.0273 | 0.000162 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9011 | 28.3494 | 462.4
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 1.9 | 462.562 | 464.5 | 0.8 | 0.7658 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0431 | 0.0442 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1010.3805 | 29 | 28.3 | 462.6
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 2 | 462.5381 | 464.5 | 0.7 | 0.7211 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.043038 | 0.0442 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1010.6188 | 29.0038 | 28.3 | 462.5
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 2 | 462.5 | 464.5 | 0.8 | 0.8 | 18.7 | 12.9 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.043 | 0.0442 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0805 | 28.3 | 462.5
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 2 | 462.3378 | 464.3 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.043 | 0.0441 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.3 | 462.3
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 2.1 | 462.2379 | 464.3 | 0.8 | 0.8 | 18.7 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.042938 | 0.0441 | 0.0273 | 0.000138 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0324 | 28.3 | 462.2
2008-07-16 | 13:44 | -0.01 | -0.001 | -10 | 2.2 | 462.2 | 464.4 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0429 | 0.044 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9104 | 28.3 | 462.2
2008-07-16 | 13:44 | -0.01 | -0.0009 | -8.6405 | 2.2 | 461.9381 | 464.1 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0429 | 0.043938 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.987 | 28.3 | 461.9
2008-07-16 | 13:44 | -0.01 | 0 | -0.45045 | 2.3 | 461.8376 | 464.1 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.0428 | 0.0439 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9363 | 28.3 | 461.8
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.3 | 461.8 | 464.1 | 0.8 | 0.8411 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0428 | 0.0439 | 0.0273 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9403 | 28.3 | 461.8
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.3 | 461.862 | 464.2 | 0.8 | 0.7827 | 18.7 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.4 | 0.0428 | 0.043838 | 0.0273 | 0.000162 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29 | 28.3 | 461.9
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.4 | 461.838 | 464.2 | 0.8 | 0.8225 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0428 | 0.0438 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29 | 28.3 | 461.8
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.4 | 461.5381 | 463.9 | 0.8 | 0.8362 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 4.6 | 195 | 59.4 | 0.0427 | 0.0438 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1401 | 28.3 | 461.5
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.4 | 461.4 | 463.8 | 0.8 | 0.811 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3624 | 0.0427 | 0.0438 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1531 | 28.3 | 461.4
2008-07-16 | 13:44 | -0.01 | 0 | 0 | 2.4 | 461.2378 | 463.6 | 0.9 | 0.8797 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3378 | 0.0427 | 0.0438 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0764 | 28.3 | 461.2
2008-07-16 | 13:44 | -0.006 | -0.0008 | -7.6063 | 2.4 | 461.0758 | 463.5 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3621 | 0.0427 | 0.043738 | 0.0273 | 0.0002 |  | n/a |  | 32.107 | 55 | 1049 | 1012.621 | 28.9997 | 28.3 | 461.1
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2.4 | 460.8761 | 463.3 | 0.9 | 0.8789 | 18.7 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0427 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32.893 | 55 | 1049 | 1013 | 28.923 | 28.3 | 460.9
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2.4 | 460.8 | 463.2 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0427 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9 | 28.3 | 460.8
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2.3 | 460.7383 | 463 | 0.8 | 0.8 | 18.7 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.042638 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9 | 28.3 | 460.7
2008-07-16 | 13:44 | 0 | -0.001 | -10 | 2.3 | 460.7623 | 463.1 | 0.8 | 0.777 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0426 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9069 | 28.3 | 460.8
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.062 | 463.4 | 0.7 | 0.7377 | 18.7 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0426 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3805 | 28.9836 | 28.3 | 461.1
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.038 | 463.4 | 0.8 | 0.7768 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0426 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.6203 | 29 | 28.3 | 461
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 460.9242 | 463.3 | 0.7 | 0.7204 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0426 | 0.0437 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29 | 28.3 | 460.9
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.0238 | 463.4 | 0.8 | 0.7707 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0426 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3812 | 29.0136 | 28.3 | 461
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.3 | 463.7 | 0.7 | 0.7099 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.0426 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0901 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.3 | 463.7 | 0.8 | 0.7837 | 18.7 | 12.9 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0426 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0332 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.3 | 463.7 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.042538 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9565 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.3 | 463.7 | 0.9 | 0.8763 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0425 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9202 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.4 | 463.8 | 0.7 | 0.7 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0425 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9969 | 28.3 | 461.4
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.3 | 463.7 | 0.8 | 0.8 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0425 | 0.0436 | 0.0273 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0735 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 2.4 | 461.2622 | 463.7 | 0.8 | 0.7769 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0425 | 0.0436 | 0.027362 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2003 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | -0.0009 | -8.75 | 2.4 | 461.2242 | 463.6 | 0.8 | 0.7897 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3621 | 0.0425 | 0.0436 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2464 | 28.3 | 461.2
2008-07-16 | 13:45 | 0 | -0.0001 | -0.73718 | 2.3 | 461.2 | 463.5 | 0.8 | 0.7906 | 18.7 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0425 | 0.0436 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3 | 461.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.3 | 461.4 | 463.7 | 0.7 | 0.7208 | 18.7 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0425 | 0.0436 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.3 | 461.4
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.3 | 461.2762 | 463.6 | 0.9 | 0.8646 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0425 | 0.0435 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6188 | 29.1 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.3 | 461.2 | 463.5 | 0.7 | 0.7205 | 18.7 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0425 | 0.0435 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0665 | 28.3 | 461.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.4 | 461.3 | 463.7 | 0.9 | 0.8779 | 18.7 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0425 | 0.0435 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0102 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.3 | 461.338 | 463.7 | 0.8 | 0.753 | 18.7 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0425 | 0.0435 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0868 | 28.3 | 461.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.4 | 461.1764 | 463.6 | 0.9 | 0.9 | 18.7 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4.4 | 195 | 59.4 | 0.042438 | 0.0435 | 0.0274 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.618 | 29.1 | 28.3 | 461.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 2.3 | 457.701 | 460 | 0.8 | 0.771 | 18.59 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4.6 | 195 | 59.4 | 0.0424 | 0.0435 | 0.027153 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0209 | 28.3 | 457.7
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.9 | 429.5125 | 430.4 | 0.8 | 0.8 | 17.58 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.041492 | 0.042716 | 0.025903 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9483 | 28.3 | 429.5
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.3 | 412.4619 | 412.7 | 0.7 | 0.6788 | 16.97 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.040314 | 0.041652 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1792 | 28.3 | 412.5
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.1 | 412.5 | 412.6 | 0.6 | 0.642 | 16.72 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039852 | 0.04119 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.13 | 28.3 | 412.5
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.4 | 412.4 | 412.8 | 0.7 | 0.6787 | 16.7 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039638 | 0.040938 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0065 | 28.3 | 412.4
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.4 | 412.4625 | 412.9 | 0.6 | 0.6228 | 16.7 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0395 | 0.0408 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.375 | 28.9234 | 28.3 | 412.5
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.4 | 412.3752 | 412.8 | 0.8 | 0.766 | 16.7 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039438 | 0.040738 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6242 | 28.9999 | 28.3 | 412.4
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.5 | 412.3 | 412.8 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039338 | 0.040638 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.377 | 28.9232 | 28.3 | 412.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.5 | 412.2 | 412.7 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0393 | 0.040538 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9 | 28.3 | 412.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.7 | 411.9759 | 412.6 | 0.6 | 0.6 | 16.6 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0392 | 0.040438 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9602 | 28.3 | 412
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.8 | 411.7762 | 412.5 | 0.7 | 0.7 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0391 | 0.040338 | 0.0245 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3 | 411.8
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.8 | 410.9 | 411.7 | 0.6 | 0.6 | 16.6 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039 | 0.0402 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3 | 410.9
2008-07-16 | 13:45 | 0 | 0 | 0 | 0.9 | 410.9622 | 411.9 | 0.7 | 0.7 | 16.6 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.039 | 0.040138 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.3 | 411
2008-07-16 | 13:45 | 0 | 0 | 0 | 1 | 411.1621 | 412.1 | 0.6 | 0.5783 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0389 | 0.0401 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 29.0632 | 28.3 | 411.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 1 | 411.2 | 412.2 | 0.5 | 0.5387 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0389 | 0.040038 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0134 | 28.3 | 411.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.1 | 411.2619 | 412.4 | 0.7 | 0.6777 | 16.6 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0388 | 0.039938 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0901 | 28.3 | 411.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.1 | 411.4 | 412.5 | 0.6 | 0.6217 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3624 | 0.0388 | 0.0399 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.3 | 411.4
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.2 | 411.4377 | 412.6 | 0.7 | 0.745 | 16.6 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0387 | 0.0399 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1434 | 28.3 | 411.4
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.2 | 411.3 | 412.5 | 0.5 | 0.52 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0387 | 0.0398 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1799 | 28.3 | 411.3
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.2 | 411.1 | 412.3 | 0.7 | 0.6789 | 16.6 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0387 | 0.0398 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1033 | 28.3 | 411.1
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.2 | 411.2 | 412.4 | 0.5 | 0.5 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3624 | 0.0386 | 0.039738 | 0.0244 | 0.000262 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.3 | 411.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.3 | 411.1624 | 412.4 | 0.7 | 0.7 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3376 | 0.0386 | 0.0397 | 0.0244 | 0.0003 |  | n/a |  | 32.1069 | 55 | 1049 | 1011 | 29.1 | 28.3 | 411.2
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.3 | 411.0622 | 412.4 | 0.5 | 0.5 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.038538 | 0.039638 | 0.0244 | 0.0003 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.1 | 28.2732 | 411.1
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.3 | 410.8379 | 412.1 | 0.6 | 0.6229 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3621 | 0.0385 | 0.0396 | 0.0244 | 0.0003 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.1069 | 28.2035 | 410.8
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.3 | 410.8 | 412.1 | 0.6 | 0.6391 | 16.6 | 12.9 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0385 | 0.0396 | 0.0244 | 0.0003 |  | n/a |  | 33 | 55 | 1049 | 1011 | 29.2603 | 28.2801 | 410.8
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.3 | 410.6763 | 412 | 0.6 | 0.6299 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0385 | 0.0396 | 0.0244 | 0.0003 |  | n/a |  | 33 | 55 | 1049 | 1011.6183 | 29.1864 | 28.3 | 410.7
2008-07-16 | 13:45 | 0 | 0 | 0 | 1.4 | 410.5619 | 412 | 0.7 | 0.6782 | 16.6 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3381 | 0.0385 | 0.039538 | 0.0244 | 0.0003 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0331 | 28.3 | 410.6
2008-07-16 | 13:45 | 0 | -0.0003 | -2.7595 | 1.4 | 410.4 | 411.8 | 0.6 | 0.6226 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0384 | 0.0395 | 0.0244 | 0.0002 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.9202 | 28.3 | 410.4
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 410.4 | 411.8 | 0.8 | 0.7679 | 16.6 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0384 | 0.0395 | 0.0244 | 0.0002 |  | n/a |  | 33 | 55 | 1049 | 1012 | 29.0736 | 28.3 | 410.4
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.5 | 410.4622 | 411.9 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.038338 | 0.0394 | 0.0244 | 0.000262 |  | n/a |  | 32.893 | 55 | 1049 | 1012 | 29.1633 | 28.3 | 410.5
2008-07-16 | 13:45 | 0 | -0.0002 | -2.4374 | 1.5 | 410.3762 | 411.9 | 0.7 | 0.7 | 16.6 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0394 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1207 | 28.2604 | 410.4
2008-07-16 | 13:45 | 0 | -0.0006 | -5.8853 | 1.5 | 410.2377 | 411.7 | 0.7 | 0.6572 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0394 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9838 | 28.2162 | 410.2
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.5 | 410.1378 | 411.6 | 0.7 | 0.7 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0394 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9068 | 28.2932 | 410.1
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.9379 | 411.3 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0394 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.04 | 28.3 | 409.9
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.8 | 411.2 | 0.7 | 0.721 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.039338 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9597 | 28.3 | 409.8
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.6 | 411 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8706 | 28.3 | 409.6
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.7619 | 411.2 | 0.6 | 0.6337 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0383 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.0995 | 28.2998 | 409.8
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.8 | 411.2 | 0.7 | 0.6757 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0382 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9463 | 28.2231 | 409.8
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 409.963 | 411.4 | 0.6 | 0.6199 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0382 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9535 | 28.2 | 410
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 410.062 | 411.5 | 0.7 | 0.6679 | 16.6 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0382 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0302 | 28.2 | 410.1
2008-07-16 | 13:45 | 0 | -0.001 | -10 | 1.4 | 410.262 | 411.7 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0382 | 0.0393 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.2 | 410.3
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.4 | 410.2381 | 411.6 | 0.7 | 0.6848 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.4 | 0.038138 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.2 | 410.2
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.4 | 410.1 | 411.5 | 0.5 | 0.5 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0381 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.2 | 410.1
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.4 | 410.0378 | 411.4 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0381 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1011.6218 | 29.1738 | 28.2369 | 410
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 410 | 411.3 | 0.6 | 0.5697 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0381 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2864 | 28.2864 | 410
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.938 | 411.2 | 0.6 | 0.6232 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.338 | 0.0381 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2097 | 28.2097 | 409.9
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.776 | 411 | 0.6 | 0.6 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0381 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.133 | 28.2 | 409.8
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.7 | 411 | 0.5 | 0.5357 | 16.6 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038038 | 0.0392 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0126 | 28.2437 | 409.7
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.6 | 410.9 | 0.6 | 0.5779 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.039138 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9406 | 28.2797 | 409.6
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.5 | 410.8 | 0.5 | 0.5205 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.094 | 28.203 | 409.5
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.3 | 409.3 | 410.6 | 0.6 | 0.6401 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.0244 | 0.0003 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.1 | 28.2 | 409.3
2008-07-16 | 13:46 | 0 | -0.0007 | -7.4035 | 1.3 | 409.2381 | 410.5 | 0.5 | 0.5 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.0244 | 0.000238 |  | n/a |  | 32 | 55 | 1049 | 1013.6192 | 29.1 | 28.2 | 409.2
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.3 | 409.2 | 410.5 | 0.7 | 0.6773 | 16.6 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1014 | 29.1 | 28.2 | 409.2
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.3 | 409.3 | 410.6 | 0.6 | 0.5554 | 16.6 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.0962 | 28.2 | 409.3
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.3 | 409.3 | 410.6 | 0.6 | 0.6234 | 16.6 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.038 | 0.0391 | 0.0244 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.0196 | 28.2 | 409.3
2008-07-16 | 13:46 | 0 | -0.0006 | -5.9383 | 1.3 | 408.2455 | 409.5 | 0.5 | 0.5 | 16.6 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.038 | 0.0391 | 0.024338 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.0571 | 28.2 | 408.2
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 0.6 | 385.12 | 385.7 | 0.4 | 0.4225 | 16.29 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.037838 | 0.039 | 0.022928 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 29.0325 | 28.2 | 385.1
2008-07-16 | 13:46 | 0 | -0.0007 | -7.3906 | 0.5 | 367.5273 | 368.1 | 0.4 | 0.3887 | 15.27 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.036274 | 0.037386 | 0.021887 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9313 | 28.2 | 367.5
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.2 | 362.1758 | 363.4 | 0.3 | 0.3368 | 14.77 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.035475 | 0.036438 | 0.021538 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.6242 | 29.1615 | 28.2 | 362.2
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.2 | 361.8622 | 363 | 0.5 | 0.4759 | 14.6 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.035151 | 0.036089 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3778 | 29.1366 | 28.2 | 361.9
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.4 | 362 | 363.4 | 0.4 | 0.4 | 14.6 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.035 | 0.0359 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.2 | 362
2008-07-16 | 13:46 | 0 | -0.0006 | -5.8512 | 1.5 | 361.838 | 363.3 | 0.5 | 0.4688 | 14.5 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0349 | 0.0358 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.2 | 361.8
2008-07-16 | 13:46 | 0 | -0.001 | -10 | 1.4 | 361.6382 | 363.1 | 0.3 | 0.3 | 14.5 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.4 | 0.0348 | 0.0357 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.2 | 361.6
2008-07-16 | 13:46 | 0 | -0.0007 | -7.2419 | 1.5 | 361.6 | 363.1 | 0.5 | 0.4786 | 14.5 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034738 | 0.035638 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.2 | 361.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.6 | 361.5 | 363.1 | 0.4 | 0.4 | 14.5 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034638 | 0.035538 | 0.0215 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9587 | 28.2 | 361.5
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.7 | 361.2137 | 362.9 | 0.5 | 0.5 | 14.5 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034538 | 0.035438 | 0.021438 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.621 | 28.8238 | 28.2 | 361.2
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 361 | 362.8 | 0.4 | 0.4 | 14.5 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0345 | 0.0354 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8994 | 28.2 | 361
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 361 | 362.8 | 0.3 | 0.3426 | 14.5 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034438 | 0.0353 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3805 | 28.8227 | 28.2 | 361
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.599 | 362.4 | 0.4 | 0.3777 | 14.5 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0344 | 0.0352 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.854 | 28.2 | 360.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.1376 | 361.9 | 0.3 | 0.3 | 14.5 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034338 | 0.0352 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9614 | 28.2 | 360.1
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.1623 | 362 | 0.4 | 0.3785 | 14.5 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.035138 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.2 | 360.2
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.3621 | 362.2 | 0.3 | 0.3207 | 14.5 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.034238 | 0.0351 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.379 | 29.1 | 28.2 | 360.4
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.3 | 362.1 | 0.5 | 0.4675 | 14.5 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0342 | 0.0351 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2216 | 28.2 | 360.3
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.7 | 360.0762 | 361.8 | 0.4 | 0.4119 | 14.5 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3619 | 0.0342 | 0.035038 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.225 | 28.2 | 360.1
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360.0379 | 361.8 | 0.5 | 0.5 | 14.5 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034138 | 0.035 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0858 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360 | 361.8 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0341 | 0.035 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0091 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 359.938 | 361.7 | 0.5 | 0.4885 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0341 | 0.0349 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.2 | 359.9
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 359.962 | 361.8 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.338 | 0.034038 | 0.0349 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9558 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360 | 361.8 | 0.4 | 0.4223 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.034 | 0.0349 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9209 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.7 | 360.0238 | 361.8 | 0.5 | 0.4778 | 14.5 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3381 | 0.034 | 0.0349 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.3808 | 28.9976 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360 | 361.8 | 0.3 | 0.3 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3622 | 0.034 | 0.034838 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0743 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 360 | 361.8 | 0.5 | 0.4559 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.033938 | 0.0348 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 359.938 | 361.7 | 0.3 | 0.3226 | 14.5 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.4 | 0.0339 | 0.0348 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0446 | 28.2 | 359.9
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.8 | 359.9 | 361.7 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0339 | 0.0347 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9044 | 28.2 | 359.9
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.7 | 359.9618 | 361.7 | 0.3 | 0.3 | 14.5 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0339 | 0.0347 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9811 | 28.2 | 360
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.7 | 359.7378 | 361.4 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.0339 | 0.0347 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6218 | 28.8845 | 28.2 | 359.7
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.6 | 359.7 | 361.3 | 0.3 | 0.3 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0338 | 0.0347 | 0.0213 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.8689 | 28.2 | 359.7
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.6 | 359.638 | 361.2 | 0.3 | 0.3192 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0338 | 0.0347 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9888 | 28.1888 | 359.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.6 | 359.4 | 361 | 0.3 | 0.3 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0338 | 0.034638 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9122 | 28.1122 | 359.4
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.6 | 359.4 | 361 | 0.2 | 0.2091 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.033738 | 0.0346 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9 | 28.1645 | 359.4
2008-07-16 | 13:46 | 0.004 | 0 | 0 | 1.7 | 359.4 | 361.1 | 0.3 | 0.3 | 14.5 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0337 | 0.0346 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9412 | 28.1588 | 359.4
2008-07-16 | 13:46 | 0.01 | 0 | 0 | 1.7 | 359.5 | 361.2 | 0.2 | 0.2141 | 14.5 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0337 | 0.0346 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.1 | 359.5
2008-07-16 | 13:46 | 0.01 | 0 | 0 | 1.7 | 359.4 | 361.1 | 0.4 | 0.378 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0337 | 0.034538 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.1 | 359.4
2008-07-16 | 13:46 | 0.01 | 0 | 0 | 1.6 | 359.5238 | 361.2 | 0.3 | 0.3144 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0337 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9293 | 28.1 | 359.5
2008-07-16 | 13:46 | 0.01 | 0 | 0 | 1.6 | 359.6 | 361.2 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3383 | 0.0337 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9475 | 28.1 | 359.6
2008-07-16 | 13:46 | 0.004 | 0 | 0 | 1.6 | 359.6 | 361.2 | 0.3 | 0.3 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0337 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.1 | 359.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1.3 | 359.6 | 360.9 | 0.5 | 0.5 | 14.5 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0336 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.1 | 359.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1 | 359.6 | 360.6 | 0.4 | 0.3532 | 14.5 | 13 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.338 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29 | 28.1 | 359.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 1 | 359.662 | 360.7 | 0.4 | 0.4225 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.3805 | 28.9454 | 28.1 | 359.7
2008-07-16 | 13:46 | 0 | 0 | 0 | 1 | 359.6 | 360.6 | 0.4 | 0.3722 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9314 | 28.1 | 359.6
2008-07-16 | 13:46 | 0 | 0 | 0 | 0.9 | 359.6624 | 360.6 | 0.3 | 0.3443 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.3758 | 29.0081 | 28.1081 | 359.7
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.6 | 360.4 | 0.5 | 0.4784 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0848 | 28.1848 | 359.6
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.6 | 360.4 | 0.3 | 0.3294 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0335 | 0.0345 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1 | 28.1385 | 359.6
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.538 | 360.3 | 0.6 | 0.5798 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.338 | 0.0334 | 0.0345 | 0.021338 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6205 | 29.0618 | 28.1 | 359.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.4 | 360.2 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3376 | 0.0334 | 0.0345 | 0.021362 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0149 | 28.1 | 359.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.4 | 360.2 | 0.5 | 0.5 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0334 | 0.034438 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0916 | 28.1 | 359.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.4379 | 360.2 | 0.4 | 0.3781 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0334 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.1 | 359.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.3 | 360.1 | 0.4 | 0.4241 | 14.5 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0334 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.1 | 359.3
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.4 | 360.2 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0334 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0782 | 28.1 | 359.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.5 | 360.3 | 0.4 | 0.4411 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0334 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0015 | 28.1 | 359.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.5376 | 360.3 | 0.4 | 0.3732 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0334 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6242 | 29 | 28.1 | 359.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.5 | 360.3 | 0.3 | 0.3442 | 14.5 | 13 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.033338 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.8963 | 28.1 | 359.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.5 | 360.3 | 0.5 | 0.4779 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0333 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.8571 | 28.1 | 359.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.438 | 360.2 | 0.3 | 0.3 | 14.5 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.362 | 0.0333 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9948 | 28.1 | 359.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.2 | 360 | 0.5 | 0.4765 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0333 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9181 | 28.1 | 359.2
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 359.3 | 360.1 | 0.3 | 0.3 | 14.5 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0333 | 0.0344 | 0.0214 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9 | 28.1 | 359.3
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.7 | 355.2584 | 356 | 0.4 | 0.4 | 14.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.2378 | 0.0333 | 0.034338 | 0.021151 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9706 | 28.1 | 355.3
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.6 | 327.4823 | 328 | 0.3 | 0.3 | 13.91 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.033014 | 0.033989 | 0.019465 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.621 | 29.0761 | 28.1 | 327.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.5 | 314.4215 | 314.9 | 0.3 | 0.3 | 12.95 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.031514 | 0.032452 | 0.018738 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9227 | 28.1 | 314.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.4 | 313.1424 | 313.6 | 0.3 | 0.2554 | 12.63 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.031052 | 0.03199 | 0.018638 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.8347 | 28.1 | 313.1
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.5 | 312 | 312.5 | 0.2 | 0.2211 | 12.5 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0308 | 0.0318 | 0.0186 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.8421 | 28.1 | 312
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.1 | 311.2376 | 311.3 | 0.3 | 0.2731 | 12.5 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.030638 | 0.0317 | 0.0185 | 0.0002 |  | n/a |  | 32.107 | 55 | 1049 | 1011 | 28.9376 | 28.1 | 311.2
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.5 | 311.0378 | 311 | 0.2 | 0.2221 | 12.48 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3622 | 0.0304 | 0.031538 | 0.0185 | 0.0002 |  | n/a |  | 32.893 | 55 | 1049 | 1011.6222 | 29.0911 | 28.1 | 311
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.5 | 310.8 | 310.8 | 0.2 | 0.2385 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0303 | 0.031438 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1722 | 28.1 | 310.8
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.4 | 310.7 | 310.7 | 0.2 | 0.2 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.030238 | 0.031338 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1511 | 28.1 | 310.7
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.3 | 310.5381 | 310.5 | 0.4 | 0.3822 | 12.4 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.030138 | 0.031238 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28.1 | 310.5
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.3 | 310.6 | 310.6 | 0.3 | 0.2787 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0301 | 0.0312 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0977 | 28.1 | 310.6
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.2 | 310.5 | 310.5 | 0.3 | 0.3 | 12.4 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.03 | 0.0311 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0211 | 28.1 | 310.5
2008-07-16 | 13:47 | 0 | 0 | 0 | -0.1 | 310.338 | 310.3 | 0.2 | 0.2 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.029938 | 0.031038 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0553 | 28.1 | 310.3
2008-07-16 | 13:47 | 0 | 0 | 0 | 0 | 310.3 | 310.3 | 0.3 | 0.3232 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0299 | 0.030938 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0361 | 28.1 | 310.3
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.1 | 310.2382 | 310.3 | 0.3 | 0.3 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3618 | 0.0298 | 0.030838 | 0.0185 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9268 | 28.1 | 310.2
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.1 | 310.2 | 310.3 | 0.3 | 0.3206 | 12.4 | 13 | 952 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.029738 | 0.030738 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.156 | 28.1 | 310.2
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.2 | 310.2 | 310.4 | 0.3 | 0.3 | 12.4 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0297 | 0.0307 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0755 | 28.1 | 310.2
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.3 | 310.1 | 310.4 | 0.2 | 0.2392 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.029638 | 0.0306 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0392 | 28.1 | 310.1
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.3 | 310.038 | 310.3 | 0.4 | 0.3778 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.029538 | 0.0305 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0521 | 28.084 | 310
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.4 | 309.8 | 310.2 | 0.2 | 0.2 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0295 | 0.0304 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8221 | 28.0074 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.5 | 309.8 | 310.3 | 0.4 | 0.3781 | 12.4 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.4 | 0.0294 | 0.030338 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.8 | 28.0693 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.5 | 309.7 | 310.2 | 0.2 | 0.2 | 12.4 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0294 | 0.0303 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8 | 28.1 | 309.7
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.6 | 309.7 | 310.3 | 0.2 | 0.2374 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0294 | 0.0303 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8 | 28.1 | 309.7
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.6 | 309.638 | 310.2 | 0.2 | 0.179 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.4 | 0.0293 | 0.0303 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.8 | 28.1 | 309.6
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.7 | 309.762 | 310.4 | 0.2 | 0.2137 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0293 | 0.0302 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.8761 | 28.1 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.7 | 309.6762 | 310.4 | 0.2 | 0.1827 | 12.4 | 13 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3381 | 0.0293 | 0.0302 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.6188 | 28.9528 | 28.1 | 309.7
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.8 | 309.7248 | 310.5 | 0.2 | 0.2195 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3624 | 0.0292 | 0.0301 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3758 | 29 | 28.1 | 309.7
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.9 | 309.8377 | 310.7 | 0.2 | 0.1746 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3377 | 0.0292 | 0.0301 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0062 | 28.1 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.9 | 309.8 | 310.7 | 0 | 0.0239 | 12.4 | 13 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0292 | 0.03 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.083 | 28.1 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 0.9 | 309.838 | 310.7 | 0.2 | 0.1763 | 12.4 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0292 | 0.03 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011.6205 | 29.1 | 28.1 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 1 | 309.9 | 310.9 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0292 | 0.03 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.0271 | 28.1 | 309.9
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.1 | 309.8 | 310.9 | 0.2 | 0.1791 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0291 | 0.03 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.9525 | 28.1 | 309.8
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.2 | 309.6378 | 310.8 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3622 | 0.0291 | 0.029938 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2592 | 28.1 | 309.6
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.2 | 309.5 | 310.7 | 0.2 | 0.2 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.0291 | 0.0299 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1004 | 28.0335 | 309.5
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.3 | 309.4 | 310.7 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0299 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9567 | 28 | 309.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.3 | 309.4 | 310.7 | 0.2 | 0.2228 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0291 | 0.0299 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.9798 | 28.0199 | 309.4
2008-07-16 | 13:47 | 0 | 0 | 0 | 1.3 | 309.4 | 310.7 | 0.2 | 0.1544 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0291 | 0.0299 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012.3758 | 29.2865 | 28.0966 | 309.4
2008-07-16 | 13:47 | 0 | -0.0003 | -2.6385 | 1.3 | 309.4 | 310.7 | 0.1 | 0.1236 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0291 | 0.029838 | 0.0184 | 0.0002 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0065 | 28.1 | 309.4
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.3 | 309.4 | 310.7 | 0.2 | 0.1718 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0291 | 0.0298 | 0.0184 | 0.000138 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.9501 | 28.1 | 309.4
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.3 | 309.5 | 310.8 | 0.1 | 0.122 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.029038 | 0.0298 | 0.0184 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0803 | 28.0732 | 309.5
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.3 | 309.6 | 310.9 | 0.3 | 0.2762 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0298 | 0.0184 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.2931 | 28.0035 | 309.6
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.4 | 309.7 | 311.1 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0298 | 0.0184 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1395 | 28.0803 | 309.7
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.4 | 309.8621 | 311.3 | 0.3 | 0.2781 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0298 | 0.0184 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.214 | 28.043 | 309.9
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.4 | 309.8 | 311.2 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0297 | 0.0184 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.2326 | 28 | 309.8
2008-07-16 | 13:47 | 0 | -0.001 | -10 | 1.4 | 309.638 | 311 | 0.2 | 0.2324 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1104 | 28.0104 | 309.6
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.4 | 309.5 | 310.9 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.029 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.1871 | 28.0871 | 309.5
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.4 | 309.4 | 310.8 | 0.2 | 0.2206 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.029 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.327 | 28.0365 | 309.4
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.4 | 309.2376 | 310.6 | 0.3 | 0.2561 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1989 | 28.0402 | 309.2
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.4 | 309.1378 | 310.5 | 0.1 | 0.1433 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.6222 | 28.9684 | 28.1 | 309.1
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.5 | 309.162 | 310.6 | 0.2 | 0.2 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0289 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012.3795 | 29.274 | 28.1 | 309.2
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.5 | 309.2619 | 310.8 | 0.1 | 0.1215 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.0289 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0885 | 28.0295 | 309.3
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.5 | 309.2623 | 310.8 | 0.3 | 0.2536 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.0943 | 28 | 309.3
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.5 | 309.3 | 310.8 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1763 | 28 | 309.3
2008-07-16 | 13:48 | 0 | -0.001 | -10 | 1.5 | 309.3 | 310.8 | 0.2 | 0.2 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.1 | 28 | 309.3
2008-07-16 | 13:48 | 0 | -0.0006 | -5.9025 | 1.5 | 309.462 | 311 | 0.1 | 0.1 | 12.4 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.3797 | 29.1 | 28 | 309.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.5 | 309.5 | 311 | 0.2 | 0.23 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.2084 | 28 | 309.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.5 | 309.4381 | 310.9 | 0.2 | 0.1897 | 12.4 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.6188 | 29.3309 | 28 | 309.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.5 | 309.4 | 310.9 | 0.1 | 0.1433 | 12.4 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.4 | 28 | 309.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.5 | 309.4 | 310.9 | 0.2 | 0.1555 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.4 | 28 | 309.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.5 | 309.4 | 310.9 | 0.1 | 0.1121 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.3389 | 28 | 309.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.6 | 309.3 | 310.9 | 0.2 | 0.1715 | 12.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0295 | 0.0183 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2622 | 28 | 309.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.9 | 290.8727 | 291.8 | 0.1 | 0.1 | 12.25 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.028676 | 0.0295 | 0.01722 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2 | 28 | 290.9
2008-07-16 | 13:48 | 0 | 0 | 0 | -0.1 | 266.8126 | 266.8 | 0.2 | 0.1768 | 11.31 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.027605 | 0.028605 | 0.015952 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1012 | 29.2 | 28 | 266.8
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.2 | 262.2026 | 262.4 | 0 | 0 | 10.65 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026638 | 0.0275 | 0.015538 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011 | 29.2679 | 28 | 262.2
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.2 | 260.9 | 261.1 | 0.2 | 0.1788 | 10.5 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026538 | 0.027338 | 0.0155 | 0.0001 |  | n/a |  | 32 | 55 | 1049 | 1011.3777 | 29.3446 | 28 | 260.9
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.3 | 260.038 | 260.3 | 0 | 0 | 10.4 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.026438 | 0.027238 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048.2665 | 1011 | 29.3786 | 28 | 260
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.8 | 259.7761 | 260.5 | 0.1 | 0.1 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.026038 | 0.026838 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3018 | 28 | 259.8
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.9 | 259.7619 | 260.6 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3619 | 0.025938 | 0.026638 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3748 | 28 | 259.8
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 259.6 | 260.6 | 0 | 0.0443 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3625 | 0.025837 | 0.0265 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.2454 | 28 | 259.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 259.5 | 260.6 | 0.2 | 0.1561 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.025738 | 0.0264 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.1566 | 28 | 259.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.2 | 259.4 | 260.6 | 0 | 0.0213 | 10.4 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3622 | 0.0257 | 0.0263 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.305 | 28 | 259.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.2 | 259.4 | 260.6 | 0.1 | 0.0758 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0256 | 0.026238 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3817 | 28 | 259.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 259.5 | 260.6 | 0 | 0.0209 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.025538 | 0.0262 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4 | 28 | 259.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 259.4 | 260.4 | 0.2 | 0.1571 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0254 | 0.0261 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.3296 | 28 | 259.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 259.4 | 260.4 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0254 | 0.026038 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.2238 | 28 | 259.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 259.276 | 260.3 | 0.1 | 0.0884 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.362 | 0.0253 | 0.026 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.3773 | 28 | 259.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 259.1 | 260.2 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0253 | 0.0259 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4 | 28 | 259.1
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 259.0374 | 260.1 | 0.1 | 0.1307 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0252 | 0.0259 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3159 | 28 | 259
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.9 | 260 | 0.1 | 0.0779 | 10.4 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.4 | 0.0252 | 0.025838 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.2376 | 28 | 258.9
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.2 | 258.7377 | 259.9 | 0 | 0.0315 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3623 | 0.025138 | 0.0258 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.3911 | 28 | 258.7
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.6622 | 259.8 | 0.1 | 0.0778 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0251 | 0.025738 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.2554 | 28 | 258.7
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.7 | 259.8 | 0 | 0.0198 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0251 | 0.0257 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.2981 | 28 | 258.7
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.6 | 259.7 | 0.1 | 0.1 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3381 | 0.0251 | 0.0257 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.5378 | 259.6 | 0 | 0.0219 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3622 | 0.025038 | 0.0257 | 0.0154 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4073 | 28 | 258.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 1.1 | 258.4 | 259.5 | 0.2 | 0.1781 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.025 | 0.0257 | 0.015338 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013.6218 | 29.6367 | 28 | 258.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 258.4 | 259.4 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3379 | 0.025 | 0.0256 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.5894 | 28 | 258.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 258.338 | 259.3 | 0.1 | 0.1 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.024938 | 0.0256 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.468 | 28 | 258.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 1 | 258.3 | 259.3 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.0249 | 0.0256 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4 | 28 | 258.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.7 | 258.2 | 258.9 | 0.1 | 0.0683 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0248 | 0.0256 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4 | 28 | 258.2
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.7 | 258.2 | 258.9 | 0.1 | 0.1 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0248 | 0.025538 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4626 | 28 | 258.2
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.2623 | 258.9 | 0.2 | 0.1603 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0248 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4607 | 28 | 258.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.3622 | 259 | 0.2 | 0.1565 | 10.4 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0248 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3837 | 28 | 258.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.5 | 258.462 | 259 | 0 | 0.0237 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0248 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3795 | 29.307 | 28 | 258.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.5 | 258.5617 | 259.1 | 0.2 | 0.1709 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0247 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.3698 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.7 | 259.3 | 0 | 0.0177 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0247 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4 | 28 | 258.7
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.7622 | 259.4 | 0.2 | 0.1766 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0247 | 0.0255 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.8
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.738 | 259.3 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0247 | 0.025438 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.7
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.6 | 259.2 | 0.2 | 0.2 | 10.4 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0247 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4767 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.6 | 259.2 | 0.1 | 0.0777 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0247 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4466 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.6 | 259.2 | 0.1 | 0.0634 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.024638 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.5623 | 259.2 | 0.1 | 0.0793 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.6
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.5 | 259.1 | 0 | 0.0211 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.5
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.438 | 259 | 0.1 | 0.0549 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0254 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.4 | 28 | 258.4
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.3 | 258.9 | 0 | 0.0107 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4 | 28 | 258.3
2008-07-16 | 13:48 | 0 | 0 | 0 | 0.6 | 258.3 | 258.9 | 0.1 | 0.0765 | 10.4 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3377 | 0.0246 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4139 | 28 | 258.3
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 258.1378 | 258.8 | 0 | 0 | 10.4 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4906 | 28 | 258.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 258.1 | 258.8 | 0.1 | 0.0614 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4327 | 28 | 258.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 258 | 258.7 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0246 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4 | 28 | 258
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 258.062 | 258.8 | 0.1 | 0.1213 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.024538 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3805 | 29.4 | 28 | 258.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 257.9376 | 258.6 | 0.1 | 0.0778 | 10.4 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.0253 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4 | 28 | 257.9
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 257.9 | 258.6 | 0.1 | 0.0589 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.025238 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4742 | 28 | 257.9
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 257.9 | 258.6 | 0.1 | 0.0771 | 10.4 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.449 | 28 | 257.9
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 257.9 | 258.6 | 0 | 0.0237 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4 | 28 | 257.9
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.9 | 258.7 | 0.1 | 0.1099 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4089 | 28 | 257.9
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.7 | 258.5 | 0 | 0.0238 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.5624 | 28 | 257.7
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.7 | 258.5 | 0.2 | 0.1727 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.3682 | 28 | 257.7
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.6379 | 258.4 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.2694 | 28 | 257.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.6 | 258.4 | 0.2 | 0.1834 | 10.4 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.0245 | 0.0252 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4229 | 28 | 257.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.6 | 258.4 | 0 | 0 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0244 | 0.025138 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.5764 | 28 | 257.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.6 | 258.4 | 0.1 | 0.1188 | 10.4 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.4 | 0.0244 | 0.0251 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4701 | 28 | 257.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 257.6 | 258.4 | 0.1 | 0.1 | 10.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3624 | 0.0244 | 0.0251 | 0.0153 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.4 | 28 | 257.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 243.0456 | 243.6 | 0 | 0.0327 | 10.19 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3623 | 0.0244 | 0.0251 | 0.014426 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3816 | 28 | 243
2008-07-16 | 13:49 | 0 | 0 | 0 | -1.1 | 213.2966 | 213.3 | 0 | 0 | 9.25 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3378 | 0.023103 | 0.024103 | 0.012878 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.3049 | 28 | 213.3
2008-07-16 | 13:49 | 0 | 0 | 0 | -0.5 | 210.2657 | 210.3 | 0 | 0.0176 | 8.63 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.022314 | 0.023052 | 0.0125 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5155 | 28 | 210.3
2008-07-16 | 13:49 | 0 | 0 | 0 | 0 | 209.1381 | 209.1 | 0 | 0 | 8.48 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.021838 | 0.022538 | 0.0124 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6192 | 29.6486 | 28 | 209.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.1 | 208.5393 | 208.6 | 0 | 0 | 8.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.021738 | 0.022438 | 0.0124 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5992 | 28 | 208.5
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.3 | 208.1759 | 208.4 | 0 | 0 | 8.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.021638 | 0.022238 | 0.0124 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6206 | 29.3065 | 28 | 208.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.4 | 208.1 | 208.5 | 0 | 0 | 8.4 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0215 | 0.022138 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6146 | 28 | 208.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.4 | 207.5028 | 207.9 | 0 | 0 | 8.38 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0214 | 0.022038 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5343 | 28 | 207.5
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 207.2 | 207.7 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0213 | 0.0219 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4319 | 28 | 207.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 207.2 | 207.7 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.021238 | 0.021838 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5177 | 28 | 207.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 207.1376 | 207.6 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0212 | 0.0218 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6706 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 207.1377 | 207.6 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3623 | 0.0212 | 0.0218 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5755 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 207.1 | 207.7 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0211 | 0.0217 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5784 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 207.1 | 207.7 | 0 | 0 | 8.3 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.4 | 0.021 | 0.0216 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6519 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 207.0381 | 207.6 | 0 | 0 | 8.3 | 12.9 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.4 | 0.021 | 0.021538 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4219 | 28 | 207
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 207.0619 | 207.8 | 0 | 0 | 8.3 | 12.9 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.020938 | 0.0215 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5389 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 207.1 | 207.9 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0209 | 0.0215 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5076 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 207.1 | 207.9 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0208 | 0.0214 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.4459 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 207.1 | 207.9 | 0 | 0 | 8.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0208 | 0.0213 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5994 | 28 | 207.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.8 | 207 | 207.8 | 0 | 0 | 8.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.020737 | 0.021237 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5235 | 28 | 207
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 207 | 207.6 | 0 | 0 | 8.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.4 | 0.0206 | 0.0212 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5533 | 28 | 207
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.8377 | 207.4 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3377 | 0.020538 | 0.021138 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.54 | 28 | 206.8
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.8 | 207.4 | 0 | 0 | 8.3 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3621 | 0.0205 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.42 | 28 | 206.8
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.8 | 207.3 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.020438 | 0.021038 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.6503 | 28 | 206.8
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.7 | 207.2 | 0 | 0 | 8.3 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0204 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.6398 | 28 | 206.7
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.7 | 207.2 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0204 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.5262 | 28 | 206.7
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.7 | 207.2 | 0 | 0 | 8.3 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3381 | 0.0203 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.441 | 28 | 206.7
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.5758 | 207.1 | 0.1 | 0.1 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.4 | 0.0203 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.6712 | 28 | 206.6
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.5 | 207 | 0 | 0 | 8.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.338 | 0.0203 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7672 | 28 | 206.5
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.5 | 207 | 0.1 | 0.1 | 8.3 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3 | 0.0202 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7561 | 28 | 206.5
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.4 | 206.9 | 0 | 0 | 8.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0202 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6794 | 28 | 206.4
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.5 | 206.3376 | 206.8 | 0 | 0.022 | 8.3 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3624 | 0.0202 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6026 | 28 | 206.3
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.2378 | 206.8 | 0 | 0 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0202 | 0.020738 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.4516 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.2 | 206.8 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.020138 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.5528 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.2 | 206.8 | 0 | 0 | 8.3 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0201 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.2 | 206.8 | 0 | 0 | 8.3 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0201 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.1378 | 206.7 | 0.1 | 0.0558 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0201 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 28 | 206.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.6 | 206.1 | 206.7 | 0 | 0.0116 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.338 | 0.0201 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 28 | 206.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.2 | 206.9 | 0.1 | 0.0738 | 8.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 4.4 | 195 | 59.4 | 0.0201 | 0.020638 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3797 | 29.7346 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.2 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 4 | 195 | 59.338 | 0.0201 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7886 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.2 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3619 | 0.0201 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7119 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.2 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.02 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 28 | 206.2
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.1377 | 206.8 | 0 | 0 | 8.3 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.02 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.623 | 29.7 | 28 | 206.1
2008-07-16 | 13:49 | 0 | 0 | 0 | 0.7 | 206.1 | 206.8 | 0 | 0 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.02 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6816 | 28 | 206.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 205.938 | 206.6 | 0 | 0.0349 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.02 | 0.0206 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.6205 | 29.6048 | 28 | 205.9
2008-07-16 | 13:50 | 0 | -0.0007 | -7.3064 | 0.7 | 206 | 206.7 | 0.1 | 0.08 | 8.3 | 13 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.02 | 0.0205 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6719 | 28 | 206
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.7 | 205.9383 | 206.6 | 0 | 0.0233 | 8.3 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3617 | 0.02 | 0.0205 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.6172 | 29.6513 | 28 | 205.9
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.7 | 206.0622 | 206.8 | 0.1 | 0.0537 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.019938 | 0.0205 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6505 | 28 | 206.1
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.7 | 206.1622 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0199 | 0.0205 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3785 | 29.7985 | 28 | 206.2
2008-07-16 | 13:50 | 0 | -0.0009 | -9.2442 | 0.8 | 206.2 | 207 | 0.1 | 0.1 | 8.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0199 | 0.0205 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7214 | 28 | 206.2
2008-07-16 | 13:50 | 0 | -0.0001 | -0.93854 | 0.8 | 206.1381 | 206.9 | 0 | 0 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.0199 | 0.0204 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8103 | 28 | 206.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 206.1 | 206.9 | 0 | 0.0222 | 8.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0199 | 0.0204 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3758 | 29.8364 | 28 | 206.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 206.1622 | 207 | 0 | 0.011 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.0199 | 0.0204 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7088 | 28 | 206.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 206.1 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.019838 | 0.0204 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7858 | 28 | 206.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 206.062 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0198 | 0.0204 | 0.0123 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3795 | 29.8 | 28 | 206.1
2008-07-16 | 13:50 | 0 | 0 | -0.36332 | 0.8 | 206 | 206.8 | 0 | 0.0433 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3618 | 0.0198 | 0.0204 | 0.0122 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7608 | 27.9608 | 206
2008-07-16 | 13:50 | 0 | -0.0009 | -9.0138 | 0.8 | 206 | 206.8 | 0 | 0 | 8.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.0198 | 0.0203 | 0.0122 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7321 | 27.9 | 206
2008-07-16 | 13:50 | 0 | -0.0003 | -2.7496 | 0.9 | 206 | 206.9 | 0 | 0 | 8.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.0198 | 0.0203 | 0.0122 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6218 | 29.8857 | 27.9 | 206
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.9 | 206 | 206.9 | 0.1 | 0.0539 | 8.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0198 | 0.0203 | 0.0122 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7608 | 27.9696 | 206
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.9 | 205.9378 | 206.8 | 0 | 0 | 8.3 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0198 | 0.0203 | 0.0122 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6536 | 28 | 205.9
2008-07-16 | 13:50 | 0 | 0 | 0 | -0.3 | 185.3715 | 185.4 | 0 | 0 | 8 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.019428 | 0.020052 | 0.011019 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6231 | 28 | 185.4
2008-07-16 | 13:50 | 0 | 0 | 0 | -0.6 | 163.9593 | 164 | 0 | 0 | 7.07 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.018329 | 0.019005 | 0.009752 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6188 | 29.6998 | 28 | 164
2008-07-16 | 13:50 | 0 | 0 | 0 | -0.6 | 158.0012 | 158 | 0 | 0 | 6.53 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3624 | 0.017875 | 0.018538 | 0.009438 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7766 | 28 | 158
2008-07-16 | 13:50 | 0 | 0 | 0 | -0.3 | 156.9016 | 156.9 | 0 | 0 | 6.4 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.0175 | 0.0181 | 0.009338 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 156.9
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.1 | 156.5 | 156.6 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3621 | 0.017138 | 0.017638 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.77 | 28 | 156.5
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.2 | 156.4626 | 156.6 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3374 | 0.017037 | 0.017537 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3735 | 29.7068 | 28 | 156.5
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.2 | 156.4381 | 156.6 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.017 | 0.0175 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6192 | 29.7836 | 28 | 156.4
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.2 | 156.3618 | 156.6 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.016838 | 0.017338 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 156.4
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.3 | 156.3 | 156.6 | 0 | 0 | 6.3 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0167 | 0.0172 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8371 | 28 | 156.3
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.3 | 156.3 | 156.6 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3379 | 0.0166 | 0.0171 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 156.3
2008-07-16 | 13:50 | 0 | -0.0005 | -5.4353 | 0.3 | 156.2378 | 156.5 | -0.1 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3622 | 0.0166 | 0.017038 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 156.2
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.4 | 156.2 | 156.6 | 0 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0165 | 0.016938 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.6979 | 28 | 156.2
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.4 | 156.2 | 156.6 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3618 | 0.0164 | 0.016838 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.6441 | 28 | 156.2
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.4 | 156 | 156.4 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.016338 | 0.0168 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 28 | 156
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.3 | 155.8754 | 156.2 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.0163 | 0.016738 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.623 | 29.7 | 28 | 155.9
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.3 | 155.8 | 156.1 | 0 | 0 | 6.3 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.016238 | 0.0167 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7744 | 28 | 155.8
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.3 | 155.7 | 156 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0162 | 0.0167 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 155.7
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.3 | 155.6 | 155.9 | -0.1 | 0 | 6.3 | 13 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0162 | 0.016638 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 155.6
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.4 | 155.5378 | 155.9 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3622 | 0.0161 | 0.016538 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 27.9954 | 155.5
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.5 | 155.5621 | 156 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.016 | 0.0165 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 27.9187 | 155.6
2008-07-16 | 13:50 | 0 | -0.001 | -10 | 0.5 | 155.562 | 156.1 | 0 | 0 | 6.3 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.016 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.742 | 27.958 | 155.6
2008-07-16 | 13:50 | 0 | -0.0003 | -2.732 | 0.6 | 155.5 | 156.1 | -0.1 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.015938 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8044 | 27.9652 | 155.5
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.6 | 155.4374 | 156 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3374 | 0.0159 | 0.016337 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9652 | 27.9116 | 155.4
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.6 | 155.4 | 156 | 0 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3376 | 0.0159 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.735 | 27.9883 | 155.4
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.3 | 156 | 0 | 0 | 6.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0159 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7651 | 27.9349 | 155.3
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.2378 | 155.9 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3378 | 0.0158 | 0.0162 | 0.009238 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7163 | 27.9 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0158 | 0.016138 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6371 | 27.9 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7906 | 27.9 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.4 | 0.015738 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 27.9721 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3378 | 0.0157 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.3 | 156.1 | -0.1 | 0 | 6.3 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0157 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 155.3
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.3 | 156 | 0 | 0 | 6.3 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0157 | 0.016038 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7954 | 28 | 155.3
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.2 | 155.9 | -0.1 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3381 | 0.015638 | 0.016 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.6426 | 28 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.1376 | 155.8 | 0 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3376 | 0.0156 | 0.016 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7666 | 28 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.1 | 155.8 | -0.1 | 0 | 6.3 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0156 | 0.016 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9323 | 27.9677 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155.1 | 155.8 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3378 | 0.0156 | 0.016 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3783 | 29.9824 | 27.9 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.1 | 155.9 | -0.1 | 0 | 6.3 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0156 | 0.016 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8285 | 27.9 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.1 | 155.9 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0156 | 0.015938 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7375 | 27.9 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | -0.1 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3377 | 0.0156 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.2 | 156 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3378 | 0.0156 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.7 | 27.9 | 155.2
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.1 | 155.9 | 0 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3379 | 0.0155 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.8 | 155.1 | 155.9 | -0.1 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0155 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7698 | 27.9 | 155.1
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 154.962 | 155.7 | 0 | 0 | 6.3 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.4 | 0.0155 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3805 | 29.8465 | 27.9465 | 155
2008-07-16 | 13:50 | 0 | 0 | 0 | 0.7 | 155 | 155.7 | -0.2 | 0 | 6.3 | 12.9 | 953 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3 | 0.0155 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8533 | 27.9767 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 155 | 155.7 | -0.1 | 0 | 6.3 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3376 | 0.0155 | 0.0159 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 155.0623 | 155.8 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0155 | 0.015838 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 155.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 155.1 | 155.8 | -0.1 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.015438 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8607 | 27.9 | 155.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155.2 | 155.8 | -0.1 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3381 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.9091 | 27.9303 | 155.2
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155.1 | 155.7 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7071 | 27.9929 | 155.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155 | 155.6 | -0.1 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7838 | 27.9162 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 154.9 | 155.5 | -0.2 | 0 | 6.3 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3379 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8606 | 27.9 | 154.9
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 154.9 | 155.5 | -0.1 | 0 | 6.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 27.9 | 154.9
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 154.9 | 155.5 | -0.2 | 0 | 6.3 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0154 | 0.0158 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 27.9 | 154.9
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 154.9619 | 155.6 | -0.1 | 0 | 6.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3381 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3812 | 29.9 | 27.9 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155 | 155.6 | -0.2 | 0 | 6.3 | 12.9 | 953 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.4 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8324 | 27.9676 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155 | 155.6 | -0.1 | 0 | 6.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3623 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8444 | 28 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155.1 | 155.7 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3622 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8578 | 27.9789 | 155.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155 | 155.6 | -0.2 | 0 | 6.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.338 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7043 | 27.9021 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 155 | 155.6 | -0.2 | 0 | 6.3 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0153 | 0.0157 | 0.0092 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 155
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.4 | 145.9713 | 146.4 | -0.2 | 0 | 6.29 | 12.9 | 953 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3376 | 0.0153 | 0.0157 | 0.00865 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8543 | 27.9514 | 146
2008-07-16 | 13:51 | 0 | 0 | 0 | -1.6 | 116.4976 | 116.5 | -0.2 | 0 | 5.42 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3379 | 0.014628 | 0.015352 | 0.006952 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30 | 27.9719 | 116.5
2008-07-16 | 13:51 | 0 | 0 | 0 | -1.1 | 108.1022 | 108.1 | -0.3 | 0 | 4.69 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.013364 | 0.013964 | 0.006438 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9855 | 27.9 | 108.1
2008-07-16 | 13:51 | 0 | 0 | 0 | -1.1 | 106.4417 | 106.4 | -0.2 | 0 | 4.4 | 12.9 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.012938 | 0.013538 | 0.006338 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7553 | 27.9 | 106.4
2008-07-16 | 13:51 | 0 | 0 | 0 | -1.1 | 105.138 | 105.1 | -0.4 | 0 | 4.3 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0128 | 0.0134 | 0.0063 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7583 | 27.9 | 105.1
2008-07-16 | 13:51 | 0 | 0 | 0 | -1.1 | 105.1 | 105.1 | -0.2 | 0 | 4.3 | 12.9 | 951 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0127 | 0.0133 | 0.0063 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7649 | 27.9351 | 105.1
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.9 | 105 | 105 | -0.3 | 0 | 4.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3376 | 0.0126 | 0.013138 | 0.0063 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7236 | 27.9882 | 105
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.8 | 104.9 | 104.9 | -0.3 | 0 | 4.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3377 | 0.012438 | 0.012938 | 0.0063 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8771 | 27.9114 | 104.9
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.6 | 104.8378 | 104.8 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3378 | 0.012338 | 0.012776 | 0.006238 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7694 | 27.9 | 104.8
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.5 | 104.862 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.338 | 0.0122 | 0.012638 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8262 | 27.9 | 104.9
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.4 | 104.8 | 104.8 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0121 | 0.012538 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9436 | 27.9 | 104.8
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.4 | 104.8 | 104.8 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0121 | 0.0125 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7134 | 27.9 | 104.8
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.3 | 104.7 | 104.7 | -0.2 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.012 | 0.0124 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8447 | 27.9 | 104.7
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.2 | 104.6379 | 104.6 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4 | 195 | 59.3379 | 0.0119 | 0.0123 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 27.9 | 104.6
2008-07-16 | 13:51 | 0 | 0 | 0 | -0.1 | 104.5 | 104.5 | -0.2 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4 | 195 | 59.338 | 0.011838 | 0.012238 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6203 | 29.8483 | 27.9 | 104.5
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.1 | 104.438 | 104.5 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4 | 195 | 59.362 | 0.0118 | 0.012138 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7026 | 27.9 | 104.4
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.3 | 104.3375 | 104.6 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0117 | 0.012 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.7794 | 27.9 | 104.3
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.3 | 104.3 | 104.6 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3 | 0.0117 | 0.012 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7442 | 27.9 | 104.3
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.4 | 104.3 | 104.7 | -0.3 | 0 | 4.2 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3 | 0.0116 | 0.0119 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.6217 | 29.7325 | 27.9 | 104.3
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.4 | 104.2 | 104.6 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.362 | 0.011538 | 0.0118 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8 | 27.9 | 104.2
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.5 | 104.1381 | 104.6 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3381 | 0.0115 | 0.011738 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8 | 27.9 | 104.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 104.1 | 104.7 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3618 | 0.011438 | 0.011638 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8 | 27.9 | 104.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.6 | 104.1 | 104.7 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3382 | 0.0114 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.618 | 29.8 | 27.9 | 104.1
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 104 | 104.7 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0114 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8161 | 27.9 | 104
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 104 | 104.7 | -0.4 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0113 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.893 | 27.9 | 104
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.7 | 104 | 104.7 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.362 | 0.0113 | 0.011538 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.7599 | 27.9 | 104
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.8 | 103.9381 | 104.7 | -0.4 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0113 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.7937 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | 0 | 0 | 0.9 | 103.9 | 104.8 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3624 | 0.0113 | 0.011438 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.9236 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | -0.0007 | -6.789 | 1 | 103.9 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3 | 0.0112 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.377 | 30 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1 | 103.9 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3 | 0.0112 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1 | 104 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0112 | 0.0113 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.3795 | 30 | 27.9 | 104
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.1 | 103.9381 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0111 | 0.0113 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011.6192 | 29.9695 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.1 | 103.9 | 105 | -0.2 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3382 | 0.0111 | 0.0113 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8927 | 27.9 | 103.9
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.1 | 103.8379 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.3621 | 0.0111 | 0.0112 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8159 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.1 | 103.8 | 104.9 | -0.2 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.4 | 0.0111 | 0.0112 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.6203 | 29.8 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 4 | 195 | 59.338 | 0.0111 | 0.0112 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.2 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4 | 195 | 59.3382 | 0.011038 | 0.0112 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013.6175 | 29.8144 | 27.9 | 103.7
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4.6 | 195 | 59.3376 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013.3758 | 29.8912 | 27.9 | 103.7
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1014 | 29.8321 | 27.9 | 103.7
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3622 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8893 | 27.9 | 103.7
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.362 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.9787 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 4.4 | 195 | 59.3 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9019 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 4 | 195 | 59.3381 | 0.011 | 0.0111 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7503 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0109 | 0.011038 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 27.9 | 103.8
2008-07-16 | 13:51 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7566 | 27.9 | 103.8
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.8 | 105 | -0.2 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.4 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 27.9 | 103.8
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.738 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.4 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8414 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.623 | 29.8 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3622 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1013 | 29.8 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.338 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012.3795 | 29.8 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3618 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8657 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3377 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8151 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.0007 | -6.5638 | 1.2 | 103.8 | 105 | -0.4 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0109 | 0.011 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 103.8
2008-07-16 | 13:52 | 0 | 0 | 0 | 1.2 | 103.738 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.338 | 0.0108 | 0.010938 | 0.0062 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8454 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | 0 | -0.11676 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0108 | 0.0109 | 0.0061 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8506 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.0008 | -8.457 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.0108 | 0.0109 | 0.0061 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8262 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.3 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3623 | 0.0108 | 0.0109 | 0.0061 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9058 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.2 | 103.7 | 104.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0108 | 0.0109 | 0.0061 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0592 | 27.9 | 103.7
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 1.1 | 99.7453 | 100.9 | -0.4 | 0 | 4.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0108 | 0.0109 | 0.005913 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8747 | 27.9 | 99.7
2008-07-16 | 13:52 | -0.009 | -0.001 | -10 | -0.4 | 72.847 | 72.8 | -0.4 | 0 | 3.67 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.338 | 0.010252 | 0.010552 | 0.004328 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.6205 | 29.7988 | 27.9 | 72.8
2008-07-16 | 13:52 | -0.003 | -0.001 | -10 | -0.2 | 56.9427 | 56.9 | -0.5 | 0 | 2.67 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.008752 | 0.009014 | 0.0034 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9812 | 27.9 | 56.9
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.7 | 53.5377 | 53.5 | -0.4 | 0 | 2.4 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.4 | 0.008438 | 0.008738 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8271 | 27.9 | 53.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.6 | 53.5 | 53.5 | -0.5 | 0 | 2.3 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.008238 | 0.008575 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8629 | 27.9 | 53.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.6 | 53.5 | 53.5 | -0.4 | 0 | 2.2 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0082 | 0.0085 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9794 | 27.9 | 53.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.5 | 53.5 | 53.5 | -0.6 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.362 | 0.0081 | 0.0084 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0502 | 27.9 | 53.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.3 | 53.5 | 53.5 | -0.5 | 0 | 2.2 | 12.9 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.008 | 0.0082 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8209 | 27.9 | 53.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.2 | 53.4382 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3618 | 0.0079 | 0.0081 | 0.0032 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.87 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | -0.001 | -10 | -0.1 | 53.4 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 4 | 195 | 59.3 | 0.0078 | 0.008 | 0.003138 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9469 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0 | 53.4 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0077 | 0.0079 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.929 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0 | 53.4 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0077 | 0.0079 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0 | 53.4 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0076 | 0.007738 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | -0.0008 | -8.2598 | 0.1 | 53.4 | 53.5 | -0.5 | 0 | 2.2 | 12.9 | 952 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3619 | 0.0075 | 0.007638 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7541 | 27.9 | 53.4
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.2 | 53.3 | 53.5 | -0.4 | 0 | 2.2 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3377 | 0.0074 | 0.007538 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9233 | 27.9 | 53.3
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.2 | 53.2378 | 53.4 | -0.6 | 0 | 2.2 | 12.9 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3622 | 0.007338 | 0.0075 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.185 | 27.9 | 53.2
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.3 | 53.2 | 53.5 | -0.4 | 0 | 2.2 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.362 | 0.0073 | 0.007438 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0315 | 27.9 | 53.2
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.3 | 53.1 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0072 | 0.0074 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.878 | 27.9 | 53.1
2008-07-16 | 13:52 | 0 | -0.0004 | -3.6415 | 0.3 | 53.1 | 53.4 | -0.3 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3619 | 0.0072 | 0.0073 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010.3812 | 29.8378 | 27.9 | 53.1
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0.4 | 53.1 | 53.5 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3624 | 0.0071 | 0.0072 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.9145 | 27.9 | 53.1
2008-07-16 | 13:52 | 0 | -0.001 | -9.9414 | 0.4 | 53.0377 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3377 | 0.0071 | 0.0072 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 29.9912 | 27.9 | 53
2008-07-16 | 13:52 | 0 | -0.0002 | -1.5732 | 0.4 | 53 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.007 | 0.0071 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 30 | 27.9 | 53
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.5 | 52.9 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.007 | 0.0071 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9553 | 27.9 | 52.9
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.5 | 52.9 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.3 | 0.0069 | 0.007 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9214 | 27.9 | 52.9
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.6 | 52.8 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0069 | 0.007 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9981 | 27.9 | 52.8
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.6 | 52.7378 | 53.3 | -0.3 | 0 | 2.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4 | 195 | 59.4 | 0.0069 | 0.006938 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.6218 | 29.9253 | 27.9 | 52.7
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.6 | 52.7 | 53.3 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3 | 0.006838 | 0.0069 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9515 | 27.9 | 52.7
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.6 | 52.662 | 53.3 | -0.3 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.362 | 0.0068 | 0.0069 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.3797 | 29.9717 | 27.9 | 52.7
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.7 | 52.638 | 53.3 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.362 | 0.0068 | 0.0068 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8949 | 27.9 | 52.6
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.7 | 52.6 | 53.3 | -0.3 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 4.6 | 195 | 59.4 | 0.0067 | 0.0068 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8182 | 27.9 | 52.6
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.7 | 52.6 | 53.3 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3624 | 0.0067 | 0.0068 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9171 | 27.9 | 52.6
2008-07-16 | 13:52 | 0 | -0.0002 | -1.729 | 0.8 | 52.6 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3623 | 0.0067 | 0.0067 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9294 | 27.9 | 52.6
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0.8 | 52.6 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.006638 | 0.0067 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8 | 27.9 | 52.6
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0.9 | 52.5 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3381 | 0.0066 | 0.0066 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8 | 27.9 | 52.5
2008-07-16 | 13:52 | 0 | -0.001 | -10 | 0.8 | 52.5 | 53.3 | -0.4 | 0 | 2.2 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0066 | 0.0066 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.931 | 27.9 | 52.5
2008-07-16 | 13:52 | 0 | -0.0005 | -4.9541 | 0.8 | 52.5 | 53.3 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0066 | 0.0066 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9578 | 27.9 | 52.5
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.9 | 52.5 | 53.4 | -0.4 | 0 | 2.2 | 12.9 | 952 | 959 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.006538 | 0.006538 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 27.9 | 52.5
2008-07-16 | 13:52 | 0 | -0.0002 | -1.729 | 0.9 | 52.5 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3 | 0.0065 | 0.0065 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.3785 | 29.9 | 27.9 | 52.5
2008-07-16 | 13:52 | 0 | -0.001 | -9.958 | 1 | 52.5 | 53.5 | -0.3 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.4 | 0.0065 | 0.0065 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9724 | 27.9724 | 52.5
2008-07-16 | 13:52 | 0 | -0.0002 | -1.5546 | 1 | 52.4 | 53.4 | -0.5 | 0 | 2.2 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4 | 195 | 59.3374 | 0.0065 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9508 | 27.9508 | 52.4
2008-07-16 | 13:52 | 0 | 0 | 0 | 1 | 52.4 | 53.4 | -0.3 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3617 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 27.9 | 52.4
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.9 | 52.4 | 53.3 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.4 | 195 | 59.3 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8973 | 27.9027 | 52.4
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.9 | 52.4 | 53.3 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 4.6 | 195 | 59.3 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.379 | 29.8206 | 27.9794 | 52.4
2008-07-16 | 13:52 | 0 | 0 | 0 | 0.9 | 52.4 | 53.3 | -0.5 | 0 | 2.2 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3374 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8 | 28 | 52.4
2008-07-16 | 13:52 | 0 | 0 | -0.034423 | 1 | 52.3381 | 53.3 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32.107 | 55 | 1048 | 1011.6192 | 29.8968 | 27.9677 | 52.3
2008-07-16 | 13:53 | 0 | -0.0009 | -8.6403 | 0.9 | 52.3 | 53.2 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.4 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32.893 | 55 | 1048 | 1012 | 30.0908 | 27.9 | 52.3
2008-07-16 | 13:53 | 0 | -0.001 | -10 | 0.9 | 52.3 | 53.2 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.0064 | 0.0064 | 0.0031 | 0 |  | n/a |  | 32.107 | 55 | 1048 | 1012 | 30.0144 | 27.9 | 52.3
2008-07-16 | 13:53 | 0 | -0.001 | -10 | 0.9 | 52.3 | 53.2 | -0.4 | 0 | 2.2 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 29.9374 | 27.9626 | 52.3
2008-07-16 | 13:53 | 0 | -0.001 | -10 | 0.9 | 52.2381 | 53.1 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3381 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 29.9396 | 27.9604 | 52.2
2008-07-16 | 13:53 | 0 | -0.0008 | -8.2739 | 0.9 | 52.2381 | 53.1 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3619 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.2 | 53.1 | -0.5 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3624 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011.3758 | 30 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.2 | 53.1 | -0.3 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3377 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 29.8602 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.2 | 53.1 | -0.4 | 0 | 2.2 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3622 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 29.8468 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.2 | 53.1 | -0.3 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 29.9235 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.2 | 53.1 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.0005 | 27.9 | 52.2
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.1382 | 53 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3618 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011 | 30.154 | 27.9 | 52.1
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.1 | 53 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3634 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011.6343 | 29.9314 | 27.9 | 52.1
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.9 | 52.1 | 53 | -0.4 | 0 | 2.2 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3621 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 29.7913 | 27.9 | 52.1
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.8 | 52.1 | 52.9 | -0.3 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.362 | 0.0063 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1011.6203 | 30 | 27.9 | 52.1
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.8 | 52.1 | 52.9 | -0.4 | 0 | 2.2 | 13 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0062 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30 | 27.9 | 52.1
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.6 | 44.9224 | 45.5 | -0.3 | 0 | 2.2 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.0062 | 0.006238 | 0.002667 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30 | 27.9606 | 44.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.7 | 24.2576 | 24.3 | -0.5 | 0 | 1.59 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.005713 | 0.0062 | 0.001425 | 0 |  | n/a |  | 32.893 | 55 | 1048 | 1012 | 29.9627 | 27.9627 | 24.3
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.9 | 5.7016 | 5.7 | -0.4 | 0 | 0.7 | 13 | 950 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3623 | 0.004251 | 0.004651 | 0.000338 | 0 |  | n/a |  | 32.107 | 55 | 1048 | 1012 | 29.9281 | 27.9 | 5.7
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.4 | 2.1378 | 2.1 | -0.6 | 0 | 0.35 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.003738 | 0.004038 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 30.0814 | 27.9 | 2.1
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.4 | 2 | 2 | -0.5 | 0 | 0.2 | 13 | 949 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0036 | 0.003876 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 29.965 | 27.9 | 2
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.3 | 2 | 2 | -0.6 | 0 | 0.2 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.0035 | 0.0038 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1048 | 1012 | 29.9 | 27.9 | 2
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.3 | 1.9 | 1.9 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.4 | 0.003438 | 0.0037 | 0.0001 | 0 |  | n/a |  | 32.893 | 55 | 1048 | 1012 | 29.9 | 27.9 | 1.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.2 | 1.8 | 1.8 | -0.6 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3622 | 0.003338 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 27.9 | 1.8
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.2 | 1.7 | 1.7 | -0.6 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.003238 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9745 | 27.9745 | 1.7
2008-07-16 | 13:53 | 0 | 0 | 0 | -1.1 | 1.6 | 1.6 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0032 | 0.003438 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9488 | 28 | 1.6
2008-07-16 | 13:53 | 0 | 0 | 0 | -1 | 1.5 | 1.5 | -0.6 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0031 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9 | 28 | 1.5
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.9 | 1.4 | 1.4 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.003 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9047 | 28 | 1.4
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.9 | 1.3376 | 1.3 | -0.6 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3376 | 0.002938 | 0.003138 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9814 | 28 | 1.3
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.8 | 1.3 | 1.3 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3377 | 0.0029 | 0.003038 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.377 | 29.9418 | 28 | 1.3
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.7 | 1.3 | 1.3 | -0.6 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0028 | 0.002938 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 1.3
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.6 | 1.2 | 1.2 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.002738 | 0.002838 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 1.2
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.5 | 1.1 | 1.1 | -0.6 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3382 | 0.002638 | 0.0028 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 1.1
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.5 | 1.1 | 1.1 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.4 | 0.0026 | 0.0027 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 27.9348 | 1.1
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.5 | 1.1 | 1.1 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3622 | 0.0026 | 0.0027 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 27.9419 | 1.1
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.5 | 1 | 1 | -0.6 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0026 | 0.002638 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8814 | 28 | 1
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.4 | 1 | 1 | -0.4 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0025 | 0.0026 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8046 | 28 | 1
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.4 | 0.9 | 0.9 | -0.6 | 0 | 0.1 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.4 | 0.0025 | 0.0026 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8721 | 27.9279 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.3 | 0.9 | 0.9 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0024 | 0.0025 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8511 | 27.9489 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.3 | 0.9 | 0.9 | -0.6 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0024 | 0.0024 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8512 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.2 | 0.9 | 0.9 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0024 | 0.0024 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.6217 | 29.9954 | 27.9977 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.2 | 0.9 | 0.9 | -0.6 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.362 | 0.0023 | 0.0024 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8419 | 27.921 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.1 | 0.9 | 0.9 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0023 | 0.0023 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9111 | 27.9555 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.1 | 0.9 | 0.9 | -0.5 | 0 | 0.1 | 13 | 950 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.002238 | 0.0023 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9678 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.1 | 0.9 | 0.9 | -0.5 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0022 | 0.0022 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9183 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | -0.1 | 0.9 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0022 | 0.0022 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0711 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | 0 | 0.9 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0022 | 0.0022 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.975 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | 0 | 0.9 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3618 | 0.0022 | 0.0022 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9 | 28 | 0.9
2008-07-16 | 13:53 | 0 | 0 | 0 | 0 | 0.83758 | 0.8 | -0.5 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.002138 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8839 | 28 | 0.8
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.1 | 0.8 | 0.9 | -0.3 | 0 | 0.1 | 13 | 950 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.0021 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8071 | 28 | 0.8
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.1 | 0.8 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.0021 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9393 | 28 | 0.8
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.1 | 0.8 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30 | 28 | 0.8
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.1 | 0.73808 | 0.8 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3619 | 0.002038 | 0.002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.9538 | 28 | 0.7
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.1 | 0.7 | 0.8 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.002 | 0.002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8003 | 28 | 0.7
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.7 | 0.9 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3622 | 0.002 | 0.002 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8765 | 28 | 0.7
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.7 | 0.9 | -0.4 | 0 | 0.1 | 13 | 952 | 960 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.002 | 0.0019 | 0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.621 | 29.9533 | 28 | 0.7
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.63797 | 0.8 | -0.4 | 0 | 0.1 | 13 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.338 | 0.0019 | 0.0019 | 3.8e-05 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.9399 | 28 | 0.6
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.0019 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 0.6
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.3 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3626 | 0.0019 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8 | 28 | 0.6
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3376 | 0.0019 | 0.0018 | 6.2e-05 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8602 | 28 | 0.6
2008-07-16 | 13:53 | 0 | 0 | 0 | 0.2 | 0.5377 | 0.7 | -0.3 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3377 | 0.0019 | 0.0018 | 3.8e-05 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.863 | 28 | 0.5
2008-07-16 | 13:54 | 0 | 0 | 0 | 0.2 | 0.5 | 0.7 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0019 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.8549 | 28 | 0.5
2008-07-16 | 13:54 | 0 | -0.0007 | -6.7723 | 0.2 | 0.4 | 0.6 | -0.4 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0019 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.1617 | 28 | 0.4
2008-07-16 | 13:54 | 0 | -0.001 | -10 | -0.2 | 0.3 | 0.3 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3383 | 0.001738 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 30.0657 | 28 | 0.3
2008-07-16 | 13:54 | 0 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8684 | 28 | 0.3
2008-07-16 | 13:54 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -0.5 | 0 | 0.1 | 13 | 951 | 959 | 947 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.379 | 29.7 | 28 | 0.2
2008-07-16 | 13:54 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -0.5 | 0 | 0.1 | 12.9 | 951 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.338 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.6203 | 29.7 | 28 | 0.2
2008-07-16 | 13:54 | 0 | -0.0001 | -1.4901 | -0.6 | 0.2 | 0.2 | -0.5 | 0 | 0.1 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.362 | 0.0017 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7741 | 28 | 0.2
2008-07-16 | 13:54 | 0 | 0 | 0 | -0.6 | 0.2 | 0.2 | -0.6 | 0 | 0.1 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3619 | 0.0016 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8508 | 28 | 0.2
2008-07-16 | 13:54 | 0 | 0 | 0 | -0.6 | 0.2 | 0.2 | -0.4 | 0 | 0.1 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.8724 | 28 | 0.2
2008-07-16 | 13:54 | 0 | 0 | 0 | -0.7 | 0.2 | 0.2 | -0.6 | 0 | 0.1 | 12.9 | 952 | 959 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.7957 | 28 | 0.2
2008-07-16 | 13:54 | 0 | 0 | 0 | -0.6 | 0.2 | 0.2 | -0.4 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3622 | 0.0016 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.719 | 28 | 0.2
2008-07-16 | 13:54 | 0 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.6 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.001738 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 29.7 | 28 | 0.1
2008-07-16 | 13:54 | -0.005 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.5 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3381 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 29.838 | 28 | 0.1
2008-07-16 | 13:54 | -0.01 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.5 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 30.0888 | 28 | 0.1
2008-07-16 | 13:54 | -0.01 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.5 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010.3782 | 30.0121 | 28 | 0.1
2008-07-16 | 13:54 | -0.01 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.5 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 30 | 28 | 0.1
2008-07-16 | 13:54 | -0.002 | 0 | 0 | -0.6 | 0.1 | 0.1 | -0.5 | 0 | 0.1 | 12.9 | 952 | 960 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0016 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1010 | 30 | 28 | 0.1
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
No Processing Errors Detected
Test Date | 2008-07-16
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.12
CO2 Span(%) | 12
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 1505
NO2 Span(ppm) | 505
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 965.4
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | White Van
License Plate
Engine Displacement | 7.3
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 13:33
Test End Time | 13:54
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 993
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.832
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 07/16/2008 13:27:51.526 - None Found
0X0000 - 07/16/2008 13:30:11.279 - None Found
0X0000 - 07/16/2008 13:32:44.484 - None Found
Warnings:
0X0000 - 07/16/2008 13:27:51.526 - None Found
0X0000 - 07/16/2008 13:30:11.279 - None Found
0X0000 - 07/16/2008 13:32:44.485 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1263
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-07-16
Time | 13:28
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
NO(ppm) | -1.3687 | -0.013333 | -1.3554
NO2(ppm) | 0.25 | -0.14667 | 0.39667
Span
InfoVer | 1
Date | 2008-07-16
Time | 13:30
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 1200 | 12 | 20.9 | 1989 | x^0: 1505; x^1: 505; x^2: 0; x^3: 965.4
Gas | Previous | Current | Difference
NO2(ppm) | 503.475 | 505.18 | -1.705
Test Information:
SEMTECH_DATA_FILE | 2008-07-16 | 13:27 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## NO2 Linearity
Test Date:
Test Time:
7/16/2008
13:32:44
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## NDUV
Test Date:
Test Time:
7/16/2008
13:32:44
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## O2 Linearity
Test Date:
Test Time:
7/16/2008
13:32:44
SEMTECH Serial Number:
D06-SDS01

## O2
Test Date:
Test Time:
7/16/2008
13:32:44
SEMTECH Serial Number:
D06-SDS01