## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 66
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 28261
NDUV Serial Number: | 131
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | NO2 NDUV Serial Number: |  |  | 131 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-12-02 |  | 518.6 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 13:52
Test software: | 2.018 66
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0.5 | 0.5 |  |  | 0 | 0.1 | 0.1
90 | 105 | 1 | 51.86 | 52 | 0.1 | 0.19 |  | 2.09 | 2 | -0.09 | -4.31
150 | 165 | 2 | 103.72 | 104.6 | 0.9 | 0.87 |  | 4.18 | 4 | -0.18 | -4.31
210 | 225 | 3 | 155.58 | 156.4 | 0.8 | 0.51 |  | 6.27 | 6 | -0.27 | -4.31
270 | 285 | 4 | 207.44 | 208.1 | 0.7 | 0.34 |  | 8.36 | 7.9 | -0.46 | -5.5
330 | 345 | 5 | 259.3 | 260.5 | 1.2 | 0.46 |  | 10.45 | 9.9 | -0.55 | -5.26
390 | 405 | 6 | 311.16 | 313.5 | 2.3 | 0.74 |  | 12.54 | 11.9 | -0.64 | -5.1
450 | 465 | 7 | 363.02 | 365.3 | 2.3 | 0.63 |  | 14.63 | 13.9 | -0.73 | -4.99
510 | 525 | 8 | 414.88 | 417.1 | 2.2 | 0.53 |  | 16.72 | 15.9 | -0.82 | -4.9
570 | 585 | 9 | 466.74 | 468.4 | 1.7 | 0.36 |  | 18.81 | 17.9 | -0.91 | -4.84
630 | 645 | 10 | 518.6 | 520.2 | 1.6 | 0.31 |  | 20.9 | 20 | -0.9 | -4.31
690 | 705 | 9 | 466.74 | 470.4 | 3.7 | 0.79 |  | 18.81 | 18 | -0.81 | -4.31
750 | 765 | 8 | 414.88 | 419.1 | 4.2 | 1.01 |  | 16.72 | 16 | -0.72 | -4.31
810 | 825 | 7 | 363.02 | 367.1 | 4.1 | 1.13 |  | 14.63 | 14 | -0.63 | -4.31
870 | 885 | 6 | 311.16 | 315.4 | 4.2 | 1.35 |  | 12.54 | 12 | -0.54 | -4.31
930 | 945 | 5 | 259.3 | 263 | 3.7 | 1.43 |  | 10.45 | 10 | -0.45 | -4.31
990 | 1005 | 4 | 207.44 | 209.9 | 2.5 | 1.21 |  | 8.36 | 8 | -0.36 | -4.31
1050 | 1065 | 3 | 155.58 | 158.1 | 2.5 | 1.61 |  | 6.27 | 6 | -0.27 | -4.31
1110 | 1125 | 2 | 103.72 | 106 | 2.3 | 2.22 |  | 4.18 | 4.1 | -0.08 | -1.91
1170 | 1185 | 1 | 51.86 | 52.4 | 0.5 | 0.96 |  | 2.09 | 2.1 | 0.01 | 0.48
1230 | 1245 | 0 | 0 | -0.7 | -0.7 |  |  | 0 | 0.1 | 0.1
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | 0.6 | 5.2 |  | Intercept | 0.06 | 0.21
|  |  |  | Slope | 1.01 | 0.99 - 1.01 |  | Slope | 0.95 | 0.99 - 1.01 | Fail
|  |  |  | SEE | 1.1374 | 5.186 |  | SEE | 0.0571 | 0.209
|  |  |  | r2 | 1 | 0.998 |  | r2 | 0.9999 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 1.7 | 2 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8335 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.3 | 28 | 1.7
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 1.6667 | 2 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1425 | 28 | 1.7
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 1.6 | 1.9 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 1.6
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 0.7 | 1 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.7
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 0.7 | 1 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0011 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0858 | 28 | 0.7
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.3 | 0.7 | 1 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.001 | 0.001 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0072 | 28 | 0.7
2008-12-02 | 13:52 | -0.003 | 0 | 0 | 0.2 | 0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8331 | 0.001 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.2136 | 28 | 0.7
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.001 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.2503 | 28 | 0.7
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.001 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.172 | 28 | 0.7
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.66621 | 0.9 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.001 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.7
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.000966 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8665 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1204 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.2 | 0.6 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.199 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0009 | 0.0009 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1226 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8331 | 0.000867 | 0.000867 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.0657 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0.0007 | 6.5459 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0.001 | 10 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0.001 | 10 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.0697 | 28 | 0.6
2008-12-02 | 13:52 | -0.01 | 0.001 | 10 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.0518 | 28 | 0.6
2008-12-02 | 13:52 | 0.003 | 0.001 | 10 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8667 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.0266 | 28 | 0.6
2008-12-02 | 13:52 | 0.01 | 0.0002 | 1.9099 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1014 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | 0.01 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8665 | 0.0008 | 0.0008 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.6
2008-12-02 | 13:52 | 0.008 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.000767 | 0.0007 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.162 | 28 | 0.6
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0007 | 0.0007 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1191 | 28 | 0.6
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.833 | 0.0007 | 0.0007 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0189 | 28 | 0.6
2008-12-02 | 13:52 | 0 | 0 | 0 | 0.1 | 0.6 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0007 | 0.0007 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0972 | 28 | 0.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.8 | 59.8662 | 0.0007 | 0.0007 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.6
2008-12-02 | 13:53 | 0 | 0.0013 | 12.9274 | 0 | 0.6 | 0.6 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0007 | 0.0007 | 6.6e-05 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013.336 | 11.1 | 28 | 0.6
2008-12-02 | 13:53 | 0 | 0.002 | 20 | 0 | 0.6 | 0.6 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013.6653 | 11.1 | 28 | 0.6
2008-12-02 | 13:53 | 0 | 0.0017 | 16.7979 | 0 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8667 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1467 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8669 | 0.0007 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1498 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.46693 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.000667 | 0.0007 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0.001 | 10 | 0 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0.0009 | 8.5108 | 0 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8329 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0604 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8338 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8336 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.1 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0006 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8667 | 0.000567 | 0.0006 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013.6667 | 11.0903 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8331 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0118 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 196 | 59.8669 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0665 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.055 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.5 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0039 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4664 | 0.5 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1608 | 28 | 0.5
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1415 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.063 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8331 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0158 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0943 | 28 | 0.4
2008-12-02 | 13:53 | -0.008 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8331 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.1 | 28 | 0.4
2008-12-02 | 13:53 | -0.01 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013 | 11.0489 | 28 | 0.4
2008-12-02 | 13:53 | -0.005 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1050 | 1013.33 | 11 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0005 | 0.0005 | 0 | 0 |  | n/a |  | 35 | 61 | 1051.9945 | 1014 | 11 | 28 | 0.4
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 1.1435 | 1.1 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0005 | 0.0005 | 3.4e-05 | 0 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | 1.1
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 14 | 14 | -0.5 | 0 | 0.3 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.0005 | 0.0005 | 0.0007 | 0 |  | n/a |  | 35 | 61 | 1052.9892 | 1014 | 11 | 28 | 14
2008-12-02 | 13:53 | 0 | 0 | 0 | -0.1 | 29.7509 | 29.8 | -0.4 | 0 | 1.34 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0006 | 0.0006 | 0.001534 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 29.8
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.2 | 50.9667 | 51.2 | -0.3 | 0 | 1.79 | 13.2 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.001533 | 0.0015 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0214 | 28 | 51
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.4 | 51.6 | 52 | -0.3 | 0 | 1.94 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.7 | 59.8332 | 0.001933 | 0.001833 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0995 | 28 | 51.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.5 | 51.6 | 52.1 | -0.3 | 0 | 2 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0022 | 0.0021 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6646 | 11.0211 | 28 | 51.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.5 | 51.6 | 52.1 | -0.3 | 0 | 2 | 13.2 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0022 | 0.0021 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 51.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.5 | 51.6 | 52.1 | -0.2 | 0 | 2 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0022 | 0.0022 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 51.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.6 | 51.6344 | 52.2 | -0.2 | 0 | 2 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8656 | 0.0023 | 0.0022 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.014 | 28 | 51.6
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.6 | 51.7 | 52.3 | -0.2 | 0 | 2 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0023 | 0.0022 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0927 | 28 | 51.7
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.6 | 51.7 | 52.3 | -0.2 | 0 | 2 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8662 | 0.002501 | 0.002401 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0288 | 28 | 51.7
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.8 | 51.7 | 52.5 | -0.2 | 0 | 2 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8664 | 0.0028 | 0.0027 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 51.7
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.8 | 51.7 | 52.5 | -0.2 | 0 | 2 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8665 | 0.0029 | 0.0027 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3347 | 11.0279 | 28 | 51.7
2008-12-02 | 13:53 | 0 | 0 | 0 | 0.9 | 51.7 | 52.6 | -0.2 | 0 | 2 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0029 | 0.0027 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0942 | 28 | 51.7
2008-12-02 | 13:53 | 0.008 | 0.0002 | 1.5149 | 0.9 | 51.7 | 52.6 | -0.2 | 0 | 2 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8668 | 0.002933 | 0.0028 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.668 | 11.015 | 28 | 51.7
2008-12-02 | 13:53 | 0.01 | 0.001 | 9.9758 | 0.9 | 51.7 | 52.6 | -0.2 | 0 | 2 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.003034 | 0.0029 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3354 | 11.0635 | 28 | 51.7
2008-12-02 | 13:53 | 0.01 | 0.0002 | 1.8982 | 0.9 | 51.8 | 52.7 | -0.2 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0031 | 0.002933 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.142 | 28 | 51.8
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 0.9 | 51.8 | 52.7 | -0.2 | 0 | 2 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8669 | 0.0032 | 0.003033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1795 | 28 | 51.8
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1 | 51.8 | 52.8 | -0.2 | 0 | 2 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0033 | 0.0031 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1011 | 28 | 51.8
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1 | 51.8 | 52.8 | -0.2 | 0 | 2 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.0033 | 0.0031 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1773 | 28 | 51.8
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.2 | 0 | 2 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.1 | 59.8 | 0.003333 | 0.0032 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1442 | 28 | 51.9
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8338 | 0.0034 | 0.0032 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0657 | 28 | 51.9
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0034 | 0.0032 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0128 | 28 | 51.9
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.003433 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0911 | 28 | 51.9
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0303 | 28 | 51.9
2008-12-02 | 13:53 | 0.01 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 51.9
2008-12-02 | 13:53 | 0.002 | 0 | 0 | 1.1 | 51.9 | 53 | -0.3 | 0 | 2 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0266 | 28 | 51.9
2008-12-02 | 13:53 | 0 | 0 | 0 | 1.2 | 51.9 | 53.1 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.095 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9 | 53.1 | -0.3 | 0 | 2 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0035 | 0.0033 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0165 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9 | 53.1 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.867 | 0.003533 | 0.0034 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.938 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9 | 53.1 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0036 | 0.0034 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.9808 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0036 | 0.0034 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0811 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9664 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.003634 | 0.003434 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0027 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9335 | 53.1 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 51.9333 | 53.1 | -0.3 | 0 | 2 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8667 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 51.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8668 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0328 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0888 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0106 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.068 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52 | 53.2 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 52
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.2 | 52.1 | 53.3 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8328 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0966 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.3 | 52.1 | 53.4 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0037 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6621 | 11.0182 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.3 | 52.1 | 53.4 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.003734 | 0.0035 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0603 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.3 | 52.1 | 53.4 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0038 | 0.003533 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0611 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.3 | 52.1 | 53.4 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0038 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0738 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8667 | 0.003833 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0476 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0039 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0312 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.0039 | 0.0036 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0904 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.4 | 52.1 | 53.5 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0039 | 0.003633 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0119 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3293 | 11.0664 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0551 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8665 | 0.0039 | 0.0037 | 0.0027 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0039 | 0.0037 | 0.0028 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.002 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8668 | 0.0039 | 0.0037 | 0.0028 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0804 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.3 | 0 | 2 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0039 | 0.0037 | 0.0028 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0416 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.4 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8667 | 0.0039 | 0.0037 | 0.0028 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0737 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 1.5 | 52.1 | 53.6 | -0.4 | 0 | 2 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.004 | 0.0037 | 0.0028 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1842 | 28 | 52.1
2008-12-02 | 13:54 | 0 | 0.0005 | 4.7696 | 1.5 | 67.9482 | 69.4 | -0.4 | 0 | 2.09 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.004 | 0.0038 | 0.003606 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1058 | 28 | 67.9
2008-12-02 | 13:54 | 0 | 0.001 | 10 | 1.8 | 90.0688 | 91.9 | -0.4 | 0 | 3.14 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.867 | 0.004232 | 0.003899 | 0.004765 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0274 | 28 | 90.1
2008-12-02 | 13:54 | 0 | 0.0009 | 8.589 | 2.6 | 100.3059 | 102.9 | -0.2 | 0 | 3.74 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.866 | 0.005338 | 0.004904 | 0.0053 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 100.3
2008-12-02 | 13:54 | 0 | 0 | 0.38556 | 2.9 | 102.9 | 105.8 | -0.2 | 0 | 3.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.006734 | 0.006234 | 0.0054 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 102.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 3 | 103 | 106 | -0.2 | 0 | 4 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.006834 | 0.006334 | 0.0054 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 103
2008-12-02 | 13:54 | 0 | 0 | 0 | 3 | 103.1 | 106.1 | -0.2 | 0 | 4 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.006933 | 0.006433 | 0.0054 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 103.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.1 | 103.2333 | 106.3 | -0.2 | 0 | 4 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8333 | 0.007 | 0.0065 | 0.005433 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 103.2
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.1 | 103.3 | 106.4 | -0.2 | 0 | 4 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.007 | 0.0065 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 103.3
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.1 | 103.3335 | 106.4 | -0.2 | 0 | 4 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8665 | 0.007034 | 0.0065 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9787 | 28 | 103.3
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.1 | 103.4331 | 106.5 | -0.2 | 0 | 4 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0071 | 0.006533 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9004 | 28 | 103.4
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.1 | 103.5331 | 106.7 | -0.2 | 0 | 4 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.007133 | 0.0066 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9788 | 28 | 103.5
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.2 | 103.7 | 106.9 | -0.2 | 0 | 4 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0072 | 0.0067 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0573 | 28 | 103.7
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.2 | 103.7 | 106.9 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0072 | 0.0067 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0645 | 28 | 103.7
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 103.8 | 107.1 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0073 | 0.0068 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0276 | 28 | 103.8
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 103.9 | 107.2 | -0.2 | 0 | 4 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.007334 | 0.0068 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1852 | 28 | 103.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 103.9335 | 107.2 | -0.2 | 0 | 4 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0074 | 0.0068 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1289 | 28 | 103.9
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 104.1 | 107.4 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8661 | 0.0074 | 0.0069 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 104.1 | 107.4 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.9 | 0.0075 | 0.0069 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1281 | 28 | 104.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 104.1 | 107.4 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8665 | 0.0075 | 0.0069 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3354 | 11.1936 | 28 | 104.1
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.3 | 104.2 | 107.5 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8333 | 0.0075 | 0.0069 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1151 | 28 | 104.2
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.4 | 104.2 | 107.6 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0075 | 0.0069 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2268 | 28 | 104.2
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.4 | 104.2 | 107.6 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0075 | 0.006933 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3306 | 11.2162 | 28 | 104.2
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.4 | 104.233 | 107.6 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.007533 | 0.007 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0798 | 28 | 104.2
2008-12-02 | 13:54 | 0 | 0.0006 | 6.4625 | 3.4 | 104.3 | 107.7 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.007 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0017 | 28 | 104.3
2008-12-02 | 13:54 | 0 | 0.0005 | 5.2071 | 3.4 | 104.3 | 107.7 | -0.2 | 0 | 4 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.007 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0771 | 28 | 104.3
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.5 | 104.3335 | 107.8 | -0.2 | 0 | 4 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.007 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6653 | 11.0443 | 28 | 104.3
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.5 | 104.4 | 107.9 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.007 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0342 | 28 | 104.4
2008-12-02 | 13:54 | 0 | 0 | 0 | 3.5 | 104.4 | 107.9 | -0.2 | 0 | 4 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.4
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.4 | 107.9 | -0.2 | 0 | 4 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0076 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.4
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.4667 | 108 | -0.2 | 0 | 4 | 13.2 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0076 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.4331 | 107.9 | -0.2 | 0 | 4 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0076 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.052 | 28 | 104.4
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.5 | 108 | -0.2 | 0 | 4 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0077 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0264 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.5 | 108 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8329 | 0.0077 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0951 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.5 | 108 | -0.2 | 0 | 4 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0077 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0166 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.5 | 108 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8338 | 0.0077 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.5 | 108 | -0.3 | 0 | 4 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.0077 | 0.0071 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0402 | 28 | 104.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.6 | 108.1 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0078 | 0.0072 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.6
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.5 | 104.6 | 108.1 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0078 | 0.0072 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.6
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.6 | 104.6 | 108.2 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.007833 | 0.0073 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0243 | 28 | 104.6
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.6 | 104.6 | 108.2 | -0.3 | 0 | 4 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0079 | 0.007334 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0541 | 28 | 104.6
2008-12-02 | 13:55 | 0 | 0.0008 | 7.875 | 3.6 | 104.6 | 108.2 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.6
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.6 | 104.7 | 108.3 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8331 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.089 | 28 | 104.7
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.6 | 104.7331 | 108.3 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0107 | 28 | 104.7
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.6 | 104.8 | 108.4 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.8
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.6 | 104.8 | 108.4 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.8
2008-12-02 | 13:55 | -0.006 | 0.001 | 10 | 3.6 | 104.8 | 108.4 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3379 | 11 | 28 | 104.8
2008-12-02 | 13:55 | -0.01 | 0.001 | 10 | 3.6 | 104.9 | 108.5 | -0.3 | 0 | 4 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8664 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | -0.01 | 0.001 | 10 | 3.6 | 104.9 | 108.5 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | -0.009 | 0.001 | 10 | 3.6 | 104.9 | 108.5 | -0.3 | 0 | 4 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0172 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8667 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3333 | 11.0956 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0.0002 | 1.9864 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0262 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8329 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8338 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0094 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.8336 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0878 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8665 | 0.008 | 0.0074 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 104.9
2008-12-02 | 13:55 | 0 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.008033 | 0.007433 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0553 | 28 | 104.9
2008-12-02 | 13:55 | -0.003 | 0 | 0 | 3.7 | 104.9 | 108.6 | -0.3 | 0 | 4 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0081 | 0.0075 | 0.0055 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 104.9
2008-12-02 | 13:55 | -0.009 | 0 | 0 | 3.7 | 108.2535 | 112 | -0.3 | 0 | 4 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.008134 | 0.0075 | 0.005701 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0018 | 28 | 108.3
2008-12-02 | 13:55 | -0.001 | 0 | 0 | 3.9 | 136.8 | 140.7 | -0.3 | 0 | 4.72 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0088 | 0.0081 | 0.0068 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0802 | 28 | 136.8
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.1 | 149.6265 | 153.8 | -0.2 | 0 | 5.58 | 13.2 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.009533 | 0.008899 | 0.007866 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0417 | 28 | 149.6
2008-12-02 | 13:55 | -0.006 | 0 | 0 | 4.3 | 154.3645 | 158.7 | -0.2 | 0 | 5.8 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.010133 | 0.0094 | 0.008133 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 154.4
2008-12-02 | 13:55 | -0.005 | 0 | 0 | 4.4 | 155.0329 | 159.5 | -0.2 | 0 | 5.9 | 13.2 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.010233 | 0.0095 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6714 | 11 | 28 | 155
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.5 | 155.1329 | 159.7 | -0.2 | 0 | 5.9 | 13.2 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8329 | 0.0103 | 0.0096 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3293 | 11 | 28 | 155.1
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.6 | 155.2338 | 159.8 | -0.2 | 0 | 5.9 | 13.1 | 922 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0104 | 0.0096 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6621 | 11 | 28 | 155.2
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.7 | 155.3 | 160 | -0.2 | 0 | 5.98 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.010734 | 0.009934 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.336 | 10.9492 | 28 | 155.3
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.8 | 155.3335 | 160.1 | -0.2 | 0 | 6 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8665 | 0.0108 | 0.01 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6653 | 10.9581 | 28 | 155.3
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.9 | 155.4332 | 160.3 | -0.2 | 0 | 6 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.010833 | 0.01 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0925 | 28 | 155.4
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.9 | 155.4335 | 160.3 | -0.2 | 0 | 6 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0109 | 0.0101 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6646 | 11.0137 | 28 | 155.4
2008-12-02 | 13:55 | 0 | 0 | 0 | 4.9 | 155.5 | 160.4 | -0.1 | 0 | 6 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0109 | 0.0101 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 155.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 5 | 155.5 | 160.5 | -0.1 | 0 | 6 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.011 | 0.010133 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 155.5
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.1 | 155.6 | 160.7 | -0.1 | 0 | 6 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.011333 | 0.0105 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 155.6
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.1 | 155.633 | 160.8 | -0.1 | 0 | 6 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0114 | 0.010533 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 155.6
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.2 | 155.8 | 161 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0115 | 0.0106 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11 | 28 | 155.8
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.3 | 155.8338 | 161.2 | -0.1 | 0 | 6 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8338 | 0.0118 | 0.0109 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.057 | 28 | 155.8
2008-12-02 | 13:55 | 0 | 0.0006 | 6.3699 | 5.4 | 155.9 | 161.3 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0119 | 0.011 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.664 | 11.1 | 28 | 155.9
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 5.4 | 156 | 161.4 | -0.1 | 0 | 6 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0119 | 0.011 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0864 | 28 | 156
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 5.4 | 156 | 161.4 | -0.1 | 0 | 6 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0119 | 0.011 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0076 | 28 | 156
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 5.4 | 156 | 161.4 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0119 | 0.011 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1418 | 28 | 156
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 5.4 | 156 | 161.4 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0119 | 0.011 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1507 | 28.0493 | 156
2008-12-02 | 13:55 | 0 | 0.001 | 10 | 5.5 | 156.0333 | 161.5 | -0.1 | 0 | 6 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8333 | 0.012 | 0.011033 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.0725 | 156
2008-12-02 | 13:55 | 0 | 0.0004 | 3.5537 | 5.5 | 156.1 | 161.6 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.012 | 0.0111 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.1
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.5 | 156.1 | 161.6 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.012 | 0.0111 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.1
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.5 | 156.2 | 161.7 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.012 | 0.0111 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1631 | 28 | 156.2
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.5 | 156.2 | 161.7 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.012 | 0.0111 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28 | 156.2
2008-12-02 | 13:55 | 0 | 0 | 0 | 5.6 | 156.2 | 161.8 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0121 | 0.0111 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1801 | 28 | 156.2
2008-12-02 | 13:56 | 0 | 0.0008 | 7.8469 | 5.6 | 156.3 | 161.9 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0121 | 0.011133 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1019 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.6 | 156.3 | 161.9 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8667 | 0.0121 | 0.0112 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.6 | 156.3 | 161.9 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0121 | 0.0112 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.4 | 162.1 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0121 | 0.0112 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0661 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.3667 | 162.1 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.012133 | 0.0112 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3333 | 11.0246 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.3 | 162 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8669 | 0.0122 | 0.0112 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1815 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.3 | 162 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0122 | 0.011233 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1308 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.3 | 162 | -0.1 | 0 | 6 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0122 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 156.3
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.4 | 162.1 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0.001 | 10 | 5.7 | 156.4 | 162.1 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8662 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0955 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0.0009 | 8.6066 | 5.8 | 156.4 | 162.2 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8336 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0173 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0 | 0.40984 | 5.8 | 156.4 | 162.2 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1227 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.8 | 156.5 | 162.3 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.12 | 28 | 156.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.8 | 156.5 | 162.3 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 156.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.8 | 156.4335 | 162.2 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6646 | 11 | 28 | 156.4
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.8 | 156.5667 | 162.4 | -0.1 | 0 | 6 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 156.6
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.8 | 156.5 | 162.3 | -0.1 | 0 | 6 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 156.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.9 | 156.5 | 162.4 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0322 | 28 | 156.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 5.9 | 156.533 | 162.4 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0123 | 0.0113 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0893 | 28 | 156.5
2008-12-02 | 13:56 | -0.003 | 0 | 0 | 5.9 | 156.6 | 162.5 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0123 | 0.011333 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0109 | 28 | 156.6
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 5.9 | 156.6 | 162.5 | -0.1 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.012334 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0676 | 28 | 156.6
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 5.9 | 156.6336 | 162.5 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.6
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 5.9 | 156.6665 | 162.6 | -0.1 | 0 | 6 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0755 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 5.9 | 156.7 | 162.6 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 156.7
2008-12-02 | 13:56 | -0.006 | 0 | 0 | 6 | 156.6 | 162.6 | -0.2 | 0 | 6 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 156.6
2008-12-02 | 13:56 | -0.003 | 0 | 0 | 6 | 156.6667 | 162.7 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.3333 | 11.0598 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6 | 156.7 | 162.7 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6 | 156.7 | 162.7 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6707 | 11.0832 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6 | 156.7 | 162.7 | -0.1 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0124 | 0.0114 | 0.0082 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0048 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6 | 156.734 | 162.7 | -0.2 | 0 | 6 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.834 | 0.0125 | 0.0115 | 0.008234 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0735 | 28 | 156.7
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6 | 156.8 | 162.8 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.012466 | 0.011466 | 0.0083 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.048 | 28 | 156.8
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6.1 | 156.9 | 163 | -0.2 | 0 | 6 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0125 | 0.011434 | 0.0083 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9694 | 28 | 156.9
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6.1 | 156.9 | 163 | -0.2 | 0 | 6 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0125 | 0.0115 | 0.0083 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9181 | 28 | 156.9
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6.1 | 159.7667 | 165.9 | -0.2 | 0 | 6 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0125 | 0.0115 | 0.008433 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.075 | 28 | 159.8
2008-12-02 | 13:56 | -0.01 | 0 | 0 | 6.1 | 172.6706 | 178.8 | -0.2 | 0 | 6.67 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0125 | 0.0115 | 0.008866 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.034 | 28 | 172.7
2008-12-02 | 13:56 | -0.004 | 0 | 0 | 6.5 | 201.2133 | 207.7 | -0.1 | 0 | 7.46 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.013435 | 0.012401 | 0.010601 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0444 | 28 | 201.2
2008-12-02 | 13:56 | 0 | 0 | 0 | 6.8 | 206.5 | 213.3 | 0 | 0 | 7.8 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.014933 | 0.0138 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 206.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 6.9 | 206.6331 | 213.5 | 0 | 0 | 7.89 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0151 | 0.013933 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 206.6
2008-12-02 | 13:56 | 0 | 0 | 0 | 7 | 206.7 | 213.7 | 0 | 0 | 7.9 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.1 | 59.9 | 0.015133 | 0.014 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 206.7
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.1 | 206.833 | 213.9 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.015233 | 0.0141 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.042 | 28 | 206.8
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.1 | 206.9 | 214 | 0 | 0 | 7.9 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0153 | 0.0141 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 206.9
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.2 | 207 | 214.2 | 0 | 0 | 7.9 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.015334 | 0.014134 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0151 | 28 | 207
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.2 | 207.1665 | 214.4 | 0 | 0 | 7.9 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0154 | 0.0142 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1013.3347 | 11.0936 | 28 | 207.2
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.3 | 207.2 | 214.5 | 0 | 0 | 7.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0154 | 0.0142 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.028 | 28 | 207.2
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.3 | 207.2668 | 214.6 | 0 | 0 | 7.9 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0154 | 0.0142 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 207.3
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.3 | 207.3 | 214.6 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8335 | 0.0155 | 0.0143 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0288 | 28 | 207.3
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.4 | 207.3 | 214.7 | 0 | 0 | 7.9 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0155 | 0.0143 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 207.3
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.4 | 207.3 | 214.7 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0155 | 0.0143 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 207.3
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.4 | 207.3331 | 214.7 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.015533 | 0.014333 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 207.3
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.5 | 207.4 | 214.9 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.015633 | 0.0144 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0575 | 28 | 207.4
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.5 | 207.4 | 214.9 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0157 | 0.0144 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 207.4
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.5 | 207.4338 | 214.9 | 0 | 0 | 7.9 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0157 | 0.0145 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 207.4
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.5 | 207.5 | 215 | 0 | 0 | 7.9 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0157 | 0.0145 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 207.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.5 | 207.5 | 215 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0157 | 0.0145 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0565 | 28 | 207.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.6 | 207.5 | 215.1 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0158 | 0.014533 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.065 | 28 | 207.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.6 | 207.5 | 215.1 | 0 | 0 | 7.9 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0158 | 0.0146 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0134 | 28 | 207.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.6 | 207.5335 | 215.2 | 0 | 0 | 7.9 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0159 | 0.0146 | 0.0109 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014.6653 | 11.0918 | 28 | 207.5
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.7 | 207.8331 | 215.5 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.6 | 59.9 | 0.015933 | 0.014633 | 0.010933 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0297 | 28 | 207.8
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.7 | 207.9 | 215.6 | 0 | 0 | 7.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.016 | 0.0147 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 207.9
2008-12-02 | 13:56 | 0 | 0 | 0 | 7.7 | 207.9329 | 215.7 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.016 | 0.014733 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.027 | 28 | 207.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 7.8 | 207.9329 | 215.8 | 0 | 0 | 7.9 | 13.1 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0161 | 0.0148 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 207.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 7.9 | 207.9662 | 215.9 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8338 | 0.0161 | 0.0148 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208
2008-12-02 | 13:57 | 0 | 0 | 0 | 7.9 | 207.9 | 215.8 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8336 | 0.0161 | 0.0148 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0374 | 28 | 207.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 7.9 | 207.9 | 215.8 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0161 | 0.0148 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.041 | 28 | 207.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 8 | 207.9333 | 215.9 | -0.1 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0162 | 0.014833 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 207.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 8 | 207.9668 | 216 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208
2008-12-02 | 13:57 | 0 | 0 | 0 | 8 | 208 | 216 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0162 | 0.0149 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1763 | 28 | 208
2008-12-02 | 13:57 | 0 | 0 | 0 | 8 | 208.1 | 216.1 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0163 | 0.0149 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1452 | 28 | 208.1
2008-12-02 | 13:57 | 0 | 0 | 0 | 8 | 208.1331 | 216.1 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0163 | 0.0149 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208.1
2008-12-02 | 13:57 | 0 | 0.0003 | 2.8916 | 8 | 208.2 | 216.2 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0163 | 0.014933 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0884 | 28 | 208.2
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.1 | 208.2 | 216.3 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0163 | 0.015 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0102 | 28 | 208.2
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.1 | 208.3 | 216.4 | 0 | 0 | 7.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.016333 | 0.015 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0685 | 28 | 208.3
2008-12-02 | 13:57 | 0 | -0.0006 | -5.751 | 8.1 | 208.3 | 216.4 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0164 | 0.015034 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0529 | 28 | 208.3
2008-12-02 | 13:57 | 0 | -0.001 | -10 | 8.1 | 208.3 | 216.4 | 0 | 0 | 7.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 208.3
2008-12-02 | 13:57 | -0.001 | -0.001 | -10 | 8.1 | 208.3 | 216.4 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0164 | 0.0151 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0039 | 208.3
2008-12-02 | 13:57 | -0.009 | -0.001 | -10 | 8.1 | 208.3667 | 216.5 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0817 | 208.4
2008-12-02 | 13:57 | -0.01 | -0.001 | -10 | 8.1 | 208.3 | 216.4 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0165 | 0.0151 | 0.011 | 3.3e-05 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0395 | 208.3
2008-12-02 | 13:57 | -0.01 | -0.001 | -10 | 8.2 | 208.3 | 216.5 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 208.3
2008-12-02 | 13:57 | -0.01 | -0.0002 | -2.1918 | 8.2 | 208.3333 | 216.5 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8667 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0177 | 28 | 208.3
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 8.2 | 208.3 | 216.5 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0962 | 28 | 208.3
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 8.2 | 208.4 | 216.6 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208.4
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 8.2 | 208.4 | 216.6 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0471 | 28.0529 | 208.4
2008-12-02 | 13:57 | -0.01 | -0.0003 | -2.9167 | 8.2 | 208.4 | 216.6 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0315 | 28.0685 | 208.4
2008-12-02 | 13:57 | -0.01 | -0.0029 | -28.9583 | 8.2 | 208.5 | 216.7 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208.5
2008-12-02 | 13:57 | -0.01 | -0.0014 | -14.2105 | 8.2 | 208.5 | 216.7 | 0 | 0 | 7.9 | 13.1 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 208.5
2008-12-02 | 13:57 | -0.01 | -0.001 | -10 | 8.2 | 208.5 | 216.7 | 0 | 0 | 7.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1664 | 28 | 208.5
2008-12-02 | 13:57 | -0.01 | -0.001 | -10 | 8.2 | 208.6 | 216.8 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1102 | 28 | 208.6
2008-12-02 | 13:57 | -0.01 | -0.0005 | -4.6598 | 8.2 | 208.6 | 216.8 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0238 | 28 | 208.6
2008-12-02 | 13:57 | -0.01 | 0.001 | 10 | 8.2 | 208.5667 | 216.8 | 0 | 0 | 7.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0165 | 0.0151 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1022 | 28 | 208.6
2008-12-02 | 13:57 | -0.004 | 0.001 | 10 | 8.2 | 208.5 | 216.7 | 0 | 0 | 7.9 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.015133 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1806 | 28 | 208.5
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.3 | 208.6 | 216.9 | 0 | 0 | 7.9 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0165 | 0.0152 | 0.011 | 0.0001 |  | n/a |  | 35 | 61 | 1053.3118 | 1014 | 11.0823 | 28 | 208.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.3 | 223.5863 | 231.9 | 0 | 0 | 8.06 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.016533 | 0.0152 | 0.011797 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 223.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.6 | 253.7454 | 262.3 | 0 | 0.0228 | 9.12 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.017731 | 0.016298 | 0.013399 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 253.7
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 8.9 | 258.2338 | 267.2 | 0.1 | 0.1 | 9.67 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.6 | 59.9 | 0.018901 | 0.017401 | 0.0136 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 258.2
2008-12-02 | 13:57 | 0 | 0.0009 | 9.1 | 9 | 258.3 | 267.3 | 0.1 | 0.1 | 9.8 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0191 | 0.0176 | 0.0136 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0729 | 28 | 258.3
2008-12-02 | 13:57 | -0.006 | 0.0001 | 0.76667 | 9.1 | 258.7012 | 267.8 | 0.1 | 0.1 | 9.8 | 13.1 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.019433 | 0.017933 | 0.013633 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.049 | 28 | 258.7
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 9.1 | 259.3333 | 268.5 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0195 | 0.018 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.0292 | 259.3
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 9.5 | 259.4332 | 268.9 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0196 | 0.018033 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.668 | 11 | 28.0918 | 259.4
2008-12-02 | 13:57 | -0.009 | -0.0009 | -8.7695 | 9.5 | 259.5335 | 269.1 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.019634 | 0.0181 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.0133 | 259.5
2008-12-02 | 13:57 | -0.001 | -0.0012 | -11.5113 | 9.6 | 259.6 | 269.2 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0197 | 0.0181 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 259.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 9.6 | 259.6 | 269.2 | 0.2 | 0.1921 | 9.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0197 | 0.0181 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.0434 | 259.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 9.6 | 259.6 | 269.2 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.019733 | 0.018133 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.0782 | 259.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 9.7 | 259.6 | 269.3 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0198 | 0.0182 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 259.6
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 9.7 | 259.7 | 269.4 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0198 | 0.0182 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 259.7
2008-12-02 | 13:57 | 0 | 0.001 | 10 | 9.7 | 259.7 | 269.4 | 0.2 | 0.1973 | 9.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.019834 | 0.018234 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 259.7
2008-12-02 | 13:57 | 0 | 0.0007 | 7.4521 | 9.7 | 259.7335 | 269.4 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0199 | 0.0183 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0356 | 28 | 259.7
2008-12-02 | 13:57 | -0.008 | 0 | 0 | 9.8 | 259.8 | 269.6 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0199 | 0.0183 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0864 | 28 | 259.8
2008-12-02 | 13:57 | -0.01 | 0 | 0 | 9.8 | 259.8 | 269.6 | 0.2 | 0.1518 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.02 | 0.0184 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0074 | 28 | 259.8
2008-12-02 | 13:57 | -0.006 | 0 | 0 | 9.8 | 259.833 | 269.6 | 0.1 | 0.1 | 9.9 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0201 | 0.0185 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28 | 259.8
2008-12-02 | 13:57 | 0 | 0 | 0 | 9.9 | 259.8 | 269.7 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0202 | 0.0185 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0496 | 28 | 259.8
2008-12-02 | 13:57 | 0 | 0 | 0 | 9.9 | 259.9 | 269.8 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0202 | 0.0185 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0721 | 28 | 259.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 9.9 | 259.9 | 269.8 | 0.2 | 0.186 | 9.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0202 | 0.0186 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0062 | 28 | 259.9
2008-12-02 | 13:57 | 0 | 0 | 0 | 10 | 260 | 270 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0202 | 0.0186 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0849 | 28 | 260
2008-12-02 | 13:57 | 0 | 0 | 0 | 10 | 260.0329 | 270.1 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0203 | 0.0186 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28 | 260
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.2 | 260.1 | 270.3 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0186 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0582 | 28 | 260.1
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.2 | 260.2 | 270.4 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0186 | 0.0137 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9798 | 28 | 260.2
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.2 | 260.2667 | 270.5 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0186 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3333 | 10.9019 | 28 | 260.3
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.2 | 260.3 | 270.5 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.020333 | 0.0187 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.054 | 28 | 260.3
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.3 | 260.3 | 270.6 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0204 | 0.0187 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0444 | 28 | 260.3
2008-12-02 | 13:57 | 0 | 0 | 0 | 10.3 | 260.3 | 270.6 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0204 | 0.0187 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.3
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.2669 | 270.6 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0204 | 0.0187 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3313 | 11 | 28.0124 | 260.3
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.3331 | 270.6 | 0.1 | 0.1 | 9.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6694 | 11 | 28.0907 | 260.3
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.4 | 270.7 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0693 | 28.0307 | 260.4
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.5 | 270.8 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0522 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.5 | 270.8 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.6 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.6 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.6 | 270.9 | 0.1 | 0.101 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0205 | 0.0188 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9384 | 28.0616 | 260.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.3 | 260.5668 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.020533 | 0.018833 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9 | 28.0599 | 260.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.4 | 260.5 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.020634 | 0.0189 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9185 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.4 | 260.5 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9969 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.4 | 260.5 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0754 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.4 | 260.5 | 270.9 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0462 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.4 | 260.534 | 271 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.5 | 260.6 | 271.1 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0108 | 28 | 260.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.5 | 260.7 | 271.2 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.336 | 11.0892 | 28 | 260.7
2008-12-02 | 13:58 | 0 | -0.0001 | -0.82835 | 10.5 | 260.7 | 271.2 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0323 | 28 | 260.7
2008-12-02 | 13:58 | 0 | -0.0009 | -9.3681 | 10.5 | 260.8339 | 271.3 | 0.1 | 0.1 | 9.9 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0207 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0461 | 28 | 260.8
2008-12-02 | 13:58 | 0 | -0.0002 | -2.1658 | 10.5 | 260.9668 | 271.5 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.020733 | 0.019 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28 | 261
2008-12-02 | 13:58 | 0.002 | 0 | 0 | 10.5 | 260.9 | 271.4 | 0.2 | 0.1998 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0208 | 0.019034 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.097 | 28 | 260.9
2008-12-02 | 13:58 | 0.01 | 0 | 0 | 10.5 | 260.8 | 271.3 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8333 | 0.0208 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0186 | 28 | 260.8
2008-12-02 | 13:58 | 0.01 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.2 | 0.1978 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.020833 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.0598 | 260.9
2008-12-02 | 13:58 | 0.002 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.2 | 0.2 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.0617 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.1 | 0.1014 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.1 | 0.1 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 261 | 271.6 | 0.2 | 0.1996 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9264 | 28 | 261
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 261 | 271.6 | 0.1 | 0.1006 | 9.9 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8664 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9521 | 28 | 261
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9665 | 271.6 | 0.1 | 0.1297 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0306 | 28 | 261
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.2 | 0.2 | 9.9 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.091 | 28.009 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.1 | 0.102 | 9.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 196 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0125 | 28.0875 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 260.9 | 271.5 | 0.2 | 0.2 | 9.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0209 | 0.0191 | 0.0138 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3354 | 11.0659 | 28.0341 | 260.9
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.6 | 262.8 | 273.4 | 0.2 | 0.2 | 9.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0209 | 0.0191 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6667 | 11.0114 | 28 | 262.8
2008-12-02 | 13:58 | 0 | 0 | 0 | 10.7 | 282.0208 | 292.8 | 0.2 | 0.2 | 10.13 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.020999 | 0.019199 | 0.014699 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9228 | 28.0228 | 282
2008-12-02 | 13:58 | 0 | 0 | 0 | 11 | 306.4944 | 317.5 | 0.2 | 0.2 | 11.19 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.021698 | 0.019898 | 0.016166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0987 | 306.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.1 | 308.626 | 319.8 | 0.3 | 0.3 | 11.7 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022199 | 0.020399 | 0.016333 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0203 | 308.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.2 | 310.5293 | 321.8 | 0.3 | 0.3 | 11.8 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022599 | 0.020732 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 310.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.3 | 311.2338 | 322.5 | 0.3 | 0.3 | 11.8 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.022834 | 0.021 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9636 | 28.0364 | 311.2
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.3 | 311.4336 | 322.8 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.023 | 0.0211 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1013 | 10.915 | 28.085 | 311.4
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.4 | 311.5 | 322.9 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023 | 0.021133 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9935 | 28.0065 | 311.5
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.5 | 311.6 | 323.1 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023133 | 0.0213 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28 | 311.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.5 | 311.6 | 323.1 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0232 | 0.0213 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0502 | 28.0502 | 311.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.6 | 311.6333 | 323.2 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023233 | 0.0213 | 0.0164 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.0714 | 311.6
2008-12-02 | 13:58 | 0 | 0 | 0 | 11.9 | 311.7 | 323.6 | 0.3 | 0.3 | 11.9 | 13.1 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023333 | 0.021433 | 0.016433 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0928 | 28 | 311.7
2008-12-02 | 13:58 | 0 | 0 | 0 | 12 | 311.8 | 323.8 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023433 | 0.0215 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0143 | 28 | 311.8
2008-12-02 | 13:58 | 0 | 0 | 0 | 12.1 | 311.833 | 323.9 | 0.3 | 0.3 | 11.9 | 13.2 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0236 | 0.0216 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0642 | 311.8
2008-12-02 | 13:58 | 0 | 0 | 0 | 12.1 | 311.934 | 324.1 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.023802 | 0.021802 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0575 | 311.9
2008-12-02 | 13:58 | 0 | -0.0001 | -0.75 | 12.2 | 311.9662 | 324.2 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.024 | 0.022 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0208 | 28.0208 | 312
2008-12-02 | 13:58 | 0 | -0.0009 | -9.0833 | 12.2 | 311.9336 | 324.2 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.024034 | 0.022 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.664 | 11.0994 | 28.0995 | 311.9
2008-12-02 | 13:58 | -0.004 | -0.001 | -10 | 12.3 | 312.0665 | 324.4 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0241 | 0.0221 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.3347 | 11.1779 | 28.0221 | 312.1
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.3 | 312.1333 | 324.4 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0241 | 0.0221 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1437 | 28 | 312.1
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.3 | 313.1 | 325.4 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0242 | 0.022133 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 313.1
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.4 | 313.267 | 325.7 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0242 | 0.0222 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1257 | 28 | 313.3
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0242 | 0.0222 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2831 | 28 | 313.2
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.024233 | 0.0222 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1598 | 28.0701 | 313.2
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0243 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0515 | 28.0515 | 313.2
2008-12-02 | 13:58 | -0.01 | -0.001 | -10 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0243 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.3293 | 11.0269 | 28 | 313.2
2008-12-02 | 13:58 | -0.009 | -0.0006 | -5.7177 | 12.4 | 313.1662 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8338 | 0.0243 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 313.2
2008-12-02 | 13:58 | -0.001 | 0 | 0 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0243 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28 | 313.2
2008-12-02 | 13:58 | 0 | 0 | 0 | 12.4 | 313.3 | 325.7 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0243 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0376 | 28 | 313.3
2008-12-02 | 13:58 | 0 | 0 | 0 | 12.4 | 313.2 | 325.6 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.024333 | 0.0223 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0408 | 28 | 313.2
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.4 | 313.2337 | 325.7 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0244 | 0.0224 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.0191 | 313.2
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.3 | 325.8 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0245 | 0.0224 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.0972 | 313.3
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4333 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0245 | 0.0224 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.3669 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0245 | 0.0224 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.3313 | 11.1546 | 28.0454 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4 | 325.9 | 0.3 | 0.3 | 11.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0245 | 0.0224 | 0.0165 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0245 | 0.0224 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1886 | 28.0114 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0245 | 0.0224 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1103 | 28.0897 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0245 | 0.0225 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.5 | 313.4336 | 325.9 | 0.3 | 0.3 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 196 | 59.9 | 0.024534 | 0.0225 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0533 | 28.1 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5335 | 326.1 | 0.3 | 0.3 | 11.9 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0246 | 0.022533 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0252 | 28.0748 | 313.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.6 | 326.2 | 0.4 | 0.4 | 11.9 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0246 | 0.0226 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.0036 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.6 | 326.2 | 0.4 | 0.4 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0246 | 0.0226 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.0816 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5667 | 326.2 | 0.4 | 0.4 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28.4 | 5 | 195.6 | 59.9 | 0.024633 | 0.0226 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1603 | 28.0397 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5669 | 326.2 | 0.4 | 0.4 | 11.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 29 | 5 | 195.2 | 59.9 | 0.0247 | 0.0226 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1221 | 28 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5 | 326.1 | 0.4 | 0.3998 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 29 | 5 | 195 | 59.9 | 0.0247 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0348 | 28 | 313.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5 | 326.1 | 0.4 | 0.3979 | 11.9 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28.6 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1917 | 28 | 313.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5 | 326.1 | 0.4 | 0.4 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1258 | 28 | 313.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.6 | 313.5 | 326.1 | 0.3 | 0.3018 | 11.9 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0474 | 28.0526 | 313.5
2008-12-02 | 13:59 | -0.004 | 0 | 0 | 12.6 | 313.4 | 326 | 0.4 | 0.4 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0623 | 28.1 | 313.4
2008-12-02 | 13:59 | -0.01 | 0 | 0 | 12.7 | 313.5 | 326.2 | 0.3 | 0.3114 | 11.9 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1809 | 28.1 | 313.5
2008-12-02 | 13:59 | -0.01 | 0 | 0 | 12.7 | 313.4667 | 326.2 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0241 | 28.1 | 313.5
2008-12-02 | 13:59 | -0.01 | 0 | 0 | 12.7 | 313.4332 | 326.1 | 0.3 | 0.3 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0661 | 28.1 | 313.4
2008-12-02 | 13:59 | -0.01 | 0 | 0 | 12.7 | 313.4 | 326.1 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0227 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.0554 | 313.4
2008-12-02 | 13:59 | -0.01 | 0 | 0 | 12.7 | 313.4 | 326.1 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0228 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.0233 | 313.4
2008-12-02 | 13:59 | -0.006 | 0 | 0 | 12.7 | 313.4331 | 326.1 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0248 | 0.0228 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0982 | 28.1 | 313.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.7 | 313.5669 | 326.3 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.024833 | 0.0228 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.3306 | 11.0198 | 28.1 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 313.6 | 326.4 | 0.4 | 0.4 | 11.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0249 | 0.0228 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 313.6 | 326.4 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.966 | 0.024934 | 0.022834 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 313.6 | 326.4 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.025 | 0.0229 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.0845 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 313.6336 | 326.4 | 0.4 | 0.4 | 11.9 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.025 | 0.0229 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.006 | 313.6
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 313.6665 | 326.5 | 0.4 | 0.4 | 11.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.025 | 0.0229 | 0.0166 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.3347 | 11 | 28.0725 | 313.7
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.8 | 315.7333 | 328.6 | 0.4 | 0.4 | 11.9 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.025 | 0.0229 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0507 | 28.1 | 315.7
2008-12-02 | 13:59 | 0 | 0 | 0 | 12.9 | 332.1853 | 345.1 | 0.4 | 0.4 | 12.19 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.025033 | 0.022933 | 0.017566 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0709 | 28.1 | 332.2
2008-12-02 | 13:59 | 0 | 0 | 0 | 13.5 | 359.5372 | 373.1 | 0.4 | 0.428 | 13.27 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.026201 | 0.024001 | 0.019034 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.0922 | 359.5
2008-12-02 | 13:59 | 0 | -0.0001 | -0.77055 | 13.7 | 362.7 | 376.4 | 0.5 | 0.5 | 13.72 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0268 | 0.0245 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.0137 | 362.7
2008-12-02 | 13:59 | 0 | -0.0009 | -9.3322 | 13.8 | 362.7331 | 376.5 | 0.5 | 0.5 | 13.8 | 13.1 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.026999 | 0.024799 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0647 | 28.0647 | 362.7
2008-12-02 | 13:59 | 0 | -0.001 | -10 | 14 | 362.9 | 376.9 | 0.5 | 0.5 | 13.82 | 13.1 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.027899 | 0.025599 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0144 | 28.1 | 362.9
2008-12-02 | 13:59 | 0 | -0.001 | -10 | 14.1 | 363.8 | 377.9 | 0.6 | 0.6 | 13.9 | 13.1 | 924 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.028133 | 0.0258 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.9209 | 28.1 | 363.8
2008-12-02 | 13:59 | 0 | -0.001 | -10 | 14.2 | 363.9 | 378.1 | 0.5 | 0.5049 | 13.9 | 13.1 | 923 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0282 | 0.0259 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 363.9
2008-12-02 | 13:59 | 0 | -0.001 | -10 | 14.2 | 363.8337 | 378 | 0.6 | 0.5961 | 13.9 | 13.1 | 923 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0282 | 0.0259 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1569 | 28.1 | 363.8
2008-12-02 | 13:59 | 0 | -0.0009 | -9.061 | 14.2 | 363.8 | 378 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0282 | 0.0259 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 363.8
2008-12-02 | 13:59 | 0 | -0.0001 | -1.0353 | 14.2 | 363.9 | 378.1 | 0.5 | 0.5131 | 13.9 | 13.1 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0282 | 0.0259 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 363.9
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.3 | 364.1 | 378.4 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.028234 | 0.0259 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1865 | 28.1 | 364.1
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.3 | 364.0671 | 378.4 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0283 | 0.026 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1079 | 28.1 | 364.1
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.3 | 364.2333 | 378.6 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0284 | 0.026 | 0.0192 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.0294 | 364.2
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.4 | 364.4331 | 378.8 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0284 | 0.0261 | 0.019233 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1491 | 28 | 364.4
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.4 | 364.5 | 378.9 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 60 | 0.0285 | 0.0261 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.0273 | 364.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.4 | 364.533 | 378.9 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0285 | 0.0261 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1947 | 28.0947 | 364.5
2008-12-02 | 13:59 | 0 | 0 | 0 | 14.5 | 364.6 | 379.1 | 0.6 | 0.6 | 13.9 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0285 | 0.0262 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1157 | 28.0157 | 364.6
2008-12-02 | 13:59 | 0 | 0.0006 | 5.905 | 14.5 | 364.5662 | 379.1 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.028534 | 0.0262 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3379 | 11.1 | 28 | 364.6
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.5 | 364.5 | 379 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0286 | 0.0262 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.0413 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.5 | 364.5335 | 379 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0286 | 0.0262 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.5 | 364.5 | 379 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0286 | 0.0262 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.5 | 364.5332 | 379.1 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0286 | 0.026233 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.6 | 364.5 | 379.1 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0286 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.155 | 28.1 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.6 | 364.5 | 379.1 | 0.6 | 0.5721 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0286 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 364.5
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.6 | 364.6331 | 379.2 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0287 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 364.6
2008-12-02 | 13:59 | 0 | 0.001 | 10 | 14.6 | 364.733 | 379.3 | 0.6 | 0.6 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0287 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 364.7
2008-12-02 | 13:59 | 0 | 0.0009 | 9.2735 | 14.6 | 364.834 | 379.4 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0287 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6601 | 11.1314 | 28.1 | 364.8
2008-12-02 | 13:59 | 0 | 0.0001 | 1.1102 | 14.6 | 364.8676 | 379.5 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0287 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6621 | 11.1 | 28.1 | 364.9
2008-12-02 | 14:00 | 0 | 0 | 0 | 14.6 | 364.9672 | 379.6 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0287 | 0.0263 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.664 | 11.1 | 28.1 | 365
2008-12-02 | 14:00 | 0 | 0 | 0 | 14.6 | 365.0665 | 379.7 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.028733 | 0.026333 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3347 | 11.1041 | 28.1 | 365.1
2008-12-02 | 14:00 | 0 | 0 | 0 | 14.7 | 365.0333 | 379.7 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0288 | 0.0264 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1824 | 28.1 | 365
2008-12-02 | 14:00 | 0 | 0.0002 | 2.2619 | 14.7 | 365.1 | 379.8 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0288 | 0.0264 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1391 | 28.0391 | 365.1
2008-12-02 | 14:00 | 0 | 0.0009 | 9.2635 | 14.7 | 365.1 | 379.8 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0288 | 0.0264 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1394 | 28.0394 | 365.1
2008-12-02 | 14:00 | 0 | 0.0001 | 1.0802 | 14.7 | 365.1333 | 379.8 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0288 | 0.0264 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6667 | 11.2 | 28.1 | 365.1
2008-12-02 | 14:00 | 0 | 0 | 0 | 14.7 | 365.2 | 379.9 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0288 | 0.0264 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1 | 365.2
2008-12-02 | 14:00 | -0.002 | 0 | 0 | 14.7 | 365.3331 | 380.1 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1256 | 28.1 | 365.3
2008-12-02 | 14:00 | -0.009 | 0 | 0 | 14.8 | 365.5 | 380.3 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.001 | 0.0002 | 2.2487 | 14.8 | 365.5 | 380.3 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.0685 | 365.5
2008-12-02 | 14:00 | 0 | 0.0009 | 8.5306 | 14.8 | 365.5 | 380.3 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.11 | 28 | 365.5
2008-12-02 | 14:00 | 0 | -0.0008 | -7.7959 | 14.8 | 365.4336 | 380.2 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.664 | 11.1884 | 28 | 365.4
2008-12-02 | 14:00 | 0 | -0.001 | -10 | 14.8 | 365.5 | 380.3 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1336 | 28 | 365.5
2008-12-02 | 14:00 | 0 | -0.001 | -10 | 14.9 | 365.5 | 380.4 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.0449 | 365.5
2008-12-02 | 14:00 | 0 | -0.001 | -10 | 14.9 | 365.4335 | 380.4 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6646 | 11.0526 | 28.1 | 365.4
2008-12-02 | 14:00 | 0 | -0.001 | -10 | 15 | 365.4667 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0289 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3333 | 10.9 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.007 | -0.001 | -10 | 15 | 365.5 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.029 | 0.0265 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.01 | -0.001 | -10 | 15 | 365.4 | 380.4 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.029 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0177 | 28.1 | 365.4
2008-12-02 | 14:00 | -0.01 | 0.0005 | 4.8415 | 15 | 365.533 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.029 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.67 | 11.1373 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.5671 | 380.6 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.029033 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3293 | 11.1841 | 28.1 | 365.6
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.5329 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6713 | 11.1056 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.5 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9664 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0272 | 28.1 | 365.5
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.6 | 380.6 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 196 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0509 | 28.0491 | 365.6
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.6 | 380.6 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28 | 365.6
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.6 | 380.6 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1081 | 28.0081 | 365.6
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15 | 365.5 | 380.5 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1866 | 28.0866 | 365.5
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.1 | 365.5 | 380.6 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0291 | 0.0266 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.07 | 28.035 | 365.5
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.1 | 365.6 | 380.7 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0291 | 0.0267 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.0434 | 365.6
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.1 | 365.7669 | 380.9 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0291 | 0.0267 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3306 | 11.0218 | 28.0782 | 365.8
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.1 | 365.7 | 380.8 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0291 | 0.0267 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.0003 | 365.7
2008-12-02 | 14:00 | -0.01 | 0.0003 | 2.7211 | 15.1 | 365.7 | 380.8 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.029133 | 0.0267 | 0.0193 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.0787 | 365.7
2008-12-02 | 14:00 | -0.01 | 0.0006 | 5.7617 | 15.1 | 365.7338 | 380.8 | 0.5 | 0.5 | 13.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0292 | 0.0267 | 0.019334 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 365.7
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.1 | 367.9158 | 383 | 0.5 | 0.5 | 13.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0292 | 0.0267 | 0.019501 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.336 | 11.1 | 28.1 | 367.9
2008-12-02 | 14:00 | -0.01 | 0.001 | 10 | 15.3 | 395.4485 | 410.7 | 0.5 | 0.5 | 14.53 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.029534 | 0.027034 | 0.020901 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 395.4
2008-12-02 | 14:00 | -0.01 | 0.0009 | 9.2665 | 15.6 | 413 | 428.6 | 0.6 | 0.5933 | 15.44 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0309 | 0.0284 | 0.0218 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 413
2008-12-02 | 14:00 | -0.01 | 0.0001 | 1.1165 | 15.8 | 413.9332 | 429.7 | 0.6 | 0.6 | 15.77 | 13.1 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0313 | 0.0288 | 0.0219 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 413.9
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16 | 414.0335 | 430 | 0.6 | 0.6 | 15.8 | 13.2 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.032134 | 0.029534 | 0.0219 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0506 | 28.1 | 414
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.1 | 414.4667 | 430.6 | 0.6 | 0.6 | 15.9 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0322 | 0.0296 | 0.0219 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0278 | 28.1 | 414.5
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.2 | 415.8669 | 432.1 | 0.6 | 0.6 | 15.9 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0323 | 0.0296 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1063 | 28.1 | 415.9
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.2 | 415.9 | 432.1 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0323 | 0.029633 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1847 | 28.1 | 415.9
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.2 | 415.9329 | 432.1 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0324 | 0.029733 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1369 | 28.1 | 415.9
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.3 | 416.1338 | 432.4 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.8 | 59.9 | 0.032434 | 0.0298 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1415 | 28.1 | 416.1
2008-12-02 | 14:00 | -0.006 | 0 | 0 | 16.4 | 416.2 | 432.6 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0325 | 0.029834 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1801 | 28.1 | 416.2
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.4 | 416.3 | 432.7 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.032533 | 0.0299 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.102 | 28.1 | 416.3
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.5 | 416.2 | 432.7 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0326 | 0.0299 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0233 | 28.1 | 416.2
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.6 | 416.3668 | 433 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0326 | 0.0299 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 416.4
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.6 | 416.4 | 433 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.032634 | 0.029934 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0675 | 28.0663 | 416.4
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.6 | 416.4333 | 433 | 0.7 | 0.6976 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9667 | 0.0327 | 0.03 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1879 | 28.0121 | 416.4
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.6 | 416.5669 | 433.2 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0327 | 0.03 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.11 | 28.09 | 416.6
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.6 | 416.5 | 433.1 | 0.6 | 0.6489 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0327 | 0.03 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2376 | 28.1 | 416.5
2008-12-02 | 14:00 | -0.001 | 0 | 0 | 16.7 | 416.633 | 433.3 | 0.6 | 0.6 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0328 | 0.0301 | 0.022 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1576 | 28.1 | 416.6
2008-12-02 | 14:00 | -0.009 | 0 | 0 | 16.7 | 416.7329 | 433.4 | 0.7 | 0.6971 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0329 | 0.030133 | 0.022033 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 416.7
2008-12-02 | 14:00 | -0.003 | 0 | 0 | 16.7 | 416.7 | 433.4 | 0.6 | 0.6025 | 15.9 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0329 | 0.0302 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0043 | 28.1 | 416.7
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.7 | 416.8336 | 433.6 | 0.7 | 0.6976 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0329 | 0.0302 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0826 | 28.1 | 416.8
2008-12-02 | 14:00 | 0 | 0.0007 | 7.3711 | 16.8 | 416.9 | 433.7 | 0.6 | 0.6008 | 15.9 | 13.2 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9665 | 0.033 | 0.0302 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 416.9
2008-12-02 | 14:00 | 0 | 0.001 | 10 | 16.8 | 417.1 | 433.9 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.033 | 0.0302 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 417.1
2008-12-02 | 14:00 | 0 | 0.0006 | 6.0806 | 16.9 | 417.0665 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.033 | 0.0302 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3354 | 11.1 | 28.1 | 417.1
2008-12-02 | 14:00 | 0 | 0 | 0 | 16.9 | 417 | 433.9 | 0.6 | 0.6002 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.033 | 0.030233 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417
2008-12-02 | 14:00 | -0.007 | 0 | 0 | 16.9 | 417 | 433.9 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.033 | 0.0303 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0254 | 28.1 | 417
2008-12-02 | 14:00 | -0.01 | 0 | 0 | 16.9 | 416.9669 | 433.9 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0331 | 0.0303 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0531 | 28.1 | 417
2008-12-02 | 14:01 | -0.006 | 0 | 0 | 16.9 | 417.1 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0331 | 0.0303 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.1
2008-12-02 | 14:01 | 0 | 0 | 0 | 16.9 | 416.9 | 433.8 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0331 | 0.0303 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 416.9
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 416.9338 | 433.9 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0331 | 0.0303 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 416.9
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 417 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8336 | 0.0331 | 0.030334 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 417.0335 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0331 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1455 | 28.1 | 417
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 417.0661 | 434.1 | 0.7 | 0.7 | 15.9 | 13.2 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.4 | 59.9 | 0.033134 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1761 | 28.1 | 417.1
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 417 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1047 | 28.1 | 417
2008-12-02 | 14:01 | 0 | 0 | 0 | 17 | 417 | 434 | 0.7 | 0.7 | 15.9 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2616 | 28.1 | 417
2008-12-02 | 14:01 | 0 | 0 | 0 | 17.1 | 417.1 | 434.2 | 0.7 | 0.7 | 15.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.241 | 28.1 | 417.1
2008-12-02 | 14:01 | 0 | 0 | 0 | 17.1 | 417.1331 | 434.2 | 0.7 | 0.7 | 15.9 | 13.2 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1625 | 28.1 | 417.1
2008-12-02 | 14:01 | -0.007 | 0 | 0 | 17.1 | 417.1 | 434.2 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1161 | 28.1 | 417.1
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.1 | 417.1659 | 434.3 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1945 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | -0.0007 | -7.2472 | 17.1 | 417.1676 | 434.3 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0332 | 0.0304 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1271 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.1 | 417.3 | 434.4 | 0.7 | 0.7 | 15.9 | 13.2 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0332 | 0.030434 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.2 | 417.4 | 434.6 | 0.6 | 0.603 | 15.9 | 13.2 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0333 | 0.0305 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 417.4
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.2 | 417.3 | 434.5 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0333 | 0.0305 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0918 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.2 | 417.4 | 434.6 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0333 | 0.0305 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0133 | 28.1 | 417.4
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.3 | 417.4 | 434.7 | 0.7 | 0.6979 | 15.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0333 | 0.0305 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0651 | 28.1 | 417.4
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.3 | 417.3667 | 434.7 | 0.6 | 0.6014 | 15.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.033333 | 0.030533 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3333 | 11.1435 | 28.1 | 417.4
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.3 | 417.3 | 434.6 | 0.7 | 0.7 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1563 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | -0.001 | -10 | 17.3 | 417.3 | 434.6 | 0.7 | 0.6886 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0008 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | -0.0008 | -7.8448 | 17.3 | 417.2 | 434.5 | 0.6 | 0.6 | 15.9 | 13.1 | 922 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1576 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.3 | 417.2 | 434.5 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.3 | 417.2338 | 434.5 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.129 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.3 | 417.3 | 434.6 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.3 | 417.3 | 434.6 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.3 | 417.3 | 434.6 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0334 | 0.0306 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0709 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.4 | 417.3332 | 434.7 | 0.6 | 0.6 | 15.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0335 | 0.030633 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.4 | 417.2 | 434.6 | 0.7 | 0.7 | 15.9 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0335 | 0.0307 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.4 | 417.2667 | 434.7 | 0.6 | 0.6014 | 15.9 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.033533 | 0.0307 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.3
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.4 | 417.2 | 434.6 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0336 | 0.0307 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1 | 28.1 | 417.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.4 | 417.4 | 434.8 | 0.6 | 0.6 | 15.9 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0336 | 0.0307 | 0.0221 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0369 | 28.1 | 417.4
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.5 | 434.2318 | 451.7 | 0.6 | 0.6 | 16.07 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.033633 | 0.030733 | 0.022697 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1246 | 28.1 | 434.2
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 17.9 | 454.9663 | 472.8 | 0.6 | 0.602 | 17.06 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.034231 | 0.031298 | 0.024032 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.6707 | 11.2602 | 28.1 | 455
2008-12-02 | 14:01 | -0.01 | 0 | 0 | 18.2 | 465.2042 | 483.4 | 0.7 | 0.7273 | 17.72 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0351 | 0.0321 | 0.0246 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1036 | 28.1 | 465.2
2008-12-02 | 14:01 | -0.004 | 0 | 0 | 18.2 | 465.8336 | 484 | 0.8 | 0.8 | 17.8 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0351 | 0.032134 | 0.0246 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 465.8
2008-12-02 | 14:01 | -0.004 | 0 | 0 | 18.2 | 466.7335 | 485 | 0.8 | 0.8 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9665 | 0.0353 | 0.0324 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0448 | 28.1 | 466.7
2008-12-02 | 14:01 | -0.008 | 0 | 0 | 18.5 | 466.8 | 485.3 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0362 | 0.0332 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0337 | 28.1 | 466.8
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.6 | 466.9332 | 485.5 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.036233 | 0.0332 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1242 | 28.1 | 466.9
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.6 | 466.9667 | 485.6 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195.1 | 59.9 | 0.0363 | 0.0333 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2802 | 28.1 | 467
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.7 | 467 | 485.7 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195 | 59.9 | 0.0364 | 0.0334 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3688 | 28.1 | 467
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.7 | 467 | 485.7 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.9 | 0.0364 | 0.0334 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 467
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.7 | 467.033 | 485.8 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.9 | 0.036433 | 0.0334 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 467
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.8 | 467.1 | 485.9 | 0.8 | 0.8 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195 | 59.9 | 0.0365 | 0.0335 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3958 | 28.1 | 467.1
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.8 | 467.2 | 486 | 0.8 | 0.8 | 17.9 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.9 | 0.0366 | 0.0335 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3173 | 28.1 | 467.2
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.8 | 467.3 | 486.1 | 0.8 | 0.8 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.9664 | 0.0366 | 0.033534 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4222 | 28.1 | 467.3
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.9 | 467.5 | 486.4 | 0.8 | 0.8 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.9 | 0.0367 | 0.0336 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5 | 28.1 | 467.5
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.9 | 467.4667 | 486.4 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195 | 59.9 | 0.0367 | 0.0336 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4821 | 28.1 | 467.5
2008-12-02 | 14:01 | 0 | 0 | 0 | 18.9 | 467.4664 | 486.4 | 0.8 | 0.8 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0367 | 0.0336 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4036 | 28.1 | 467.5
2008-12-02 | 14:01 | 0 | 0 | 0 | 19 | 467.5 | 486.5 | 0.7 | 0.7009 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0367 | 0.033634 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 467.5
2008-12-02 | 14:01 | 0 | 0 | 0 | 19 | 467.6 | 486.6 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0367 | 0.0337 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5062 | 28.1 | 467.6
2008-12-02 | 14:01 | 0 | 0 | 0 | 19.1 | 467.7331 | 486.8 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0368 | 0.0337 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.5366 | 28.1317 | 467.7
2008-12-02 | 14:01 | 0 | 0.0001 | 0.68438 | 19.1 | 467.8 | 486.9 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0368 | 0.0337 | 0.0247 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.2 | 467.8
2008-12-02 | 14:01 | 0 | 0.0009 | 8.7359 | 19.1 | 467.866 | 487 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.036833 | 0.033733 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.2 | 467.9
2008-12-02 | 14:01 | 0 | 0.001 | 10 | 19.1 | 467.9 | 487 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.966 | 0.0369 | 0.0338 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.2 | 467.9
2008-12-02 | 14:01 | 0 | 0.001 | 10 | 19.2 | 467.9 | 487.1 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9336 | 0.0369 | 0.0338 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.155 | 467.9
2008-12-02 | 14:01 | 0 | 0.001 | 10 | 19.3 | 468.0335 | 487.3 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037 | 0.0339 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 468
2008-12-02 | 14:01 | 0 | 0.0008 | 7.8856 | 19.3 | 468.1333 | 487.4 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037 | 0.0339 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3978 | 28.1 | 468.1
2008-12-02 | 14:01 | 0 | 0 | 0 | 19.3 | 468.2 | 487.5 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037 | 0.0339 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3193 | 28.1 | 468.2
2008-12-02 | 14:01 | 0 | 0 | 0 | 19.3 | 468.2 | 487.5 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037 | 0.0339 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4181 | 28.1 | 468.2
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.4 | 468.1667 | 487.6 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0371 | 0.033933 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4626 | 28.1 | 468.2
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.1669 | 487.7 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037133 | 0.034 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 468.2
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.2 | 487.7 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.034 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.4 | 28.1 | 468.2
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.3 | 487.8 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.034 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3272 | 28.1 | 468.3
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.3 | 487.8 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.034 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3 | 28.1 | 468.3
2008-12-02 | 14:02 | 0.006 | 0 | 0 | 19.5 | 468.3 | 487.8 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.034034 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3293 | 28.1 | 468.3
2008-12-02 | 14:02 | 0.01 | 0 | 0 | 19.5 | 468.3 | 487.8 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9336 | 0.0372 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.392 | 28.108 | 468.3
2008-12-02 | 14:02 | 0.01 | 0 | 0 | 19.5 | 468.4 | 487.9 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3136 | 28.1864 | 468.4
2008-12-02 | 14:02 | 0.01 | 0 | 0 | 19.5 | 468.4 | 487.9 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3 | 28.1351 | 468.4
2008-12-02 | 14:02 | 0.01 | 0 | 0 | 19.5 | 468.5 | 488 | 0.7 | 0.7 | 17.9 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9663 | 0.0372 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.3431 | 28.1431 | 468.5
2008-12-02 | 14:02 | 0.01 | 0 | 0 | 19.5 | 468.5 | 488 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0372 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.3788 | 28.2 | 468.5
2008-12-02 | 14:02 | 0.005 | 0 | 0 | 19.5 | 468.5 | 488 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.037233 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2998 | 28.2 | 468.5
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.5 | 488 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2213 | 28.2 | 468.5
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.5 | 468.633 | 488.2 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.000133 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1429 | 468.6
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.634 | 488.2 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1354 | 468.6
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.8 | 488.4 | 0.7 | 0.7 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1863 | 28.2 | 468.8
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.8336 | 488.4 | 0.7 | 0.7 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.000166 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1077 | 28.2 | 468.8
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.9 | 488.5 | 0.6 | 0.603 | 17.9 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0341 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1707 | 28.1293 | 468.9
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.8667 | 488.5 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2492 | 28.1 | 468.9
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.6664 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0373 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2449 | 28.1275 | 468.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.7329 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0374 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3354 | 11.1056 | 28.1944 | 468.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.5667 | 488.2 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0374 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1844 | 28.1156 | 468.6
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0374 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1371 | 28.1629 | 468.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.7 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0374 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1413 | 28.1587 | 468.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.733 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0374 | 0.0342 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2196 | 28.1196 | 468.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 19.6 | 468.6 | 488.2 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0374 | 0.0343 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2974 | 28.1974 | 468.6
2008-12-02 | 14:02 | 0 | -0.0004 | -3.8235 | 19.6 | 468.6676 | 488.3 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0375 | 0.0343 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3 | 28.1236 | 468.7
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 19.6 | 468.8 | 488.4 | 0.6 | 0.6 | 17.9 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9336 | 0.0375 | 0.0343 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.19 | 28.1 | 468.8
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 19.7 | 468.8665 | 488.6 | 0.6 | 0.6 | 17.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0376 | 0.0344 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1334 | 28.1 | 468.9
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 19.7 | 468.8 | 488.5 | 0.6 | 0.6 | 17.9 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0376 | 0.034433 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1881 | 28.1119 | 468.8
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 19.7 | 468.9 | 488.6 | 0.6 | 0.6 | 17.9 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0377 | 0.0345 | 0.0248 | 0.0001 |  | n/a |  | 35 | 61 | 1054.1904 | 1015 | 11.1099 | 28.1901 | 468.9
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 19.7 | 468.8671 | 488.6 | 0.6 | 0.6 | 17.9 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.037666 | 0.0345 | 0.0248 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0313 | 28.1313 | 468.9
2008-12-02 | 14:02 | -0.001 | -0.0006 | -6.369 | 19.7 | 474.4333 | 494.1 | 0.6 | 0.6 | 17.92 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0377 | 0.0345 | 0.0251 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0471 | 28.1471 | 474.4
2008-12-02 | 14:02 | -0.009 | 0 | 0 | 20 | 503.0943 | 523.1 | 0.6 | 0.6 | 18.81 | 13.2 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.038599 | 0.035299 | 0.026499 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1 | 28.1744 | 503.1
2008-12-02 | 14:02 | -0.01 | -0.0001 | -0.60338 | 20.3 | 517.4952 | 537.8 | 0.7 | 0.6976 | 19.62 | 13.2 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.039633 | 0.036333 | 0.027433 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015.3306 | 11.104 | 28.104 | 517.5
2008-12-02 | 14:02 | -0.01 | -0.0009 | -8.6484 | 20.3 | 518.7 | 539 | 0.7 | 0.7 | 19.86 | 13.1 | 924 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.039733 | 0.036433 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1016 | 11.182 | 28.182 | 518.7
2008-12-02 | 14:02 | -0.006 | -0.0003 | -2.7965 | 20.4 | 518.734 | 539.2 | 0.8 | 0.8 | 19.9 | 13.2 | 925 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.039834 | 0.036534 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1016 | 11.0787 | 28.2 | 518.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.5 | 518.9338 | 539.4 | 0.8 | 0.8 | 19.9 | 13.2 | 925 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0402 | 0.0369 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0784 | 28.1608 | 518.9
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.6 | 519 | 539.6 | 0.7 | 0.7027 | 19.9 | 13.2 | 924 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.040334 | 0.037 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1647 | 28.1177 | 519
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.7 | 519.0667 | 539.8 | 0.7 | 0.7 | 19.9 | 13.2 | 924 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.040433 | 0.0371 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0077 | 28.1962 | 519.1
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.7 | 519.3 | 540 | 0.7 | 0.7 | 19.9 | 13.2 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0405 | 0.0372 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 519.3
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.7 | 519.1 | 539.8 | 0.7 | 0.7 | 20 | 13.2 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0406 | 0.037234 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0526 | 28.1474 | 519.1
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.8 | 519.3 | 540.1 | 0.7 | 0.7 | 20 | 13.2 | 923 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0407 | 0.037333 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1 | 28.1314 | 519.3
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.8 | 519.4 | 540.2 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0407 | 0.0374 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1 | 28.2 | 519.4
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.8 | 519.5 | 540.3 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0408 | 0.0374 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1 | 28.2 | 519.5
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.9 | 519.5 | 540.4 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0409 | 0.0375 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1665 | 28.1335 | 519.5
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.9 | 519.6 | 540.5 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0409 | 0.0375 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1551 | 28.1 | 519.6
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.9 | 519.5673 | 540.5 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0409 | 0.0375 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1235 | 28.1 | 519.6
2008-12-02 | 14:02 | 0 | 0 | 0 | 20.9 | 519.5 | 540.4 | 0.7 | 0.7 | 20 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0409 | 0.0375 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1962 | 28.1 | 519.5
2008-12-02 | 14:02 | 0 | -0.0006 | -5.5379 | 20.9 | 519.7 | 540.6 | 0.7 | 0.7 | 20 | 13.2 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.040933 | 0.0376 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0392 | 28.1 | 519.7
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 20.9 | 519.6333 | 540.6 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.041 | 0.037633 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0585 | 28.1585 | 519.6
2008-12-02 | 14:02 | 0 | -0.001 | -10 | 21 | 519.6663 | 540.7 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0411 | 0.0377 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0631 | 28.2 | 519.7
2008-12-02 | 14:02 | 0 | -0.001 | -9.7513 | 21.1 | 519.7335 | 540.8 | 0.7 | 0.7 | 20 | 13.2 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0412 | 0.037734 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0468 | 28.1844 | 519.7
2008-12-02 | 14:02 | 0 | -0.0001 | -1.175 | 21.1 | 519.7667 | 540.9 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0412 | 0.0378 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.2821 | 28.106 | 519.8
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.2 | 519.7 | 540.9 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9331 | 0.0412 | 0.0378 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1551 | 28.1725 | 519.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.2 | 519.7 | 540.9 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.041233 | 0.0378 | 0.0275 | 0.000167 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0495 | 28.2 | 519.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.2 | 519.7 | 540.9 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.966 | 0.041334 | 0.037934 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9712 | 28.2 | 519.7
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.3 | 519.9 | 541.2 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0414 | 0.038 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9233 | 28.2 | 519.9
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.4 | 520.0328 | 541.4 | 0.7 | 0.7 | 20 | 13.2 | 921 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0415 | 0.038 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1584 | 28.2 | 520
2008-12-02 | 14:02 | 0 | 0 | 0 | 21.4 | 519.9669 | 541.4 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0415 | 0.038 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0062 | 28.2 | 520
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.4 | 520 | 541.4 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0415 | 0.0381 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9429 | 28.2 | 520
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.4 | 520.1 | 541.5 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9332 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0212 | 28.2 | 520.1
2008-12-02 | 14:03 | 0 | -0.0004 | -3.5986 | 21.5 | 520.1 | 541.6 | 0.7 | 0.7 | 20 | 13.2 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0999 | 28.2 | 520.1
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.6 | 520 | 541.6 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.1784 | 28.2 | 520
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.6 | 520.1669 | 541.8 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0865 | 28.2 | 520.2
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.6 | 520.1 | 541.7 | 0.7 | 0.7 | 20 | 13.2 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0349 | 28.1651 | 520.1
2008-12-02 | 14:03 | 0 | -0.0003 | -2.8756 | 21.6 | 520.1 | 541.7 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0869 | 28.1131 | 520.1
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.6 | 520.2329 | 541.8 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.008 | 28.192 | 520.2
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.6 | 520.3 | 541.9 | 0.7 | 0.7 | 20 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0416 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.1296 | 520.3
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.6 | 520.3 | 541.9 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0417 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.1489 | 520.3
2008-12-02 | 14:03 | -0.007 | 0 | 0 | 21.7 | 520.3335 | 542 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0417 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.3
2008-12-02 | 14:03 | -0.01 | 0 | 0 | 21.7 | 520.4337 | 542.1 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0417 | 0.0382 | 0.0275 | 0.0001 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.4
2008-12-02 | 14:03 | -0.014 | -0.0007 | -6.8526 | 21.7 | 520.5 | 542.2 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9665 | 0.041734 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.5
2008-12-02 | 14:03 | -0.02 | -0.001 | -10 | 21.7 | 520.4 | 542.1 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.4
2008-12-02 | 14:03 | -0.019 | -0.001 | -10 | 21.8 | 520.4 | 542.2 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.4
2008-12-02 | 14:03 | -0.003 | -0.0008 | -7.9547 | 21.8 | 520.4 | 542.2 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9806 | 28.2 | 520.4
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.4 | 542.2 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9025 | 28.2 | 520.4
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.4329 | 542.2 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.0275 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9762 | 28.2 | 520.4
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.5 | 542.3 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0418 | 0.0383 | 0.027534 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.5
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.5672 | 542.4 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0418 | 0.0383 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.6
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.5668 | 542.4 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0419 | 0.038333 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9885 | 28.2 | 520.6
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.7333 | 542.5 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9667 | 0.0419 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9105 | 28.2 | 520.7
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.8 | 520.6664 | 542.5 | 0.7 | 0.7 | 20 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0419 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 10.9683 | 28.2 | 520.7
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.8 | 542.7 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0419 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0469 | 28.2 | 520.8
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.8 | 542.7 | 0.7 | 0.7 | 20 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0419 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0747 | 28.2 | 520.8
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.7 | 542.6 | 0.7 | 0.7 | 20 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.041933 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.7
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.7331 | 542.6 | 0.7 | 0.7 | 20 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.042 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.7
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.8 | 542.7 | 0.7 | 0.7 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.042 | 0.0384 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.0605 | 28.2 | 520.8
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.9 | 542.8 | 0.6 | 0.6025 | 20 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9671 | 0.042 | 0.038433 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11.061 | 28.2 | 520.9
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.8662 | 542.8 | 0.6 | 0.6 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.042 | 0.0385 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1055 | 1015 | 11 | 28.2 | 520.9
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.8336 | 542.7 | 0.7 | 0.6985 | 20 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.042 | 0.0385 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.2 | 520.8
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.9 | 520.9669 | 542.9 | 0.6 | 0.6 | 20 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.042 | 0.0385 | 0.0276 | 0.0002 |  | n/a |  | 35 | 61 | 1053.4204 | 1014.6653 | 11 | 28.2 | 521
2008-12-02 | 14:03 | 0 | -0.0004 | -3.8033 | 21.9 | 517.5 | 539.4 | 0.7 | 0.6891 | 20 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9333 | 0.042 | 0.0385 | 0.0274 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9475 | 28.2 | 517.5
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.9 | 493.0946 | 515 | 0.6 | 0.6 | 19.46 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.041967 | 0.0384 | 0.026135 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9311 | 28.2 | 493.1
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.8 | 474.0335 | 495.8 | 0.6 | 0.6 | 18.49 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9665 | 0.0416 | 0.038066 | 0.0251 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 474
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.7 | 473.1667 | 494.8 | 0.5 | 0.5 | 18.14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0412 | 0.0377 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 473.2
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21.5 | 472.7018 | 494.2 | 0.5 | 0.5 | 18.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.040567 | 0.037067 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 472.7
2008-12-02 | 14:03 | 0 | -0.0005 | -4.6988 | 21.4 | 472 | 493.4 | 0.5 | 0.5 | 18.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.9 | 0.0402 | 0.0368 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 472
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.2 | 471.5672 | 492.8 | 0.5 | 0.5 | 18.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195.2 | 59.9 | 0.039867 | 0.0364 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 471.6
2008-12-02 | 14:03 | 0 | 0 | 0 | 21.2 | 471.5 | 492.7 | 0.5 | 0.5 | 18 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.039766 | 0.036366 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 471.5
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.4 | 492.4 | 0.5 | 0.5 | 18 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0396 | 0.0362 | 0.025 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 471.4
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.4669 | 492.5 | 0.6 | 0.5951 | 18 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0396 | 0.0362 | 0.024967 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9415 | 28.1415 | 471.5
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.5 | 492.5 | 0.5 | 0.5018 | 18 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.039567 | 0.036167 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9369 | 28.1369 | 471.5
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.5332 | 492.5 | 0.6 | 0.5914 | 18 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0154 | 28.2 | 471.5
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.6 | 492.6 | 0.5 | 0.5 | 18 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0938 | 28.2 | 471.6
2008-12-02 | 14:03 | 0 | 0 | 0 | 21 | 471.3667 | 492.4 | 0.6 | 0.6 | 18 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1278 | 471.4
2008-12-02 | 14:03 | 0 | 0 | -0.42465 | 21 | 471.3 | 492.3 | 0.5 | 0.5 | 18 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0496 | 28.1504 | 471.3
2008-12-02 | 14:03 | 0 | -0.0009 | -8.751 | 21 | 471.3 | 492.3 | 0.5 | 0.5021 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9714 | 28.2 | 471.3
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21 | 471.3 | 492.3 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.907 | 28.2 | 471.3
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 21 | 471.2671 | 492.3 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9859 | 28.2 | 471.3
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 471.2 | 492.1 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9662 | 0.039466 | 0.0361 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 471.2
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 471.3 | 492.2 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0394 | 0.036066 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 471.3
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 471.1 | 492 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0394 | 0.036 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0416 | 28.2 | 471.1
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 471.1 | 492 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0394 | 0.036 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1994 | 28.2 | 471.1
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 471.0665 | 492 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0394 | 0.036 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1219 | 28.2 | 471.1
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 470.9667 | 491.9 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0394 | 0.036 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0435 | 28.1435 | 471
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 470.8337 | 491.7 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9669 | 0.039367 | 0.036 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1349 | 470.8
2008-12-02 | 14:03 | -0.008 | -0.001 | -10 | 20.9 | 470.6 | 491.5 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0132 | 28.2 | 470.6
2008-12-02 | 14:03 | -0.003 | -0.001 | -10 | 20.9 | 470.567 | 491.5 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.967 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0918 | 28.2 | 470.6
2008-12-02 | 14:03 | 0 | -0.001 | -10 | 20.9 | 470.5659 | 491.5 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8894 | 28.2 | 470.6
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.9 | 470.5662 | 491.5 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8487 | 28.1513 | 470.6
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.9 | 470.4664 | 491.4 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.927 | 28.127 | 470.5
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.9 | 470.4 | 491.3 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.995 | 28.2 | 470.4
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.3333 | 491.1 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.039267 | 0.0359 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.6667 | 10.9161 | 28.2 | 470.3
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.4668 | 491.3 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9668 | 0.0392 | 0.0358 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.2 | 470.5
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.4 | 491.2 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0358 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8592 | 28.2 | 470.4
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.4 | 491.2 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.9 | 0.0391 | 0.035767 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8384 | 28.2 | 470.4
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.5 | 491.3 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0391 | 0.0357 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9953 | 28.2 | 470.5
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.6 | 491.4 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.039 | 0.0357 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9239 | 28.2 | 470.6
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.8 | 470.567 | 491.3 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.039 | 0.035667 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.2 | 470.6
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.7 | 470.3662 | 491.1 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.038966 | 0.0356 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.2 | 470.4
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.7 | 470.3 | 491 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0389 | 0.0356 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.2 | 470.3
2008-12-02 | 14:04 | 0 | -0.001 | -10 | 20.7 | 470.2665 | 491 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0389 | 0.0356 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.2 | 470.3
2008-12-02 | 14:04 | 0 | -0.0005 | -4.9374 | 20.7 | 470.2 | 490.9 | 0.5 | 0.5 | 18 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0389 | 0.0356 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9681 | 28.2 | 470.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.1 | 490.8 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0389 | 0.035567 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.9067 | 28.1533 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.1335 | 490.8 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0389 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.6646 | 10.8502 | 28.1251 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.2 | 490.9 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0389 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 470.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.2663 | 491 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0389 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 470.3
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.2 | 490.9 | 0.5 | 0.5 | 18 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0389 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 470.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.7 | 470.2 | 490.9 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.933 | 0.038867 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.2 | 470.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470.2 | 490.8 | 0.5 | 0.5 | 18 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0355 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9828 | 28.2 | 470.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470 | 490.6 | 0.5 | 0.5 | 18 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.9043 | 28.2 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470 | 490.6 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.826 | 28.126 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470 | 490.6 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.9049 | 28.1 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470 | 490.6 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.969 | 28.131 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470 | 490.6 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.8906 | 28.2 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.6 | 470.1 | 490.7 | 0.5 | 0.5 | 18 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8122 | 28.2 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.5 | 470.0333 | 490.5 | 0.5 | 0.5 | 18 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.6667 | 10.866 | 28.2 | 470
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.5 | 470.1 | 490.6 | 0.5 | 0.5 | 18 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8556 | 28.1556 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.5 | 470.1 | 490.6 | 0.6 | 0.599 | 18 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8231 | 28.1231 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.5 | 470.1 | 490.6 | 0.6 | 0.6 | 18 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.0249 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8985 | 28.1985 | 470.1
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.5 | 460.5715 | 481.1 | 0.6 | 0.6 | 18 | 13.2 | 913 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0354 | 0.024335 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.82 | 28.12 | 460.6
2008-12-02 | 14:04 | 0 | 0 | 0 | 20.2 | 437.1832 | 457.3 | 0.6 | 0.6 | 17.06 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.037927 | 0.034661 | 0.023162 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.9163 | 28.1582 | 437.2
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.7 | 423.6281 | 443.3 | 0.5 | 0.5 | 16.29 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.036866 | 0.0337 | 0.022466 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1634 | 423.6
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.5 | 421.4855 | 441 | 0.5 | 0.5 | 16.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.036398 | 0.033263 | 0.022366 | 0.000166 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 421.5
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.4 | 420.4 | 439.8 | 0.5 | 0.5 | 16.04 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0361 | 0.033 | 0.0223 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 420.4
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.4 | 419.4332 | 438.8 | 0.5 | 0.4998 | 16 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0361 | 0.0329 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.8558 | 28.1721 | 419.4
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.4 | 419.6 | 439 | 0.5 | 0.4943 | 16 | 13.2 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.036 | 0.0329 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9009 | 28.1496 | 419.6
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.3 | 419.5669 | 438.9 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.036 | 0.032867 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1 | 419.6
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.3 | 419.5 | 438.8 | 0.5 | 0.5 | 16 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0359 | 0.0328 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1074 | 419.5
2008-12-02 | 14:04 | 0 | 0 | 0 | 19.3 | 419.5 | 438.8 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0359 | 0.0328 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1858 | 419.5
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19.3 | 419.5 | 438.8 | 0.5 | 0.5 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0359 | 0.0327 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0642 | 28.1358 | 419.5
2008-12-02 | 14:04 | -0.02 | 0 | 0 | 19.2 | 419.4662 | 438.7 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.5 | 59.9 | 0.0359 | 0.0327 | 0.0222 | 0.0001 |  | n/a |  | 35 | 61 | 1054 | 1014.3379 | 11.0575 | 28.1 | 419.5
2008-12-02 | 14:04 | -0.018 | 0 | 0 | 19.2 | 419.4 | 438.6 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0358 | 0.0327 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 419.4
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19.1 | 419.3665 | 438.5 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0358 | 0.0327 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 419.4
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19.1 | 419.3 | 438.4 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0358 | 0.0327 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 419.3
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19.1 | 419.4 | 438.5 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.035767 | 0.032667 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1 | 419.4
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19.1 | 419.4 | 438.5 | 0.5 | 0.5 | 16 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0357 | 0.0326 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1 | 419.4
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19 | 419.3 | 438.3 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0356 | 0.0325 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1128 | 419.3
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19 | 419.3 | 438.3 | 0.5 | 0.5 | 16 | 13.2 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0355 | 0.032467 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1916 | 419.3
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 19 | 419.4 | 438.4 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0355 | 0.0324 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.2 | 419.4
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 18.9 | 419.267 | 438.2 | 0.5 | 0.5 | 16 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0354 | 0.032367 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.2 | 419.3
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 18.9 | 419.2 | 438.1 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0354 | 0.0323 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1732 | 419.2
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 18.9 | 419.2 | 438.1 | 0.5 | 0.5 | 16 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0353 | 0.0323 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0051 | 28.1051 | 419.2
2008-12-02 | 14:04 | -0.01 | 0 | 0 | 18.9 | 419.1 | 438 | 0.4 | 0.4014 | 16 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0353 | 0.0322 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0838 | 28.1838 | 419.1
2008-12-02 | 14:04 | -0.01 | -0.0005 | -4.975 | 18.9 | 419.1667 | 438 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0353 | 0.0322 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0378 | 28.2 | 419.2
2008-12-02 | 14:04 | -0.01 | -0.001 | -10 | 18.8 | 419.1 | 437.9 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0352 | 0.0322 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.2 | 419.1
2008-12-02 | 14:04 | -0.01 | -0.001 | -10 | 18.8 | 419.1 | 437.9 | 0.5 | 0.4996 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0352 | 0.0322 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1811 | 419.1
2008-12-02 | 14:04 | -0.01 | -0.001 | -10 | 18.8 | 418.9 | 437.7 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 196 | 59.9 | 0.0352 | 0.0321 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1031 | 418.9
2008-12-02 | 14:04 | -0.01 | -0.001 | -10 | 18.8 | 419.1 | 437.9 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0352 | 0.0321 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 419.1
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 18.8 | 419.0339 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0352 | 0.0321 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 419
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 18.8 | 418.9329 | 437.7 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0351 | 0.0321 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1327 | 418.9
2008-12-02 | 14:05 | -0.01 | -0.0006 | -6.4895 | 18.8 | 419 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9671 | 0.0351 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1889 | 419
2008-12-02 | 14:05 | -0.002 | 0 | 0 | 18.7 | 419.1662 | 437.9 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0351 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1104 | 419.2
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9336 | 0.035066 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1681 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419 | 437.7 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1535 | 419
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0499 | 28.1249 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1933 | 28.1966 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1335 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0369 | 28.1185 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1333 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1656 | 437.9 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1 | 419.2
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1 | 437.8 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.983 | 28.117 | 419.1
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.7 | 419.1671 | 437.9 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9045 | 28.1955 | 419.2
2008-12-02 | 14:05 | 0 | -0.0005 | -4.9376 | 18.6 | 419.1 | 437.7 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9737 | 28.1263 | 419.1
2008-12-02 | 14:05 | 0 | -0.0007 | -6.7333 | 18.6 | 419 | 437.6 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11 | 28.1522 | 419
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.6 | 419 | 437.6 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9334 | 0.035 | 0.032 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0307 | 28.1693 | 419
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.6 | 419 | 437.6 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9333 | 0.0349 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 419
2008-12-02 | 14:05 | -0.005 | 0 | 0 | 18.5 | 419 | 437.5 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0349 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1 | 419
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.5 | 418.8669 | 437.4 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0349 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.9681 | 28.1659 | 418.9
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.5 | 418.8 | 437.3 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0349 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.9444 | 28.2 | 418.8
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.5 | 418.7 | 437.2 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.034867 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1771 | 418.7
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.4 | 418.7669 | 437.2 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0319 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 418.8
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.4 | 418.7 | 437.1 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.031867 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 418.7
2008-12-02 | 14:05 | -0.005 | 0 | 0 | 18.4 | 418.7 | 437.1 | 0.4 | 0.4 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 418.7
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.4 | 418.7662 | 437.2 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1 | 418.8
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.4 | 418.7336 | 437.1 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1151 | 418.7
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.4 | 418.8 | 437.2 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11 | 28.1935 | 418.8
2008-12-02 | 14:05 | 0 | 0 | 0 | 18.4 | 418.9 | 437.3 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0719 | 28.1281 | 418.9
2008-12-02 | 14:05 | -0.007 | 0 | 0 | 18.4 | 418.9 | 437.3 | 0.4 | 0.4017 | 16 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 10.9995 | 28.1502 | 418.9
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 18.4 | 418.7 | 437.1 | 0.5 | 0.494 | 16 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0348 | 0.0318 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 10.9286 | 28.2 | 418.7
2008-12-02 | 14:05 | -0.01 | -0.0007 | -6.6559 | 18.4 | 418.8667 | 437.3 | 0.5 | 0.5 | 16 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0347 | 0.031767 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0071 | 28.1929 | 418.9
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 18.4 | 418.8 | 437.2 | 0.4 | 0.4 | 16 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0347 | 0.0317 | 0.0222 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0856 | 28.1144 | 418.8
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 18.4 | 403.985 | 422.4 | 0.4 | 0.4 | 15.9 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.034562 | 0.031562 | 0.021422 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.6561 | 11.036 | 28.1 | 404
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.9 | 376.771 | 394.7 | 0.4 | 0.4 | 14.81 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.033074 | 0.030174 | 0.019969 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0424 | 28.1424 | 376.8
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.5 | 369.3329 | 386.8 | 0.4 | 0.4 | 14.2 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.031667 | 0.0288 | 0.0196 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0792 | 28.1792 | 369.3
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.4 | 368.4 | 385.8 | 0.3 | 0.3 | 14.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0316 | 0.0287 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0006 | 28.1006 | 368.4
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.4 | 368.3664 | 385.7 | 0.3 | 0.3 | 14 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0315 | 0.028666 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 368.4
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.3 | 368.3341 | 385.6 | 0.3 | 0.3 | 14 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0314 | 0.0286 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1562 | 368.3
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 17.1 | 368.3 | 385.4 | 0.4 | 0.3991 | 14 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0313 | 0.0285 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0345 | 28.1655 | 368.3
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.7 | 368.2665 | 385 | 0.3 | 0.3 | 14 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0312 | 0.0284 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0872 | 28.1 | 368.3
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.7 | 368.0667 | 384.8 | 0.3 | 0.3 | 14 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9333 | 0.0311 | 0.0284 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0086 | 28.1 | 368.1
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.7 | 368 | 384.7 | 0.4 | 0.4 | 14 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0311 | 0.0284 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 368
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.7 | 368 | 384.7 | 0.4 | 0.4 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0311 | 0.0284 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 368
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.6 | 367.9 | 384.5 | 0.4 | 0.4 | 14 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.031067 | 0.028367 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.9
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.6 | 367.8671 | 384.5 | 0.4 | 0.4 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.031 | 0.028267 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.9
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.6 | 367.8 | 384.4 | 0.4 | 0.4 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0309 | 0.0282 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.8
2008-12-02 | 14:05 | -0.01 | -0.001 | -10 | 16.6 | 367.7 | 384.3 | 0.4 | 0.4 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0309 | 0.0282 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.062 | 28.1 | 367.7
2008-12-02 | 14:05 | -0.01 | -0.0002 | -1.8932 | 16.5 | 367.6 | 384.1 | 0.4 | 0.3709 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0309 | 0.0282 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0191 | 28.1 | 367.6
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 16.5 | 367.6333 | 384.1 | 0.3 | 0.3 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0308 | 0.0282 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9189 | 28.1 | 367.6
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 16.5 | 367.6337 | 384.1 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0308 | 0.0281 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9972 | 28.1 | 367.6
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 16.4 | 367.7 | 384.1 | 0.3 | 0.3 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0308 | 0.0281 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1757 | 367.7
2008-12-02 | 14:05 | -0.01 | 0 | 0 | 16.4 | 367.6 | 384 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.9 | 0.030767 | 0.0281 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9459 | 28.1459 | 367.6
2008-12-02 | 14:05 | -0.003 | 0 | 0 | 16.4 | 367.6 | 384 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0307 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9 | 28.1325 | 367.6
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.467 | 383.8 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9219 | 28.189 | 367.5
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.4 | 383.7 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0787 | 28.1107 | 367.4
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.4 | 383.7 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0321 | 28.1 | 367.4
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.4336 | 383.7 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9536 | 28.1 | 367.4
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.4 | 383.7 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.028 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9248 | 28.1 | 367.4
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.5 | 383.8 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.5
2008-12-02 | 14:05 | 0 | 0 | 0 | 16.3 | 367.5 | 383.8 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.5
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.3 | 367.5 | 383.8 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 367.5
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.3 | 367.3667 | 383.7 | 0.2 | 0.2003 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0306 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9616 | 28.1 | 367.4
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.2 | 367.3 | 383.5 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.030567 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9169 | 28.1 | 367.3
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.2 | 367.2331 | 383.4 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9953 | 28.1 | 367.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.2 | 367.167 | 383.4 | 0.3 | 0.3 | 14 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.967 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0736 | 28.1736 | 367.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.2 | 367 | 383.2 | 0.3 | 0.2515 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.0479 | 28.1479 | 367
2008-12-02 | 14:06 | 0 | 0.0007 | 6.543 | 16.2 | 367.0338 | 383.2 | 0.3 | 0.3 | 14 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.6621 | 10.9694 | 28.1 | 367
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 367.0664 | 383.2 | 0.3 | 0.3 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.909 | 28.1 | 367.1
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 367.1 | 383.2 | 0.2 | 0.2 | 14 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9875 | 28.1 | 367.1
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 367.0333 | 383.1 | 0.3 | 0.3 | 14 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9342 | 28.1658 | 367
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 367.1 | 383.2 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.1557 | 367.1
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 367.0667 | 383.2 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0305 | 0.0279 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9227 | 28.1 | 367.1
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 196 | 59.9 | 0.0305 | 0.027867 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9988 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0.0005 | 5.16 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9204 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.842 | 28.158 | 366.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8728 | 28.1636 | 366.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9851 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9067 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.7335 | 382.8 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.6653 | 10.7565 | 28.1718 | 366.7
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.2 | 0.2 | 14 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8003 | 28.1498 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8716 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8335 | 382.9 | 0.2 | 0.2481 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.9 | 383 | 0.2 | 0.2 | 14 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0305 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8671 | 383 | 0.3 | 0.3 | 14 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7576 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0.0005 | 4.7456 | 16.1 | 366.9 | 383 | 0.2 | 0.2022 | 14 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7208 | 28.1 | 366.9
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 16.1 | 366.8 | 382.9 | 0.3 | 0.3 | 14 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7992 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0.0008 | 8.2519 | 16.1 | 366.8 | 382.9 | 0.3 | 0.3 | 14 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9662 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8776 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.3 | 0.3 | 14 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8439 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.3 | 0.3 | 14 | 13.2 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.4 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 16.1 | 366.8 | 382.9 | 0.3 | 0.3 | 14 | 13.2 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.7 | 59.9 | 0.0304 | 0.0278 | 0.0195 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.8
2008-12-02 | 14:06 | 0.002 | 0 | 0 | 16.1 | 366.136 | 382.2 | 0.3 | 0.2757 | 14 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.9 | 59.9332 | 0.0304 | 0.0278 | 0.019467 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 366.1
2008-12-02 | 14:06 | 0.01 | 0 | 0 | 16 | 339.5105 | 355.5 | 0.3 | 0.3 | 13.48 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0303 | 0.0277 | 0.017999 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 339.5
2008-12-02 | 14:06 | 0.002 | 0 | 0 | 15.7 | 320.7 | 336.4 | 0.2 | 0.2083 | 12.47 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0292 | 0.026667 | 0.017 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8482 | 28.1 | 320.7
2008-12-02 | 14:06 | 0 | 0 | 0 | 15.4 | 317.9669 | 333.3 | 0.1 | 0.104 | 12.1 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.028468 | 0.025901 | 0.0169 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8734 | 28.1 | 318
2008-12-02 | 14:06 | 0 | 0.0008 | 8.3656 | 15.3 | 317 | 332.3 | 0.1 | 0.1021 | 12.05 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.027967 | 0.025467 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.795 | 28.105 | 317
2008-12-02 | 14:06 | 0 | 0.001 | 10 | 15.2 | 316.9 | 332.1 | 0.1 | 0.1 | 12 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.027867 | 0.0254 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7164 | 28.1836 | 316.9
2008-12-02 | 14:06 | -0.001 | 0.001 | 10 | 15.1 | 316.9 | 332 | 0.2 | 0.2 | 12 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0278 | 0.025366 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.762 | 28.138 | 316.9
2008-12-02 | 14:06 | -0.01 | 0.001 | 10 | 15.1 | 316.7336 | 331.8 | 0.1 | 0.1029 | 12 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0278 | 0.0253 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7596 | 28.1 | 316.7
2008-12-02 | 14:06 | -0.002 | 0.0009 | 8.5543 | 15.1 | 316.7 | 331.8 | 0.2 | 0.2 | 12 | 13.2 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0277 | 0.0253 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7188 | 28.1 | 316.7
2008-12-02 | 14:06 | 0 | 0 | 0 | 15.1 | 316.6 | 331.7 | 0.2 | 0.2 | 12 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0277 | 0.0252 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7972 | 28.1 | 316.6
2008-12-02 | 14:06 | 0 | 0 | 0 | 15 | 316.5 | 331.5 | 0.2 | 0.2 | 12 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0276 | 0.025167 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9513 | 28.1 | 316.5
2008-12-02 | 14:06 | 0 | 0 | 0 | 15 | 316.4329 | 331.4 | 0.2 | 0.2 | 12 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.027566 | 0.0251 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9459 | 28.1 | 316.4
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.9 | 316.3 | 331.2 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0275 | 0.0251 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8675 | 28.1 | 316.3
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.9 | 316.2 | 331.1 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0275 | 0.025067 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.811 | 28.1 | 316.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.8 | 316.2 | 331 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.027467 | 0.025 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8893 | 28.1 | 316.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.8 | 316.167 | 330.9 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.027367 | 0.024967 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9679 | 28.1 | 316.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.7 | 316.1329 | 330.8 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.027267 | 0.024867 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 316.1
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.7 | 316.1662 | 330.9 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.027166 | 0.024766 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9752 | 28.1 | 316.2
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.7 | 316.0664 | 330.8 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0271 | 0.0247 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8969 | 28.1 | 316.1
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.7 | 316.0335 | 330.7 | 0.3 | 0.3 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0271 | 0.0247 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8185 | 28.1 | 316
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.6 | 316.1 | 330.7 | 0.3 | 0.3 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.027067 | 0.0247 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9199 | 28.1 | 316.1
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.6 | 315.9663 | 330.5 | 0.3 | 0.3 | 12 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.026966 | 0.0246 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9616 | 28.1 | 316
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.5 | 315.9335 | 330.4 | 0.3 | 0.2754 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0269 | 0.0245 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9169 | 28.1169 | 315.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.5 | 316 | 330.5 | 0.3 | 0.3 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.026867 | 0.0245 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9953 | 28.1953 | 316
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.5 | 315.9 | 330.4 | 0.3 | 0.3 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0268 | 0.024467 | 0.0168 | 0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9264 | 28.2 | 315.9
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.5 | 315.8343 | 330.3 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0268 | 0.0244 | 0.016767 | 0.000167 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9521 | 28.1479 | 315.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.4 | 315.8 | 330.2 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0268 | 0.0244 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 315.8
2008-12-02 | 14:06 | 0 | 0 | 0 | 14.4 | 315.8 | 330.2 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0268 | 0.0244 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9909 | 28.1091 | 315.8
2008-12-02 | 14:06 | -0.005 | 0 | 0 | 14.3 | 315.8 | 330.1 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.026766 | 0.0244 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1013.664 | 10.9125 | 28.1875 | 315.8
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.3 | 315.7 | 330 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0267 | 0.0244 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9659 | 28.1341 | 315.7
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.3 | 315.6667 | 330 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0267 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9115 | 28.1 | 315.7
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.3 | 315.5336 | 329.8 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9668 | 0.026667 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8227 | 28.1 | 315.5
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.3 | 315.4 | 329.7 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8976 | 28.1 | 315.4
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.2 | 315.4 | 329.6 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.7408 | 28.1 | 315.4
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.2 | 315.3 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.816 | 28.1 | 315.3
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.2 | 315.3 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 10.9 | 28.1 | 315.3
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.2 | 315.3 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.67 | 10.9 | 28.1 | 315.3
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.1 | 315.3 | 329.4 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.1 | 315.3
2008-12-02 | 14:07 | -0.002 | 0 | 0 | 14.1 | 315.3663 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0266 | 0.0243 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9718 | 28.1 | 315.4
2008-12-02 | 14:07 | -0.006 | 0 | 0 | 14.1 | 315.3 | 329.4 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8996 | 28.1 | 315.3
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 14.1 | 315.4 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8 | 28.1 | 315.4
2008-12-02 | 14:07 | -0.007 | 0 | 0 | 14.1 | 315.4 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8071 | 28.1 | 315.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 14.1 | 315.4 | 329.5 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8856 | 28.1 | 315.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 14.1 | 315.3 | 329.4 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9667 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.964 | 28.1 | 315.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 14 | 315.3 | 329.3 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 315.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 14 | 315.3669 | 329.4 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0265 | 0.0242 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9791 | 28.1 | 315.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 14 | 315.3 | 329.3 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.026467 | 0.024167 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9007 | 28.1 | 315.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 14 | 315.3 | 329.3 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0264 | 0.0241 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.8223 | 28.1 | 315.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315.3662 | 329.3 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0261 | 0.0238 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9123 | 28.1 | 315.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315.2 | 329.1 | 0.2 | 0.2 | 12 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.026 | 0.023766 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9655 | 28.1 | 315.2
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315.0665 | 329 | 0.2 | 0.2 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.1 | 315.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315.1 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.1 | 315.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315.0668 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9 | 28.1 | 315.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.9 | 315 | 328.9 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9483 | 28.1 | 315
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.8 | 315.0333 | 328.8 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.9733 | 28.1 | 315
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.8 | 315.2 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 10.905 | 28.1 | 315.2
2008-12-02 | 14:07 | 0 | 0 | 0 | 13.8 | 315.3 | 329.1 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9669 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1013 | 10.9836 | 28.1 | 315.3
2008-12-02 | 14:07 | 0 | 0 | -0.12019 | 13.8 | 315.2 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 315.2
2008-12-02 | 14:07 | 0 | -0.0024 | -24.1587 | 13.8 | 315.2 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9595 | 28.1 | 315.2
2008-12-02 | 14:07 | 0 | -0.003 | -30 | 13.8 | 315.1662 | 329 | 0.1 | 0.1 | 12 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9338 | 0.026 | 0.0237 | 0.0167 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9 | 28.1 | 315.2
2008-12-02 | 14:07 | -0.001 | -0.0016 | -16.0982 | 13.7 | 307.2123 | 320.9 | 0.1 | 0.1 | 11.96 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.026 | 0.0237 | 0.016266 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9 | 28.1 | 307.2
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 13.5 | 284.1719 | 297.6 | 0.1 | 0.1 | 11.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.025666 | 0.023499 | 0.015032 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.8243 | 28.1 | 284.2
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 12.9 | 266.3333 | 279.2 | 0 | 0.001 | 10.35 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.024467 | 0.0223 | 0.0141 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9084 | 28.1 | 266.3
2008-12-02 | 14:07 | -0.01 | 0 | 0 | 12.6 | 265.4 | 278 | 0 | 0 | 10.11 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0236 | 0.0215 | 0.0141 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6626 | 11 | 28.1 | 265.4
2008-12-02 | 14:07 | -0.005 | 0 | 0 | 12.6 | 264.4 | 277 | 0 | 0 | 10 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0235 | 0.0215 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 10.9669 | 28.1 | 264.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.6 | 264.3333 | 276.9 | 0 | 0 | 10 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0235 | 0.021467 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3333 | 10.7319 | 28.1 | 264.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.6 | 264.0669 | 276.7 | 0 | 0 | 10 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.023467 | 0.0214 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9038 | 28.1 | 264.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.5 | 264 | 276.5 | 0 | 0 | 10 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.023301 | 0.021201 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 264
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.4 | 263.9341 | 276.3 | 0 | 0 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0231 | 0.021 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.9
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.3 | 263.8 | 276.1 | 0 | 0 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.023066 | 0.021 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9904 | 28.1032 | 263.8
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.3 | 263.6664 | 276 | 0 | 0 | 10 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 196 | 59.9 | 0.022966 | 0.0209 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 10.7554 | 28.1815 | 263.7
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.3 | 263.6335 | 275.9 | 0.1 | 0.0992 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0229 | 0.0209 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.6653 | 10.88 | 28.14 | 263.6
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.2 | 263.6 | 275.8 | 0.1 | 0.0988 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022867 | 0.020867 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.6
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.2 | 263.5668 | 275.8 | 0 | 0.0482 | 10 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0228 | 0.0208 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.6
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.2 | 263.567 | 275.8 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0227 | 0.0207 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.6
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.1 | 263.5667 | 275.7 | 0.1 | 0.1 | 10 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0227 | 0.0207 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.6
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.1 | 263.5 | 275.6 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.022667 | 0.0207 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.5
2008-12-02 | 14:07 | 0 | 0 | 0 | 12.1 | 263.5 | 275.6 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0226 | 0.020667 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9693 | 28.1 | 263.5
2008-12-02 | 14:07 | 0 | 0 | 0 | 12 | 263.5 | 275.5 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0226 | 0.0206 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9092 | 28.1 | 263.5
2008-12-02 | 14:07 | 0 | 0 | 0 | 12 | 263.466 | 275.5 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0225 | 0.0206 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9876 | 28.1 | 263.5
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.9 | 263.3662 | 275.3 | 0.1 | 0.1 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0225 | 0.0206 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.4
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.9 | 263.3 | 275.2 | 0 | 0 | 10 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0225 | 0.0206 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.9 | 263.2667 | 275.1 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0458 | 28.1 | 263.3
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.1668 | 275 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1973 | 28.1 | 263.2
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.1665 | 275 | 0 | 0 | 10 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0405 | 28.1 | 263.2
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.1 | 274.9 | 0 | 0 | 10 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9419 | 28.1 | 263.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.1 | 274.9 | 0 | 0 | 10 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9366 | 28.1 | 263.1
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.1669 | 275 | 0 | 0 | 10 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0225 | 0.0205 | 0.014 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263.2
2008-12-02 | 14:07 | 0 | 0 | 0 | 11.8 | 263.133 | 274.9 | 0 | 0 | 10 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.022467 | 0.0205 | 0.013967 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.67 | 11 | 28.1 | 263.1
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.8 | 263.1671 | 275 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3293 | 11.0719 | 28.1 | 263.2
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.8 | 263.1662 | 274.9 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 10.9995 | 28.1 | 263.2
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.7 | 263.1 | 274.8 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 10.9286 | 28.1 | 263.1
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.7 | 263.1 | 274.8 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11 | 28.1 | 263.1
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.7 | 263 | 274.7 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 263
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.7 | 263 | 274.7 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0224 | 0.0205 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9359 | 28.1 | 263
2008-12-02 | 14:08 | -0.003 | 0 | 0 | 11.7 | 263 | 274.7 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0224 | 0.020466 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9425 | 28.1 | 263
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.7 | 263 | 274.7 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.021 | 28.1 | 263
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 263 | 274.6 | 0 | 0 | 10 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0995 | 28.1 | 263
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 262.9 | 274.5 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0221 | 28.1 | 262.9
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 263 | 274.6 | 0 | 0 | 10 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11 | 28.1 | 263
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 262.9 | 274.5 | 0 | 0 | 10 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11 | 28.1 | 262.9
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 262.9 | 274.5 | 0 | 0 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0223 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.336 | 10.9871 | 28.1 | 262.9
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.6 | 262.8 | 274.4 | 0 | 0 | 10 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.022267 | 0.0204 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9084 | 28.1 | 262.8
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.5 | 262.7667 | 274.3 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9 | 28.1 | 262.8
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.4 | 262.6668 | 274.1 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.8515 | 28.1 | 262.7
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.4 | 262.5665 | 274 | 0 | 0 | 10 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.8 | 28.1 | 262.6
2008-12-02 | 14:08 | -0.005 | 0 | 0 | 11.4 | 262.5 | 273.9 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.8101 | 28.1 | 262.5
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.5 | 273.8 | 0 | 0 | 10 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9676 | 28.1 | 262.5
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.5 | 273.8 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 262.5
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.4 | 273.7 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 262.4
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.4 | 273.7 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0222 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.981 | 28.1 | 262.4
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.4671 | 273.8 | 0 | 0 | 10 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.022167 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9027 | 28.1 | 262.5
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.4 | 273.7 | 0 | 0 | 10 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0221 | 0.0203 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9759 | 28.1 | 262.4
2008-12-02 | 14:08 | 0 | 0 | 0 | 11.3 | 262.4 | 273.7 | 0 | 0 | 10 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0221 | 0.020267 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 262.4
2008-12-02 | 14:08 | -0.001 | 0 | 0 | 11.3 | 262.4 | 273.7 | 0 | 0 | 10 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.022067 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 262.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.022 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9888 | 28.1 | 262.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.022 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9106 | 28.1 | 262.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.022 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0361 | 28.1 | 262.4
2008-12-02 | 14:08 | -0.001 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0534 | 28.1 | 262.4
2008-12-02 | 14:08 | 0.01 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022 | 0.0202 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.075 | 262.4
2008-12-02 | 14:08 | 0.009 | 0 | 0 | 11.2 | 262.4 | 273.6 | 0 | 0 | 10 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.022 | 0.020167 | 0.0139 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.0034 | 262.4
2008-12-02 | 14:08 | 0 | 0 | 0 | 11 | 257.8377 | 268.8 | 0 | 0 | 10 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.021966 | 0.0201 | 0.013663 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.0818 | 257.8
2008-12-02 | 14:08 | 0 | 0 | 0 | 10.6 | 234.6715 | 245.3 | 0 | 0 | 9.34 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 196 | 59.9 | 0.021365 | 0.019598 | 0.012432 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 234.7
2008-12-02 | 14:08 | 0 | 0 | 0 | 10.2 | 213.9 | 224.1 | -0.1 | 0 | 8.41 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0198 | 0.0181 | 0.011367 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0387 | 28.1 | 213.9
2008-12-02 | 14:08 | 0 | 0 | 0 | 10 | 211.8 | 221.8 | -0.1 | 0 | 8.16 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0192 | 0.0176 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0828 | 28.1 | 211.8
2008-12-02 | 14:08 | 0 | 0 | 0 | 10 | 211.8 | 221.8 | -0.2 | 0 | 8.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0192 | 0.017567 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0044 | 28.1 | 211.8
2008-12-02 | 14:08 | 0 | 0 | 0 | 9.9 | 211.7665 | 221.7 | -0.2 | 0 | 8.06 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.019066 | 0.0174 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11 | 28.1 | 211.8
2008-12-02 | 14:08 | 0 | 0 | 0 | 9.9 | 210.8 | 220.7 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.018967 | 0.0173 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6667 | 11 | 28.1 | 210.8
2008-12-02 | 14:08 | 0 | 0 | 0 | 9.8 | 210.6669 | 220.5 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0189 | 0.0173 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3313 | 11 | 28.1 | 210.7
2008-12-02 | 14:08 | 0 | 0 | 0 | 9.8 | 210.6 | 220.4 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0189 | 0.0173 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 210.6
2008-12-02 | 14:08 | 0 | -0.0008 | -7.8577 | 9.8 | 210.567 | 220.4 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.018867 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 210.6
2008-12-02 | 14:08 | 0 | -0.001 | -10 | 9.8 | 210.5 | 220.3 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0188 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.066 | 28.034 | 210.5
2008-12-02 | 14:08 | 0 | -0.0015 | -14.6613 | 9.8 | 210.5 | 220.3 | -0.1 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0188 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.0444 | 210.5
2008-12-02 | 14:08 | 0 | -0.002 | -20 | 9.8 | 210.4664 | 220.3 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0188 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1231 | 28.1 | 210.5
2008-12-02 | 14:08 | 0 | -0.002 | -20 | 9.8 | 210.4665 | 220.3 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0188 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1969 | 28.1 | 210.5
2008-12-02 | 14:08 | -0.006 | -0.002 | -20 | 9.8 | 210.4 | 220.2 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0188 | 0.0172 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.04 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | -0.0012 | -11.964 | 9.8 | 210.4 | 220.2 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0187 | 0.0171 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.668 | 11.0583 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | -0.0004 | -3.5696 | 9.8 | 210.4 | 220.2 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0187 | 0.0171 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.7 | 210.4 | 220.1 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0187 | 0.0171 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1154 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.7 | 210.4 | 220.1 | -0.2 | 0 | 8 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.018667 | 0.017067 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1937 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.6 | 210.4 | 220 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0186 | 0.017 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.6 | 210.4 | 220 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0186 | 0.017 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.6 | 210.4 | 220 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0185 | 0.016966 | 0.0112 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1426 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | -0.0006 | -6.3699 | 9.6 | 210.4 | 220 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0185 | 0.0169 | 0.011166 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0075 | 28.1 | 210.4
2008-12-02 | 14:08 | -0.01 | -0.001 | -10 | 9.6 | 210.3 | 219.9 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0859 | 28.1 | 210.3
2008-12-02 | 14:08 | -0.01 | -0.0007 | -7.1046 | 9.5 | 210.3 | 219.8 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0185 | 0.0169 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2286 | 28.1 | 210.3
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.5 | 210.3 | 219.8 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0184 | 0.0169 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2573 | 28.1 | 210.3
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.5 | 210.2665 | 219.8 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0184 | 0.0169 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1789 | 28.1 | 210.3
2008-12-02 | 14:08 | -0.01 | 0 | 0 | 9.5 | 210.1667 | 219.7 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0184 | 0.0168 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3333 | 11.1003 | 28.1 | 210.2
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.5 | 210.1 | 219.6 | -0.2 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0184 | 0.0168 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 210.1
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.4 | 210.0656 | 219.5 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0168 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0435 | 28.1 | 210.1
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.4 | 209.967 | 219.4 | -0.2 | 0 | 8 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0348 | 28.1 | 210
2008-12-02 | 14:09 | -0.002 | 0 | 0 | 9.4 | 209.9 | 219.3 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.4 | 209.9338 | 219.3 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6621 | 11.1 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.005 | 0 | 0 | 9.4 | 210 | 219.4 | -0.2 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1703 | 28.1 | 210
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.4 | 210 | 219.4 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1513 | 28.0513 | 210
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.4 | 210 | 219.4 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0183 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3333 | 11.1 | 28.027 | 210
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.4 | 209.9669 | 219.4 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0182 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.105 | 28.1 | 210
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.9667 | 219.3 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0182 | 0.0167 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.184 | 28.1 | 210
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.8331 | 219.1 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0182 | 0.016667 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1375 | 28.1 | 209.8
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.9 | 219.2 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9331 | 0.0182 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.8658 | 219.2 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9658 | 0.018166 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1193 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.9 | 219.2 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1974 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.8 | 219.1 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1239 | 28.1 | 209.8
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.8 | 219.1 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.8
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.9 | 219.2 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.3 | 209.8667 | 219.2 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.9
2008-12-02 | 14:09 | -0.007 | 0 | 0 | 9.3 | 209.8 | 219.1 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0181 | 0.0166 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.3 | 209.8 | 219.1 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.018066 | 0.016566 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.0318 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8 | 219 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.017967 | 0.016467 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.0469 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8 | 219 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0179 | 0.0164 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8669 | 219.1 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0179 | 0.0164 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3306 | 11.1037 | 28.1 | 209.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8 | 219 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.182 | 28.1 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.9 | 219.1 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1395 | 28.1 | 209.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8664 | 219.1 | -0.4 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.139 | 28.1 | 209.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8 | 219 | -0.4 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3347 | 11.1826 | 28.1 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.8 | 219 | -0.3 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3333 | 11.1042 | 28.1 | 209.8
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.7 | 218.9 | -0.4 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.5 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1739 | 28.1 | 209.7
2008-12-02 | 14:09 | 0 | 0 | 0 | 9.2 | 209.7 | 218.9 | -0.4 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.095 | 28.1 | 209.7
2008-12-02 | 14:09 | -0.003 | 0 | 0 | 9.1 | 209.7 | 218.8 | -0.4 | 0 | 8 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0178 | 0.0163 | 0.0111 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 209.7
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9.1 | 208.0103 | 217.1 | -0.3 | 0 | 8 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0178 | 0.0163 | 0.011034 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0096 | 28.1 | 208
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 9 | 180.1992 | 189.2 | -0.4 | 0 | 7.57 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.017368 | 0.015901 | 0.009569 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0881 | 28.1 | 180.2
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 8.5 | 163.2811 | 171.7 | -0.4 | 0 | 6.52 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.016302 | 0.014902 | 0.008701 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0335 | 28.1 | 163.3
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 8.1 | 160.0601 | 168.1 | -0.5 | 0 | 6.22 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.015766 | 0.014466 | 0.0085 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0449 | 28.1 | 160.1
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 8 | 159.3662 | 167.3 | -0.5 | 0 | 6.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.015599 | 0.014299 | 0.0085 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28.1 | 159.4
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.9 | 159.2664 | 167.2 | -0.4 | 0 | 6.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.015366 | 0.014066 | 0.0085 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0982 | 28.1 | 159.3
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.9 | 159.2 | 167.1 | -0.5 | 0 | 6.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.015267 | 0.013967 | 0.008467 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0198 | 28.1 | 159.2
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.8 | 159.1 | 166.9 | -0.4 | 0 | 6.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.015067 | 0.0138 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0584 | 28.1 | 159.1
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.7 | 159.1668 | 166.9 | -0.4 | 0 | 6.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.015 | 0.0137 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0631 | 28.1 | 159.2
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.7 | 159.0665 | 166.8 | -0.4 | 0 | 6.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.015 | 0.0137 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0156 | 28.1 | 159.1
2008-12-02 | 14:09 | -0.006 | 0 | 0 | 7.7 | 158.9667 | 166.7 | -0.4 | 0 | 6.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0149 | 0.0136 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.094 | 28.1 | 159
2008-12-02 | 14:09 | 0 | 0 | 0 | 7.7 | 158.9 | 166.6 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.014867 | 0.0136 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1725 | 28.1 | 158.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 7.7 | 158.9 | 166.6 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0148 | 0.0135 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1495 | 28.1 | 158.9
2008-12-02 | 14:09 | 0 | 0 | 0 | 7.6 | 158.9 | 166.5 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.014767 | 0.0135 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1577 | 28.1 | 158.9
2008-12-02 | 14:09 | -0.008 | 0 | 0 | 7.6 | 158.9 | 166.5 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0147 | 0.0134 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2845 | 28.1 | 158.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.5 | 158.9 | 166.4 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 196 | 59.9 | 0.0146 | 0.0134 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6621 | 11.1276 | 28.1 | 158.9
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.5 | 158.8 | 166.3 | -0.4 | 0 | 6.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0146 | 0.0134 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9707 | 28.1 | 158.8
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.5 | 158.8 | 166.3 | -0.4 | 0 | 6.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0146 | 0.0133 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.943 | 28.1 | 158.8
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.5 | 158.7 | 166.2 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.014567 | 0.0133 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 158.7
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.5 | 158.7 | 166.2 | -0.4 | 0 | 6 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0145 | 0.0133 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 158.7
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.4 | 158.6 | 166 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0145 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0784 | 28.1 | 158.6
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.4 | 158.6 | 166 | -0.4 | 0 | 6 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1569 | 28.1 | 158.6
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.4 | 158.5 | 165.9 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.165 | 28.1 | 158.5
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.4 | 158.5 | 165.9 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 158.5
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.4 | 158.5 | 165.9 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 158.5
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.3 | 158.5 | 165.8 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1706 | 28.1 | 158.5
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.3 | 158.5 | 165.8 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0144 | 0.0132 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.151 | 28.1 | 158.5
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.2 | 158.4 | 165.6 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0143 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1274 | 28.1 | 158.4
2008-12-02 | 14:09 | -0.01 | 0 | 0 | 7.2 | 158.3667 | 165.6 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0143 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 158.4
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.2 | 158.4 | 165.6 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0143 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 158.4
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.2 | 158.3 | 165.5 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0143 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 158.3
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.2 | 158.2667 | 165.5 | -0.4 | 0 | 6 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0143 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1588 | 28.1 | 158.3
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.2 | 158.2 | 165.4 | -0.4 | 0 | 6 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0142 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0805 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.1 | 158.2 | 165.3 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.9 | 0.0142 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0024 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.1 | 158.2 | 165.3 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0142 | 0.0131 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.67 | 11.0764 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.1 | 158.2 | 165.3 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0142 | 0.013 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1549 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.1 | 158.2 | 165.3 | -0.5 | 0 | 6 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.014166 | 0.013 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7.1 | 158.1664 | 165.3 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0141 | 0.013 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.336 | 11.2 | 28.1 | 158.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7 | 158.1 | 165.1 | -0.5 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0141 | 0.0129 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 158.1
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7 | 158.0668 | 165.1 | -0.5 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.014 | 0.0129 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1316 | 28.1 | 158.1
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 7 | 158 | 165 | -0.5 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0139 | 0.0128 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 158
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.9667 | 164.9 | -0.5 | 0 | 6 | 13.2 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0139 | 0.0127 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1256 | 28.1 | 158
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.9 | 164.8 | -0.5 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0138 | 0.0127 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.9
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.9329 | 164.8 | -0.4 | 0 | 6 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0138 | 0.012667 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.9
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.9 | 164.8 | -0.2 | 0 | 6 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.967 | 0.0138 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.9
2008-12-02 | 14:10 | -0.016 | 0 | 0 | 6.9 | 157.9 | 164.8 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2786 | 28.1 | 157.9
2008-12-02 | 14:10 | -0.015 | 0 | 0 | 6.9 | 157.8 | 164.7 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3468 | 28.1 | 157.8
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.7 | 164.6 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1115 | 28.1 | 157.7
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.7 | 164.6 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1741 | 28.0259 | 157.7
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.7 | 164.6 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 196 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1473 | 28.0527 | 157.7
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.9 | 157.7 | 164.6 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1315 | 28.1 | 157.7
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.8 | 157.7 | 164.5 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.21 | 28.1 | 157.7
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 6.8 | 157.7 | 164.5 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9667 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2884 | 28.1 | 157.7
2008-12-02 | 14:10 | -0.007 | 0 | 0 | 6.8 | 157.7 | 164.5 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2333 | 28.1 | 157.7
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.8 | 157.667 | 164.5 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.7
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.8 | 157.6 | 164.4 | -0.2 | 0 | 6 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.6
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.8 | 157.7 | 164.5 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.7
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.8 | 157.6 | 164.4 | -0.2 | 0 | 6 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 157.6
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.8 | 157.6 | 164.4 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1414 | 28.1 | 157.6
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.7 | 157.6 | 164.3 | -0.2 | 0 | 6 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0137 | 0.0126 | 0.0084 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.137 | 28.1 | 157.6
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.7 | 155.5125 | 162.2 | -0.3 | 0 | 6 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0137 | 0.0126 | 0.008301 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1842 | 28.1 | 155.5
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.7 | 138.5422 | 145.2 | -0.2 | 0 | 5.69 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8665 | 0.013499 | 0.012433 | 0.007399 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3354 | 11.1057 | 28.1 | 138.5
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.4 | 115.7333 | 122.1 | -0.3 | 0 | 4.64 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0124 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6667 | 11.1 | 28.1 | 115.7
2008-12-02 | 14:10 | 0 | 0 | 0 | 6.1 | 109.0018 | 115.1 | -0.4 | 0 | 4.22 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.011767 | 0.0107 | 0.0058 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 109
2008-12-02 | 14:10 | 0 | 0.0008 | 7.8649 | 5.7 | 106.6669 | 112.4 | -0.4 | 0 | 4.17 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0113 | 0.010367 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106.7
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.7 | 106.6 | 112.3 | -0.4 | 0 | 4.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0112 | 0.010267 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.092 | 28.1 | 106.6
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.7 | 106.6 | 112.3 | -0.4 | 0 | 4.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0111 | 0.0102 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0136 | 28.1 | 106.6
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.6 | 106.5 | 112.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.011066 | 0.0101 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0649 | 28.1 | 106.5
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.6 | 106.5664 | 112.2 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.010966 | 0.01 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1431 | 28.1 | 106.6
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.6 | 106.5 | 112.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0109 | 0.01 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1786 | 28.1 | 106.5
2008-12-02 | 14:10 | 0 | 0.001 | 10 | 5.5 | 106.4667 | 112 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.010867 | 0.009967 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0998 | 28.1 | 106.5
2008-12-02 | 14:10 | 0 | -0.0002 | -2.4937 | 5.5 | 106.4 | 111.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0108 | 0.0099 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0213 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.5 | 106.4 | 111.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0107 | 0.0098 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.5 | 106.4 | 111.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0107 | 0.0098 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.5 | 106.4 | 111.9 | -0.4 | 0 | 4.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0107 | 0.0098 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0133 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.4 | 106.4 | 111.8 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.010667 | 0.0097 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0923 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.4 | 106.4 | 111.8 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0106 | 0.0097 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.3 | 106.4 | 111.7 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0105 | 0.0097 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1985 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.3 | 106.4 | 111.7 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0105 | 0.0097 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2724 | 28.1 | 106.4
2008-12-02 | 14:10 | 0 | -0.001 | -10 | 5.3 | 106.4 | 111.7 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0105 | 0.0096 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1941 | 28.1 | 106.4
2008-12-02 | 14:10 | -0.008 | -0.001 | -10 | 5.2 | 106.3 | 111.5 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0104 | 0.009567 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1154 | 28.1 | 106.3
2008-12-02 | 14:10 | -0.01 | -0.001 | -10 | 5.2 | 106.2668 | 111.5 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0103 | 0.0095 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1631 | 28.1 | 106.3
2008-12-02 | 14:10 | -0.005 | -0.001 | -10 | 5.2 | 106.2 | 111.4 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0103 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1585 | 28.1 | 106.2
2008-12-02 | 14:10 | -0.003 | -0.0004 | -3.8462 | 5.2 | 106.2 | 111.4 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0103 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1396 | 28.1 | 106.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 5.2 | 106.1669 | 111.4 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0103 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2955 | 28.1 | 106.2
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 5.2 | 106.1 | 111.3 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0102 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1467 | 28.1 | 106.1
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 5.2 | 106.1 | 111.3 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0102 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106.1
2008-12-02 | 14:10 | -0.01 | 0 | 0 | 5.1 | 106.0671 | 111.2 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0102 | 0.0094 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106.1
2008-12-02 | 14:10 | -0.01 | -0.0008 | -7.7931 | 5.1 | 106 | 111.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0102 | 0.009366 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | -0.0004 | -4.0048 | 5.1 | 106 | 111.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0102 | 0.0093 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 5.1 | 106 | 111.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0101 | 0.0093 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1689 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 5.1 | 106 | 111.1 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8333 | 0.0101 | 0.0093 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1525 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 5 | 106 | 111 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0101 | 0.009267 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 5 | 106 | 111 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0101 | 0.0092 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1043 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106.0667 | 111 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.01 | 0.0092 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1823 | 28.1 | 106.1
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106 | 110.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.01 | 0.0092 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106 | 110.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.01 | 0.0092 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1208 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106 | 110.9 | -0.4 | 0 | 4.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.01 | 0.009167 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.0361 | 28.082 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106 | 110.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1929 | 28.0035 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.9 | 106 | 110.9 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1251 | 28.0749 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.8 | 105.9665 | 110.8 | -0.4 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.9 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3347 | 11.1 | 28.1 | 106
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1318 | 28.1 | 105.9
2008-12-02 | 14:11 | -0.009 | 0 | 0 | 4.8 | 106 | 110.8 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8668 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 106
2008-12-02 | 14:11 | -0.001 | 0 | 0 | 4.8 | 105.9335 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.6646 | 11.2 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.8 | 105.9667 | 110.8 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0099 | 0.0091 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.3333 | 11.2668 | 28.1 | 106
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.0098 | 0.009 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1641 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0098 | 0.009 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.024 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0098 | 0.009 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0976 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | -0.0004 | -4.1583 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0098 | 0.009 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0191 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0098 | 0.009 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0588 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.9 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.009766 | 0.008966 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.8666 | 110.7 | -0.5 | 0 | 4.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1162 | 28.1 | 105.9
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.8333 | 110.6 | -0.5 | 0 | 4.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1947 | 28.1 | 105.8
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.8 | 110.6 | -0.5 | 0 | 4.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 105.8
2008-12-02 | 14:11 | 0 | -0.001 | -10 | 4.8 | 105.8667 | 110.7 | -0.5 | 0 | 4.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1486 | 28.0486 | 105.9
2008-12-02 | 14:11 | 0 | -0.0008 | -7.5187 | 4.8 | 105.8 | 110.6 | -0.5 | 0 | 4.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.0297 | 105.8
2008-12-02 | 14:11 | 0 | -0.0001 | -0.80833 | 4.8 | 105.8 | 110.6 | -0.5 | 0 | 4.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1084 | 28.1 | 105.8
2008-12-02 | 14:11 | 0 | -0.0009 | -9.1417 | 4.8 | 105.8 | 110.6 | -0.5 | 0 | 4.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0097 | 0.0089 | 0.0057 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1868 | 28.1 | 105.8
2008-12-02 | 14:11 | 0 | -0.0002 | -2.2203 | 4.8 | 105.7671 | 110.5 | -0.5 | 0 | 4.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0097 | 0.0089 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 105.8
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 105.7 | 110.4 | -0.5 | 0 | 4.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0097 | 0.0089 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1566 | 28.1 | 105.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 105.7 | 110.4 | -0.5 | 0 | 4.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0097 | 0.0089 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 105.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 105.7 | 110.4 | -0.5 | 0 | 4.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.5 | 59.9 | 0.009667 | 0.0089 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 105.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 105.7 | 110.4 | -0.5 | 0 | 4.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0096 | 0.0089 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 105.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 105.7 | 110.4 | -0.5 | 0 | 4.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0096 | 0.008866 | 0.0056 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1575 | 28.0425 | 105.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.7 | 96.9253 | 101.6 | -0.5 | 0 | 4.08 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0096 | 0.0088 | 0.005165 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1642 | 28.0358 | 96.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.6 | 71.9 | 76.5 | -0.5 | 0 | 3.13 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0095 | 0.0087 | 0.0039 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1283 | 28.1 | 71.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.4 | 60.6127 | 65 | -0.6 | 0 | 2.37 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8669 | 0.008869 | 0.008102 | 0.003301 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2855 | 28.1 | 60.6
2008-12-02 | 14:11 | 0 | 0 | 0 | 4.1 | 55.0694 | 59.1 | -0.7 | 0 | 2.2 | 13.2 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.007801 | 0.007101 | 0.002967 | 0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1577 | 28.1 | 55.1
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.9 | 54.4 | 58.3 | -0.7 | 0 | 2.18 | 13.2 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0075 | 0.006867 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1496 | 28.1 | 54.4
2008-12-02 | 14:11 | -0.006 | 0 | 0 | 3.8 | 54.2671 | 58 | -0.7 | 0 | 2.1 | 13.2 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.007367 | 0.006767 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 54.3
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.6 | 54.2 | 57.8 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8338 | 0.006966 | 0.006366 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.188 | 28.1 | 54.2
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.6 | 54.1 | 57.7 | -0.7 | 0 | 2.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0069 | 0.0063 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0303 | 28.1 | 54.1
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.5 | 54.1 | 57.6 | -0.7 | 0 | 2.1 | 13.1 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.006867 | 0.006267 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0634 | 28.1 | 54.1
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.5 | 54.1 | 57.6 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0068 | 0.0062 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1418 | 28.1 | 54.1
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.4 | 54 | 57.4 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0067 | 0.0062 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1595 | 28.1 | 54
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.4 | 54 | 57.4 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0067 | 0.006166 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0027 | 28.1 | 54
2008-12-02 | 14:11 | -0.01 | 0 | 0 | 3.4 | 53.9667 | 57.4 | -0.7 | 0 | 2.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.006667 | 0.0061 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1542 | 28.1 | 54
2008-12-02 | 14:11 | -0.008 | 0 | 0 | 3.3 | 53.9 | 57.2 | -0.7 | 0 | 2.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0066 | 0.006 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2555 | 28.0445 | 53.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.3 | 53.7669 | 57.1 | -0.7 | 0 | 2.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0065 | 0.005967 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28.034 | 53.8
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.3 | 52.9 | 56.2 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.006467 | 0.0059 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2753 | 28.1 | 52.9
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.3 | 52.8 | 56.1 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.5 | 59.9 | 0.0064 | 0.0059 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.6707 | 11.1189 | 28.1 | 52.8
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.3 | 52.8 | 56.1 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0064 | 0.0059 | 0.0029 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1692 | 28.1 | 52.8
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.2 | 52.7665 | 56 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0064 | 0.0058 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.8
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.2 | 52.7 | 55.9 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0063 | 0.0058 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.2 | 52.7 | 55.9 | -0.7 | 0 | 2.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0063 | 0.0058 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.7
2008-12-02 | 14:11 | 0 | 0 | 0 | 3.1 | 52.6 | 55.7 | -0.7 | 0 | 2.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0063 | 0.0058 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.6
2008-12-02 | 14:11 | 0 | 0 | 0 | 3 | 52.6 | 55.6 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.006267 | 0.0057 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.3333 | 11.2 | 28.1 | 52.6
2008-12-02 | 14:11 | 0 | 0 | 0 | 3 | 52.5 | 55.5 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.006167 | 0.005667 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2399 | 28.1 | 52.5
2008-12-02 | 14:11 | 0 | 0 | 0 | 3 | 52.5 | 55.5 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0061 | 0.0056 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28.1 | 52.5
2008-12-02 | 14:12 | 0 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.2 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.006067 | 0.005567 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28.1 | 52.4
2008-12-02 | 14:12 | 0 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.006 | 0.0055 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.225 | 28.1 | 52.4
2008-12-02 | 14:12 | 0 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.005966 | 0.005466 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1464 | 28.1 | 52.4
2008-12-02 | 14:12 | 0 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0059 | 0.0054 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1322 | 28.1 | 52.4
2008-12-02 | 14:12 | 0 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0059 | 0.0054 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2106 | 28.1 | 52.4
2008-12-02 | 14:12 | -0.004 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0059 | 0.0054 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.289 | 28.1 | 52.4
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0059 | 0.0054 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2325 | 28.1 | 52.4
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 2.9 | 52.3335 | 55.2 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8665 | 0.0058 | 0.0054 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.6646 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0058 | 0.0053 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.4
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 2.9 | 52.4 | 55.3 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8331 | 0.0058 | 0.0053 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2027 | 28.1 | 52.4
2008-12-02 | 14:12 | -0.004 | 0 | 0 | 2.8 | 52.367 | 55.2 | -0.7 | 0 | 2.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0058 | 0.0053 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2812 | 28.1 | 52.4
2008-12-02 | 14:12 | 0 | -0.0007 | -7.3833 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0057 | 0.005267 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.6707 | 11.3 | 28.1 | 52.3
2008-12-02 | 14:12 | 0 | -0.001 | -10 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0056 | 0.0052 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2619 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.001 | -0.001 | -10 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0056 | 0.005166 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.009 | -0.001 | -10 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0056 | 0.0051 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.003 | -0.001 | -10 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0056 | 0.0051 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | 0 | -0.001 | -10 | 2.8 | 52.3 | 55.1 | -0.7 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0055 | 0.0051 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | 0 | -0.001 | -10 | 2.7 | 52.3 | 55 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8665 | 0.0055 | 0.0051 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.002 | -0.001 | -10 | 2.7 | 52.3 | 55 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.005467 | 0.005 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1913 | 28.0913 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.7 | 52.3 | 55 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0054 | 0.0049 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1129 | 28.0129 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.7 | 52.3 | 55 | -0.8 | 0 | 2.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0054 | 0.0049 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2311 | 28.0656 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.6 | 52.3 | 54.9 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.005367 | 0.0049 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.256 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.6 | 52.3 | 54.9 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0049 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1776 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.6 | 52.3 | 54.9 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1019 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.6 | 52.3 | 54.9 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2587 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.6 | 52.3 | 54.9 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.6667 | 11.1844 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 52.3 | 54.8 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1362 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 52.3 | 54.8 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 52.3 | 54.8 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8667 | 0.005267 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 52.3 | 54.8 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0053 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1285 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 52.2656 | 54.8 | -0.8 | 0 | 2.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0052 | 0.0048 | 0.0028 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1499 | 28.1 | 52.3
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.5 | 49.1642 | 51.7 | -0.8 | 0 | 2.08 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0052 | 0.0048 | 0.002635 | 0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.33 | 11.1719 | 28.1 | 49.2
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.3 | 22.8281 | 25.2 | -0.8 | 0 | 1.23 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.004996 | 0.004596 | 0.001224 | 6.6e-05 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1122 | 28.1 | 22.8
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 2.1 | 3.1621 | 5.3 | -0.9 | 0 | 0.45 | 13.1 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.004196 | 0.003896 | 0.0002 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.6621 | 11.2704 | 28.1 | 3.2
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 1.8 | 2.4 | 4.2 | -0.9 | 0 | 0.2 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.003466 | 0.003166 | 0.0002 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.664 | 11.3 | 28.1 | 2.4
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 1.7 | 1.4 | 3.1 | -0.9 | 0 | 0.18 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8341 | 0.003166 | 0.002866 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2156 | 28.0578 | 1.4
2008-12-02 | 14:12 | -0.01 | -0.001 | -10 | 1.7 | 1.4 | 3.1 | -0.9 | 0 | 0.1 | 13.1 | 905 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.003 | 0.0028 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1413 | 28.0206 | 1.4
2008-12-02 | 14:12 | -0.01 | -0.0006 | -5.8015 | 1.6 | 1.4 | 3 | -1 | 0 | 0.1 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8668 | 0.003 | 0.0027 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2981 | 28.0991 | 1.4
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 1.6 | 1.3 | 2.9 | -0.9 | 0 | 0.1 | 13.1 | 905 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8335 | 0.0029 | 0.002666 | 0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3775 | 28.1 | 1.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 1.6 | 0.3 | 1.9 | -1 | 0 | 0.1 | 13.1 | 905 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.002867 | 0.0026 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3441 | 28.1 | 0.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 1.5 | 0.2 | 1.7 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.002767 | 0.0025 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.1 | 0.2
2008-12-02 | 14:12 | -0.006 | 0 | 0 | 1.5 | 0.2 | 1.7 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.002667 | 0.0024 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3128 | 28.1 | 0.2
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.4 | 0.2 | 1.6 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0026 | 0.0024 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3912 | 28.1 | 0.2
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.4 | 0.16707 | 1.5 | -0.9 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.002468 | 0.002201 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28.1 | 0.2
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.3 | 0.1 | 1.4 | -0.9 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0022 | 0.002 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3518 | 28.1 | 0.1
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.3 | 0 | 1.3 | -0.9 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0022 | 0.001966 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2734 | 28.1 | 0
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.3 | -0.033465 | 1.3 | -0.9 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0021 | 0.0019 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.21 | 28.095 | 0
2008-12-02 | 14:12 | 0 | -0.0007 | -7.3819 | 1.2 | -0.13333 | 1.2 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.002067 | 0.001867 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3669 | 28.0165 | -0.1
2008-12-02 | 14:12 | 0 | 0.0002 | 1.5835 | 1.2 | -0.2 | 1.2 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8332 | 0.002 | 0.0018 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.332 | 11.3381 | 28.0619 | -0.2
2008-12-02 | 14:12 | 0 | 0.0006 | 5.8657 | 1.2 | -0.2 | 1.2 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.002 | 0.0018 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2597 | 28.1 | -0.2
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.1 | -0.2 | 1.1 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.002 | 0.0018 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.6667 | 11.2187 | 28.1 | -0.2
2008-12-02 | 14:12 | 0 | 0 | 0 | 1.1 | -0.2 | 1.1 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.001967 | 0.0018 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.6694 | 11.2972 | 28.1 | -0.2
2008-12-02 | 14:12 | -0.006 | 0 | 0 | 0.9 | -0.3 | 0.9 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0019 | 0.0017 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28.1 | -0.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.9 | -0.3 | 0.9 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.001867 | 0.0017 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.6707 | 11.3 | 28.1 | -0.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.9 | -0.3 | 0.9 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.001766 | 0.0016 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2675 | 28.1 | -0.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.8 | -0.3336 | 0.8 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8336 | 0.0017 | 0.0015 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.222 | 28.1 | -0.3
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.8 | -0.4 | 0.8 | -1 | 0 | 0.1 | 13.1 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0016 | 0.0015 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3788 | 28.1 | -0.4
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.8 | -0.5 | 0.8 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8333 | 0.0016 | 0.0015 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3322 | 28.1 | -0.5
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.8 | -0.5 | 0.8 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0016 | 0.0015 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2537 | 28.1 | -0.5
2008-12-02 | 14:12 | -0.01 | 0 | 0 | 0.6 | -0.5 | 0.6 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.001566 | 0.0015 | 0 | 0 |  | n/a |  | 35 | 61 | 1051 | 1012.6646 | 11.2247 | 28.1 | -0.5
2008-12-02 | 14:13 | -0.01 | 0 | 0 | 0.4 | -0.6 | 0.4 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0015 | 0.0014 | -3.3e-05 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3031 | 28.0969 | -0.6
2008-12-02 | 14:13 | -0.009 | 0 | 0 | 0.4 | -0.63313 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0015 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3815 | 28.0185 | -0.6
2008-12-02 | 14:13 | -0.001 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0015 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2801 | 28.06 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.7 | 59.8342 | 0.0015 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0014 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2168 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8662 | 0.0014 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2954 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8664 | 0.0014 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3736 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.4 | -0.7 | 0.4 | -1.1 | 0 | 0.1 | 13.1 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8335 | 0.0014 | 0.0014 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3479 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.7 | 0.3 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8332 | 0.0014 | 0.001367 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2387 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.7 | 0.3 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.001366 | 0.0013 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1182 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.73333 | 0.3 | -1.1 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0013 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2751 | 28.1 | -0.7
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.8 | 0.3 | -1.1 | 0 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8331 | 0.0013 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.0341 | -0.8
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.8 | 0.3 | -1 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0013 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3443 | 28.0443 | -0.8
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.3 | -0.8 | 0.3 | -1 | 0 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.833 | 0.001267 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3315 | 28.1 | -0.8
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.8 | 0.2 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0012 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28.0987 | -0.8
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.8 | 0.2 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8338 | 0.0012 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28.0202 | -0.8
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.9 | 0.2 | -1 | 0 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8336 | 0.0012 | 0.0012 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1581 | 28.0581 | -0.9
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.9 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0012 | 0.001167 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.127 | 28.1 | -0.9
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.9 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8667 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28.1 | -0.9
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -0.9 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8332 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28.1 | -0.9
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8669 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1438 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8667 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1497 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.5 | 59.8331 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0928 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0 | 0 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0143 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0.0004 | 3.9456 | 0.2 | -1 | 0.2 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8662 | 0.0012 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0642 | 28.1 | -1
2008-12-02 | 14:13 | 0 | 0.001 | 10 | 0.1 | -1 | 0.1 | -1 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8664 | 0.0011 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1425 | 28.1 | -1
2008-12-02 | 14:13 | -0.007 | 0.001 | 10 | 0.1 | -1 | 0.1 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0011 | 0.0011 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013.6653 | 11.2 | 28.1 | -1
2008-12-02 | 14:13 | -0.01 | 0.001 | 10 | 0.1 | -1.0333 | 0.1 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8667 | 0.0011 | 0.001067 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.1 | -1
2008-12-02 | 14:13 | -0.01 | 0.001 | 10 | 0 | -1.1 | 0 | -1 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.000767 | 0.000767 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.0221 | -1.1
2008-12-02 | 14:13 | -0.01 | 0.0016 | 15.5675 | 0 | -1.1 | 0 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8665 | 0.0007 | 0.0007 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.0563 | -1.1
2008-12-02 | 14:13 | -0.009 | 0.0004 | 4.5 | 0 | -1.1 | 0 | -1.1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8333 | 0.0007 | 0.0007 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1653 | 28.1 | -1.1
2008-12-02 | 14:13 | -0.001 | -0.002 | -20 | 0 | -1.1 | 0 | -1 | 0 | 0.1 | 13.2 | 906 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0007 | 0.0007 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28.1 | -1.1
2008-12-02 | 14:13 | 0 | -0.0019 | -19.2653 | 0 | -1.1 | 0 | -1 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0007 | 0.0007 | -0.0001 | 0 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28.1 | -1.1
Summary Information:
Post Processor DLL Version | 6.01
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-12-02
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 66
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 28261
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 2005
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 131
Version | 1.611
NO Span(ppm) | 302
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.00
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | Backhoe J Deere MN 700212 310G
License Plate
Engine Displacement | 4.5
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 13:52
Test End Time | 14:13
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 990.989
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iAMBII_RPM
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
Faults:
0X5009 - 12/02/2008 13:45:49.983 - FID battery backed RAM corrupt
Warnings:
0X0000 - 12/02/2008 13:52:01.347 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1263
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-12-02 | 13:52 | [RELEASE_VER=2.018 BUILD=66 BDATE=10/21/2008 IP=10.10.1.55]
No Information Available

## NO2 Linearity
Test Date:
Test Time:
12/2/2008
13:52:01
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
131

## NDUV
Test Date:
Test Time:
12/2/2008
13:52:01
SEMTECH Serial Number:
G05-SDS02
NDUV Serial Number:
131

## O2 Linearity
Test Date:
Test Time:
12/2/2008
13:52:01
SEMTECH Serial Number:
G05-SDS02

## O2
Test Date:
Test Time:
12/2/2008
13:52:01
SEMTECH Serial Number:
G05-SDS02