## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 321.8 | Scott gas | ALM054255 | ±1% | 2010-06-16 | 965.4
1200 ppm CO | 1205
12% CO2 | 12
1500 ppm NO | 1505
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.12 | Scott gas | ALM004744 | ±1% | 2010-06-04 | 93.36
200 ppm CO | 200.7
6% CO2 | 6.009
300 ppm NO | 297.2
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 37249 |  |  | CO - AMBII Serial Number: |  |  | 37249 |  | HC - AMBII Serial Number: |  |  | 37249 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-07-16 |  | 6.009 | X | 0.1 | 3 |  |  | 200.7 |  | 50 | 3 |  | 31.12 |  | 4 | 3 |  | 297.2 |  | 15 | 3 |  | 93.36 | 100 | 5 | 2
Test Time: | 10:00
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0.003 | 0.003 |  |  | 0 | 0 | 0 |  |  | 0 | 1012.5 | 1012.5 |  |  | 0 | -0.2 | -0.2 |  |  | 0 | -0.09 | -0.09
90 | 105 | 1 | 0.6009 | 0.597 | 0.585 | -0.012 | -2.01 |  | 20.07 | 28 | 8 | 39.86 |  | 3.112 | 1012.9 | 1009.8 | 32448.59 | Fail | 29.72 | 29.7 | 0 | 0 |  | 9.336 | 9.2 | -0.14 | -1.5
150 | 165 | 2 | 1.2018 | 1.195 | 1.144 | -0.051 | -4.27 |  | 40.14 | 40 | 0 | 0 |  | 6.224 | 1013.4 | 1007.2 | 16182.52 | Fail | 59.44 | 59.1 | -0.3 | -0.5 |  | 18.672 | 18.45 | -0.22 | -1.18
210 | 225 | 3 | 1.8027 | 1.794 | 1.723 | -0.071 | -3.96 |  | 60.21 | 60 | 0 | 0 |  | 9.336 | 1012.6 | 1003.3 | 10746.57 | Fail | 89.16 | 89.2 | 0 | 0 |  | 28.008 | 27.79 | -0.22 | -0.79
270 | 285 | 4 | 2.4036 | 2.393 | 2.327 | -0.066 | -2.76 |  | 80.28 | 90 | 10 | 12.46 |  | 12.448 | 1012.7 | 1000.3 | 8035.83 | Fail | 118.88 | 118.6 | -0.3 | -0.25 |  | 37.344 | 37.06 | -0.28 | -0.75
330 | 345 | 5 | 3.0045 | 2.994 | 2.93 | -0.064 | -2.14 |  | 100.35 | 108 | 8 | 7.97 |  | 15.56 | 1013.5 | 997.9 | 6413.24 | Fail | 148.6 | 148.8 | 0.2 | 0.13 |  | 46.68 | 46.46 | -0.22 | -0.47
390 | 405 | 6 | 3.6054 | 3.595 | 3.552 | -0.043 | -1.2 |  | 120.42 | 121 | 1 | 0.83 |  | 18.672 | 1013.3 | 994.6 | 5326.69 | Fail | 178.32 | 179.3 | 1 | 0.56 |  | 56.016 | 55.7 | -0.32 | -0.57
450 | 465 | 7 | 4.2063 | 4.197 | 4.169 | -0.028 | -0.67 |  | 140.49 | 145 | 5 | 3.56 |  | 21.784 | 1013.2 | 991.4 | 4551.05 | Fail | 208.04 | 207.3 | -0.7 | -0.34 |  | 65.352 | 64.7 | -0.65 | -0.99
510 | 525 | 8 | 4.8072 | 4.8 | 4.78 | -0.02 | -0.42 |  | 160.56 | 160 | -1 | -0.62 |  | 24.896 | 1012.5 | 987.6 | 3966.9 | Fail | 237.76 | 237.2 | -0.6 | -0.25 |  | 74.688 | 74.02 | -0.67 | -0.9
570 | 585 | 9 | 5.4081 | 5.404 | 5.399 | -0.005 | -0.09 |  | 180.63 | 180 | -1 | -0.55 |  | 28.008 | 1012.1 | 984.1 | 3513.64 | Fail | 267.48 | 268.2 | 0.7 | 0.26 |  | 84.024 | 83.48 | -0.54 | -0.64
630 | 645 | 10 | 6.009 | 6.009 | 6.008 | -0.001 | -0.02 |  | 200.7 | 196 | -5 | -2.49 |  | 31.12 | 1013.6 | 982.5 | 3157.13 | Fail | 297.2 | 297.2 | 0 | 0 |  | 93.36 | 92.8 | -0.56 | -0.6
690 | 705 | 9 | 5.4081 | 5.404 | 5.41 | 0.006 | 0.11 |  | 180.63 | 180 | -1 | -0.55 |  | 28.008 | 1012.9 | 984.9 | 3516.5 | Fail | 267.48 | 268.7 | 1.2 | 0.45 |  | 84.024 | 83.3 | -0.72 | -0.86
750 | 765 | 8 | 4.8072 | 4.8 | 4.78 | -0.02 | -0.42 |  | 160.56 | 160 | -1 | -0.62 |  | 24.896 | 1012.3 | 987.4 | 3966.1 | Fail | 237.76 | 239.4 | 1.6 | 0.67 |  | 74.688 | 73.93 | -0.76 | -1.02
810 | 825 | 7 | 4.2063 | 4.197 | 4.162 | -0.035 | -0.83 |  | 140.49 | 141 | 1 | 0.71 |  | 21.784 | 1013.6 | 991.8 | 4552.88 | Fail | 208.04 | 209.4 | 1.4 | 0.67 |  | 65.352 | 64.59 | -0.76 | -1.16
870 | 885 | 6 | 3.6054 | 3.595 | 3.552 | -0.043 | -1.2 |  | 120.42 | 120 | 0 | 0 |  | 18.672 | 1012.6 | 993.9 | 5322.94 | Fail | 178.32 | 180 | 1.7 | 0.95 |  | 56.016 | 55.29 | -0.73 | -1.3
930 | 945 | 5 | 3.0045 | 2.994 | 2.933 | -0.061 | -2.04 |  | 100.35 | 100 | 0 | 0 |  | 15.56 | 1012.8 | 997.2 | 6408.74 | Fail | 148.6 | 150.3 | 1.7 | 1.14 |  | 46.68 | 45.9 | -0.78 | -1.67
990 | 1005 | 4 | 2.4036 | 2.393 | 2.33 | -0.063 | -2.63 |  | 80.28 | 80 | 0 | 0 |  | 12.448 | 1012.7 | 1000.3 | 8035.83 | Fail | 118.88 | 120.2 | 1.3 | 1.09 |  | 37.344 | 36.58 | -0.76 | -2.04
1050 | 1065 | 3 | 1.8027 | 1.794 | 1.723 | -0.071 | -3.96 |  | 60.21 | 60 | 0 | 0 |  | 9.336 | 1012.2 | 1002.9 | 10742.29 | Fail | 89.16 | 90.6 | 1.4 | 1.57 |  | 28.008 | 27.4 | -0.61 | -2.18
1110 | 1125 | 2 | 1.2018 | 1.195 | 1.14 | -0.055 | -4.6 |  | 40.14 | 40 | 0 | 0 |  | 6.224 | 1012.5 | 1006.3 | 16168.06 | Fail | 59.44 | 61.4 | 2 | 3.36 |  | 18.672 | 18.1 | -0.57 | -3.05
1170 | 1185 | 1 | 0.6009 | 0.597 | 0.58 | -0.017 | -2.85 |  | 20.07 | 20 | 0 | 0 |  | 3.112 | 1011.9 | 1008.8 | 32416.45 | Fail | 29.72 | 31.6 | 1.9 | 6.39 |  | 9.336 | 8.8 | -0.54 | -5.78
1230 | 1245 | 0 | 0 | 0 | 0.004 | 0.004 |  |  | 0 | -10 | -10 |  |  | 0 | 1012.2 | 1012.2 |  |  | 0 | 1.6 | 1.6 |  |  | 0 | -0.54 | -0.54
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.042 | 0.06 |  |  |  | Intercept | 1 | 2 |  |  | Intercept | 1013 | 0 | Fail |  | Intercept | 0.9 | 3 |  |  | Intercept | -0.32 | 0.93
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 0.02 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0273 | 0.0601 |  |  |  | SEE | 4.3143 | 2.007 | Fail |  | SEE | 0.4932 | 0.3112 | Fail |  | SEE | 0.9292 | 2.972 |  |  | SEE | 0.2019 | 0.9336
|  |  | r2 | 0.9998 | 0.998 |  |  |  | r2 | 0.9955 | 0.998 | Fail |  | r2 | 0.1061 | 0.998 | Fail |  | r2 | 0.9999 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-07-16 | 10:00 | 0.01 | 0 | 0 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.00058 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 12.9 | 951 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0005 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 27.9872 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0.1 | 0.1 | 0 | 0.0202 | 0.1 | 12.9 | 951 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 27.9104 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0005 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 55 | 1048 | 1012 | 27.8336 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0.1 | 0.1 | 0 | 0.0184 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3206 | 0.0005 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32.897 | 55 | 1048 | 1012 | 27.9294 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 28.0265 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1011 | 28.0503 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1011 | 28.0728 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.1 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1048 | 1012 | 28.0118 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0.08008 | 0.1 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3801 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2422 | 27.7 | 0.1
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2425 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.4 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1657 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3802 | 0.0004 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0004 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.2293 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.00038 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1756 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0183 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0951 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0282 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0487 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.3 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.1023 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1791 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3204 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2559 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2346 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0906 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0137 | 27.7 | 0
2008-07-16 | 10:00 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0626 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | 0 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0607 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3802 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.198 | 28.0163 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0934 | 27.7 | 0
2008-07-16 | 10:01 | 0.004 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2396 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2534 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1527 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0014 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.154 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0923 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.2 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.1996 | 28.0616 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.2 | 28.1925 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1155 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2228 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3803 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2236 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.07 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0003 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 27.9163 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0002 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1747 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1636 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0223 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | -0.1 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.099 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0242 | 27.7 | 0
2008-07-16 | 10:01 | 0 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0526 | 27.7 | 0
2008-07-16 | 10:01 | 0.004 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0706 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 2e-05 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1197 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1268 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.1 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.7956 | 28.0267 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0 | 0 | 0 | -0.2 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1805 | 27.7 | 0
2008-07-16 | 10:01 | 0.01 | 0 | 0 | 0.1 | 0 | 0.1 | -0.2 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.7972 | 28.1329 | 27.7 | 0
2008-07-16 | 10:01 | 0.085 | 0.0008 | 8.3828 | 0.1 | 0 | 0.1 | -0.2 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0002 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1877 | 27.7 | 0
2008-07-16 | 10:01 | 0.536 | 0.0018 | 17.5063 | 2.4 | 0 | 2.4 | 0.2 | 0.2384 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.000656 | 0.0002 | 0.0001 | 8e-05 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2587 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 20.2 | 0 | 20.2 | 3.9 | 3.892 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.00428 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.8004 | 28.1051 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 24.4 | 0 | 24.4 | 7.1 | 7.1032 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0051 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 27.3 | 0 | 27.3 | 8.6 | 8.5832 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.00572 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1988 | 28.1 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 28.5 | 0 | 28.5 | 9 | 8.9721 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.006 | 0.0002 | 0.0001 | 2e-05 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0722 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 28.6 | 0 | 28.6 | 9.1 | 9.1 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.006 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0047 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 28.6 | 0 | 28.6 | 9.2 | 9.2 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.006 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0814 | 27.7 | 0
2008-07-16 | 10:01 | 0.56 | 0.001 | 10 | 28.7 | 0 | 28.7 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.006 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2164 | 27.7 | 0
2008-07-16 | 10:01 | 0.577 | 0.001 | 10 | 29 | 0 | 29 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.00602 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.335 | 27.7 | 0
2008-07-16 | 10:01 | 0.59 | 0.0018 | 17.5342 | 29.3 | 0.1 | 29.4 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0002 | 0.0001 | 8e-05 |  | n/a |  | 33 | 55 | 1049 | 1011.7988 | 28.3524 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0452 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.589 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0655 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.198 | 28.1422 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0002 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.181 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.00018 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1041 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.002 | 20 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.8016 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.0032 | 31.9293 | 29.4 | 0.1 | 29.5 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0011 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.004 | 40 | 29.5 | 0.1 | 29.6 | 9.3 | 9.3 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3796 | 0.0061 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 27.9 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.0035 | 35.0712 | 29.5 | 0.1 | 29.6 | 9.3 | 9.3 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 2e-05 | 2e-05 |  | n/a |  | 33 | 55 | 1049 | 1011 | 27.9061 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.58 | 0.002 | 20 | 29.5 | 0.1 | 29.6 | 9.2 | 9.2007 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0598 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.586 | 0.002 | 20 | 29.5 | 0.1 | 29.6 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.5 | 0.1 | 29.6 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1011.2 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.00612 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1012.1992 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3196 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1013.2063 | 28.135 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3794 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2886 | 27.7 | 0.1
2008-07-16 | 10:01 | 0.59 | 0.002 | 20 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0.0001 | 0.0001 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0866 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 8e-05 | 8e-05 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0959 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.0023 | 22.6889 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1507 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.003 | 30 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0024 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.003 | 30 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7988 | 28.1552 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.003 | 30 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2545 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.59 | 0.003 | 30 | 29.6 | 0.1 | 29.7 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2055 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.589 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0085 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.581 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0855 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3206 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2241 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3798 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.145 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.7972 | 27.9628 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1996 | 28.2693 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8004 | 28.3 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1617 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.1996 | 28.0231 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.58 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8008 | 28.0999 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.576 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2535 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.003 | 30 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1929 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.0026 | 25.8911 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1304 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1856 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.032 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.7 | 0.1 | 29.8 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1217 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1624 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8004 | 28.1289 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2825 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 27.9598 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 27.8897 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8044 | 28.0217 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.57 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.0985 | 27.7 | 0.1
2008-07-16 | 10:02 | 0.679 | 0.002 | 20 | 29.8 | 0.1 | 29.9 | 9.2 | 9.2 | 0.1 | 13 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1753 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.03 | 0.0024 | 24.1021 | 30 | 0.1 | 30.1 | 9.2 | 9.2 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.006241 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7952 | 28.2521 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.13 | 0.003 | 30 | 45.3 | 0.1 | 45.4 | 11.6 | 11.5769 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.009243 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2134 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.003 | 30 | 54.4 | 0.1 | 54.5 | 15.4 | 15.3994 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.01106 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0171 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.003 | 30 | 55.7 | 0.1 | 55.8 | 17.4 | 17.4389 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.01132 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2475 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.0037 | 37.35 | 57.6 | 0.1 | 57.7 | 18.1 | 18.0722 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0117 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2407 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 57.8 | 0.1 | 57.9 | 18.3 | 18.2784 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.01172 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1639 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.2 | 0.1 | 58.3 | 18.4 | 18.3769 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.0871 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.3 | 0.1 | 58.4 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.0103 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.4 | 0.1 | 58.5 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.1995 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.4 | 0.1 | 58.5 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.7 | 0.1 | 58.8 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3206 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2598 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.8 | 0.1 | 58.9 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1061 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0262 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1518 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2452 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1346 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1771 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.0235 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1952 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 58.9 | 0.1 | 59 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2581 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7988 | 28.2187 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2955 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2277 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1017 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0028 | 27.7 | 0.1
2008-07-16 | 10:02 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3195 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1952 | 28.0796 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2127 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.29 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3202 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2132 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.1 | 59.1 | 18.4 | 18.4209 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1364 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59 | 0.08 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.4781 | 0.1 | 12.9 | 956 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.1173 | 27.7 | 0
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.1972 | 28.1942 | 27.7 | 0
2008-07-16 | 10:03 | 1.14 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2 | 27.7 | 0
2008-07-16 | 10:03 | 1.146 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2478 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2268 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0548 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0684 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.024 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2552 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.4 | 18.4216 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1758 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1388 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.4791 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2153 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2924 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | 0 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3694 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.020635 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.215 | 27.7 | 0
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.023 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0999 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3798 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.1996 | 28.0464 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.15 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0073 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.266 | 0.004 | 40 | 59.1 | -0.1 | 59.1 | 18.5 | 18.5 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0841 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.599 | 0.0048 | 48.0945 | 59.1 | -0.1 | 59.1 | 18.4 | 18.4 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.161 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.679 | 0.006 | 60 | 74.2 | -0.1 | 74.2 | 20.8 | 20.8318 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.014914 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2378 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.69 | 0.006 | 60 | 84.5 | -0.1 | 84.5 | 24.8 | 24.8266 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.016921 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.2562 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.701 | 0.006 | 60 | 85.2 | -0.1 | 85.2 | 26.8 | 26.77 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.017 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0258 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.714 | 0.006 | 60 | 85.7 | -0.1 | 85.7 | 27.5 | 27.4562 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.01712 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0683 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 86.3 | -0.1 | 86.3 | 27.6 | 27.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.01722 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0549 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 86.6 | -0.1 | 86.6 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0173 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 86.6 | -0.1 | 86.6 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049.7197 | 1013 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 86.7 | -0.1 | 86.7 | 27.8 | 27.7843 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.151 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 86.7 | -0.1 | 86.7 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0431 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87 | -0.1 | 87 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 27.9291 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.1 | -0.1 | 87.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.1 | -0.1 | 87.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.1 | -0.1 | 87.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1014 | 28 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.2 | -0.1 | 87.2 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3204 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.1094 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.2 | -0.1 | 87.2 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3799 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1049.2616 | 1014 | 28.2601 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.2 | -0.1 | 87.2 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.0295 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.2 | -0.1 | 87.2 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.067 | 27.7 | -0.1
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.3 | 0 | 87.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0562 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.3 | 0 | 87.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3198 | 0.0174 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0206 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.5 | 0 | 87.5 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3803 | 0.01742 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0974 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 87.8 | 0 | 87.8 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.01752 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.0257 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.3 | 0 | 88.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3802 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.802 | 27.9489 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3206 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0115 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3796 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.2859 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3794 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.0553 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1168 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 4.6 | 195 | 59.3 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2353 | 27.7 | 0
2008-07-16 | 10:03 | 1.72 | 0.006 | 60 | 88.4 | 0 | 88.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 4.4 | 195 | 59.3 | 0.0176 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.2759 | 27.7 | 0
2008-07-16 | 10:03 | 1.722 | 0.006 | 60 | 88.8 | 0 | 88.8 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.32 | 0.0177 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1223 | 27.7 | 0
2008-07-16 | 10:04 | 1.729 | 0.006 | 60 | 89.1 | 0 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.01772 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1011 | 28.0343 | 27.7 | 0
2008-07-16 | 10:04 | 1.721 | 0.006 | 60 | 89.1 | 0 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.1988 | 28.085 | 27.7 | 0
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.1 | -0.1 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.1 | -0.1 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.1 | -0.1 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1049 | 1012 | 28.2 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.722 | 0.006 | 60 | 89.1 | -0.1 | 89.1 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4 | 195 | 59.3797 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1502 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.729 | 0.006 | 60 | 89.2 | -0.1 | 89.2 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.6734 | -0.1
2008-07-16 | 10:04 | 1.721 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.7 | 27.7213 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.6034 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.6803 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1571 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.724 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1661 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.73 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049.7415 | 1013.1988 | 28.1108 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.729 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3208 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.7921 | 28.1876 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.721 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0068 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0237 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.182 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.4 | -0.1 | 89.4 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1051 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2433 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2515 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4.6 | 195 | 59.38 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.2 | 28.2253 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.8 | 27.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 4.4 | 195 | 59.32 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2978 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.006 | 60 | 89.3 | -0.1 | 89.3 | 27.7 | 27.711 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.221 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.0067 | 67.2627 | 89.3 | -0.11972 | 89.3 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2 | 27.7 | -0.1
2008-07-16 | 10:04 | 1.72 | 0.0064 | 64.2578 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3209 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1674 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.72 | 0.0064 | 64.0129 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3206 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1189 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.72 | 0.007 | 70 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2726 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.727 | 0.007 | 70 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.2044 | 28.3 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.724 | 0.007 | 70 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3203 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1805 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.72 | 0.007 | 70 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0164 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.729 | 0.007 | 70 | 89.4 | -0.2 | 89.4 | 27.7 | 27.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0928 | 27.7 | -0.2
2008-07-16 | 10:04 | 1.786 | 0.0074 | 74.0032 | 89.4 | -0.2 | 89.4 | 27.8 | 27.782 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0178 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.802 | 28.1698 | 27.6302 | -0.2
2008-07-16 | 10:04 | 2.261 | 0.0078 | 77.8833 | 96 | -0.2 | 96 | 28 | 28.0389 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.019059 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1063 | 27.6468 | -0.2
2008-07-16 | 10:04 | 2.31 | 0.007 | 70 | 111.7 | -0.2 | 111.7 | 31.7 | 31.7466 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.021258 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0239 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.31 | 0.007 | 70 | 114.7 | -0.2 | 114.7 | 35 | 35.038 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0226 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.0991 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.31 | 0.0077 | 77.3503 | 116.8 | -0.2 | 116.8 | 36.4 | 36.3822 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.02292 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0227 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.31 | 0.008 | 80 | 117.3 | -0.2 | 117.3 | 36.8 | 36.7762 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.023021 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0541 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.31 | 0.008 | 80 | 117.3 | -0.3 | 117.3 | 37 | 36.9775 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0231 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0689 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.31 | 0.008 | 80 | 117.4 | -0.3 | 117.4 | 37 | 37 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0231 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.31 | 0.008 | 80 | 117.9 | -0.3 | 117.9 | 37 | 37.0145 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.02312 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.327 | 0.008 | 80 | 118.3 | -0.3 | 118.3 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0232 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8008 | 28.123 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.34 | 0.008 | 80 | 118.4 | -0.28 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.2 | 28.123 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.34 | 0.008 | 80 | 118.4 | -0.2 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0463 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.34 | 0.008 | 80 | 118.4 | -0.2 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2767 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.331 | 0.008 | 80 | 118.4 | -0.2 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2309 | 27.7 | -0.2
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.4 | -0.2 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.246 | 27.654 | -0.2
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.4 | -0.3 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2772 | 27.6228 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.4 | -0.3 | 118.4 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7956 | 28.2004 | 27.6996 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.5 | -0.3 | 118.5 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3203 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.5 | -0.3 | 118.5 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2533 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.32 | 0.0233 | 0.00012 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1050 | 1012 | 28.2253 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.3 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3805 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3794 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0321 | 27.7 | -0.4
2008-07-16 | 10:04 | 2.33 | 0.008 | 80 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0895 | 27.7 | -0.4
2008-07-16 | 10:04 | 2.33 | 0.0087 | 87.2379 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7972 | 28.1568 | 27.7 | -0.4
2008-07-16 | 10:04 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7988 | 28.0032 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 12.9 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1996 | 28.2256 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37 | 37 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37 | 37 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2886 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.326 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37 | 37 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.135 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.32 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37 | 37 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2186 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.32 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37 | 37 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2639 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.32 | 0.009 | 90 | 118.7 | -0.4 | 118.7 | 37.1 | 37.0764 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1871 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.32 | 0.009 | 90 | 118.7 | -0.4 | 118.7 | 37.1 | 37.1 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2028 | 28.1103 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.326 | 0.009 | 90 | 118.7 | -0.4 | 118.7 | 37 | 37 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.7 | -0.4 | 118.7 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1406 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2942 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2493 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1724 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.6 | -0.4 | 118.6 | 37.1 | 37.1 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0233 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 118.7 | -0.4 | 118.7 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3794 | 0.023321 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 119.2 | -0.4 | 119.2 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.216 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.009 | 90 | 119.3 | -0.4 | 119.3 | 37.1 | 37.1 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3201 | 0.0235 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.323 | 0.009 | 89.6417 | 119.3 | -0.4 | 119.3 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3798 | 0.0235 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.326 | 0.0081 | 81.3083 | 119.3 | -0.4 | 119.3 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0235 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.3 | -0.4 | 119.3 | 37.1 | 37.1 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.02348 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1039 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.2 | -0.4 | 119.2 | 37 | 37.0236 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0843 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.2 | -0.4 | 119.2 | 37.1 | 37.1 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3197 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.162 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.2 | -0.4 | 119.2 | 37 | 37.0172 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0083 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.2 | -0.4 | 119.2 | 37.1 | 37.0616 | 0.1 | 13 | 955 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0727 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.2 | -0.4 | 119.2 | 37.1 | 37.0772 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1050 | 1013 | 28.1989 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.33 | 0.008 | 80 | 119.3 | -0.4 | 119.3 | 37 | 37 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.441 | 0.008 | 80 | 119.3 | -0.4 | 119.3 | 37 | 37 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192.4 | 28 | 5 | 195 | 59.3 | 0.0234 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.2028 | 28.2939 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.823 | 0.01 | 99.5356 | 121.1 | -0.4 | 121.1 | 37 | 37 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 193 | 28 | 5 | 195 | 59.3 | 0.023761 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1401 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.892 | 0.0128 | 128.0116 | 137.5 | -0.4 | 137.5 | 39.5 | 39.4822 | 0.1 | 13 | 956 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192.6 | 28 | 5 | 195 | 59.38 | 0.026139 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.907 | 0.0119 | 119.3823 | 144.2 | -0.4 | 144.2 | 43.4 | 43.4065 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.02798 | 0.0002 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012.2 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.91 | 0.0103 | 102.6878 | 146.6 | -0.4 | 146.6 | 45.3 | 45.2992 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.91 | 0.01 | 100 | 147.3 | -0.4 | 147.3 | 46 | 45.9584 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.02844 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.91 | 0.01 | 100 | 148.2 | -0.4 | 148.2 | 46.2 | 46.1814 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.02872 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2282 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.91 | 0.01 | 100 | 148.4 | -0.4 | 148.4 | 46.2 | 46.2 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3792 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2181 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.919 | 0.01 | 100 | 148.5 | -0.4 | 148.5 | 46.3 | 46.3 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1355 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.92 | 0.01 | 100 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3794 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2891 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.92 | 0.01 | 100 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.92 | 0.01 | 100 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2518 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.92 | 0.01 | 100 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1499 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.92 | 0.01 | 100 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.929 | 0.0102 | 102.1195 | 148.7 | -0.4 | 148.7 | 46.3 | 46.3 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.6 | -0.4 | 148.6 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0555 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.7 | -0.4 | 148.7 | 46.3 | 46.3 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.7 | -0.4 | 148.7 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1184 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2714 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.7 | -0.4 | 148.7 | 46.3 | 46.3 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3199 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2214 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.7 | -0.4 | 148.7 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1316 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3799 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2857 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1602 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.3792 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.193 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.4 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2307 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.8008 | 28.0005 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2316 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1383 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3792 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0306 | 27.7 | -0.4
2008-07-16 | 10:05 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1152 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.3 | 46.3 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2691 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.3842 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.177 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.0109 | 109.3585 | 148.7 | -0.4 | 148.7 | 46.5 | 46.4782 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28.603 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.8008 | 28.1767 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.0101 | 101.3392 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2546 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.0107 | 107.2059 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.1996 | 28.0242 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28.397 | 191 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.0687 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.0106 | 106.2541 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28.603 | 191 | 28 | 5 | 195 | 59.3801 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.8012 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.01 | 100 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 28.397 | 191 | 28 | 5 | 195 | 59.3198 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0776 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.01 | 100.2667 | 148.7 | -0.4 | 148.7 | 46.5 | 46.5 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.0008 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.0109 | 108.6 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.2281 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3205 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.5 | 46.5 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3797 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.7972 | 28.2406 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.4152 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.5 | 46.4892 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.38 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.93 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2203 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.936 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1050 | 1013 | 28.2261 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 148.8 | -0.4 | 148.8 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8008 | 28.1138 | 27.7 | -0.4
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 148.9 | -0.5 | 148.9 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3197 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1906 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 148.9 | -0.5 | 148.9 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1326 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1442 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3796 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7956 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3799 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1253 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.2031 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 110 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011.1996 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.011 | 109.7833 | 149 | -0.5 | 149 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3801 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.0101 | 101.45 | 148.9 | -0.5 | 148.9 | 46.4 | 46.4 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.4 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.94 | 0.01 | 100 | 148.9 | -0.5 | 148.9 | 46.4 | 46.3861 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2644 | 27.7 | -0.5
2008-07-16 | 10:06 | 2.995 | 0.01 | 100 | 148.8 | -0.5 | 148.8 | 46.3 | 46.3 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.802 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.236 | 0.01 | 100 | 149 | -0.5 | 149 | 46.3 | 46.3 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.486 | 0.01 | 100 | 156.8 | -0.5 | 156.8 | 47.1 | 47.0576 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.030305 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1339 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.53 | 0.01 | 100.1833 | 168.6 | -0.5 | 168.6 | 51.3 | 51.2739 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.032403 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1429 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.54 | 0.0109 | 108.5167 | 174.3 | -0.5 | 174.3 | 53.9 | 53.9368 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.03334 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.011 | 110 | 175.2 | -0.5 | 175.2 | 55 | 55.0385 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03352 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.011 | 110 | 176.1 | -0.5 | 176.1 | 55.4 | 55.3698 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03372 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2733 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.011 | 110 | 176.9 | -0.5 | 176.9 | 55.5 | 55.5 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03382 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.198 | 28.1996 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.011 | 110 | 177.5 | -0.5 | 177.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.03392 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1541 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.543 | 0.011 | 110 | 178.4 | -0.5 | 178.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0341 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.545 | 0.011 | 110 | 178.6 | -0.5 | 178.6 | 55.6 | 55.6121 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0341 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.0117 | 116.8443 | 178.6 | -0.5 | 178.6 | 55.6 | 55.6211 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03412 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 178.7 | -0.5 | 178.7 | 55.7 | 55.6776 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2657 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 178.8 | -0.5 | 178.8 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.7972 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.559 | 0.012 | 120 | 178.9 | -0.5 | 178.9 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.56 | 0.012 | 120 | 179 | -0.5 | 179 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2647 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.555 | 0.012 | 120 | 179 | -0.5 | 179 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179 | -0.5 | 179 | 55.7 | 55.6769 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8004 | 28.2633 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.1 | -0.5 | 179.1 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1095 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.1 | -0.51988 | 179.1 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2441 | 27.7 | -0.5
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.1 | -0.6 | 179.1 | 55.6 | 55.6393 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011.8016 | 28.3 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.2 | -0.6 | 179.2 | 55.6 | 55.6109 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3208 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.2487 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.1049 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.2585 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.03422 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1878 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.2012 | 28.1329 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.61996 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.32 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.1996 | 28.2098 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.68 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2866 | 27.7 | -0.7
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.2 | -0.6 | 179.2 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1099 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3199 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0402 | 27.7 | -0.6
2008-07-16 | 10:06 | 3.55 | 0.012 | 120 | 179.2 | -0.6 | 179.2 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.2 | -0.6 | 179.2 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3206 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0296 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1419 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2759 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.6 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.7988 | 28.2005 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.62 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.2775 | 27.7 | -0.6
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.3 | -0.7 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1364 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.558 | 0.012 | 120 | 179.3 | -0.7 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.094 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.56 | 0.012 | 120 | 179.3 | -0.7 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2767 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.56 | 0.012 | 120 | 179.3 | -0.7 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0455 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.556 | 0.0127 | 126.8182 | 179.3 | -0.7 | 179.3 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1237 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.013 | 130 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.0126 | 126.3468 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3208 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7919 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3203 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1539 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.7 | 55.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3798 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1534 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2302 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.293 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2162 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2144 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2912 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.55 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.57 | 0.012 | 120 | 179.5 | -0.7 | 179.5 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2137 | 27.7 | -0.7
2008-07-16 | 10:07 | 3.683 | 0.012 | 120 | 179.4 | -0.7 | 179.4 | 55.6 | 55.6 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3803 | 0.0343 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.9065 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.092 | 0.0123 | 123.4195 | 183.2 | -0.7 | 183.2 | 55.7 | 55.6967 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.034978 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.0504 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 201.6 | -0.6 | 201.6 | 59.5 | 59.5318 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.038159 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 204.7 | -0.62059 | 204.7 | 62.9 | 62.9022 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.038641 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 205.9 | -0.7 | 205.9 | 64.1 | 64.1454 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0389 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 206.4 | -0.67968 | 206.4 | 64.6 | 64.5808 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3203 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 206.5 | -0.6 | 206.5 | 64.7 | 64.6894 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.2779 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 206.5 | -0.6 | 206.5 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.3638 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.163 | 0.014 | 140 | 206.6 | -0.6 | 206.6 | 64.8 | 64.7901 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.7 | -0.6 | 206.7 | 64.8 | 64.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.8004 | 28.3 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.8 | -0.6 | 206.8 | 64.8 | 64.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012.1976 | 28.1668 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.8 | -0.6 | 206.8 | 64.8 | 64.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3801 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1434 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.8 | -0.6 | 206.8 | 64.8 | 64.7924 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.039 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.9 | -0.6 | 206.9 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3196 | 0.03902 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8036 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 206.9 | -0.6 | 206.9 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0391 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207 | -0.6 | 207 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0391 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.1 | -0.62036 | 207.1 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0391 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0391 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2042 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.6 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.0391 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2811 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.6 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.03912 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.6 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0392 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.6
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 958 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2885 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2116 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.1348 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.16 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.168 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3204 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1011.2044 | 28.1376 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3201 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012 | 28.2912 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1551 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.32 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1996 | 28.1 | 27.7 | -0.7
2008-07-16 | 10:07 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1521 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3198 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.2435 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3208 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2333 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.0147 | 146.7143 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3794 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.015 | 150 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3204 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.0146 | 146.4309 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3799 | 0.0392 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.2012 | 28.2362 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.162 | 0.0142 | 141.6653 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.2012 | 28.2406 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.015 | 149.9917 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013.8004 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.165 | 0.015 | 150 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.015 | 150 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.1996 | 28.229 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.015 | 150 | 207.1 | -0.7 | 207.1 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.2478 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.015 | 150 | 207.1 | -0.7 | 207.1 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3803 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.226 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.0148 | 148.28 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0046 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2341 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.17 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3206 | 0.0393 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2451 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.166 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3796 | 0.0393 | 0.00042 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3201 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1837 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.202 | 28.0296 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1867 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2221 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1154 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1925 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1305 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.4 | -0.7 | 207.4 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3206 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1463 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.177 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.2044 | 28.1001 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3203 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.1768 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.2537 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.014 | 140 | 207.3 | -0.7 | 207.3 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1050 | 1014 | 28.3 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.16 | 0.0144 | 143.5525 | 207.2 | -0.7 | 207.2 | 64.7 | 64.7 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0393 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2926 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.61 | 0.0148 | 148.3788 | 209.2 | -0.6802 | 209.2 | 64.6 | 64.6213 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.039359 | 0.0005 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1014 | 28.2158 | 27.7 | -0.7
2008-07-16 | 10:08 | 4.741 | 0.014 | 140.1075 | 222.3 | -0.6 | 222.3 | 66.6 | 66.6011 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.041877 | 0.0005 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1014 | 28.139 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.014 | 140 | 230.1 | -0.6 | 230.1 | 70.8 | 70.8208 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.04316 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1756 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.0147 | 146.8425 | 231.6 | -0.6 | 231.6 | 72.8 | 72.8345 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.04342 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2854 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.015 | 150 | 232.5 | -0.6 | 232.5 | 73.5 | 73.5445 | 0.1 | 13 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0436 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2086 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.015 | 150 | 233.5 | -0.6 | 233.5 | 73.8 | 73.8 | 0.1 | 13 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.04382 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014.2044 | 28.2 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.015 | 150 | 234.4 | -0.6 | 234.4 | 73.9 | 73.9 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0439 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.155 | 27.7 | -0.6
2008-07-16 | 10:08 | 4.77 | 0.015 | 150 | 234.5 | -0.5 | 234.5 | 74 | 73.979 | 0.1 | 13 | 957 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.04392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1 | 27.7 | -0.5
2008-07-16 | 10:08 | 4.77 | 0.015 | 150 | 234.7 | -0.5 | 234.7 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3201 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1 | 27.7 | -0.5
2008-07-16 | 10:08 | 4.773 | 0.0157 | 156.6807 | 234.9 | -0.5 | 234.9 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3798 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1755 | 27.7 | -0.5
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 235.5 | -0.48 | 235.5 | 74 | 74 | 0.1 | 13 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1477 | 27.7 | -0.5
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 235.8 | -0.4 | 235.8 | 74 | 74 | 0.1 | 13 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 235.9 | -0.4 | 235.9 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1118 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.1 | -0.4 | 236.1 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2654 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.2 | -0.4 | 236.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.1984 | 28.181 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.2 | -0.4 | 236.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3794 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7941 | 28.1727 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.4 | -0.4 | 236.4 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2606 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.5 | -0.4 | 236.5 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.03 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.6 | -0.4 | 236.6 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.04422 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0668 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 236.8 | -0.4 | 236.8 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0563 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.785 | 0.016 | 160 | 236.9 | -0.4 | 236.9 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.041 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 236.9 | -0.4 | 236.9 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1946 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 236.9 | -0.4 | 236.9 | 74 | 74 | 0.1 | 13 | 957 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 236.9 | -0.4 | 236.9 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013.1972 | 28.2509 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 236.9 | -0.4 | 236.9 | 74 | 74 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2445 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 237 | -0.4 | 237 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3206 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1045 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.79 | 0.016 | 160 | 237 | -0.4 | 237 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1813 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.786 | 0.016 | 160 | 237 | -0.4 | 237 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1419 | 27.7 | -0.4
2008-07-16 | 10:08 | 4.78 | 0.016 | 160 | 237.1 | -0.4 | 237.1 | 74 | 74 | 0.1 | 13 | 957 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0651 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0353 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0443 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2657 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.04432 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.3 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2154 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0809 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1984 | 28.0041 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1455 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3203 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2028 | 28.2263 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2937 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74.1 | 74.1 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1401 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74 | 74 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2135 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.2 | -0.4 | 237.2 | 74.1 | 74.0824 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.2329 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.3 | -0.4 | 237.3 | 74 | 74 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1104 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.3 | -0.4 | 237.3 | 74.1 | 74.0799 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1872 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.4 | -0.38012 | 237.4 | 74.1 | 74.1 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3801 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1359 | 27.7 | -0.4
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74.1 | 74.1 | 0.1 | 13 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74 | 74.0105 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3794 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.78 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74 | 74 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.2028 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.772 | 0.016 | 160 | 237.3 | -0.3 | 237.3 | 74.1 | 74.086 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0286 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74 | 74 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3201 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.32 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0499 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.1965 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.4 | -0.3 | 237.4 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0429 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.5 | -0.3 | 237.5 | 74 | 74.0166 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 237.7 | -0.3 | 237.7 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.04442 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0641 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 238.3 | -0.3 | 238.3 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.04452 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.775 | 0.016 | 160 | 238.7 | -0.2802 | 238.7 | 74.1 | 74.1 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0446 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:09 | 4.77 | 0.016 | 160 | 238.8 | -0.2 | 238.8 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3206 | 0.0446 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2626 | 27.7 | -0.2
2008-07-16 | 10:09 | 4.766 | 0.016 | 160 | 238.8 | -0.2 | 238.8 | 74.1 | 74.1 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0446 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.2215 | 27.7 | -0.2
2008-07-16 | 10:09 | 4.85 | 0.0163 | 163.39 | 238.8 | -0.2 | 238.8 | 74.1 | 74.1 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3799 | 0.0446 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.1316 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.268 | 0.0168 | 168.1863 | 241.4 | -0.2 | 241.4 | 74.2 | 74.1756 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3799 | 0.045041 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.2857 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.373 | 0.016 | 160.0163 | 257.9 | -0.2 | 257.9 | 77.4 | 77.3533 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.38 | 0.04778 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011.1996 | 28.2301 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.0168 | 168.1938 | 263.4 | -0.2 | 263.4 | 81.1 | 81.1364 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.04866 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.107 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 264.9 | -0.2 | 264.9 | 82.7 | 82.7014 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3 | 0.04892 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 265.7 | -0.2 | 265.7 | 83.2 | 83.2098 | 0.1 | 12.9 | 959 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0002 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 265.8 | -0.2 | 265.8 | 83.4 | 83.3818 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0772 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 265.9 | -0.2 | 265.9 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3802 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1984 | 28.1 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 265.8 | -0.2 | 265.8 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1307 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 265.9 | -0.2 | 265.9 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1924 | 27.7 | -0.2
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 266 | -0.1 | 266 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3203 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1155 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 266 | -0.1 | 266 | 83.5 | 83.4863 | 0.1 | 12.9 | 959 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9158 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 266.2 | -0.1 | 266.2 | 83.4 | 83.4211 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9147 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 266.2 | -0.1 | 266.2 | 83.5 | 83.4998 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0849 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.38 | 0.017 | 170 | 266.3 | -0.1 | 266.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0081 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.385 | 0.017 | 170 | 266.7 | -0.1 | 266.7 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0492 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9313 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.0177 | 176.6071 | 266.7 | -0.1 | 266.7 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0492 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0819 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.1 | -0.1 | 267.1 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2331 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.2 | -0.1 | 267.2 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0025 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.2 | -0.1 | 267.2 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012 | 28 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.2 | -0.1 | 267.2 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.8945 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.3 | -0.1 | 267.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012.2028 | 27.9184 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.3 | -0.1 | 267.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1871 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.2 | -0.1 | 267.2 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0493 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0334 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.39 | 0.018 | 180 | 267.4 | -0.1 | 267.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.04932 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.8 | 28.1804 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.398 | 0.018 | 180 | 267.7 | -0.1 | 267.7 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.38 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.1892 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.4 | 0.018 | 180 | 267.7 | -0.1 | 267.7 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3801 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28 | 27.7 | -0.1
2008-07-16 | 10:09 | 5.4 | 0.018 | 180 | 267.8 | 0 | 267.8 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.8012 | 28 | 27.7 | 0
2008-07-16 | 10:09 | 5.4 | 0.018 | 180 | 267.9 | 0 | 267.9 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3198 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 27.9326 | 27.7 | 0
2008-07-16 | 10:09 | 5.4 | 0.018 | 180 | 268 | 0 | 268 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3198 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 27.9442 | 27.7 | 0
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268 | 0 | 268 | 83.4 | 83.4193 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3796 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2044 | 28 | 27.7 | 0
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.1 | 0 | 268.1 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3797 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28 | 27.7 | 0
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.1 | 0 | 268.1 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1493 | 27.7 | 0
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.38 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1485 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.1 | 0.1 | 268.2 | 83.4 | 83.4 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0433 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.5 | 83.4998 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.04942 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9051 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011.8024 | 27.982 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.4 | 0.1 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0588 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32.8969 | 55 | 1050 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.4 | 0.1 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3204 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.4 | 0.1 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3796 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012 | 28.0339 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.4 | 0.1 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3201 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012 | 28 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.392 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.0197 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.39 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.0965 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.39 | 0.018 | 180 | 268.3 | 0.1 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.32 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1733 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.39 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3801 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2502 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.391 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1648 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3792 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.8077 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.1 | 268.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3802 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.9614 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.12059 | 268.3 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3206 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1151 | 27.7 | 0.1
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.2 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7956 | 28.1313 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.3 | 0.2 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.9889 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.3 | 0.2 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.912 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.3 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 27.9648 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.3 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 27.9584 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.2 | 0.3 | 268.5 | 83.5 | 83.5 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3802 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.198 | 27.9737 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.1 | 0.3 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.281 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.4 | 0.018 | 180 | 268.1 | 0.3 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.1557 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.476 | 0.018 | 180 | 268.1 | 0.3 | 268.4 | 83.5 | 83.5 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3803 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.849 | 0.0183 | 183.0782 | 268.1 | 0.3 | 268.4 | 83.5 | 83.5332 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3794 | 0.0495 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0742 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.899 | 0.019 | 190 | 280.9 | 0.2 | 281.1 | 85.3 | 85.2508 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.051524 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.93 | 0.019 | 190 | 292.2 | 0.2 | 292.4 | 89.2 | 89.225 | 0.1 | 12.9 | 959 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.05332 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7972 | 28 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.945 | 0.019 | 190 | 294.1 | 0.2 | 294.3 | 91.6 | 91.6076 | 0.1 | 12.9 | 960 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.05362 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2028 | 28.1126 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.953 | 0.019 | 190 | 294.7 | 0.2 | 294.9 | 92.5 | 92.4511 | 0.1 | 12.9 | 960 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0537 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.962 | 0.019 | 190 | 294.7 | 0.2 | 294.9 | 92.8 | 92.7669 | 0.1 | 12.9 | 960 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0537 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.8004 | 28.1802 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.97 | 0.019 | 190 | 294.8 | 0.2 | 295 | 92.8 | 92.8 | 0.1 | 12.9 | 960 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0537 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012 | 28.0266 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.978 | 0.019 | 190 | 294.8 | 0.2 | 295 | 92.9 | 92.9 | 0.1 | 12.9 | 960 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0537 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 27.9364 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.987 | 0.019 | 190 | 295.6 | 0.2 | 295.8 | 92.9 | 92.9 | 0.1 | 12.9 | 960 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.05382 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.8008 | 28.0212 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.019 | 190 | 296.3 | 0.2 | 296.5 | 92.9 | 92.9 | 0.1 | 12.9 | 959 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.2398 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012 | 28.1828 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.019 | 190 | 296.4 | 0.2 | 296.6 | 92.9 | 92.9 | 0.1 | 12.9 | 959 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1050 | 1012 | 28.1059 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.0195 | 194.888 | 296.4 | 0.2 | 296.6 | 92.9 | 92.9 | 0.1 | 12.9 | 959 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1050 | 1012 | 28.0291 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.02 | 200 | 296.4 | 0.2 | 296.6 | 92.9 | 92.9 | 0.1 | 12.9 | 959 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012 | 28.1433 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.02 | 200 | 296.4 | 0.2 | 296.6 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.3 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.02 | 200 | 296.5 | 0.2 | 296.7 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.054 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2972 | 27.7 | 0.2
2008-07-16 | 10:10 | 5.99 | 0.02 | 200 | 296.7 | 0.3 | 297 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.05402 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1445 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.99 | 0.02 | 200 | 296.7 | 0.3 | 297 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.3
2008-07-16 | 10:10 | 5.991 | 0.02 | 200 | 296.7 | 0.3 | 297 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.1171 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1012 | 28.2699 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1144 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.38 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0778 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2158 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3205 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2048 | 28.2923 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.8 | 0.3 | 297.1 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0923 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 296.9 | 0.3 | 297.2 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1387 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 297 | 0.3 | 297.3 | 92.9 | 92.9 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2073 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 297.1 | 0.28 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 27.9003 | 27.7 | 0.3
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 27.9 | 27.7 | 0.2
2008-07-16 | 10:10 | 6 | 0.02 | 200 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3801 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0075 | 27.7 | 0.2
2008-07-16 | 10:11 | 6 | 0.02 | 200 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8095 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3208 | 0.0541 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0694 | 27.6694 | 0.2
2008-07-16 | 10:11 | 6 | 0.02 | 200 | 297.3 | 0.2 | 297.5 | 92.8 | 92.819 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3196 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0074 | 27.6075 | 0.2
2008-07-16 | 10:11 | 6.007 | 0.02 | 200 | 297.4 | 0.2 | 297.6 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0843 | 27.6843 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.02 | 200 | 297.4 | 0.2 | 297.6 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1611 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.02 | 200 | 297.3 | 0.2 | 297.5 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.02 | 200 | 297.3 | 0.2 | 297.5 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3799 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1852 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.02 | 200 | 297.3 | 0.2 | 297.5 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013.1996 | 28.1084 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.0199 | 198.5 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1014 | 28.2369 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190.0254 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 4.6 | 195 | 59.3198 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1014.198 | 28.2547 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 4.4 | 195 | 59.3199 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014.8008 | 28.2221 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1015 | 28.2989 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1485 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3794 | 0.0542 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1014 | 28.2051 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.3 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3797 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013.7972 | 28.2876 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.2028 | 28.1339 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.1 | 0.2 | 297.3 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0401 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.38 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1102 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2198 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 958 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3198 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.3 | 0.3 | 297.6 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.3 | 27.7 | 0.3
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.3 | 297.5 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.1972 | 28.2327 | 27.7 | 0.3
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3198 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1118 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 12.9 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0627 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2931 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.2 | 0.2 | 297.4 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3203 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.2255 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.01 | 0.019 | 190 | 297.3 | 0.2 | 297.5 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3797 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1487 | 27.7 | 0.2
2008-07-16 | 10:11 | 6.004 | 0.019 | 190 | 297.4 | 0.2 | 297.6 | 92.9 | 92.8777 | 0.1 | 13 | 957 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3799 | 0.0542 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.7988 | 28.1282 | 27.7 | 0.2
2008-07-16 | 10:11 | 6 | 0.019 | 190 | 297.5 | 0.2 | 297.7 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0543 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 6 | 0.019 | 190 | 297.5 | 0.2 | 297.7 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.0543 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.96 | 0.0195 | 194.7664 | 297.5 | 0.2 | 297.7 | 92.8 | 92.8 | 0.1 | 13 | 956 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0543 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.1996 | 28.2586 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.731 | 0.0197 | 196.757 | 297.5 | 0.2 | 297.7 | 92.8 | 92.8 | 0.1 | 13 | 956 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0543 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2645 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.489 | 0.0192 | 191.5 | 292.1 | 0.2 | 292.3 | 92.1 | 92.0836 | 0.1 | 13 | 957 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.053442 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.43 | 0.02 | 199.9746 | 276.2 | 0.2 | 276.4 | 87.8 | 87.7687 | 0.1 | 13 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.050961 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.43 | 0.02 | 200 | 272.2 | 0.2 | 272.4 | 85.1 | 85.1492 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.050261 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.8036 | 28.2 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.427 | 0.0194 | 193.7354 | 271.3 | 0.2 | 271.5 | 84 | 83.9863 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0501 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1572 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.019 | 190 | 271 | 0.2 | 271.2 | 83.6 | 83.6315 | 0.1 | 13 | 951 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0501 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1196 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.019 | 190 | 270.8 | 0.2 | 271 | 83.5 | 83.5 | 0.1 | 13 | 951 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0501 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.7972 | 28.1965 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.0188 | 188.4987 | 270.2 | 0.2 | 270.4 | 83.5 | 83.5 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3202 | 0.05 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.2016 | 28.1267 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.018 | 180.017 | 269.7 | 0.2 | 269.9 | 83.5 | 83.5 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.04988 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.8004 | 28.1 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.0172 | 171.9904 | 269.4 | 0.2 | 269.6 | 83.5 | 83.5 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1269 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.017 | 170 | 269.2 | 0.2 | 269.4 | 83.4 | 83.4 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.1962 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.017 | 170 | 269.3 | 0.2 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1194 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.42 | 0.0173 | 173.0277 | 269.2 | 0.2 | 269.4 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.1972 | 28.1574 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.412 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.4 | 83.4 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2343 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013.7956 | 28.2668 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0362 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3203 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1295 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2416 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3202 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1050 | 1013 | 28.2449 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 952 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0145 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.3 | 269.5 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2161 | 27.7 | 0.3
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2512 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.1 | 0.2 | 269.3 | 83.4 | 83.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2513 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3902 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0825 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.3 | 83.301 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192.4 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1012 | 28.0562 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.1 | 0.2 | 269.3 | 83.3 | 83.3195 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192.6 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1661 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.2 | 0.2 | 269.4 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192.4 | 28 | 5 | 195 | 59.3203 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.2028 | 28.2803 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.1 | 0.2 | 269.3 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 193 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1267 | 27.7 | 0.2
2008-07-16 | 10:11 | 5.41 | 0.018 | 180 | 269.1 | 0.1 | 269.2 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 193 | 28 | 5 | 195 | 59.4 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.227 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.9 | 0.1 | 269 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 193 | 28 | 5 | 195 | 59.3 | 0.0498 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2194 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.9 | 0.11992 | 269 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192.6 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1171 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.9 | 0.2 | 269.1 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1939 | 27.7 | 0.2
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.8 | 0.18016 | 269 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.059 | 27.7 | 0.2
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.8 | 0.1 | 268.9 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0948 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.8 | 0.1 | 268.9 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1516 | 27.6758 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.7 | 0.1 | 268.8 | 83.3 | 83.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3796 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012.796 | 28.0006 | 27.6006 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.8 | 0.1 | 268.9 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1012 | 28.0777 | 27.6777 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.7 | 0.1 | 268.8 | 83.3 | 83.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1547 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.7 | 0.1 | 268.8 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1371 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0159 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.17 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.076 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0497 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.0388 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3208 | 0.0497 | 0.000721 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1153 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.41 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1923 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.407 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.2 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1538 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.018 | 180 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1013 | 28.1 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.018 | 180 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1050 | 1013 | 27.9463 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.018 | 180 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 27.9536 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.0174 | 173.7277 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1014 | 28.0305 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28.0926 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1050 | 1014 | 28.0158 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3208 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.061 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28.0622 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.5 | 0.1 | 268.6 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1014 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.4 | 0.017 | 170 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1050 | 1013 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 5.323 | 0.017 | 170 | 268.6 | 0.1 | 268.7 | 83.3 | 83.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1050 | 1013 | 28.1366 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.887 | 0.0167 | 166.8129 | 268.7 | 0.1 | 268.8 | 83.3 | 83.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0497 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.1996 | 28.1097 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.852 | 0.016 | 160 | 261.1 | 0.1 | 261.2 | 81.7 | 81.6769 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.38 | 0.04848 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.022 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.831 | 0.016 | 160 | 245.6 | 0.1 | 245.7 | 77.2 | 77.204 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3802 | 0.046001 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0988 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.82 | 0.016 | 160 | 242.4 | 0.1 | 242.5 | 75 | 75.0438 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.04548 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1757 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.808 | 0.016 | 160 | 241.8 | 0.1 | 241.9 | 74.4 | 74.3562 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0454 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1475 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 241.5 | 0.1 | 241.6 | 74.1 | 74.1204 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3802 | 0.0453 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1293 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 241.3 | 0.1 | 241.4 | 74 | 74.0205 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.045279 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7937 | 28.1816 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.7 | 0.1 | 240.8 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3794 | 0.0452 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9512 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.6 | 0.1 | 240.7 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3796 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 27.9598 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.6 | 0.1 | 240.7 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.5 | 0.1 | 240.6 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.0315 | 1012 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.5 | 0.1 | 240.6 | 73.9 | 73.9108 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1992 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.4 | 0.1 | 240.5 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.4 | 0.1 | 240.5 | 74 | 73.9783 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3801 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0879 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 240.3 | 0.1 | 240.4 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3803 | 0.0451 | 0.0008 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.1792 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.771 | 0.016 | 160 | 239.5 | 0.1 | 239.6 | 73.9 | 73.9169 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.04498 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1023 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.763 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 74 | 73.9795 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.249 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3794 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1974 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 74 | 74 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1563 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.285 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0546 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9196 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1172 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1646 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0877 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.8024 | 28.0109 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.3 | 0.1 | 239.4 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1319 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.77 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3805 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1144 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.774 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3802 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0393 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1929 | 27.7 | 0.1
2008-07-16 | 10:12 | 4.78 | 0.016 | 160 | 239.2 | 0.1 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2733 | 27.7 | 0.1
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.1 | 0.1 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1014 | 28.3 | 27.7 | 0.1
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.1 | 0.1 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3798 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2191 | 27.7 | 0.1
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.1 | 0.1 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.0113 | 27.7 | 0.1
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.04492 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2419 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.045 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2426 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.04498 | 0.0008 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.1657 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.4 | 0 | 239.4 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.5 | 0 | 239.5 | 74 | 73.9826 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.00078 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.5 | 0 | 239.5 | 74 | 74 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3205 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2296 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.5 | 0 | 239.5 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2584 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.6 | 0 | 239.6 | 74 | 74 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32.8961 | 55 | 1049 | 1011 | 28.1631 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.5 | 0 | 239.5 | 74 | 74 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0094 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.5 | 0 | 239.5 | 73.9 | 73.9026 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3202 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.202 | 28 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.4 | 0 | 239.4 | 73.9 | 73.9176 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.4 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1011 | 28.0978 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.4 | 0 | 239.4 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.198 | 28.1485 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0052 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1588 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 951 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.2044 | 28.2 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.2 | 0 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1802 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.2 | 0 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0264 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.2 | 0 | 239.2 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3799 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.0636 | 27.7 | 0
2008-07-16 | 10:13 | 4.777 | 0.016 | 160 | 239.1 | 0 | 239.1 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.32 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.1405 | 27.7 | 0
2008-07-16 | 10:13 | 4.771 | 0.016 | 160 | 239.1 | 0 | 239.1 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011.2 | 28.1482 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239.1 | 0 | 239.1 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011.198 | 27.9177 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 239 | 0 | 239 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.0419 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 238.9 | 0 | 238.9 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3197 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0522 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 238.8 | 0 | 238.8 | 73.9 | 73.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3802 | 0.0449 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 238.8 | 0 | 238.8 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3792 | 0.0448 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0023 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 238.8 | 0 | 238.8 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2334 | 27.7 | 0
2008-07-16 | 10:13 | 4.78 | 0.016 | 160 | 238.7 | 0 | 238.7 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7956 | 28.1903 | 27.7 | 0
2008-07-16 | 10:13 | 4.392 | 0.016 | 160 | 238.7 | 0 | 238.7 | 73.9 | 73.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.04478 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011.7972 | 28.1 | 27.7 | 0
2008-07-16 | 10:13 | 4.226 | 0.016 | 160 | 231.3 | 0 | 231.3 | 72.6 | 72.6245 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3799 | 0.043599 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.1081 | 27.7 | 0
2008-07-16 | 10:13 | 4.19 | 0.016 | 160 | 216 | 0 | 216 | 68.7 | 68.6919 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.041 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.1851 | 27.7 | 0
2008-07-16 | 10:13 | 4.19 | 0.016 | 160 | 212 | 0 | 212 | 66.1 | 66.1169 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.38 | 0.04038 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1011 | 28.2 | 27.7 | 0
2008-07-16 | 10:13 | 4.19 | 0.016 | 160 | 211.7 | 0 | 211.7 | 65.2 | 65.1596 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.6886 | 1011.1996 | 28.1222 | 27.7 | 0
2008-07-16 | 10:13 | 4.19 | 0.016 | 160 | 210.9 | 0 | 210.9 | 64.8 | 64.8252 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0402 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.8151 | 1012 | 27.9846 | 27.7 | 0
2008-07-16 | 10:13 | 4.19 | 0.0159 | 158.9573 | 210.8 | 0 | 210.8 | 64.7 | 64.7 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.04018 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.6927 | 1012 | 27.9076 | 27.7 | 0
2008-07-16 | 10:13 | 4.184 | 0.015 | 150.4103 | 210.7 | 0 | 210.7 | 64.6 | 64.6223 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3803 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 27.9695 | 27.7 | 0
2008-07-16 | 10:13 | 4.18 | 0.015 | 150 | 210.7 | 0 | 210.7 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3208 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 27.9537 | 27.7 | 0
2008-07-16 | 10:13 | 4.18 | 0.015 | 150 | 210.6 | 0 | 210.6 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 27.969 | 27.7 | 0
2008-07-16 | 10:13 | 4.179 | 0.015 | 150 | 210.7 | 0 | 210.7 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3795 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.7952 | 28.1998 | 27.7 | 0
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.6 | 0 | 210.6 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.0463 | 27.7 | 0
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.6 | -0.01996 | 210.6 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.1075 | 27.7 | 0
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.7 | -0.1 | 210.7 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.2306 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.7 | -0.1 | 210.7 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1010 | 28.2925 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.7 | -0.1 | 210.7 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.2156 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.6 | -0.1 | 210.6 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.2613 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.17 | 0.015 | 150 | 210.5 | -0.1 | 210.5 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3803 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.1968 | 28.2238 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.178 | 0.0149 | 149.0017 | 210.4 | -0.1 | 210.4 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1964 | 28.1149 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.18 | 0.014 | 140.3959 | 210.2 | -0.1 | 210.2 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.2044 | 28.1917 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.18 | 0.014 | 140 | 210.1 | -0.12044 | 210.1 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 27.9943 | 27.7 | -0.1
2008-07-16 | 10:13 | 4.172 | 0.014 | 140 | 210.1 | -0.2 | 210.1 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 27.9 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.15 | 0.0144 | 144.2369 | 210.1 | -0.2 | 210.1 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.04008 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 27.9444 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.15 | 0.015 | 150 | 210 | -0.2 | 210 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1014 | 28.098 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.158 | 0.015 | 150 | 210 | -0.2 | 210 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1014 | 28.1759 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.16 | 0.015 | 150 | 210.1 | -0.2 | 210.1 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013.802 | 28.2 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.16 | 0.015 | 150 | 210 | -0.2 | 210 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.2 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.16 | 0.015 | 150 | 210 | -0.2 | 210 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.2 | 27.7 | -0.2
2008-07-16 | 10:13 | 4.16 | 0.015 | 150 | 210 | -0.2 | 210 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.0147 | 147.3895 | 209.7 | -0.2 | 209.7 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.04 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0799 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.5 | -0.2 | 209.5 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1106 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.5 | -0.2 | 209.5 | 64.5 | 64.5182 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1049 | 1013 | 28.2726 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.5 | -0.2 | 209.5 | 64.6 | 64.5779 | 0.1 | 12.9 | 955 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013.2012 | 28.1189 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1014 | 28.1674 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.32 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.802 | 28.158 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0043 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.4 | -0.2 | 209.4 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1545 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.16 | 0.014 | 140 | 209.3 | -0.2 | 209.3 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3198 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2716 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.161 | 0.014 | 140 | 209.3 | -0.2 | 209.3 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3794 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1948 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 209.2 | -0.2 | 209.2 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014.2028 | 28.118 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1015 | 28.1588 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1643 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1125 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.38 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1893 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.0016 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.2 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0289 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.2 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3805 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2198 | 27.6802 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.2 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3196 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.2966 | 27.6034 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.2 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1529 | 27.6735 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.2 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.2 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3797 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.2028 | 28.1543 | 27.7 | -0.2
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.3 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.3 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.3 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.32 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013.2 | 28.127 | 27.7 | -0.3
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.8 | -0.3 | 208.8 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3802 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:14 | 4.17 | 0.014 | 140 | 208.9 | -0.3 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:14 | 4.121 | 0.014 | 140 | 208.9 | -0.3 | 208.9 | 64.6 | 64.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.8028 | 28 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.66 | 0.0135 | 134.7314 | 206.6 | -0.3 | 206.6 | 64.5 | 64.5188 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3792 | 0.039426 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1949 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.625 | 0.012 | 120 | 192.8 | -0.3 | 192.8 | 61.8 | 61.7682 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.037123 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2164 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.602 | 0.012 | 120 | 183.8 | -0.3 | 183.8 | 58.1 | 58.1186 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3796 | 0.035559 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.59 | 0.012 | 120 | 182.6 | -0.3 | 182.6 | 56.2 | 56.2354 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0353 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.587 | 0.012 | 120 | 182.3 | -0.3 | 182.3 | 55.6 | 55.6485 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.03528 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.58 | 0.0124 | 124.2053 | 182 | -0.3 | 182 | 55.4 | 55.4187 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0352 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.58 | 0.013 | 130 | 180.9 | -0.3 | 180.9 | 55.4 | 55.4 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.035 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.1521 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.572 | 0.013 | 130 | 180.5 | -0.3 | 180.5 | 55.3 | 55.3 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.035 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2886 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.2 | -0.3 | 180.2 | 55.3 | 55.3 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3802 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 27.9814 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.2 | -0.3 | 180.2 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0695 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3801 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1667 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3197 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1304 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3804 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012.1964 | 28.361 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 948 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3204 | 0.0349 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2723 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.57 | 0.013 | 130 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0349 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1594 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.566 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3799 | 0.0349 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0825 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.56 | 0.013 | 130 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0056 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.559 | 0.0126 | 125.8043 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2136 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.198 | 28.2039 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.558 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1498 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.554 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2966 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.554 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.143 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.557 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2104 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2682 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0348 | 0.0007 | -8e-05 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2082 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3203 | 0.0348 | 0.0007 | -2e-05 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2853 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3201 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.113 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.554 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0787 | 27.7 | -0.3
2008-07-16 | 10:14 | 3.557 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.6837 | -0.3
2008-07-16 | 10:14 | 3.55 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.6073 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3197 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2 | 27.6696 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1066 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 179.9 | -0.3 | 179.9 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.802 | 28.0471 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3796 | 0.0348 | 0.0007 | -8e-05 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1231 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.551 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.56 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.1308 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.56 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3202 | 0.0348 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2075 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.56 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.2845 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.556 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.3 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.2617 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1847 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.108 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.198 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.2808 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.0542 | 27.6542 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2186 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.7972 | 28 | 27.6226 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012.2012 | 28 | 27.6995 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1061 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.1996 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180 | -0.3 | 180 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3198 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1865 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0329 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.3 | 55.3 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3792 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.1 | -0.3 | 180.1 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3204 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.796 | 28.0142 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3796 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.091 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.8269 | 1012.2028 | 28.1679 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1013 | 28.1106 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1050 | 1012 | 28.043 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049.7324 | 1012 | 28.1968 | 27.7 | -0.3
2008-07-16 | 10:15 | 3.55 | 0.012 | 120 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2 | 27.6248 | -0.3
2008-07-16 | 10:15 | 3.2 | 0.0119 | 119.0894 | 180.2 | -0.3 | 180.2 | 55.2 | 55.2 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0348 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0959 | 27.652 | -0.3
2008-07-16 | 10:15 | 3.023 | 0.0111 | 110.5787 | 171.8 | -0.3 | 171.8 | 54.3 | 54.3404 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.033602 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 27.9422 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.97 | 0.011 | 110 | 157 | -0.3 | 157 | 50.3 | 50.2672 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3206 | 0.030697 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 27.8171 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.97 | 0.011 | 110 | 152.9 | -0.3 | 152.9 | 47.4 | 47.4332 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.029979 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0477 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.968 | 0.0106 | 105.8672 | 152.6 | -0.3 | 152.6 | 46.5 | 46.5082 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1594 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.952 | 0.01 | 100 | 152.6 | -0.3 | 152.6 | 46.1 | 46.1192 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0299 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2363 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 152.1 | -0.3 | 152.1 | 46 | 46 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3201 | 0.0298 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2606 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.8 | -0.3 | 151.8 | 45.9 | 45.9 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0298 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0302 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.8 | -0.3 | 151.8 | 45.9 | 45.9 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.38 | 0.0298 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0668 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.8 | -0.3 | 151.8 | 45.9 | 45.9 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0298 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0563 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.7 | -0.3 | 151.7 | 45.9 | 45.9 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0298 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.5 | -0.3 | 151.5 | 45.9 | 45.9 | 0.1 | 13 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0297 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.3 | -0.3 | 151.3 | 45.9 | 45.9 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0297 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.2 | -0.3 | 151.2 | 45.9 | 45.9 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0297 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.2 | -0.3 | 151.2 | 45.9 | 45.9 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.4 | 0.0297 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0556 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.2 | -0.3 | 151.2 | 45.9 | 45.9 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0297 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1906 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.2 | -0.3 | 151.2 | 45.9 | 45.9 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0369 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151.1 | -0.3 | 151.1 | 45.9 | 45.9 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.0584 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 151 | -0.3 | 151 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.7 | -0.3 | 150.7 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.7 | -0.3 | 150.7 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.8008 | 28.1658 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.8 | -0.3 | 150.8 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1574 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.7 | -0.3 | 150.7 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3803 | 0.0296 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1968 | 28.1194 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.6 | -0.3 | 150.6 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.02958 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1963 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.5 | -0.3 | 150.5 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3206 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1268 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.5 | -0.3 | 150.5 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.4 | -0.3 | 150.4 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2732 | 27.7 | -0.3
2008-07-16 | 10:15 | 2.95 | 0.01 | 100 | 150.4 | -0.3 | 150.4 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1962 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.95 | 0.01 | 100 | 150.4 | -0.3 | 150.4 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1194 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.942 | 0.01 | 100 | 150.4 | -0.3 | 150.4 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1014 | 28.0426 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 27.9314 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 27.8334 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.935 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0638 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.92 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 953 | 960 | 951 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0352 | 27.6352 | -0.3
2008-07-16 | 10:16 | 2.92 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.0833 | 27.6416 | -0.3
2008-07-16 | 10:16 | 2.92 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.2184 | 27.6816 | -0.3
2008-07-16 | 10:16 | 2.926 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.2953 | 27.6047 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.0113 | 27.6722 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.047 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2258 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2945 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1407 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3805 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 27.9306 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.934 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 27.9331 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.7956 | 28.1797 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.2044 | 28.026 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.3 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0295 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 27.9362 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.2 | -0.3 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 27.9813 | 27.7 | -0.3
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.2 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.135 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.94 | 0.01 | 100 | 150.3 | -0.2 | 150.3 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2888 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.938 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1996 | 28.0864 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.05 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.1982 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1968 | 28.1214 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.93 | 0.01 | 100 | 150.2 | -0.2 | 150.2 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.7941 | 28.0445 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.932 | 0.01 | 100 | 150.1 | -0.2 | 150.1 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0323 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.911 | 0.01 | 100 | 150.1 | -0.2 | 150.1 | 45.9 | 45.9 | 0.1 | 12.9 | 955 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1 | 27.6911 | -0.2
2008-07-16 | 10:16 | 2.65 | 0.0097 | 97.3432 | 150.1 | -0.2 | 150.1 | 45.9 | 45.9 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.2012 | 28.1 | 27.6147 | -0.2
2008-07-16 | 10:16 | 2.429 | 0.009 | 90 | 145.8 | -0.2 | 145.8 | 45.5 | 45.5158 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.02864 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.0377 | 27.6623 | -0.2
2008-07-16 | 10:16 | 2.36 | 0.009 | 90 | 130.9 | -0.2 | 130.9 | 41.4 | 41.4 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3198 | 0.026042 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.6607 | -0.2
2008-07-16 | 10:16 | 2.36 | 0.009 | 90 | 124.6 | -0.2 | 124.6 | 38.5 | 38.4598 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3198 | 0.02478 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28 | 27.6163 | -0.2
2008-07-16 | 10:16 | 2.358 | 0.009 | 90 | 123.3 | -0.2 | 123.3 | 37.3 | 37.2727 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28 | 27.6933 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.009 | 90 | 123.2 | -0.2 | 123.2 | 36.9 | 36.9223 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.009 | 90 | 122.9 | -0.2 | 122.9 | 36.8 | 36.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0244 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0931 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.009 | 90 | 121.7 | -0.2 | 121.7 | 36.7 | 36.7 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.1764 | 27.6764 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.009 | 90 | 121.6 | -0.2 | 121.6 | 36.7 | 36.7 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.0994 | 27.6006 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.009 | 90 | 121.6 | -0.2 | 121.6 | 36.7 | 36.7 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3203 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.0223 | 27.6777 | -0.2
2008-07-16 | 10:16 | 2.35 | 0.0078 | 78.4941 | 121.6 | -0.17988 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.4 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.1627 | 27.7 | -0.2
2008-07-16 | 10:16 | 2.349 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2691 | 27.6691 | -0.1
2008-07-16 | 10:16 | 2.341 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2079 | 27.6 | -0.1
2008-07-16 | 10:16 | 2.34 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.285 | 27.6 | -0.1
2008-07-16 | 10:16 | 2.351 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.2384 | 27.6616 | -0.1
2008-07-16 | 10:16 | 2.36 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 32.1031 | 55 | 1049 | 1011 | 28.1232 | 27.6616 | -0.1
2008-07-16 | 10:16 | 2.358 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.1984 | 28.0306 | 27.6153 | -0.1
2008-07-16 | 10:16 | 2.35 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1846 | 27.6923 | -0.1
2008-07-16 | 10:16 | 2.35 | 0.007 | 70 | 121.6 | -0.1 | 121.6 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2 | 27.6307 | -0.1
2008-07-16 | 10:16 | 2.335 | 0.007 | 70 | 121.5 | -0.1 | 121.5 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.2044 | 28.1076 | 27.6 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.007 | 70 | 121.4 | -0.1 | 121.4 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.023 | 27.623 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.007 | 70 | 121.3 | -0.1 | 121.3 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0999 | 27.6999 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.007 | 70 | 121.2 | -0.1 | 121.2 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2536 | 27.6232 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.007 | 70 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1928 | 27.6536 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.007 | 70 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0005 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.0696 | 27.6696 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.0071 | 70.894 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.02408 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28 | 27.6073 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.0079 | 79.1722 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1972 | 28 | 27.6842 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.008 | 80 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1221 | 27.7 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.008 | 80 | 121.1 | -0.1 | 121.1 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1242 | 27.7 | -0.1
2008-07-16 | 10:16 | 2.33 | 0.008 | 80 | 120.9 | -0.1 | 120.9 | 36.5 | 36.5242 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.0443 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.9 | -0.1 | 120.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2749 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.8 | -0.1 | 120.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.163 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.7 | -0.1 | 120.7 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.024 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.6547 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.6 | -0.1 | 120.6 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.6 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.3 | -0.1 | 120.3 | 36.5 | 36.5 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0239 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.6 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 120.2 | -0.1 | 120.2 | 36.6 | 36.5792 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0239 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.6759 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | -0.1 | 119.9 | 36.5 | 36.5126 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3197 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2055 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | -0.1 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2408 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | -0.1 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1065 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | -0.1 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1833 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | -0.079563 | 119.9 | 36.5 | 36.5195 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.7956 | 28.2 | 27.7 | -0.1
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | 0 | 119.9 | 36.6 | 36.5812 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1261 | 27.7 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | 0 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3201 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.0415 | 27.6862 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | 0 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1996 | 28.2723 | 27.6092 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.9 | 0 | 119.9 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3801 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2555 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.1787 | 27.6213 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.1018 | 27.6982 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.4 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.7 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3206 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.1 | 27.6481 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.8969 | 55 | 1049 | 1013 | 28.0713 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.0166 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.2472 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.1815 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.5 | 36.5 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.1723 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | 0 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3802 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.8024 | 28.3 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | -0.01988 | 119.8 | 36.5 | 36.519 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3801 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.2333 | 27.6 | 0
2008-07-16 | 10:17 | 2.33 | 0.008 | 80 | 119.8 | -0.1 | 119.8 | 36.5 | 36.5123 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.1984 | 28.2436 | 27.6 | -0.1
2008-07-16 | 10:17 | 2.287 | 0.0072 | 71.9156 | 119.8 | -0.1 | 119.8 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.881 | 0.0062 | 62.2852 | 119.2 | -0.1 | 119.2 | 36.6 | 36.6 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.023698 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.2044 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.766 | 0.0071 | 70.8361 | 107.7 | -0.1 | 107.7 | 34.9 | 34.87 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3797 | 0.021677 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.2258 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.74 | 0.0079 | 79.0328 | 95.6 | -0.1 | 95.6 | 30.7 | 30.7323 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.01926 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1014 | 28.3021 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.736 | 0.0073 | 72.6805 | 93.3 | -0.1 | 93.3 | 28.6 | 28.5578 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.01886 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3721 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92.3 | -0.1 | 92.3 | 27.8 | 27.7639 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3047 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92.2 | -0.1 | 92.2 | 27.5 | 27.5 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3816 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92.2 | -0.1 | 92.2 | 27.4 | 27.4 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.8008 | 28.3416 | 27.6584 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92.1 | -0.1 | 92.1 | 27.4 | 27.4 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.3 | 27.6647 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92.1 | -0.1 | 92.1 | 27.4 | 27.4 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92 | -0.1 | 92 | 27.3 | 27.3317 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.007 | 70 | 92 | -0.1 | 92 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.32 | 0.01858 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2341 | 27.6659 | -0.1
2008-07-16 | 10:17 | 1.731 | 0.007 | 70 | 91.7 | -0.1 | 91.7 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0185 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2427 | 27.6573 | -0.1
2008-07-16 | 10:17 | 1.739 | 0.007 | 70 | 91.7 | -0.1 | 91.7 | 27.4 | 27.3837 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0185 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.74 | 0.007 | 70 | 91.6 | -0.1 | 91.6 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0185 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.734 | 0.007 | 70 | 91.6 | -0.1 | 91.6 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.32 | 0.0185 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.0063 | 62.6367 | 91.5 | -0.1 | 91.5 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0185 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6501 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91.5 | -0.1 | 91.5 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3198 | 0.0184 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.354 | 27.673 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91.4 | -0.1 | 91.4 | 27.3 | 27.3 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0184 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4923 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91.1 | -0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0184 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3386 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91.1 | -0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0184 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1848 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91 | -0.1 | 91 | 27.3 | 27.3136 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3794 | 0.0184 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.169 | 27.6345 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 91 | -0.1 | 91 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.018379 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.7941 | 28.3 | 27.7 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.9 | -0.1 | 90.9 | 27.3 | 27.3097 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.7 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.9 | -0.1 | 90.9 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.635 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.9 | -0.1 | 90.9 | 27.3 | 27.3211 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3799 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.9 | -0.1 | 90.9 | 27.4 | 27.3787 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.38 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.8 | -0.1 | 90.8 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.8 | -0.1 | 90.8 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.8 | -0.1 | 90.8 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:17 | 1.73 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.73 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.73 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.73 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6568 | -0.1
2008-07-16 | 10:18 | 1.726 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6664 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2895 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.7 | -0.1 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2127 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.6 | -0.1 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3925 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.6 | -0.1 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.459 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.6 | -0.1 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3787 | 0.0183 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.6 | -0.1 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3197 | 0.01828 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.6 | -0.1 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3284 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2515 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2253 | 27.6253 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.2044 | 28.3 | 27.6978 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6209 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3798 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.1 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.38 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.2 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.2 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.2 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.72 | 0.006 | 60 | 90.5 | -0.2 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.721 | 0.006 | 60 | 90.5 | -0.2 | 90.5 | 27.4 | 27.4 | 0.1 | 13 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2292 | 27.6708 | -0.2
2008-07-16 | 10:18 | 1.738 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2475 | 27.6525 | -0.2
2008-07-16 | 10:18 | 1.74 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.7988 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.74 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.202 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.74 | 0.006 | 60 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.74 | 0.006 | 60 | 90.3 | -0.2 | 90.3 | 27.4 | 27.4 | 0.1 | 13 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.802 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.717 | 0.0056 | 55.9774 | 90.4 | -0.2 | 90.4 | 27.4 | 27.4 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 4.6 | 195 | 59.3803 | 0.0182 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3315 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.284 | 0.005 | 50 | 90 | -0.18016 | 90 | 27.4 | 27.4 | 0.1 | 13 | 954 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4 | 195 | 59.3 | 0.018121 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4086 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.169 | 0.005 | 50 | 77.6 | -0.1 | 77.6 | 25.4 | 25.3942 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 4.4 | 195 | 59.3802 | 0.015705 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4856 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 65.7 | -0.1 | 65.7 | 21.3 | 21.2515 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.013418 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3758 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 63.1 | -0.1 | 63.1 | 19.2 | 19.225 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3388 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 62.3 | -0.1 | 62.3 | 18.6 | 18.5835 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0128 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3526 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 62.2 | -0.1 | 62.2 | 18.3 | 18.3198 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0128 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.1215 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 62.1 | -0.1 | 62.1 | 18.2 | 18.2 | 0.1 | 13 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3798 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2398 | 27.6 | -0.1
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 61.9 | -0.2 | 61.9 | 18.2 | 18.2 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.393 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.005 | 50 | 61.8 | -0.2 | 61.8 | 18.2 | 18.2 | 0.1 | 13 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.32 | 0.0127 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4538 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.0043 | 42.6475 | 61.8 | -0.2 | 61.8 | 18.2 | 18.2 | 0.1 | 13 | 952 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3801 | 0.01268 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2998 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.8 | -0.2 | 61.8 | 18.2 | 18.2 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2228 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.7 | -0.2 | 61.7 | 18.2 | 18.2 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3079 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.6 | -0.2 | 61.6 | 18.2 | 18.2 | 0.1 | 13 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.802 | 28.4 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.6 | -0.2 | 61.6 | 18.2 | 18.2 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3794 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.3924 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.6 | -0.2 | 61.6 | 18.2 | 18.1615 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.3154 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.6 | -0.2 | 61.6 | 18.1 | 18.1126 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3616 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.148 | 0.004 | 40 | 61.6 | -0.2 | 61.6 | 18.2 | 18.2 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.00028 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.3615 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.5 | -0.2 | 61.5 | 18.2 | 18.1511 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3153 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.5 | -0.2 | 61.5 | 18.2 | 18.184 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3923 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.5 | -0.2 | 61.5 | 18.2 | 18.2 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.2 | -0.3 | 61.2 | 18.2 | 18.2 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.3
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.2 | -0.3 | 61.2 | 18.1 | 18.1131 | 0.1 | 13 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3801 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3771 | 27.6 | -0.3
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.2 | -0.2 | 61.2 | 18.1 | 18.1 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.8016 | 28.3003 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.2 | -0.2 | 61.2 | 18.1 | 18.1 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3766 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.149 | 0.004 | 40 | 61.2 | -0.2 | 61.2 | 18.1 | 18.1 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3466 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1011.7956 | 28.3303 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.2028 | 28.3929 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.15 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3798 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.316 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.148 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.2 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:18 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3198 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3802 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190.6 | 28 | 5 | 195 | 59.3198 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.8016 | 28.3683 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.3548 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3794 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3204 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.5274 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4422 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.38 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.38 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.24 | 27.66 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.2738 | 27.7 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.3 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1013 | 28.3725 | 27.6862 | -0.3
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.3 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3794 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2188 | 27.6094 | -0.3
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.2044 | 28.335 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.6 | 28 | 5 | 195 | 59.3797 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.4 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3788 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3019 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.3 | -0.2 | 61.3 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1014 | 28.2482 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2287 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3197 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3802 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2871 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.14 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3797 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.2028 | 28.2102 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.142 | 0.004 | 40 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2667 | 27.6 | -0.2
2008-07-16 | 10:19 | 1.132 | 0.0036 | 36.2747 | 61.4 | -0.2 | 61.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.701 | 0.0028 | 28.025 | 60.4 | -0.2 | 60.4 | 18.1 | 18.1 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.611 | 0.002 | 20 | 48.4 | -0.2 | 48.4 | 15.2 | 15.214 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.009801 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.595 | 0.002 | 20 | 35.5 | -0.2 | 35.5 | 11.6 | 11.5772 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.00726 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1988 | 28.2258 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 33.3 | -0.2 | 33.3 | 9.8 | 9.794 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3198 | 0.0068 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.251 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.9 | -0.21984 | 32.9 | 9.2 | 9.1669 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.00676 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.2
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.3 | -0.3 | 32.3 | 9 | 9.0141 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3204 | 0.0066 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.2 | -0.3 | 32.2 | 8.9 | 8.9 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0066 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.1 | -0.3 | 32.1 | 8.9 | 8.9 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0066 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.2028 | 28.2416 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.1 | -0.3 | 32.1 | 8.9 | 8.9 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0066 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.6353 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32.1 | -0.3 | 32.1 | 8.9 | 8.9 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3798 | 0.0066 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.6878 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 32 | -0.3 | 32 | 8.9 | 8.8887 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0065 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.2 | 27.6109 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.2659 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.59 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.575 | 0.002 | 20 | 32 | -0.3 | 32 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 8e-05 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3392 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.57 | 0.002 | 20 | 32 | -0.3 | 32 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4929 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.57 | 0.002 | 20 | 32 | -0.3 | 32 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.28 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.57 | 0.002 | 20 | 32 | -0.3 | 32 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.2502 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.57 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.6 | 28 | 5 | 195 | 59.3203 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.579 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3797 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188 | 28 | 5 | 195 | 59.3201 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 188.4 | 28 | 5 | 195 | 59.38 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0065 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3689 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.4772 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3235 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.3 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3651 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.31968 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.358 | 27.6 | -0.3
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.4 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.9 | -0.4 | 31.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:19 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3798 | 0.0064 | 0 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | -2e-05 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.8 | -0.4 | 31.8 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.7 | -0.4 | 31.7 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0064 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.6 | -0.4 | 31.6 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.006379 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3643 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.364 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3598 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011 | 28.2103 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.38 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3434 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.32 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.4 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.4 | -0.4 | 31.4 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3199 | 0.0063 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.1988 | 28.3977 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31.2 | -0.4 | 31.2 | 8.8 | 8.8 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.00628 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3209 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 31 | -0.4 | 31 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013.8016 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3796 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32.897 | 55 | 1049 | 1013 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012.7956 | 28.3 | 27.6098 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.3 | 27.6866 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3799 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1012 | 28.3 | 27.6366 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.8004 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.3343 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.488 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011.1992 | 28.3583 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 189.4 | 28 | 5 | 195 | 59.3199 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.2523 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3197 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 55 | 1049 | 1011 | 28.2244 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32.103 | 55 | 1049 | 1011.2075 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.58 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3206 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.571 | 0.002 | 20 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3549 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.222 | 0.001 | 10.0794 | 30.9 | -0.4 | 30.9 | 8.8 | 8.8 | 0.1 | 12.9 | 954 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0062 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3681 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.052 | -0.0029 | -29.272 | 24.4 | -0.4 | 24.4 | 7.8 | 7.8256 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3202 | 0.004895 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048.7802 | 1012.202 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.0021 | -21.272 | 7 | -0.4 | 7 | 3.7 | 3.6624 | 0.1 | 12.9 | 954 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.32 | 0.00138 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 4 | -0.4 | 4 | 1 | 0.9914 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.00078 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 3 | -0.4 | 3 | 0 | 0 | 0.1 | 12.9 | 952 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2.7 | -0.4 | 2.7 | -0.3 | 0 | 0.1 | 12.9 | 952 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.00048 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.284 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2.4 | -0.4 | 2.4 | -0.5 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0004 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.8016 | 28.2076 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2 | -0.4 | 2 | -0.5 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3388 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2 | -0.4 | 2 | -0.5 | 0 | 0.1 | 12.9 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3535 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2 | -0.4 | 2 | -0.5 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3235 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 2 | -0.4 | 2 | -0.5 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3203 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.4 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.9 | -0.4 | 1.9 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3799 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.4 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.9 | -0.4 | 1.9 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3202 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3462 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.9 | -0.4 | 1.9 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011 | 28.3 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.8 | -0.4 | 1.8 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190 | 28 | 5 | 195 | 59.3 | 0.0003 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.2921 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.7 | -0.4 | 1.7 | -0.6 | 0 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.2151 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.4 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3199 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.3233 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.4 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3792 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.4384 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.4 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.2802 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.5 | 27.6 | -0.4
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1013 | 28.5 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.7956 | 28.5 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.4078 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1011.7988 | 28.346 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.4997 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.5 | -0.5 | 1.5 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3466 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.5 | -0.5 | 1.5 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.5 | -0.5 | 1.5 | -0.6 | 0 | 0.1 | 12.9 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.00012 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.1988 | 28.3912 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.5 | -0.5 | 1.5 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012.8016 | 28.5856 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.002 | -20 | 1.5 | -0.5 | 1.5 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3196 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.4318 | 27.6 | -0.5
2008-07-16 | 10:20 | 0.01 | -0.0016 | -16.1762 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.339 | 27.6 | -0.5
2008-07-16 | 10:21 | 0.01 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | -0.0001 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.3378 | 27.6 | -0.5
2008-07-16 | 10:21 | 0.01 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1048 | 1012 | 28.4147 | 27.6 | -0.5
2008-07-16 | 10:21 | 0.01 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1011.7988 | 28.4915 | 27.6 | -0.5
2008-07-16 | 10:21 | 0.01 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 191.4 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3632 | 27.6 | -0.5
2008-07-16 | 10:21 | 0.008 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.4 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3221 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.399 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3241 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 952 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3296 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3794 | 0.0001 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.7941 | 28.4129 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3797 | 0.0001 | -0.00022 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.7972 | 28.5665 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4197 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.3415 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 952 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.5721 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 952 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3199 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4649 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.0006 | -6.1097 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 961 | 948 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3197 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.4444 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | 0 | 0 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3802 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1013 | 28.4787 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | 0 | -0.27586 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4018 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.0009 | -8.8966 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.55 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.5 | -0.5 | 1.5 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.5481 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4713 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.5 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3407 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.6 | -0.5 | 1.6 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.3361 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.001 | -10 | 1.7 | -0.5 | 1.7 | -0.5 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 994 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.0015 | -15.4773 | 1.7 | -0.5 | 1.7 | -0.6 | 0 | 0.1 | 13 | 953 | 960 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.4 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.002 | -20 | 1.7 | -0.5 | 1.7 | -0.6 | 0 | 0.1 | 13 | 953 | 960 | 950 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.5334 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.002 | -20 | 1.7 | -0.5 | 1.7 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012 | 28.5564 | 27.6 | -0.5
2008-07-16 | 10:21 | 0 | -0.002 | -19.7412 | 1.7 | -0.5 | 1.7 | -0.6 | 0 | 0.1 | 13 | 953 | 961 | 949 | 0 | 0 | 0 | 993 | -40 | 0 | 29 | 192 | 28 | 5 | 195 | 59.3 | 0 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 55 | 1049 | 1012.204 | 28.5204 | 27.6 | -0.5
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
No Processing Errors Detected
Test Date | 2008-07-16
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 297
NO2 Span(ppm) | 505
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.36
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | White Van
License Plate
Engine Displacement | 7.3
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 10:00
Test End Time | 10:21
Test Duration (s) | 1262
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 993.626
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.832
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 07/16/2008 09:53:44.257 - None Found
0X0000 - 07/16/2008 09:57:13.753 - None Found
0X0000 - 07/16/2008 10:00:00.274 - None Found
Warnings:
0X0000 - 07/16/2008 09:53:44.257 - None Found
0X0000 - 07/16/2008 09:57:13.754 - None Found
0X0000 - 07/16/2008 10:00:00.274 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 123
Brake Specific Dropout Limit Count | 1262
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-07-16
Time | 9:54
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
CO(ppm) | 0 | 0 | 0
CO2(%) | 0 | 0 | 0
NO(ppm) | 0.04375 | 0.99333 | -0.94958
NO2(ppm) | 0.125 | -0.38667 | 0.51167
THC(ppmC) | -1.1 | -0.14667 | -0.95333
Span
InfoVer | 1
Date | 2008-07-16
Time | 9:57
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 200 | 6 | 20.9 | 1989 | x^0: 297; x^1: 505; x^2: 0; x^3: 93.36
Gas | Previous | Current | Difference
CO(ppm) | 210 | 204 | 6
CO2(%) | 6.1519 | 6.0073 | 0.14454
NO(ppm) | 298.6625 | 297.1133 | 1.5492
THC(ppmC) | 93.5375 | 93.42 | 0.1175
Test Information:
SEMTECH_DATA_FILE | 2008-07-16 | 9:53 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## THC Lin
Test Date:
Test Time:
7/16/2008
10:00:00
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
7/16/2008
10:00:00
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
7/16/2008
10:00:00
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
7/16/2008
10:00:00
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
7/16/2008
10:00:00
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
7/16/2008
10:00:00
NDUV Serial No:
208
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6.009 | Rm = | 0.99276 | 0.99276 | 0.99276 | 0.99276 | 0.99276 | 0.99276 | 0.99276 | 0.99276 | 0.99276
| C.F.= | 0.099348 | 0.19884 | 0.29848 | 0.39826 | 0.49818 | 0.59825 | 0.69847 | 0.79883 | 0.89934
| gas = | 0.59698 | 1.1948 | 1.7935 | 2.3931 | 2.9936 | 3.5949 | 4.1971 | 4.8002 | 5.4042
Ro for CO2= | 0.8795