## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-710C | GDU-002 | 2008-01-01 | 2009-01-01
0 | 0 | 0 | 0
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 41304
NDUV Serial Number: | 217
FID Serial Number: | 177
Heated Filter Serial number: | 18185
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALMO23797 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.1 | Scott gas | ALMO24749 | ±1% | 2010-08-25 | 93.3
200 ppm CO | 200.2
6% CO2 | 6.07
300 ppm NO | 301.3
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6.01
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALMO31297 | ±1% | 2010-10-06
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | CO2 - AMBII Serial Number: |  |  | 41304 |  |  | CO - AMBII Serial Number: |  |  | 41304 |  | HC - AMBII Serial Number: |  |  | 41304 |  | NO - NDUV Serial Number: |  |  | 217 |  | THC - FID Serial Number: |  |  | 177
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-11-05 |  | 6.07 | X | 0.1 | 3 |  |  | 200.2 |  | 50 | 3 |  | 31.1 |  | 4 | 3 |  | 301.3 |  | 15 | 3 |  | 93.3 | 100 | 5 | 2
Test Time: | 10:29
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0.002 | 0.002 |  |  | 0 | -3 | -3 |  |  | 0 | 1016 | 1016 |  |  | 0 | 0.2 | 0.2 |  |  | 0 | 0.02 | 0.02
90 | 105 | 1 | 0.607 | 0.603 | 0.579 | -0.024 | -3.98 |  | 20.02 | 20 | 0 | 0 |  | 3.11 | 1016.1 | 1013 | 32572.35 | Fail | 30.13 | 30.4 | 0.3 | 1 |  | 9.33 | 9.45 | 0.12 | 1.29
150 | 165 | 2 | 1.214 | 1.207 | 1.14 | -0.067 | -5.55 |  | 40.04 | 49 | 9 | 22.48 |  | 6.22 | 1016.4 | 1010.2 | 16241.16 | Fail | 60.26 | 60.7 | 0.4 | 0.66 |  | 18.66 | 18.54 | -0.12 | -0.64
210 | 225 | 3 | 1.821 | 1.812 | 1.723 | -0.089 | -4.91 |  | 60.06 | 64 | 4 | 6.66 |  | 9.33 | 1016.5 | 1007.2 | 10795.28 | Fail | 90.39 | 91.1 | 0.7 | 0.77 |  | 27.99 | 27.48 | -0.51 | -1.82
270 | 285 | 4 | 2.428 | 2.417 | 2.328 | -0.089 | -3.68 |  | 80.08 | 90 | 10 | 12.49 |  | 12.44 | 1016.1 | 1003.7 | 8068.33 | Fail | 120.52 | 121.3 | 0.8 | 0.66 |  | 37.32 | 36.81 | -0.51 | -1.37
330 | 345 | 5 | 3.035 | 3.024 | 2.938 | -0.086 | -2.84 |  | 100.1 | 102 | 2 | 2 |  | 15.55 | 1016.8 | 1001.3 | 6439.23 | Fail | 150.65 | 150.9 | 0.3 | 0.2 |  | 46.65 | 46.09 | -0.56 | -1.2
390 | 405 | 6 | 3.642 | 3.631 | 3.554 | -0.077 | -2.12 |  | 120.12 | 120 | 0 | 0 |  | 18.66 | 1016.4 | 997.7 | 5346.73 | Fail | 180.78 | 180.5 | -0.3 | -0.17 |  | 55.98 | 55.41 | -0.57 | -1.02
450 | 465 | 7 | 4.249 | 4.24 | 4.181 | -0.059 | -1.39 |  | 140.14 | 132 | -8 | -5.71 |  | 21.77 | 1016.8 | 995 | 4570.51 | Fail | 210.91 | 210.9 | 0 | 0 |  | 65.31 | 64.56 | -0.75 | -1.15
510 | 525 | 8 | 4.856 | 4.849 | 4.835 | -0.014 | -0.29 |  | 160.16 | 150 | -10 | -6.24 |  | 24.88 | 1016.9 | 992 | 3987.14 | Fail | 241.04 | 240.4 | -0.6 | -0.25 |  | 74.64 | 73.78 | -0.86 | -1.15
570 | 585 | 9 | 5.463 | 5.459 | 5.449 | -0.01 | -0.18 |  | 180.18 | 166 | -14 | -7.77 |  | 27.99 | 1016.8 | 988.8 | 3532.69 | Fail | 271.17 | 271 | -0.2 | -0.07 |  | 83.97 | 83.91 | -0.06 | -0.07
630 | 645 | 10 | 6.07 | 6.07 | 6.072 | 0.002 | 0.03 |  | 200.2 | 184 | -16 | -7.99 |  | 31.1 | 1016.2 | 985.1 | 3167.52 | Fail | 301.3 | 301.1 | -0.2 | -0.07 |  | 93.3 | 93.01 | -0.29 | -0.31
690 | 705 | 9 | 5.463 | 5.459 | 5.452 | -0.007 | -0.13 |  | 180.18 | 163 | -17 | -9.44 |  | 27.99 | 1016.3 | 988.3 | 3530.9 | Fail | 271.17 | 272.6 | 1.4 | 0.52 |  | 83.97 | 83.08 | -0.89 | -1.06
750 | 765 | 8 | 4.856 | 4.849 | 4.813 | -0.036 | -0.74 |  | 160.16 | 153 | -7 | -4.37 |  | 24.88 | 1016.1 | 991.2 | 3983.92 | Fail | 241.04 | 242.7 | 1.7 | 0.71 |  | 74.64 | 73.91 | -0.73 | -0.98
810 | 825 | 7 | 4.249 | 4.24 | 4.183 | -0.057 | -1.34 |  | 140.14 | 134 | -6 | -4.28 |  | 21.77 | 1016.1 | 994.3 | 4567.29 | Fail | 210.91 | 212.4 | 1.5 | 0.71 |  | 65.31 | 64.67 | -0.64 | -0.98
870 | 885 | 6 | 3.642 | 3.631 | 3.564 | -0.067 | -1.85 |  | 120.12 | 117 | -3 | -2.5 |  | 18.66 | 1016.1 | 997.4 | 5345.12 | Fail | 180.78 | 182.5 | 1.7 | 0.94 |  | 55.98 | 55.03 | -0.95 | -1.7
930 | 945 | 5 | 3.035 | 3.024 | 2.934 | -0.09 | -2.98 |  | 100.1 | 100 | 0 | 0 |  | 15.55 | 1016.1 | 1000.6 | 6434.73 | Fail | 150.65 | 152.3 | 1.7 | 1.13 |  | 46.65 | 45.7 | -0.95 | -2.04
990 | 1005 | 4 | 2.428 | 2.417 | 2.331 | -0.086 | -3.56 |  | 80.08 | 83 | 3 | 3.75 |  | 12.44 | 1016 | 1003.6 | 8067.52 | Fail | 120.52 | 122.6 | 2.1 | 1.74 |  | 37.32 | 36.53 | -0.79 | -2.12
1050 | 1065 | 3 | 1.821 | 1.812 | 1.719 | -0.093 | -5.13 |  | 60.06 | 69 | 9 | 14.99 |  | 9.33 | 1016.2 | 1006.9 | 10792.07 | Fail | 90.39 | 92.7 | 2.3 | 2.54 |  | 27.99 | 27.48 | -0.51 | -1.82
1110 | 1125 | 2 | 1.214 | 1.207 | 1.14 | -0.067 | -5.55 |  | 40.04 | 42 | 2 | 5 |  | 6.22 | 1015.7 | 1009.5 | 16229.9 | Fail | 60.26 | 62.2 | 1.9 | 3.15 |  | 18.66 | 18.01 | -0.65 | -3.48
1170 | 1185 | 1 | 0.607 | 0.603 | 0.58 | -0.023 | -3.81 |  | 20.02 | 20 | 0 | 0 |  | 3.11 | 1015.9 | 1012.8 | 32565.92 | Fail | 30.13 | 31.6 | 1.5 | 4.98 |  | 9.33 | 8.11 | -1.22 | -13.08
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -20 | -20 |  |  | 0 | 1015.6 | 1015.6 |  |  | 0 | 0.2 | 0.2 |  |  | 0 | -1.38 | -1.38
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.057 | 0.061 |  |  |  | Intercept | 3 | 2 | Fail |  | Intercept | 1016 | 0 | Fail |  | Intercept | 1.1 | 3 |  |  | Intercept | -0.6 | 0.93
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.93 | 0.99 - 1.01 | Fail |  | Slope | 0.02 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.036 | 0.0607 |  |  |  | SEE | 7.8038 | 2.002 | Fail |  | SEE | 0.3073 | 0.311 |  |  | SEE | 0.9079 | 3.013 |  |  | SEE | 0.3983 | 0.933
|  |  | r2 | 0.9997 | 0.998 |  |  |  | r2 | 0.9832 | 0.998 | Fail |  | r2 | 0.2968 | 0.998 | Fail |  | r2 | 0.9999 | 0.998 |  |  | r2 | 0.9998 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0 | 0.1 | 13 | 910 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0003 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0 | 0.0323 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8359 | 0.0003 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1 | 0.1 | 12.8 | 912 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0003 | 0.000136 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.3598 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0002 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.3486 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8639 | 0.0002 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4271 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.003 | 0 | 0 | 0.6 | 0.2 | 0.8 | 0.1 | 0.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.8639 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4945 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.01 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.2 | 0.2 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.416 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.01 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.1 | 0.134 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.8639 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.01 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.2 | 0.2 | 0.1 | 12.8 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4407 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.01 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.2 | 0.2 | 0.1 | 12.7 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0.005 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.2 | 0.2 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.5 | 0.2 | 0.7 | 0.1 | 0.1 | 0.1 | 12.7 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.0009 | -9.4991 | 0.5 | 0.2 | 0.7 | 0.1 | 0.1 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1017 | 20.4451 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.002 | -20 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.002 | -20 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1053 | 1017 | 20.4116 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.002 | -20 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.8 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4903 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.002 | -20 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.8 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4313 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.0014 | -13.7645 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0001 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 12.9 | 912 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4257 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.8 | 59.9 | 0.0001 | 0 | 0 | -6.4e-05 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.4 | 0.2 | 0.6 | 0.1 | 0.1 | 0.1 | 13 | 911 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.4 | 59.8361 | 0.0001 | 0 | -6.4e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.3 | 0.2 | 0.5 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.8361 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4389 | 27.9 | 0.2
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.3 | 0.136 | 0.4 | 0.1 | 0.1 | 0.1 | 13 | 910 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 195 | 59.8 | 3.6e-05 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.3 | 0.1 | 0.4 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195 | 59.8358 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.3 | 0.1 | 0.4 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.1 | 59.836 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.0007 | -6.864 | 0.3 | 0.1 | 0.4 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.4 | 59.9 | 0 | 0 | -3.6e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.64 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.3 | 0.1 | 0.4 | 0.1 | 0.1 | 0.1 | 13 | 910 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.8 | 59.836 | 0 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4535 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.2 | 0.1 | 0.3 | 0.1 | 0.1 | 0.1 | 12.8 | 912 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.2 | 0.1 | 0.3 | 0.1 | 0.1 | 0.1 | 12.7 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.0006 | -6.3454 | 0.2 | 0.1 | 0.3 | 0 | 0.036 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | 0.1 | 0.0686 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.9 | 0 | 0 | -6.4e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4326 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | 0 | 0 | 0.1 | 12.7 | 916 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0 | 0 | -3.6e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4458 | 27.9 | 0.1
2008-11-05 | 10:29 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | 0 | 0.046 | 0.1 | 12.7 | 915 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0 | -6.4e-05 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4758 | 27.9 | 0.1
2008-11-05 | 10:29 | 0.006 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | 0 | 0 | 0.1 | 12.7 | 916 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 195.7 | 59.8 | 0 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4024 | 27.9 | 0.1
2008-11-05 | 10:29 | 0.01 | -0.0002 | -2.1013 | 0.2 | 0.1 | 0.3 | 0.1 | 0.0706 | 0.1 | 12.7 | 916 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.3 | 59.836 | 0 | -0.0001 | 0 | -3.6e-05 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4812 | 27.9 | 0.1
2008-11-05 | 10:29 | 0.01 | 0 | 0 | 0.2 | 0.036 | 0.2 | 0 | 0 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.1 | 59.864 | -6.4e-05 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:29 | 0.009 | 0 | 0 | 0.2 | 0 | 0.2 | 0.1 | 0.0649 | 0.1 | 12.8 | 915 | 907 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 195.4 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 195.8 | 59.8361 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:29 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.8 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8642 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4264 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.8 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.409 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 13 | 912 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4868 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.2 | 0 | 0.2 | 0 | 0 | 0.1 | 13 | 910 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.864 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.1 | 0 | 0.1 | 0 | 0 | 0.1 | 13.1 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | 0 | 0 | 0.1 | 0 | 0.1 | 0 | 0 | 0.1 | 13 | 909 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.1 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.0006 | -6.4772 | 0.1 | 0 | 0.1 | 0 | 0 | 0.1 | 13.1 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.4 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0 | 0 | 0.1 | 13 | 909 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.8 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0 | 0.049 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.0694 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4632 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0 | 0.034 | 0.1 | 13 | 910 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4152 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 12.8 | 912 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4937 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4495 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 195.7 | 59.8 | -3.6e-05 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0
2008-11-05 | 10:30 | 0 | -0.001 | -10 | 0.2 | 0 | 0.2 | 0.1 | 0.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 195.3 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4075 | 27.9 | 0
2008-11-05 | 10:30 | -0.006 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.1 | 0.1 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 195 | 59.8 | -6.4e-05 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.486 | 27.9 | 0
2008-11-05 | 10:30 | -0.005 | -0.001 | -10 | 0.1 | 0 | 0.1 | 0.2 | 0.2 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 195 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4359 | 27.9 | 0
2008-11-05 | 10:30 | 0.155 | 0.0015 | 15.1926 | 0.1 | 0 | 0.1 | 0.2 | 0.2 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0
2008-11-05 | 10:30 | 0.525 | 0.0021 | 20.9416 | 6.6 | 0 | 6.6 | 0.2 | 0.2 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.1 | 59.836 | 0.000996 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4214 | 27.9 | 0
2008-11-05 | 10:30 | 0.568 | 0.001 | 10 | 21.4 | 0 | 21.4 | 0.8 | 0.7656 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.5 | 59.8641 | 0.003526 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4999 | 27.9 | 0
2008-11-05 | 10:30 | 0.57 | 0.001 | 10 | 25.8 | 0 | 25.8 | 6.6 | 6.5765 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 195.9 | 59.9 | 0.004292 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.57 | 0.001 | 10 | 27.6 | 0 | 27.6 | 9 | 9.0228 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.7 | 59.8 | 0.004564 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4433 | 27.9 | 0
2008-11-05 | 10:30 | 0.57 | 0.001 | 10 | 29.4 | 0 | 29.4 | 9.5 | 9.5273 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.4 | 59.864 | 0.004828 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4351 | 27.9 | 0
2008-11-05 | 10:30 | 0.573 | 0.001 | 10 | 29.8 | 0 | 29.8 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.8 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.001 | 10 | 29.8 | 0 | 29.8 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.3 | 59.864 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.001 | 10 | 29.8 | 0 | 29.8 | 9.6 | 9.6 | 0.1 | 12.8 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.8 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4293 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.001 | 10 | 29.8 | 0 | 29.8 | 9.6 | 9.6407 | 0.1 | 12.9 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.836 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4491 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.001 | 10 | 29.8 | 0 | 29.8 | 9.7 | 9.6856 | 0.1 | 13 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.6 | 59.8639 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.3613 | 20.4724 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.0018 | 17.8767 | 29.8 | 0 | 29.8 | 9.6 | 9.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.2 | 59.9 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4061 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 29.9 | 0 | 29.9 | 9.6 | 9.6 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.9 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4845 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 29.9 | 0 | 29.9 | 9.6 | 9.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.8 | 0.0049 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.9 | 0.0049 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195 | 59.9 | 0.0049 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.6393 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.8 | 0.0049 | -0.000136 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.864 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.836 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4447 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30 | 0 | 30 | 9.6 | 9.6 | 0.1 | 12.8 | 917 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195 | 59.836 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4337 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.1 | 0 | 30.1 | 9.6 | 9.6 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 195.1 | 59.864 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.1 | 0 | 30.1 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.4 | 59.9 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.1 | 0 | 30.1 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 195.8 | 59.8 | 0.004964 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0 | 30.2 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4524 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0 | 30.2 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8361 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.426 | 27.9 | 0
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0.1 | 30.3 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0.1 | 30.3 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8361 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.3613 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0.1 | 30.3 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.864 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.2 | 0.1 | 30.3 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.0028 | 27.8639 | 30.2 | 0.1 | 30.3 | 9.6 | 9.6 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8639 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.003 | 30 | 30.2 | 0.1 | 30.3 | 9.6 | 9.5703 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.003 | 30 | 30.2 | 0.1 | 30.3 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 195.9 | 59.8 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.003 | 30 | 30.2 | 0.1 | 30.3 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.6 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4461 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.003 | 30 | 30.3 | 0.1 | 30.4 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.2 | 59.9 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4322 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.003 | 30 | 30.3 | 0.1 | 30.4 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.0022 | 22.133 | 30.4 | 0.1 | 30.5 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.5 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.4 | 0.1 | 30.5 | 9.5 | 9.5 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.8 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 12.9 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4538 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.58 | 0.002 | 20 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:30 | 0.572 | 0.002 | 20 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.575 | 0.0012 | 12.1416 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.8 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4385 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.6 | 59.864 | 0.005 | -0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.4 | 0.1 | 30.5 | 9.4 | 9.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.8639 | 0.005 | -0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.5 | 0.1 | 30.6 | 9.4 | 9.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195 | 59.9 | 0.005 | -0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4262 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.5 | 0.1 | 30.6 | 9.3 | 9.3318 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.8 | 0.005 | -0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4524 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.0018 | 17.8565 | 30.5 | 0.1 | 30.6 | 9.4 | 9.4 | 0.1 | 12.9 | 915 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.8639 | 0.005 | -0.000136 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.0014 | 13.955 | 30.5 | 0.1 | 30.6 | 9.4 | 9.4 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.1 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.5 | 0.1 | 30.6 | 9.4 | 9.4 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 195.5 | 59.836 | 0.005 | -0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.6 | 0.1 | 30.7 | 9.4 | 9.3656 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.9 | 59.9 | 0.005 | -0.000136 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.58 | 0.001 | 10 | 30.6 | 0.1 | 30.7 | 9.4 | 9.4 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.574 | 0.001 | 10 | 30.6 | 0.1 | 30.7 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.8641 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4768 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.0018 | 17.851 | 30.6 | 0.1 | 30.7 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.36 | 20.4016 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.6 | 0.1 | 30.7 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4801 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8358 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8644 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.57 | 0.002 | 20 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.864 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.621 | 0.002 | 20 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.836 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 0.802 | 0.0044 | 44.1413 | 30.7 | 0.1 | 30.8 | 9.5 | 9.4643 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.005 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.11 | 0.0056 | 55.7082 | 39.7 | 0.1 | 39.8 | 9.5 | 9.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.006594 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.11 | 0.005 | 50 | 55.8 | 0.1 | 55.9 | 12.4 | 12.3838 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.836 | 0.00922 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.118 | 0.005 | 50 | 58.4 | 0.1 | 58.5 | 17.5 | 17.48 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0097 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.126 | 0.005 | 50 | 58.6 | 0.1 | 58.7 | 18.4 | 18.4144 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0097 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.134 | 0.0042 | 42.2933 | 58.7 | 0.1 | 58.8 | 18.6 | 18.6 | 0.1 | 12.9 | 919 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0097 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.004 | 40 | 58.8 | 0.1 | 58.9 | 18.7 | 18.7 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8 | 0.0097 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.0044 | 44.2918 | 58.9 | 0.1 | 59 | 18.7 | 18.7 | 0.1 | 13 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0098 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4998 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.2 | 0.1 | 59.3 | 18.7 | 18.7 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0098 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4214 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.3 | 0.1 | 59.4 | 18.7 | 18.7 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0098 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.457 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.5 | 0.1 | 59.6 | 18.7 | 18.7 | 0.1 | 13 | 916 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0099 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.7 | 0.1 | 59.8 | 18.7 | 18.7 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0099 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.8 | 0.1 | 59.9 | 18.7 | 18.7 | 0.1 | 13 | 916 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0099 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.8 | 0.1 | 59.9 | 18.7 | 18.7 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0099 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.429 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.9 | 0.1 | 60 | 18.7 | 18.7 | 0.1 | 13 | 916 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.009964 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4494 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 59.9 | 0.1 | 60 | 18.7 | 18.7 | 0.1 | 12.9 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.8 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60 | 0.1 | 60.1 | 18.6 | 18.6348 | 0.1 | 12.8 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.1 | 0.1 | 60.2 | 18.7 | 18.7 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.8361 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.2 | 0.1 | 60.3 | 18.7 | 18.7 | 0.1 | 12.7 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.8 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.2 | 0.1 | 60.3 | 18.7 | 18.7 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.2 | 0.1 | 60.3 | 18.6 | 18.6367 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.01 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.2 | 0.1 | 60.3 | 18.6 | 18.6 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.836 | 0.01 | -0.0001 | 0 | -6.4e-05 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.3 | 0.1 | 60.4 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.01 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.14 | 0.005 | 50 | 60.3 | 0.1 | 60.4 | 18.6 | 18.6 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.01 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.137 | 0.005 | 50 | 60.3 | 0.1 | 60.4 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.01 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.3 | 0.1 | 60.4 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.010064 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.3 | 0.1 | 60.4 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.4 | 0.1 | 60.5 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.864 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.4 | 0.1 | 60.5 | 18.6 | 18.6 | 0.1 | 12.7 | 921 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.4 | 0.1 | 60.5 | 18.5 | 18.5494 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.8 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.474 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.4 | 0.1 | 60.5 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.64 | 20.4049 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.005 | 50 | 60.5 | 0.1 | 60.6 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4834 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.0047 | 47.2659 | 60.6 | 0.1 | 60.7 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 909 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.13 | 0.0041 | 41.0607 | 60.6 | 0.1 | 60.7 | 18.5 | 18.5 | 0.1 | 12.9 | 919 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.142 | 0.0049 | 49.1579 | 60.6 | 0.1 | 60.7 | 18.5 | 18.5 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5184 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.15 | 0.005 | 50 | 60.6 | 0.1 | 60.7 | 18.6 | 18.6 | 0.1 | 13 | 917 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5972 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.15 | 0.005 | 50 | 60.7 | 0.1 | 60.8 | 18.5 | 18.5 | 0.1 | 13 | 916 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8639 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4486 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.15 | 0.005 | 50 | 60.7 | 0.1 | 60.8 | 18.6 | 18.5639 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0101 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4542 | 27.9 | 0.1
2008-11-05 | 10:31 | 1.15 | 0.005 | 50 | 60.7 | 0.036 | 60.7 | 18.5 | 18.5 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0101 | -0.0001 | 0 | -3.6e-05 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0
2008-11-05 | 10:31 | 1.15 | 0.005 | 50 | 60.7 | 0 | 60.7 | 18.5 | 18.5 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0101 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4891 | 27.9 | 0
2008-11-05 | 10:32 | 1.132 | 0.005 | 50 | 60.7 | 0 | 60.7 | 18.5 | 18.5 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0101 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4105 | 27.9 | 0
2008-11-05 | 10:32 | 1.128 | 0.005 | 50 | 60.7 | 0 | 60.7 | 18.6 | 18.6 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.8 | 191 | 5 | 196 | 59.8 | 0.0101 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4681 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.005 | 50 | 60.8 | 0 | 60.8 | 18.6 | 18.6 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.2 | 191 | 5 | 196 | 59.836 | 0.0102 | 0 | 0 | -6.4e-05 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4535 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.005 | 50 | 60.8 | 0 | 60.8 | 18.6 | 18.6 | 0.1 | 12.9 | 917 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4248 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.005 | 50 | 60.8 | 0 | 60.8 | 18.6 | 18.5705 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.0049 | 48.9546 | 60.8 | 0 | 60.8 | 18.5 | 18.5 | 0.1 | 12.8 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.0041 | 40.8509 | 60.8 | 0 | 60.8 | 18.5 | 18.5 | 0.1 | 12.8 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 32.685 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.004 | 40 | 60.8 | 0 | 60.8 | 18.5 | 18.5 | 0.1 | 12.7 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 32 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.151 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.8 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 32 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.18 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 32.315 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.18 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 920 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.864 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.165 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.16 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.16 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190.8 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4903 | 27.9 | 0
2008-11-05 | 10:32 | 1.16 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4119 | 27.9 | 0
2008-11-05 | 10:32 | 1.158 | 0.004 | 40 | 60.9 | 0 | 60.9 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4666 | 27.9 | 0
2008-11-05 | 10:32 | 1.142 | 0.0043 | 42.6478 | 61 | 0 | 61 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.005 | 50 | 61 | 0 | 61 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0102 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.14 | 0.005 | 50 | 61 | 0 | 61 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0103 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.149 | 0.005 | 50 | 61.1 | 0 | 61.1 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0103 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.308 | 0.0062 | 62.0643 | 61.1 | 0 | 61.1 | 18.5 | 18.5 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.0103 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.72 | 0.0061 | 61.1794 | 68.6 | 0 | 68.6 | 18.5 | 18.5 | 0.1 | 12.9 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.011568 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.72 | 0.005 | 50 | 83.4 | 0 | 83.4 | 20.7 | 20.6684 | 0.1 | 12.9 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.013996 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.72 | 0.0051 | 51.0645 | 88.8 | 0 | 88.8 | 26.1 | 26.0505 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0149 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.72 | 0.0059 | 59.129 | 90 | 0 | 90 | 27.4 | 27.3676 | 0.1 | 13 | 919 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.015064 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 90.3 | 0 | 90.3 | 27.6 | 27.6142 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0151 | 6.4e-05 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.728 | 0.006 | 60 | 90.3 | 0 | 90.3 | 27.7 | 27.7 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0151 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5299 | 27.9 | 0
2008-11-05 | 10:32 | 1.75 | 0.006 | 60 | 90.4 | 0 | 90.4 | 27.7 | 27.7 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0151 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5917 | 27.9 | 0
2008-11-05 | 10:32 | 1.75 | 0.006 | 60 | 90.5 | 0 | 90.5 | 27.7 | 27.7 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8361 | 0.0151 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5133 | 27.9 | 0
2008-11-05 | 10:32 | 1.743 | 0.006 | 60 | 90.5 | 0 | 90.5 | 27.7 | 27.7 | 0.1 | 13.1 | 918 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0151 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.74 | 0.006 | 60 | 90.6 | 0 | 90.6 | 27.7 | 27.7 | 0.1 | 13.1 | 917 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.015164 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.736 | 0.006 | 60 | 90.6 | 0 | 90.6 | 27.7 | 27.7 | 0.1 | 13.1 | 918 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0152 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.7 | 0.1 | 90.8 | 27.7 | 27.7 | 0.1 | 12.9 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.8 | 191 | 5 | 196 | 59.9 | 0.0152 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.7 | 0.1 | 90.8 | 27.7 | 27.7 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 193 | 190.2 | 5 | 196 | 59.9 | 0.0152 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.8 | 0.1 | 90.9 | 27.7 | 27.6693 | 0.1 | 12.8 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 193 | 190 | 5 | 196 | 59.8 | 0.0152 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.8 | 0.1 | 90.9 | 27.6 | 27.6 | 0.1 | 12.7 | 921 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.2 | 190 | 5 | 196 | 59.8 | 0.0152 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.036 | 90.9 | 27.6 | 27.6 | 0.1 | 12.8 | 921 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0152 | 6.4e-05 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0 | 90.9 | 27.6 | 27.6 | 0.1 | 12.7 | 921 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0152 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.8 | 0 | 90.8 | 27.6 | 27.6 | 0.1 | 12.8 | 921 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0152 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0 | 90.9 | 27.5 | 27.5 | 0.1 | 12.7 | 921 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.063928 | 91 | 27.5 | 27.5 | 0.1 | 12.8 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.1 | 91 | 27.5 | 27.4644 | 0.1 | 12.8 | 921 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.1 | 91 | 27.4 | 27.4 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.836 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.1 | 91 | 27.4 | 27.4 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.1 | 91 | 27.3 | 27.3 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.8361 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.6387 | 20.458 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 90.9 | 0.1 | 91 | 27.4 | 27.3703 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4209 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4994 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.7 | 922 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.4 | 27.4 | 0.1 | 12.7 | 922 | 908 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4654 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 12.9 | 920 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4133 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91 | 0.1 | 91.1 | 27.3 | 27.3 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.015364 | 0.000164 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4917 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.4 | 27.4 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4298 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.73 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.4 | 27.4 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.8 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4487 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.724 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.4 | 27.4 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.4 | 27.4 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.5 | 27.5 | 0.1 | 13 | 917 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.5 | 27.5 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.5 | 27.5 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4592 | 27.9 | 0.1
2008-11-05 | 10:32 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.5 | 27.5 | 0.1 | 12.8 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.72 | 0.006 | 60 | 91.1 | 0.1 | 91.2 | 27.6 | 27.6 | 0.1 | 12.8 | 919 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8639 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.72 | 0.0067 | 67.3902 | 91.1 | 0.1 | 91.2 | 27.6 | 27.6 | 0.1 | 12.7 | 920 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0153 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.72 | 0.0076 | 75.6894 | 91.1 | 0.1 | 91.2 | 27.6 | 27.6 | 0.1 | 12.8 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.8641 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.6414 | 20.4549 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.72 | 0.008 | 80 | 91.2 | 0.1 | 91.3 | 27.6 | 27.5676 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.72 | 0.0077 | 77.3857 | 91.2 | 0.1 | 91.3 | 27.6 | 27.6 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.728 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.5 | 27.5 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0154 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.5 | 27.5 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.015464 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.5 | 27.5 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0155 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.5 | 27.5 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.0155 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.5 | 27.5 | 0.1 | 12.8 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0155 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5042 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.5 | 27.5 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.836 | 0.0155 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5827 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.5 | 27.5 | 0.1 | 12.8 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0155 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5388 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.4 | 27.4 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.8 | 0.0155 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.4 | 27.4 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0155 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4819 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.4 | 27.4406 | 0.1 | 12.9 | 920 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0155 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4033 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.5 | 27.5 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0156 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4752 | 27.9 | 0.1
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.2 | 91.6 | 27.5 | 27.5 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0156 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.2 | 91.6 | 27.5 | 27.5 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0156 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4679 | 27.9 | 0.2
2008-11-05 | 10:33 | 1.73 | 0.007 | 70 | 91.4 | 0.26393 | 91.7 | 27.5 | 27.5 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8361 | 0.0156 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4105 | 27.9 | 0.3
2008-11-05 | 10:33 | 1.796 | 0.007 | 70 | 91.4 | 0.3 | 91.7 | 27.5 | 27.5 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.0156 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4891 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.007 | 0.007 | 70 | 91.4 | 0.3 | 91.7 | 27.6 | 27.6 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.015564 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 105 | 0.3 | 105.3 | 27.6 | 27.6 | 0.1 | 13 | 917 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.017712 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 116.7 | 0.3 | 117 | 30.2 | 30.2083 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.019592 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5243 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 118.7 | 0.3 | 119 | 35.4 | 35.3982 | 0.1 | 13 | 919 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.019964 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5976 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 119.5 | 0.3 | 119.8 | 36.6 | 36.5514 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.020064 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5186 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.322 | 0.0085 | 85.1443 | 119.9 | 0.3 | 120.2 | 36.7 | 36.7398 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8 | 0.0201 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.313 | 0.009 | 90 | 119.9 | 0.3 | 120.2 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.836 | 0.020164 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:33 | 2.321 | 0.0086 | 85.7761 | 120 | 0.236 | 120.2 | 36.8 | 36.8 | 0.1 | 12.7 | 923 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | 0.0202 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.329 | 0.008 | 80 | 120.1 | 0.2 | 120.3 | 36.8 | 36.8472 | 0.1 | 12.8 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.836 | 0.0202 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.0079 | 79.0523 | 120.1 | 0.2 | 120.3 | 36.9 | 36.8733 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.8 | 0.0202 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.0071 | 70.8824 | 120.6 | 0.2 | 120.8 | 36.9 | 36.9 | 0.1 | 12.8 | 922 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.020264 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121 | 0.2 | 121.2 | 36.9 | 36.9 | 0.1 | 12.8 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4692 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121 | 0.2 | 121.2 | 36.9 | 36.9 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8361 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4092 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 12.8 | 921 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4876 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4338 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 12.8 | 921 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4448 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.9 | 36.8736 | 0.1 | 12.8 | 921 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4983 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.007 | 70 | 121.1 | 0.2 | 121.3 | 36.8 | 36.7723 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.836 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4204 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.329 | 0.0074 | 73.9268 | 121.2 | 0.2 | 121.4 | 36.7 | 36.7 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4583 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.321 | 0.008 | 80 | 121.2 | 0.2 | 121.4 | 36.7 | 36.7316 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5371 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.008 | 80 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 12.9 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5844 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.008 | 80 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 919 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0203 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.506 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.008 | 80 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.008 | 80 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.836 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4489 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.0084 | 83.968 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4296 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.1 | 0.2 | 121.3 | 36.8 | 36.8 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0202 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0202 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.8 | 36.8 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 195.9 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.5 | 59.8361 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.32 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.8 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.328 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 195.2 | 59.8361 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.6 | 59.8 | 0.0203 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5359 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.836 | 0.0203 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5856 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | 0.0203 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5071 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.9 | 36.8675 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 196 | 59.8 | 0.0203 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4289 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.9 | 59.864 | 0.0203 | 0.0002 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:33 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.8 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.6 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4283 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.2 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.8 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.8 | 0.020236 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.3 | 0.2 | 121.5 | 36.8 | 36.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.5 | 59.9 | 0.0202 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.8 | 36.8 | 0.1 | 12.8 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.8 | 36.8 | 0.1 | 12.9 | 919 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.836 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.8 | 36.8 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.864 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.8 | 36.8 | 0.1 | 13 | 917 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7278 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.4 | 59.8 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.7 | 59.8359 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0203 | 0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.43 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13 | 917 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0204 | 0.000264 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4483 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.4 | 0.3 | 121.7 | 36.7 | 36.7 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.836 | 0.0204 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:34 | 2.33 | 0.0084 | 84.3272 | 121.5 | 0.236 | 121.7 | 36.7 | 36.7 | 0.1 | 12.8 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0204 | 0.0003 | 0 | -6.4e-05 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.008 | 80 | 121.5 | 0.2 | 121.7 | 36.7 | 36.7 | 0.1 | 12.7 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0204 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.0082 | 82.175 | 121.6 | 0.2 | 121.8 | 36.7 | 36.7 | 0.1 | 12.7 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 196 | 59.864 | 0.0204 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.33 | 0.009 | 90 | 121.6 | 0.2 | 121.8 | 36.7 | 36.7 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0204 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.319 | 0.009 | 90 | 121.6 | 0.2 | 121.8 | 36.7 | 36.7 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0204 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.489 | 0.009 | 90 | 121.6 | 0.2 | 121.8 | 36.7 | 36.7 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0204 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.809 | 0.0107 | 106.9324 | 122 | 0.2 | 122.2 | 36.7 | 36.7 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.9 | 0.020464 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4237 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.0116 | 116.175 | 132 | 0.236 | 132.2 | 36.8 | 36.7978 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.021984 | 0.000236 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4547 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.011 | 110 | 146.3 | 0.2 | 146.5 | 41.2 | 41.2244 | 0.1 | 12.7 | 923 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.02422 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.533 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.011 | 110 | 148.4 | 0.2 | 148.6 | 44.8 | 44.8386 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.024564 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5888 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.011 | 110 | 148.8 | 0.2 | 149 | 45.6 | 45.6305 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0246 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5098 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.011 | 110 | 148.9 | 0.2 | 149.1 | 45.7 | 45.7 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190.8 | 5 | 196 | 59.9 | 0.0247 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.95 | 0.011 | 110 | 149 | 0.2 | 149.2 | 45.7 | 45.7296 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0247 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4528 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.949 | 0.011 | 110 | 149.1 | 0.2 | 149.3 | 45.8 | 45.8 | 0.1 | 12.8 | 921 | 907 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0247 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4256 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.941 | 0.011 | 110 | 149.2 | 0.2 | 149.4 | 45.8 | 45.8475 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8 | 0.0247 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.3 | 0.2 | 149.5 | 45.8 | 45.8 | 0.1 | 12.8 | 920 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.9 | 0.024764 | 0.000264 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.4 | 0.2 | 149.6 | 45.8 | 45.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.836 | 0.0248 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.5 | 0.2 | 149.7 | 45.9 | 45.9 | 0.1 | 12.9 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.9 | 0.0248 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.6 | 0.2 | 149.8 | 45.9 | 45.9 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.8 | 0.0248 | 0.0003 | 0 | -3.6e-05 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.7 | 0.2 | 149.9 | 45.9 | 45.9 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.8 | 59.836 | 0.0249 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.8 | 0.2 | 150 | 45.9 | 45.9 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.9 | 0.0249 | 0.0003 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.7 | 0.2 | 149.9 | 45.9 | 45.9 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.5 | 59.8 | 0.0249 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4466 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 149.8 | 0.2 | 150 | 45.9 | 45.9 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.836 | 0.0249 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4319 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 150.2 | 0.2 | 150.4 | 45.9 | 45.9 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.2 | 59.836 | 0.024964 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4898 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.0104 | 104.3545 | 150.5 | 0.2 | 150.7 | 45.9 | 45.9 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.6 | 59.864 | 0.025 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4112 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.6 | 0.2 | 150.8 | 45.9 | 45.9 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.025 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4673 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.6 | 0.2 | 150.8 | 45.9 | 45.9 | 0.1 | 12.8 | 919 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.8 | 0.025 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.7 | 0.2 | 150.9 | 45.9 | 45.9 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4758 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.7 | 0.2 | 150.9 | 45.9 | 45.9 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8639 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4027 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.7 | 0.2 | 150.9 | 45.9 | 45.9 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0251 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4812 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.7 | 0.2 | 150.9 | 45.9 | 45.9 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190.2 | 5 | 196 | 59.864 | 0.0252 | 0.000464 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46 | 46 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.934 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.93 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.93 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.93 | 0.01 | 100 | 150.8 | 0.2 | 151 | 46.1 | 46.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.7 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.491 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.93 | 0.0102 | 102.3652 | 150.8 | 0.2 | 151 | 46 | 46.0301 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.3 | 59.9 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4128 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.937 | 0.011 | 110 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.864 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4659 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 150.8 | 0.2 | 151 | 46.1 | 46.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.5 | 59.836 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.8 | 59.9 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.011 | 110 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 12.9 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.0102 | 102.4604 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.009 | 90 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8361 | 0.0253 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:34 | 2.94 | 0.009 | 90.4427 | 150.8 | 0.2 | 151 | 46.1 | 46.1 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.8 | 59.9 | 0.025236 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.0099 | 99.1233 | 150.8 | 0.2 | 151 | 46.1 | 46.1 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13.1 | 916 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.836 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.1 | 0.1 | 13.1 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.0693 | 0.1 | 13 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4923 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46 | 46.0262 | 0.1 | 12.9 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4138 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.936 | 0.0093 | 92.7124 | 150.9 | 0.2 | 151.1 | 46.1 | 46.0511 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.9 | 59.836 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4645 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.932 | 0.0095 | 95.4442 | 150.9 | 0.2 | 151.1 | 46 | 46 | 0.1 | 12.8 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 195.5 | 59.8641 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6413 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.939 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.0606 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.1 | 59.9 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.931 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 195.3 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.937 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46 | 46 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.6 | 59.864 | 0.0252 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0252 | 0.000436 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 150.9 | 0.2 | 151.1 | 46.1 | 46.0706 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.9 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4646 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151 | 0.2 | 151.2 | 46 | 46.0308 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4139 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151.1 | 0.2 | 151.3 | 46.1 | 46.1 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4924 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151.1 | 0.2 | 151.3 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8639 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151.1 | 0.2 | 151.3 | 46 | 46 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.836 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4506 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151 | 0.2 | 151.2 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4278 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151 | 0.2 | 151.2 | 46 | 46 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.94 | 0.01 | 100 | 151 | 0.2 | 151.2 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.864 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 2.962 | 0.0107 | 107.4074 | 151 | 0.2 | 151.2 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0252 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.466 | 0.0118 | 117.9174 | 160.4 | 0.2 | 160.6 | 46 | 46 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.026605 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4583 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.527 | 0.011 | 110 | 174.2 | 0.2 | 174.4 | 46.4 | 46.41 | 0.1 | 12.9 | 920 | 908 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.028684 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4202 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.53 | 0.011 | 110 | 177.4 | 0.2 | 177.6 | 52.8 | 52.7911 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.02922 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4986 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.534 | 0.011 | 110 | 178.8 | 0.2 | 179 | 54.9 | 54.8925 | 0.1 | 13 | 919 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.029464 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4229 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.54 | 0.011 | 110 | 178.9 | 0.2 | 179.1 | 55.2 | 55.2402 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0295 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.54 | 0.011 | 110 | 179.1 | 0.2 | 179.3 | 55.3 | 55.3 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0295 | 0.000464 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4339 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.54 | 0.0114 | 113.9033 | 179.2 | 0.2 | 179.4 | 55.3 | 55.3 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.54 | 0.013 | 129.7049 | 179.2 | 0.2 | 179.4 | 55.3 | 55.3489 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.54 | 0.0121 | 121.2732 | 179.2 | 0.2 | 179.4 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.544 | 0.012 | 120 | 179.3 | 0.2 | 179.5 | 55.4 | 55.4 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 179.3 | 0.2 | 179.5 | 55.5 | 55.5 | 0.1 | 12.9 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 179.4 | 0.2 | 179.6 | 55.5 | 55.5 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 179.7 | 0.2 | 179.9 | 55.6 | 55.6 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.029664 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 179.9 | 0.2 | 180.1 | 55.6 | 55.6 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.864 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 179.9 | 0.2 | 180.1 | 55.6 | 55.6 | 0.1 | 12.8 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.0106 | 105.7458 | 179.9 | 0.2 | 180.1 | 55.6 | 55.6 | 0.1 | 12.8 | 921 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 195.9 | 59.836 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.01 | 100 | 180 | 0.2 | 180.2 | 55.6 | 55.6 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.6 | 59.9 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.0104 | 103.9038 | 180.1 | 0.2 | 180.3 | 55.7 | 55.7 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.2 | 59.8 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4243 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.0112 | 112.069 | 180.2 | 0.2 | 180.4 | 55.7 | 55.7 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 195.1 | 59.864 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4542 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.7 | 55.7 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 195.5 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.6 | 55.6 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.9 | 59.8 | 0.029736 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4893 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.554 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.6 | 55.6279 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.864 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4105 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.6 | 55.6 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4681 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.6 | 55.6 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.6 | 55.5849 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.5 | 55.5 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.5 | 55.5 | 0.1 | 12.9 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.029764 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.56 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 919 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3613 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.562 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.569 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.552 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.557 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.555 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.864 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.554 | 0.012 | 120 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13 | 918 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:35 | 3.558 | 0.012 | 120 | 180.4 | 0.3 | 180.7 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 180.4 | 0.3 | 180.7 | 55.4 | 55.4 | 0.1 | 12.8 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 180.5 | 0.3 | 180.8 | 55.4 | 55.4 | 0.1 | 12.8 | 920 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8356 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:35 | 3.55 | 0.012 | 120 | 180.6 | 0.2 | 180.8 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4548 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.5 | 0.2 | 180.7 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.6 | 0.2 | 180.8 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4021 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.6 | 0.2 | 180.8 | 55.4 | 55.4 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.8 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4806 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.6 | 0.2 | 180.8 | 55.5 | 55.469 | 0.1 | 12.8 | 921 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8644 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6439 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.029864 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0299 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.0299 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.8 | 0.2 | 181 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0299 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.5 | 55.4686 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.864 | 0.0299 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4485 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.8 | 0.2 | 181 | 55.4 | 55.4 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8644 | 0.0299 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.8 | 0.2 | 181 | 55.4 | 55.4 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.9 | 59.836 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4084 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.3 | 55.3303 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.6 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4868 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.3 | 55.3 | 0.1 | 12.9 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.2 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.0127 | 127.1184 | 180.7 | 0.2 | 180.9 | 55.3 | 55.3 | 0.1 | 13 | 919 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.013 | 130 | 180.7 | 0.2 | 180.9 | 55.3 | 55.3 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 195.5 | 59.9 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4778 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.55 | 0.013 | 130 | 180.7 | 0.2 | 180.9 | 55.3 | 55.3315 | 0.1 | 13 | 918 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.9 | 59.864 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4007 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.559 | 0.013 | 130 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 13 | 917 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.8 | 59.864 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4789 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.56 | 0.0129 | 129.4127 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.4 | 59.8361 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.56 | 0.0121 | 121.3677 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.1 | 59.8 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.56 | 0.012 | 120 | 180.7 | 0.2 | 180.9 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.3 | 59.864 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 3.685 | 0.0125 | 125.3431 | 180.8 | 0.2 | 181 | 55.4 | 55.4 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.7 | 59.8641 | 0.0298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 4.2 | 0.013 | 130 | 190.3 | 0.2 | 190.5 | 55.4 | 55.4 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.031239 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 4.191 | 0.013 | 130 | 203.8 | 0.2 | 204 | 57 | 56.955 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.033248 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 207.7 | 0.26393 | 208 | 62.8 | 62.7576 | 0.1 | 12.7 | 924 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0338 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4715 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 208.5 | 0.3 | 208.8 | 64.2 | 64.2063 | 0.1 | 12.7 | 924 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.8 | 0.033964 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4138 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 209.5 | 0.3 | 209.8 | 64.6 | 64.6 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.9 | 0.0341 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.6452 | 20.5708 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 210 | 0.3 | 210.3 | 64.6 | 64.6 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5361 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 210.1 | 0.3 | 210.4 | 64.7 | 64.7 | 0.1 | 12.7 | 923 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4576 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 210.1 | 0.3 | 210.4 | 64.7 | 64.7 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4209 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.183 | 0.013 | 130 | 210.1 | 0.3 | 210.4 | 64.7 | 64.7 | 0.1 | 12.8 | 922 | 907 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4993 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.1 | 0.3 | 210.4 | 64.7 | 64.7 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 195.7 | 59.836 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4222 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.184 | 0.013 | 130 | 210.2 | 0.3 | 210.5 | 64.8 | 64.7709 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4559 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 210.2 | 0.3 | 210.5 | 64.8 | 64.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.19 | 0.013 | 130 | 210.3 | 0.3 | 210.6 | 64.8 | 64.8 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.173 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.8 | 64.7792 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.8 | 59.9 | 0.0342 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.7 | 64.7 | 0.1 | 12.8 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.7 | 64.7 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.7 | 64.7 | 0.1 | 12.9 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.6387 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.7 | 64.6691 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.6 | 64.6 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.5 | 0.3 | 210.8 | 64.6 | 64.6 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5624 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.4 | 0.3 | 210.7 | 64.6 | 64.6 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5188 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.5 | 0.3 | 210.8 | 64.6 | 64.6 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4188 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.5 | 0.3 | 210.8 | 64.6 | 64.5672 | 0.1 | 13 | 919 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.64 | 20.4978 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.6 | 0.3 | 210.9 | 64.5 | 64.5294 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055.4028 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.6 | 0.3 | 210.9 | 64.5 | 64.5 | 0.1 | 13 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4453 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.6 | 0.3 | 210.9 | 64.5 | 64.5 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4331 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0343 | 0.000564 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.6 | 64.5733 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 196 | 59.9 | 0.0343 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.836 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0343 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.36 | 20.4744 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.17 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.7 | 59.9 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.177 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64.5 | 64.5 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.3 | 59.8359 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.8 | 0.3 | 211.1 | 64.5 | 64.5 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195 | 59.864 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.8 | 0.3 | 211.1 | 64.5 | 64.5 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 195 | 59.9 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4394 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.8 | 0.3 | 211.1 | 64.5 | 64.5297 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 195 | 59.9 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4822 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.8 | 0.3 | 211.1 | 64.7 | 64.7461 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.836 | 0.0344 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4036 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 210.9 | 0.3 | 211.2 | 64.6 | 64.6 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.5 | 59.836 | 0.0345 | 0.000664 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:36 | 4.18 | 0.013 | 130 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 12.8 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4533 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.013 | 130 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.187 | 0.013 | 130 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.864 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4898 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.2 | 0.0135 | 135.1626 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.8 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4113 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.2 | 0.014 | 140 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.183 | 0.014 | 140 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4456 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211 | 0.3 | 211.3 | 64.6 | 64.6412 | 0.1 | 13 | 917 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.836 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211 | 0.3 | 211.3 | 64.7 | 64.6675 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.6 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211 | 0.3 | 211.3 | 64.6 | 64.6 | 0.1 | 13 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.2 | 59.836 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.6 | 64.6499 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.1 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4405 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.5 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.3574 | 20.438 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.9 | 59.864 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.6387 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.3 | 211.4 | 64.7 | 64.7 | 0.1 | 12.7 | 922 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.864 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.236 | 211.3 | 64.7 | 64.7 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.64 | 20.5087 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.36 | 20.5867 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.469 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7295 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4442 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.2 | 211.3 | 64.8 | 64.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.18 | 0.014 | 140 | 211.1 | 0.2 | 211.3 | 64.8 | 64.8 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0345 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.188 | 0.0151 | 150.7382 | 211.1 | 0.2 | 211.3 | 64.9 | 64.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.9 | 0.0346 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.752 | 0.0166 | 165.7824 | 212.7 | 0.2 | 212.9 | 64.9 | 64.9 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.7 | 59.864 | 0.034792 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.829 | 0.015 | 150.463 | 231.7 | 0.2 | 231.9 | 65.7 | 65.6805 | 0.1 | 12.7 | 922 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.4 | 59.9 | 0.037568 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5365 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.0159 | 158.5865 | 235.8 | 0.2 | 236 | 71.3 | 71.2945 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 195 | 59.9 | 0.0381 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.57 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 238.3 | 0.2 | 238.5 | 73.7 | 73.6623 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.3 | 59.9 | 0.0385 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4132 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 238.5 | 0.2 | 238.7 | 74 | 74.0299 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.7 | 59.8641 | 0.0385 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.0156 | 156.4427 | 238.7 | 0.2 | 238.9 | 74.1 | 74.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0386 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 238.9 | 0.2 | 239.1 | 74.1 | 74.1 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0386 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4289 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.015 | 150.1095 | 238.9 | 0.2 | 239.1 | 74.1 | 74.1 | 0.1 | 12.9 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0386 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.0159 | 158.5341 | 238.9 | 0.2 | 239.1 | 74.1 | 74.1 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0386 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:37 | 4.82 | 0.0153 | 153.196 | 239.1 | 0.264 | 239.3 | 74.1 | 74.1 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0386 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.0155 | 155.0958 | 239.2 | 0.3 | 239.5 | 74.1 | 74.1 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.038664 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 239.2 | 0.3 | 239.5 | 74.1 | 74.0726 | 0.1 | 13 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0387 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 239.3 | 0.3 | 239.6 | 74 | 74.0331 | 0.1 | 13.1 | 917 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0387 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 239.4 | 0.3 | 239.7 | 74.1 | 74.0599 | 0.1 | 13.1 | 917 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0387 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.016 | 160 | 239.5 | 0.3 | 239.8 | 74 | 74.0327 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0387 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4435 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.0153 | 153.0936 | 239.5 | 0.3 | 239.8 | 74.1 | 74.1 | 0.1 | 13.1 | 917 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8641 | 0.0387 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4349 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 239.6 | 0.3 | 239.9 | 74.1 | 74.1 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4865 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 239.7 | 0.3 | 240 | 74.1 | 74.1 | 0.1 | 12.9 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.408 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 239.7 | 0.3 | 240 | 74 | 74 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 239.8 | 0.3 | 240.1 | 74 | 74.0479 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190.2 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055.9694 | 1017 | 20.4486 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.815 | 0.015 | 150 | 239.9 | 0.3 | 240.2 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.81 | 0.015 | 150 | 239.9 | 0.3 | 240.2 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.81 | 0.015 | 150 | 240 | 0.3 | 240.3 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.8361 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017.6387 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.81 | 0.015 | 150 | 240.1 | 0.3 | 240.4 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4373 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.819 | 0.0152 | 152.0241 | 240.2 | 0.3 | 240.5 | 73.9 | 73.9471 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8361 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4412 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.813 | 0.016 | 160 | 240.2 | 0.3 | 240.5 | 73.9 | 73.9301 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.81 | 0.016 | 160 | 240.2 | 0.3 | 240.5 | 73.9 | 73.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0388 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.814 | 0.0146 | 145.9881 | 240.3 | 0.3 | 240.6 | 73.9 | 73.9 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.0145 | 145.3164 | 240.3 | 0.3 | 240.6 | 73.9 | 73.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5551 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 240.3 | 0.3 | 240.6 | 73.9 | 73.9 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.9 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5329 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.82 | 0.015 | 150 | 240.3 | 0.3 | 240.6 | 73.9 | 73.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.412 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.834 | 0.015 | 150 | 240.3 | 0.3 | 240.6 | 73.8 | 73.8 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4905 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.835 | 0.015 | 150 | 240.3 | 0.3 | 240.6 | 73.8 | 73.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4311 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.834 | 0.015 | 150 | 240.4 | 0.3 | 240.7 | 73.8 | 73.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0389 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.8 | 73.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4258 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 12.9 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8361 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:37 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 13 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8361 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.015 | 150 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.015 | 150 | 240.6 | 0.3 | 240.9 | 73.8 | 73.8 | 0.1 | 13.1 | 917 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.0152 | 152.0622 | 240.5 | 0.3 | 240.8 | 73.7 | 73.7305 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4465 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.016 | 160 | 240.6 | 0.3 | 240.9 | 73.8 | 73.8 | 0.1 | 13.1 | 917 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4319 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.016 | 160 | 240.7 | 0.3 | 241 | 73.8 | 73.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.84 | 0.016 | 160 | 240.7 | 0.3 | 241 | 73.9 | 73.9 | 0.1 | 12.9 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.836 | 0.016 | 160 | 240.8 | 0.3 | 241.1 | 73.9 | 73.9 | 0.1 | 12.8 | 920 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.8361 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4326 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.82 | 0.016 | 160 | 240.9 | 0.3 | 241.2 | 73.9 | 73.9 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.836 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4459 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.82 | 0.016 | 160 | 240.9 | 0.3 | 241.2 | 73.9 | 73.9 | 0.1 | 12.7 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1015.36 | 20.4757 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.813 | 0.016 | 160 | 240.9 | 0.3 | 241.2 | 73.9 | 73.9416 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055.995 | 1015.64 | 20.4027 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.9 | 0.3 | 241.2 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4812 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.8 | 0.3 | 241.1 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.8 | 0.3 | 241.1 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.8361 | 0.039 | 0.000836 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.8 | 0.3 | 241.1 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8361 | 0.039 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.8 | 0.36387 | 241.2 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.039 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.9 | 0.4 | 241.3 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 240.9 | 0.4 | 241.3 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 4.81 | 0.016 | 160 | 241 | 0.4 | 241.4 | 74 | 74 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.039 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 4.934 | 0.0165 | 165.128 | 241 | 0.4 | 241.4 | 74.1 | 74.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.039 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4911 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.449 | 0.0174 | 173.605 | 250.3 | 0.4 | 250.7 | 74.1 | 74.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.040124 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4128 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.441 | 0.018 | 180 | 262.4 | 0.4 | 262.8 | 75.6 | 75.6354 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.042064 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4657 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.44 | 0.018 | 180 | 268.4 | 0.4 | 268.8 | 81.2 | 81.1808 | 0.1 | 12.7 | 924 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.042928 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.44 | 0.0171 | 171.4846 | 269.5 | 0.4 | 269.9 | 83.1 | 83.0562 | 0.1 | 12.7 | 924 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.443 | 0.017 | 170 | 269.5 | 0.4 | 269.9 | 83.4 | 83.3515 | 0.1 | 12.9 | 922 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.6 | 0.4 | 270 | 83.4 | 83.4 | 0.1 | 13 | 920 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8361 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.6 | 0.4 | 270 | 83.4 | 83.4 | 0.1 | 13 | 920 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.7 | 0.4 | 270.1 | 83.5 | 83.5 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.8 | 0.4 | 270.2 | 83.5 | 83.5296 | 0.1 | 13.1 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.836 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.8 | 0.4 | 270.2 | 83.6 | 83.6 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.8 | 0.4 | 270.2 | 83.6 | 83.6295 | 0.1 | 13.1 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8361 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1056 | 1017 | 20.5718 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 269.9 | 0.4 | 270.3 | 83.7 | 83.7 | 0.1 | 13 | 919 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0431 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5498 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270 | 0.4 | 270.4 | 83.7 | 83.7 | 0.1 | 13.1 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.043164 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.1 | 0.4 | 270.5 | 83.7 | 83.7 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.3 | 0.4 | 270.7 | 83.7 | 83.7312 | 0.1 | 12.9 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.638 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.4 | 0.4 | 270.8 | 83.7 | 83.7478 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.043264 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.5 | 0.4 | 270.9 | 83.7 | 83.7298 | 0.1 | 12.8 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4578 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.5 | 0.4 | 270.9 | 83.8 | 83.8 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1056 | 1016.64 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.5 | 0.336 | 270.8 | 83.8 | 83.8 | 0.1 | 12.8 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.5 | 0.3 | 270.8 | 83.8 | 83.7729 | 0.1 | 12.8 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.8 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.6 | 0.3 | 270.9 | 83.7 | 83.7 | 0.1 | 12.7 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.836 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4563 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.6 | 0.3 | 270.9 | 83.8 | 83.8 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055.4203 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.6 | 0.3 | 270.9 | 83.8 | 83.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4867 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.6 | 0.4 | 271 | 83.8 | 83.8 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4082 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.6 | 0.4 | 271 | 83.8 | 83.8 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.7 | 0.4 | 271.1 | 83.8 | 83.8286 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190.8 | 5 | 196 | 59.864 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.7 | 0.4 | 271.1 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0432 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4267 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.8 | 0.4 | 271.2 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.9 | 0.4 | 271.3 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.9 | 0.4 | 271.3 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 270.9 | 0.4 | 271.3 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 271 | 0.4 | 271.4 | 83.9 | 83.9 | 0.1 | 12.9 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1017 | 20.4808 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 271.1 | 0.4 | 271.5 | 83.9 | 83.9 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.36 | 20.4022 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 271 | 0.4 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 919 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4763 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.017 | 170 | 271 | 0.4 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.0165 | 164.9053 | 271 | 0.4 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 919 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.45 | 0.016 | 160 | 271 | 0.4 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 919 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:38 | 5.433 | 0.016 | 160 | 271 | 0.3 | 271.3 | 83.9 | 83.9 | 0.1 | 13 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:38 | 5.443 | 0.016 | 160 | 271.1 | 0.3 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 918 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:39 | 5.45 | 0.016 | 160 | 271.1 | 0.3 | 271.4 | 83.9 | 83.9 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:39 | 5.45 | 0.016 | 160 | 271.1 | 0.3 | 271.4 | 83.9 | 83.9 | 0.1 | 13.1 | 918 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4746 | 27.9 | 0.3
2008-11-05 | 10:39 | 5.45 | 0.0165 | 164.9206 | 271.1 | 0.3 | 271.4 | 83.9 | 83.9 | 0.1 | 12.9 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4037 | 27.9 | 0.3
2008-11-05 | 10:39 | 5.45 | 0.017 | 170 | 271.1 | 0.4 | 271.5 | 83.9 | 83.9483 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.36 | 20.4823 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.45 | 0.017 | 170 | 271.1 | 0.4 | 271.5 | 83.9 | 83.9154 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5608 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.45 | 0.017 | 170 | 271.1 | 0.4 | 271.5 | 84 | 84 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.8 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.36 | 20.5608 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.455 | 0.017 | 170 | 271.2 | 0.4 | 271.6 | 84 | 84 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 193 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.2 | 0.4 | 271.6 | 84 | 84 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192.2 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.2 | 0.4 | 271.6 | 84 | 84 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 84 | 84 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.447 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 84 | 84 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.463 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.59 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.8 | 921 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5116 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 196 | 59.9 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.8 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0433 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1016.36 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.2 | 0.4 | 271.6 | 83.9 | 83.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.46 | 0.017 | 170 | 271.3 | 0.4 | 271.7 | 83.9 | 83.9 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.441 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.61 | 0.017 | 170 | 271.2 | 0.4 | 271.6 | 83.9 | 83.9 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0434 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 5.929 | 0.017 | 170 | 273.6 | 0.4 | 274 | 83.9 | 83.9 | 0.1 | 12.9 | 920 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.043719 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.068 | 0.0173 | 173.3052 | 286.5 | 0.4 | 286.9 | 84 | 83.9625 | 0.1 | 13 | 919 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.045511 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 297.5 | 0.4 | 297.9 | 88 | 88.0346 | 0.1 | 13 | 921 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.046964 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4729 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 299 | 0.4 | 299.4 | 92.4 | 92.3733 | 0.1 | 13 | 922 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0472 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.06 | 0.018 | 180 | 299.6 | 0.4 | 300 | 93.1 | 93.126 | 0.1 | 13 | 923 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0473 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4703 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.05 | 0.0187 | 186.6102 | 299.7 | 0.3 | 300 | 93.3 | 93.2552 | 0.1 | 13 | 922 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0473 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4082 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.05 | 0.019 | 190 | 299.8 | 0.3 | 300.1 | 93.3 | 93.2709 | 0.1 | 13 | 922 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0473 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.4866 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.05 | 0.019 | 190 | 299.9 | 0.364 | 300.2 | 93.2 | 93.2294 | 0.1 | 13 | 922 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0473 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.059 | 0.019 | 190 | 300 | 0.4 | 300.4 | 93.3 | 93.3 | 0.1 | 13 | 921 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.9 | 59.8 | 0.0473 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.067 | 0.019 | 190 | 300.3 | 0.4 | 300.7 | 93.3 | 93.3 | 0.1 | 13 | 921 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.5 | 59.864 | 0.047364 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1056 | 1016.36 | 20.478 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.065 | 0.019 | 190 | 300.6 | 0.4 | 301 | 93.3 | 93.3 | 0.1 | 12.8 | 922 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.2 | 59.8 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4005 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.06 | 0.019 | 190 | 300.6 | 0.4 | 301 | 93.3 | 93.3487 | 0.1 | 12.7 | 923 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4788 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.062 | 0.019 | 190 | 300.6 | 0.4 | 301 | 93.4 | 93.4 | 0.1 | 12.8 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 195.5 | 59.8 | 0.0474 | 0.000964 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4427 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.019 | 190 | 300.6 | 0.4 | 301 | 93.3 | 93.3 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 195.9 | 59.836 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.019 | 190 | 300.7 | 0.4 | 301.1 | 93.4 | 93.4 | 0.1 | 12.8 | 922 | 908 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.4142 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.019 | 190 | 300.6 | 0.4 | 301 | 93.2 | 93.2269 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4927 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.019 | 190 | 300.7 | 0.4 | 301.1 | 93.4 | 93.3531 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.0183 | 183.3088 | 300.7 | 0.4 | 301.1 | 93.4 | 93.4 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.4505 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.8 | 0.4 | 301.2 | 93.3 | 93.3357 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.3 | 93.3 | 0.1 | 12.7 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1015.3613 | 20.4064 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.062 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.3 | 93.3 | 0.1 | 12.8 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4849 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.053 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.3 | 93.2686 | 0.1 | 12.8 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1015.6393 | 20.4366 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.06 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.2 | 93.2 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4418 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.2 | 93.2 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.8 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.8 | 0.4 | 301.2 | 93.2 | 93.1581 | 0.1 | 12.7 | 921 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.8 | 0.4 | 301.2 | 93.1 | 93.1 | 0.1 | 12.7 | 922 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5772 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.1 | 93.1 | 0.1 | 12.7 | 922 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.836 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5446 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.336 | 301.2 | 93.1 | 93.1 | 0.1 | 12.9 | 920 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.36 | 20.4662 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.3 | 301.2 | 93.1 | 93.1 | 0.1 | 12.9 | 920 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4125 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.077 | 0.018 | 180 | 300.9 | 0.3 | 301.2 | 93.1 | 93.1 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.4909 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.087 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93.1 | 93.1 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0474 | 0.000936 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4306 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93 | 93 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.4479 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.07 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93 | 93 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.3613 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.083 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93 | 93 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.64 | 20.4952 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.09 | 0.018 | 180 | 300.9 | 0.4 | 301.3 | 93 | 93 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4168 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.09 | 0.018 | 180 | 301 | 0.4 | 301.4 | 93 | 93 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.6387 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.09 | 0.018 | 180 | 301.1 | 0.3 | 301.4 | 93 | 93 | 0.1 | 13 | 919 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.36 | 20.4399 | 27.9 | 0.3
2008-11-05 | 10:39 | 6.09 | 0.0183 | 183.2075 | 301.1 | 0.36393 | 301.5 | 93 | 93 | 0.1 | 12.8 | 921 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.09 | 0.019 | 190 | 301.1 | 0.4 | 301.5 | 93 | 93 | 0.1 | 12.8 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.836 | 0.0474 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.3597 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:39 | 6.09 | 0.019 | 190 | 301.2 | 0.4 | 301.6 | 93 | 93 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.9 | 0.0475 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4245 | 27.9 | 0.4
2008-11-05 | 10:40 | 6.07 | 0.019 | 190 | 301.2 | 0.4 | 301.6 | 93 | 93 | 0.1 | 12.7 | 923 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0475 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4539 | 27.9 | 0.4
2008-11-05 | 10:40 | 6.047 | 0.019 | 190 | 301.3 | 0.4 | 301.7 | 93 | 93 | 0.1 | 12.7 | 922 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0475 | 0.000964 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 6.04 | 0.019 | 190 | 301.3 | 0.336 | 301.6 | 93 | 93 | 0.1 | 12.8 | 922 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.836 | 0.047564 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4892 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.04 | 0.019 | 190 | 301.3 | 0.3 | 301.6 | 93 | 93 | 0.1 | 12.7 | 921 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0476 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4107 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.048 | 0.0188 | 188.2944 | 301.3 | 0.3 | 301.6 | 93 | 93 | 0.1 | 12.8 | 921 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0476 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.05 | 0.018 | 180.2494 | 301.3 | 0.3 | 301.6 | 93 | 92.9685 | 0.1 | 12.8 | 921 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0476 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.3538 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.05 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.9 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0476 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.3491 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.05 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.9 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0476 | 0.001064 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.052 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.9 | 0.1 | 12.7 | 921 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 195.7 | 59.8639 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.06 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.9 | 0.1 | 12.8 | 921 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.864 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.076 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.9 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195 | 59.836 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4616 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.08 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.9 | 92.8868 | 0.1 | 12.8 | 921 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.4 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4168 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.075 | 0.018 | 180 | 301.4 | 0.3 | 301.7 | 92.8 | 92.8 | 0.1 | 12.7 | 921 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.7 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016 | 20.4953 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.077 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.8 | 92.8 | 0.1 | 12.8 | 921 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1016.6393 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.09 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.8 | 92.8 | 0.1 | 12.9 | 920 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4478 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.09 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.8 | 92.8 | 0.1 | 13 | 919 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.4307 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.09 | 0.018 | 180 | 301.4 | 0.3 | 301.7 | 92.8 | 92.8 | 0.1 | 13 | 919 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.09 | 0.018 | 180 | 301.3 | 0.3 | 301.6 | 92.8 | 92.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.864 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.085 | 0.018 | 180 | 301.4 | 0.3 | 301.7 | 92.8 | 92.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 6.083 | 0.018 | 180 | 301.4 | 0.3 | 301.7 | 92.8 | 92.8 | 0.1 | 13 | 918 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0476 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1056 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.981 | 0.018 | 180 | 301.5 | 0.36393 | 301.8 | 92.8 | 92.8 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.0476 | 0.001036 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.476 | 0.0183 | 183.2941 | 294.7 | 0.4 | 295.1 | 92.8 | 92.8 | 0.1 | 13 | 919 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.046704 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4988 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.0188 | 188.363 | 279.5 | 0.4 | 279.9 | 92 | 91.9737 | 0.1 | 13 | 918 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.044816 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054.5324 | 1016 | 20.4209 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180.2593 | 275.3 | 0.4 | 275.7 | 86.1 | 86.1077 | 0.1 | 12.8 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.8 | 0.044036 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 274.8 | 0.4 | 275.2 | 84 | 83.953 | 0.1 | 12.8 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.044 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 274.1 | 0.4 | 274.5 | 83.4 | 83.4 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0439 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.9 | 0.4 | 274.3 | 83.3 | 83.3348 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.9 | 0.0438 | 0.001064 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.8 | 0.4 | 274.2 | 83.3 | 83.3 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.8639 | 0.0438 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4717 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.9 | 0.3 | 274.2 | 83.3 | 83.3 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0438 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.8 | 0.3 | 274.1 | 83.3 | 83.3 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0438 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4711 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.7 | 0.3 | 274 | 83.3 | 83.3 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0438 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.5 | 0.3 | 273.8 | 83.2 | 83.2 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190.8 | 5 | 196 | 59.864 | 0.043736 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.2 | 0.3 | 273.5 | 83.2 | 83.1684 | 0.1 | 12.7 | 917 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8361 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4637 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 273.2 | 0.3 | 273.5 | 83.2 | 83.2 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4578 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.0178 | 178.3387 | 273 | 0.3 | 273.3 | 83.1 | 83.1 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4211 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.017 | 170.2742 | 272.9 | 0.4 | 273.3 | 83.1 | 83.1 | 0.1 | 12.8 | 917 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4996 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.017 | 170 | 272.9 | 0.4 | 273.3 | 83.1 | 83.073 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.017 | 170 | 272.9 | 0.4 | 273.3 | 83 | 83 | 0.1 | 12.8 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.017 | 170 | 272.8 | 0.4 | 273.2 | 83 | 83 | 0.1 | 12.9 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4653 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.0173 | 173.3249 | 272.8 | 0.4 | 273.2 | 82.9 | 82.9 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.8 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4133 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 272.8 | 0.4 | 273.2 | 82.9 | 82.9 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195.4 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4918 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 272.8 | 0.4 | 273.2 | 82.9 | 82.9 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 272.8 | 0.4 | 273.2 | 83 | 83 | 0.1 | 13.1 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.3 | 59.864 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.018 | 180 | 272.8 | 0.4 | 273.2 | 82.9 | 82.9 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.7 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4734 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.0175 | 175.1877 | 272.8 | 0.4 | 273.2 | 82.9 | 82.9 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4056 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 82.9 | 82.9 | 0.1 | 13 | 915 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0437 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6393 | 20.4841 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.436 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 82.9 | 82.9 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.8361 | 0.043636 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4375 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.443 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 82.9 | 82.8703 | 0.1 | 12.8 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8361 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4408 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 83 | 83 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8359 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4808 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 83.1 | 83.0645 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4022 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 83 | 83.0307 | 0.1 | 12.7 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.8 | 0.3 | 273.1 | 83 | 83 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.7 | 0.3 | 273 | 83 | 83 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.64 | 20.4329 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.017 | 170 | 272.6 | 0.3 | 272.9 | 83.1 | 83.0684 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.864 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4889 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.45 | 0.0162 | 161.99 | 272.6 | 0.4 | 273 | 83 | 83 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8641 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.41 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.453 | 0.016 | 160 | 272.6 | 0.4 | 273 | 83 | 83 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.4
2008-11-05 | 10:40 | 5.46 | 0.016 | 160 | 272.5 | 0.336 | 272.8 | 83 | 83 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.447 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.46 | 0.016 | 160 | 272.6 | 0.3 | 272.9 | 83.1 | 83.0692 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190.8 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4746 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.452 | 0.016 | 160 | 272.6 | 0.3 | 272.9 | 83 | 83.0499 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4036 | 27.9 | 0.3
2008-11-05 | 10:40 | 5.456 | 0.016 | 160 | 272.6 | 0.3 | 272.9 | 83.1 | 83.0856 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4822 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.455 | 0.016 | 160 | 272.5 | 0.3 | 272.8 | 83 | 83 | 0.1 | 12.7 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.5 | 0.3 | 272.8 | 83.1 | 83.1 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8361 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4217 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.5 | 0.3 | 272.8 | 83.1 | 83.1 | 0.1 | 12.9 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.3352 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.4 | 0.3 | 272.7 | 83.1 | 83.1 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4923 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.4 | 0.3 | 272.7 | 83.3 | 83.3 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.4 | 0.3 | 272.7 | 83.4 | 83.4 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.016 | 160 | 272.4 | 0.3 | 272.7 | 83.4 | 83.4 | 0.1 | 13.1 | 914 | 904 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.0166 | 166.4983 | 272.4 | 0.3 | 272.7 | 83.4 | 83.4271 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0436 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.017 | 170 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 13.1 | 914 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8361 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.0163 | 163.4007 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 13.1 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.015 | 150 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 13.1 | 914 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.015 | 150 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.9 | 916 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4764 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.455 | 0.015 | 150 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.402 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.46 | 0.015 | 150 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4806 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.46 | 0.0155 | 154.7766 | 272.4 | 0.3 | 272.7 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4409 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.46 | 0.016 | 160 | 272.4 | 0.3 | 272.7 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.7 | 59.8639 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4375 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.452 | 0.0161 | 161.2927 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5312 | 0.1 | 12.7 | 920 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.3 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.017 | 169.633 | 272.5 | 0.3 | 272.8 | 83.6 | 83.5613 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.1 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.017 | 170 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.4 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.45 | 0.017 | 170 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.8 | 59.8 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.382 | 0.017 | 170 | 272.5 | 0.3 | 272.8 | 83.5 | 83.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4704 | 27.9 | 0.3
2008-11-05 | 10:41 | 5.031 | 0.0163 | 163.4503 | 272.2 | 0.23613 | 272.4 | 83.5 | 83.5 | 0.1 | 12.7 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.9 | 0.0437 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6387 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:41 | 4.861 | 0.015 | 150 | 256.7 | 0.264 | 256.9 | 83.5 | 83.5 | 0.1 | 12.7 | 920 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.04152 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 248 | 0.3 | 248.3 | 80.6 | 80.5713 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.040244 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 244.9 | 0.3 | 245.2 | 75.5 | 75.5018 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.039836 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4435 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.823 | 0.015 | 150 | 244.4 | 0.3 | 244.7 | 74.4 | 74.4 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.039736 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.827 | 0.015 | 150 | 244.1 | 0.3 | 244.4 | 74.2 | 74.2 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.7 | 59.864 | 0.0397 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.811 | 0.015 | 150 | 244.1 | 0.3 | 244.4 | 74.1 | 74.1 | 0.1 | 12.9 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.3 | 59.8359 | 0.0397 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.81 | 0.015 | 150 | 243.9 | 0.3 | 244.2 | 74.1 | 74.1 | 0.1 | 13 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.9 | 0.0397 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4427 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.81 | 0.015 | 150 | 243.9 | 0.3 | 244.2 | 74.1 | 74.1 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.4 | 59.836 | 0.0397 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4357 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.83 | 0.015 | 150 | 243.8 | 0.3 | 244.1 | 74 | 74.0303 | 0.1 | 13 | 913 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.8 | 59.8361 | 0.0396 | 0.001036 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.84 | 0.015 | 150 | 243.8 | 0.3 | 244.1 | 74 | 74 | 0.1 | 13.1 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0396 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 243.8 | 0.3 | 244.1 | 74 | 74 | 0.1 | 13.1 | 913 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0396 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 243.4 | 0.3 | 243.7 | 74 | 74 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.039536 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.836 | 0.015 | 150 | 243.3 | 0.3 | 243.6 | 74 | 74 | 0.1 | 13.1 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.001064 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4721 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.84 | 0.015 | 150 | 243.2 | 0.3 | 243.5 | 74 | 74 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.835 | 0.015 | 150 | 243.1 | 0.3 | 243.4 | 74 | 74 | 0.1 | 12.9 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6393 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.83 | 0.015 | 150 | 243.1 | 0.3 | 243.4 | 74 | 73.9673 | 0.1 | 12.8 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4633 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.83 | 0.0145 | 145.0751 | 243 | 0.3 | 243.3 | 73.9 | 73.9 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190.2 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4583 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.83 | 0.0143 | 143.2803 | 243 | 0.3 | 243.3 | 73.9 | 73.9 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.822 | 0.015 | 150 | 242.9 | 0.3 | 243.2 | 73.9 | 73.9285 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 242.9 | 0.3 | 243.2 | 74 | 74 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.477 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.0158 | 157.9142 | 242.9 | 0.3 | 243.2 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.9 | 0.3 | 243.2 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4662 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.9 | 0.3 | 243.2 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.864 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4124 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.8 | 0.3 | 243.1 | 74 | 74 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4909 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.8 | 0.3 | 243.1 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.7 | 0.3 | 243 | 74 | 74 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4522 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.7 | 0.3 | 243 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.817 | 0.016 | 160 | 242.7 | 0.3 | 243 | 74 | 74 | 0.1 | 12.7 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4046 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.811 | 0.016 | 160 | 242.7 | 0.3 | 243 | 74 | 74 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4831 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.8 | 0.3 | 243.1 | 74 | 73.9701 | 0.1 | 12.7 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.8 | 0.3 | 243.1 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8639 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.4602 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.016 | 160 | 242.7 | 0.3 | 243 | 73.9 | 73.9 | 0.1 | 12.9 | 917 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4181 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.0152 | 151.9175 | 242.7 | 0.3 | 243 | 73.9 | 73.9 | 0.1 | 13 | 916 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4968 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.817 | 0.015 | 150 | 242.8 | 0.3 | 243.1 | 73.9 | 73.9 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.811 | 0.015 | 150 | 242.8 | 0.3 | 243.1 | 73.9 | 73.9 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5538 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 242.8 | 0.3 | 243.1 | 73.9 | 73.9 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5679 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 242.8 | 0.3 | 243.1 | 73.9 | 73.9 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8361 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:41 | 4.82 | 0.015 | 150 | 242.7 | 0.3 | 243 | 73.9 | 73.9 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.815 | 0.015 | 150 | 242.7 | 0.3 | 243 | 73.9 | 73.9 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.807 | 0.015 | 150 | 242.7 | 0.3 | 243 | 73.8 | 73.8346 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.7 | 0.2 | 242.9 | 73.9 | 73.9 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0395 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4757 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.7 | 0.2 | 242.9 | 73.9 | 73.9 | 0.1 | 12.9 | 916 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0395 | 0.001036 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4026 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.808 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.8 | 917 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4813 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.804 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.8688 | 0.1 | 12.8 | 918 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.6 | 0.26387 | 242.9 | 73.9 | 73.9 | 0.1 | 12.8 | 918 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.797 | 0.0156 | 156.2046 | 242.6 | 0.3 | 242.9 | 73.9 | 73.9 | 0.1 | 12.8 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4835 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.792 | 0.0156 | 155.5335 | 242.6 | 0.3 | 242.9 | 73.9 | 73.9 | 0.1 | 12.8 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.8361 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4055 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.799 | 0.015 | 150 | 242.6 | 0.3 | 242.9 | 73.9 | 73.9 | 0.1 | 12.8 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8639 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4734 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.6 | 0.3 | 242.9 | 73.9 | 73.9 | 0.1 | 12.7 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.836 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.3
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.8361 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4696 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.7 | 0.2 | 242.9 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.8 | 0.015 | 150 | 242.7 | 0.2 | 242.9 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.836 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.64 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.801 | 0.015 | 150 | 242.7 | 0.2 | 242.9 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4657 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.809 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.81 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.81 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4989 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.828 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.8697 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.421 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.772 | 0.015 | 150 | 242.6 | 0.2 | 242.8 | 73.9 | 73.9 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4577 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.516 | 0.0146 | 145.5381 | 242.5 | 0.2 | 242.7 | 73.9 | 73.9 | 0.1 | 12.9 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0394 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4635 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.173 | 0.0143 | 142.756 | 231.2 | 0.2 | 231.4 | 73.9 | 73.9 | 0.1 | 13 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.037816 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4149 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.15 | 0.015 | 150 | 217.4 | 0.2 | 217.6 | 71.7 | 71.7409 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.035816 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4933 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.162 | 0.015 | 150 | 214 | 0.2 | 214.2 | 66.4 | 66.3699 | 0.1 | 13 | 915 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.035336 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.015 | 150 | 213.5 | 0.2 | 213.7 | 65 | 65.0056 | 0.1 | 13 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0352 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.015 | 150 | 213.5 | 0.2 | 213.7 | 64.7 | 64.7 | 0.1 | 13 | 915 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8361 | 0.0352 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6387 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.015 | 150 | 213.4 | 0.2 | 213.6 | 64.7 | 64.7 | 0.1 | 13 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0352 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4929 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.015 | 150 | 213.3 | 0.2 | 213.5 | 64.7 | 64.7 | 0.1 | 13 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0352 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4144 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.0149 | 148.6095 | 213.2 | 0.2 | 213.4 | 64.7 | 64.7 | 0.1 | 13.1 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.035136 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4636 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.17 | 0.0141 | 140.5255 | 213.1 | 0.2 | 213.3 | 64.7 | 64.6673 | 0.1 | 12.9 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4579 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.174 | 0.014 | 140 | 213.1 | 0.2 | 213.3 | 64.7 | 64.7 | 0.1 | 12.8 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4209 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.014 | 140 | 213.1 | 0.2 | 213.3 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4994 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.014 | 140 | 213.1 | 0.2 | 213.3 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.4222 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.014 | 140 | 213 | 0.2 | 213.2 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4561 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.0137 | 137.3398 | 213.1 | 0.2 | 213.3 | 64.6 | 64.6306 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4654 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.0121 | 121.1192 | 213 | 0.2 | 213.2 | 64.7 | 64.7 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4132 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.0128 | 127.8965 | 213 | 0.2 | 213.2 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4916 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.9 | 0.2 | 213.1 | 64.7 | 64.7 | 0.1 | 12.8 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.836 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.9 | 0.2 | 213.1 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.9 | 0.2 | 213.1 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.9 | 0.2 | 213.1 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.205 | 0.013 | 130 | 212.9 | 0.1 | 213 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.211 | 0.013 | 130 | 212.9 | 0.1 | 213 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.36 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.9 | 0.1 | 213 | 64.6 | 64.6 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8361 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015 | 20.4594 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.9 | 0.1 | 213 | 64.6 | 64.6 | 0.1 | 12.9 | 917 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0351 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.8 | 0.1 | 212.9 | 64.6 | 64.6 | 0.1 | 13 | 916 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.7 | 0.2 | 212.9 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.7 | 0.2 | 212.9 | 64.6 | 64.6 | 0.1 | 13 | 915 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4545 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.6 | 0.2 | 212.8 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.2 | 0.013 | 130 | 212.6 | 0.13587 | 212.7 | 64.6 | 64.6 | 0.1 | 13 | 916 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.197 | 0.013 | 130 | 212.6 | 0.1 | 212.7 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.181 | 0.013 | 130 | 212.5 | 0.1 | 212.6 | 64.6 | 64.6 | 0.1 | 13 | 915 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.7 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.5 | 0.1 | 212.6 | 64.6 | 64.6 | 0.1 | 13 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 195.4 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.4 | 0.1 | 212.5 | 64.6 | 64.6 | 0.1 | 12.8 | 917 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.4 | 0.1 | 212.5 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.3 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4965 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.18 | 0.013 | 130 | 212.4 | 0.2 | 212.6 | 64.6 | 64.6 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.7 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.418 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.181 | 0.013 | 130 | 212.4 | 0.2 | 212.6 | 64.7 | 64.7 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190.2 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4605 | 27.9 | 0.2
2008-11-05 | 10:42 | 4.189 | 0.013 | 130 | 212.4 | 0.13613 | 212.5 | 64.7 | 64.7 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4611 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.19 | 0.013 | 130 | 212.4 | 0.1 | 212.5 | 64.7 | 64.7 | 0.1 | 12.7 | 919 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4173 | 27.9 | 0.1
2008-11-05 | 10:42 | 4.184 | 0.013 | 130 | 212.4 | 0.1 | 212.5 | 64.7 | 64.7 | 0.1 | 12.7 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8639 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4954 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.0138 | 137.75 | 212.4 | 0.1 | 212.5 | 64.7 | 64.7 | 0.1 | 12.7 | 919 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.8 | 64.7633 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.8 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.8 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4338 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.7 | 918 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.7 | 0.1 | 12.8 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4234 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.181 | 0.014 | 140 | 212.3 | 0.1 | 212.4 | 64.7 | 64.6726 | 0.1 | 12.7 | 918 | 907 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.189 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 12.8 | 918 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.864 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.19 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 916 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.9 | 59.8 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.19 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.5 | 59.836 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.194 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 914 | 904 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.1 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.194 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191.8 | 5 | 195.3 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191.2 | 5 | 195.6 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4281 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.18 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 195.9 | 59.8639 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.4506 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.219 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.224 | 0.014 | 140 | 212.2 | 0.1 | 212.3 | 64.6 | 64.6 | 0.1 | 13 | 916 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.194 | 0.014 | 140 | 212.1 | 0.1 | 212.2 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.64 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 4.078 | 0.014 | 140 | 212.1 | 0.1 | 212.2 | 64.6 | 64.6 | 0.1 | 13 | 915 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4355 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.829 | 0.0131 | 131.0041 | 212.1 | 0.1 | 212.2 | 64.6 | 64.6 | 0.1 | 12.8 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.035 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.443 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.579 | 0.0123 | 122.758 | 199.6 | 0.1 | 199.7 | 64.6 | 64.6 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.03312 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.013 | 130 | 186.1 | 0.1 | 186.2 | 61.7 | 61.6759 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.8646 | 0.031112 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.013 | 130 | 184 | 0.1 | 184.1 | 56.7 | 56.6623 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0308 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.013 | 130 | 183.8 | 0.1 | 183.9 | 55.6 | 55.568 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0307 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.36 | 20.3862 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.013 | 130 | 183.8 | 0.1 | 183.9 | 55.4 | 55.4 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.8 | 0.0307 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3708 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0126 | 125.6355 | 183.8 | 0.1 | 183.9 | 55.3 | 55.3 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.030764 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.8 | 0.1 | 183.9 | 55.3 | 55.3 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | 0.030736 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.6 | 0.1 | 183.7 | 55.2 | 55.2 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0307 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.5 | 0.1 | 183.6 | 55.3 | 55.3 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.8361 | 0.0307 | 0.000936 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4508 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.4 | 0.1 | 183.5 | 55.2 | 55.2 | 0.1 | 12.7 | 916 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8639 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4276 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.3 | 0.1 | 183.4 | 55.2 | 55.2 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.8361 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.6393 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183.2 | 0.1 | 183.3 | 55.2 | 55.2 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.6387 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.2 | 55.2 | 0.1 | 12.8 | 916 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.963 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.2 | 55.2 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4586 | 27.9586 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.1 | 55.1 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.1 | 55.1 | 0.1 | 12.9 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.1 | 55.1 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55.1 | 55.1 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4552 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55 | 55 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0119 | 118.6688 | 183 | 0.1 | 183.1 | 55 | 55.0399 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4879 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0111 | 110.6496 | 183 | 0.1 | 183.1 | 55.1 | 55.0679 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.4095 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0118 | 117.8749 | 182.9 | 0.1 | 183 | 55 | 55.0312 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.6393 | 20.5379 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 183 | 0.1 | 183.1 | 55 | 55 | 0.1 | 13 | 914 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.36 | 20.5526 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0116 | 115.6544 | 182.9 | 0.1 | 183 | 55 | 55 | 0.1 | 13 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.8645 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.011 | 110 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0306 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4957 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.0111 | 111.0069 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.000936 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1017 | 20.4172 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.566 | 0.012 | 119.6127 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016.3607 | 20.4613 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 196 | 59.9 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0306 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8639 | 0.030536 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.8 | 0.1 | 182.9 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.7 | 0.1 | 182.8 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.7 | 0.1 | 182.8 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.36 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.57 | 0.012 | 120 | 182.7 | 0.1 | 182.8 | 55.1 | 55.0511 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.8 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4898 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.567 | 0.012 | 120 | 182.7 | 0.1 | 182.8 | 55 | 55.0284 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4112 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 182.6 | 0.1 | 182.7 | 55.1 | 55.0666 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4673 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 182.6 | 0.1 | 182.7 | 55 | 55.0318 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4542 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 182.5 | 0.1 | 182.6 | 55.1 | 55.0692 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4242 | 27.9 | 0.1
2008-11-05 | 10:43 | 3.56 | 0.012 | 120 | 182.4 | 0.1 | 182.5 | 55 | 55 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.012 | 120 | 182.4 | 0.1 | 182.5 | 55.1 | 55.0695 | 0.1 | 12.9 | 915 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.012 | 120 | 182.4 | 0.1 | 182.5 | 55 | 55 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.012 | 120 | 182.3 | 0.1 | 182.4 | 55 | 55 | 0.1 | 13.1 | 912 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.462 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.0115 | 114.6325 | 182.3 | 0.1 | 182.4 | 55 | 55 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0305 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.01 | 100 | 182.2 | 0.1 | 182.3 | 55 | 55 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1055 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.01 | 100 | 182.2 | 0.1 | 182.3 | 55 | 55 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016.64 | 20.4732 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.0108 | 107.9134 | 182.2 | 0.1 | 182.3 | 55 | 55 | 0.1 | 13.1 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8357 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.563 | 0.011 | 110 | 182.2 | 0.1 | 182.3 | 55 | 55 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4697 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.569 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55 | 55.0283 | 0.1 | 12.9 | 915 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.561 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55 | 55 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55 | 55 | 0.1 | 12.8 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4655 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55 | 55.0333 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.864 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1015.64 | 20.456 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55 | 55.0298 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182.1 | 0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 12.7 | 918 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054.2044 | 1016 | 20.4009 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 919 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4789 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 919 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.64 | 20.4637 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4147 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.011 | 110 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.64 | 20.4932 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.56 | 0.0116 | 116.0499 | 182 | 0.1 | 182.1 | 55.1 | 55.1 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4284 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.501 | 0.0116 | 115.6937 | 182 | 0.1 | 182.1 | 55.1 | 55.0861 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.0304 | 0.0009 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1016 | 20.45 | 27.9 | 0.1
2008-11-05 | 10:44 | 3.07 | 0.0105 | 104.6393 | 181.7 | 0.1 | 181.8 | 55 | 55 | 0.1 | 12.7 | 918 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0303 | 0.000836 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4715 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.928 | 0.009 | 90 | 164.2 | 0.1 | 164.3 | 55.1 | 55.07 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.027625 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.407 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.918 | 0.009 | 90 | 156.2 | 0.1 | 156.3 | 51.2 | 51.1793 | 0.1 | 12.9 | 916 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.026381 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4854 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.009 | 90 | 154 | 0.1 | 154.1 | 46.9 | 46.9018 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.026036 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4365 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.009 | 90 | 153.6 | 0.1 | 153.7 | 46 | 45.9683 | 0.1 | 13 | 913 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30.8278 | 190.8 | 191 | 5 | 196 | 59.9 | 0.026 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.009 | 90 | 153.2 | 0.1 | 153.3 | 45.9 | 45.8693 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 191 | 191 | 5 | 196 | 59.836 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4207 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.009 | 90 | 153.1 | 0.1 | 153.2 | 45.8 | 45.8 | 0.1 | 13 | 913 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 191 | 191 | 5 | 196 | 59.864 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4992 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.938 | 0.009 | 90 | 153.1 | 0.1 | 153.2 | 45.8 | 45.8 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30.172 | 191 | 191 | 5 | 196 | 59.9 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.009 | 90 | 153.1 | 0.1 | 153.2 | 45.7 | 45.7 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.836 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3879 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.0098 | 97.8528 | 153 | 0.1 | 153.1 | 45.7 | 45.7 | 0.1 | 13.1 | 912 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3345 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.9 | 0.1 | 153 | 45.7 | 45.7 | 0.1 | 13 | 912 | 904 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0259 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.413 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.9 | 0.1 | 153 | 45.7 | 45.7 | 0.1 | 13.1 | 912 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.025836 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4915 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.8 | 0.1 | 152.9 | 45.7 | 45.7 | 0.1 | 12.9 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0258 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.7 | 0.1 | 152.8 | 45.7 | 45.7 | 0.1 | 12.8 | 915 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8358 | 0.0258 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.7 | 0.1 | 152.8 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.864 | 0.0258 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.6 | 0.1 | 152.7 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0258 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.0094 | 93.8107 | 152.6 | 0.1 | 152.7 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0258 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.009 | 90 | 152.6 | 0.1 | 152.7 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.8639 | 0.025736 | 0.000736 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.0093 | 92.6716 | 152.6 | 0.1 | 152.7 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4595 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.6 | 0.1 | 152.7 | 45.7 | 45.7 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.836 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.419 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.5 | 0.1 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 196 | 59.9 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4975 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.5 | 0.1 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.937 | 0.01 | 100 | 152.5 | 0.1 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.5 | 0.1 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4674 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.5 | 0.2 | 152.7 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190.8 | 5 | 196 | 59.9 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.5 | 0.2 | 152.7 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.5 | 0.2 | 152.7 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0257 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4681 | 27.9 | 0.2
2008-11-05 | 10:44 | 2.934 | 0.01 | 100 | 152.5 | 0.136 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.025636 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4534 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.5 | 0.1 | 152.6 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.375 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.3069 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 12.9 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4639 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.000636 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 914 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4612 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4172 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.94 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4953 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.939 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.931 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4474 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.864 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4311 | 27.9 | 0.1
2008-11-05 | 10:44 | 2.93 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13 | 912 | 904 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9094 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.4 | 0.1 | 152.5 | 45.7 | 45.7 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.5 | 27.9872 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.3 | 0.1 | 152.4 | 45.7 | 45.7 | 0.1 | 12.9 | 914 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0256 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.434 | 27.934 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.2 | 0.1 | 152.3 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.025536 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.2 | 0.1 | 152.3 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.36 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.2 | 0.1 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 195.4 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.2 | 0.1 | 152.3 | 45.7 | 45.7 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 195 | 59.864 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1015.64 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:45 | 2.93 | 0.01 | 100 | 152.2 | 0.1 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 195.3 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4585 | 27.9 | 0.1
2008-11-05 | 10:45 | 2.933 | 0.01 | 100 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.8 | 917 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 195.7 | 59.8 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1055 | 1016 | 20.4259 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3155 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.8 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.864 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.394 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152 | 0.2 | 152.2 | 45.6 | 45.6344 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8643 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152 | 0.2 | 152.2 | 45.6 | 45.6 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.9 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190.8 | 5 | 195.5 | 59.836 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4293 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.0104 | 104.2429 | 152.1 | 0.2 | 152.3 | 45.7 | 45.7 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.2 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.011 | 110 | 152.1 | 0.2 | 152.3 | 45.6 | 45.6 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 195.6 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.0109 | 109.2083 | 152 | 0.2 | 152.2 | 45.7 | 45.6526 | 0.1 | 12.7 | 918 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.9 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.0101 | 100.875 | 152 | 0.2 | 152.2 | 45.6 | 45.6 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4786 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.94 | 0.01 | 100 | 152 | 0.2 | 152.2 | 45.6 | 45.6 | 0.1 | 12.9 | 916 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4001 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.824 | 0.01 | 100 | 152 | 0.2 | 152.2 | 45.6 | 45.6 | 0.1 | 12.9 | 915 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.9 | 59.9 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9784 | 0.2
2008-11-05 | 10:45 | 2.486 | 0.0092 | 91.5266 | 151.9 | 0.2 | 152.1 | 45.6 | 45.6 | 0.1 | 13 | 914 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 195.6 | 59.836 | 0.0255 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9436 | 0.2
2008-11-05 | 10:45 | 2.336 | 0.0083 | 82.6516 | 139.6 | 0.2 | 139.8 | 45.6 | 45.5692 | 0.1 | 13 | 914 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.2 | 59.9 | 0.02358 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4348 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.335 | 0.009 | 90 | 128.5 | 0.2 | 128.7 | 42.4 | 42.4146 | 0.1 | 13 | 913 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.9 | 0.021772 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4863 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.34 | 0.009 | 90 | 125.6 | 0.2 | 125.8 | 37.6 | 37.5503 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.5 | 59.8355 | 0.021335 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4078 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.34 | 0.009 | 90 | 124.8 | 0.2 | 125 | 36.7 | 36.6645 | 0.1 | 13 | 912 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.9 | 59.836 | 0.021236 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4706 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.335 | 0.009 | 90 | 124.1 | 0.2 | 124.3 | 36.5 | 36.5 | 0.1 | 13 | 912 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0211 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4511 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 124.1 | 0.2 | 124.3 | 36.4 | 36.4355 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0211 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4273 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 124 | 0.2 | 124.2 | 36.4 | 36.4 | 0.1 | 12.9 | 914 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.021036 | 0.000536 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.3613 | 20.4941 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 124 | 0.2 | 124.2 | 36.4 | 36.4 | 0.1 | 12.8 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.864 | 0.021 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4156 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.4 | 0.2 | 123.6 | 36.4 | 36.4 | 0.1 | 12.8 | 915 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0209 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4628 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.2 | 0.2 | 123.4 | 36.4 | 36.4 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.8361 | 0.0209 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.2 | 0.2 | 123.4 | 36.3 | 36.3293 | 0.1 | 12.8 | 915 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0209 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9197 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.2 | 0.2 | 123.4 | 36.4 | 36.372 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8639 | 0.0209 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9981 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.1 | 0.2 | 123.3 | 36.4 | 36.3516 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0208 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4235 | 27.9235 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123.1 | 0.2 | 123.3 | 36.4 | 36.3688 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0208 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.455 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.32 | 0.009 | 90 | 123 | 0.2 | 123.2 | 36.4 | 36.4 | 0.1 | 12.8 | 916 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0208 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.324 | 0.009 | 90 | 123 | 0.2 | 123.2 | 36.4 | 36.4 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.864 | 0.0208 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4881 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.325 | 0.009 | 90 | 122.9 | 0.2 | 123.1 | 36.4 | 36.4 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0208 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4097 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.31 | 0.009 | 90 | 122.9 | 0.2 | 123.1 | 36.4 | 36.4 | 0.1 | 12.8 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0208 | 0.000436 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4688 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.31 | 0.009 | 90 | 122.9 | 0.2 | 123.1 | 36.4 | 36.4 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4527 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.347 | 0.009 | 90 | 122.9 | 0.2 | 123.1 | 36.5 | 36.5 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.36 | 0.009 | 90 | 122.9 | 0.2 | 123.1 | 36.5 | 36.5 | 0.1 | 12.9 | 915 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4042 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.356 | 0.0083 | 82.5042 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30.828 | 190 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4826 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.35 | 0.008 | 80 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 190 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.35 | 0.008 | 80 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 189.2 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.35 | 0.008 | 80 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30.172 | 189 | 191 | 5 | 196 | 59.8639 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4821 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.35 | 0.008 | 80 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4037 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.344 | 0.008 | 80 | 122.8 | 0.2 | 123 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.344 | 0.0087 | 87.4769 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 13 | 913 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.35 | 0.009 | 90 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.8 | 915 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4317 | 27.9317 | 0.2
2008-11-05 | 10:45 | 2.349 | 0.0086 | 85.8212 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9898 | 0.2
2008-11-05 | 10:45 | 2.332 | 0.0083 | 82.609 | 122.7 | 0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 12.7 | 917 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8643 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.6429 | 20.5 | 27.9114 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.009 | 90 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5344 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.433 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.009 | 90 | 122.7 | 0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190.2 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4454 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.0082 | 82.4407 | 122.7 | 0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4761 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.008 | 80 | 122.7 | 0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4024 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.008 | 80 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4808 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.008 | 80 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 906 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4408 | 27.9 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.008 | 80 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1054 | 1016 | 20.4376 | 27.9376 | 0.2
2008-11-05 | 10:45 | 2.33 | 0.008 | 80 | 122.7 | 0.2 | 122.9 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8361 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4839 | 27.9839 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.0088 | 87.6206 | 122.6 | 0.2 | 122.8 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4055 | 27.9055 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.0084 | 83.9747 | 122.6 | 0.2 | 122.8 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.473 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.6 | 0.2 | 122.8 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0207 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4701 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 12.8 | 916 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4083 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 12.9 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016.64 | 20.4867 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.334 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1017 | 20.4349 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.34 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4436 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.339 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 13.1 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.864 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.478 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.331 | 0.008 | 80 | 122.5 | 0.2 | 122.7 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.4 | 0.2 | 122.6 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.4 | 0.2 | 122.6 | 36.5 | 36.5 | 0.1 | 13 | 913 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4573 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.4 | 0.2 | 122.6 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.836 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4641 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.6 | 36.5682 | 0.1 | 13 | 913 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4144 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 13 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4928 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 12.8 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30.828 | 189 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 189 | 191 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4505 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 31 | 189 | 190.2 | 5 | 196 | 59.8361 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.33 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30.172 | 189 | 190 | 5 | 196 | 59.8361 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.3613 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.317 | 0.008 | 80 | 122.3 | 0.2 | 122.5 | 36.5 | 36.5 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.9 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 2.23 | 0.0072 | 71.7444 | 122.2 | 0.2 | 122.4 | 36.5 | 36.5 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8639 | 0.0206 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.6393 | 20.4633 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.765 | 0.0065 | 65.153 | 114.1 | 0.2 | 114.3 | 36.5 | 36.5 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.019256 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4582 | 27.9418 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0078 | 78.027 | 100.4 | 0.2 | 100.6 | 35.8 | 35.8201 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0.017096 | 0.000336 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4203 | 27.9797 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0053 | 52.7319 | 95.4 | 0.2 | 95.6 | 29.6 | 29.5761 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 195.7 | 59.9 | 0.016244 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.36 | 20.4987 | 27.9013 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.005 | 50 | 94.3 | 0.2 | 94.5 | 27.7 | 27.7073 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.3 | 59.9 | 0.016036 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0056 | 55.7571 | 93.8 | 0.2 | 94 | 27.4 | 27.4 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.1 | 59.864 | 0.016 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.006 | 60 | 93.8 | 0.2 | 94 | 27.3 | 27.3 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 195.4 | 59.9 | 0.016 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.36 | 20.4662 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0057 | 57.3878 | 93.7 | 0.2 | 93.9 | 27.3 | 27.3 | 0.1 | 12.7 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.8 | 59.8 | 0.0159 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4125 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0051 | 50.6771 | 93.7 | 0.2 | 93.9 | 27.3 | 27.3 | 0.1 | 12.8 | 913 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0159 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4909 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.0059 | 59.3576 | 93.7 | 0.2 | 93.9 | 27.3 | 27.3 | 0.1 | 12.7 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0159 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.738 | 0.006 | 60 | 93.6 | 0.2 | 93.8 | 27.3 | 27.3 | 0.1 | 12.8 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0159 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4522 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.747 | 0.006 | 60 | 93.5 | 0.2 | 93.7 | 27.3 | 27.3153 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0159 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.738 | 0.006 | 60 | 93.4 | 0.2 | 93.6 | 27.4 | 27.4 | 0.1 | 12.9 | 912 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.015836 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.722 | 0.006 | 60 | 93.4 | 0.2 | 93.6 | 27.4 | 27.4 | 0.1 | 13 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.006 | 60 | 93.4 | 0.2 | 93.6 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9616 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.006 | 60 | 93.3 | 0.2 | 93.5 | 27.4 | 27.4 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 28 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.006 | 60 | 93.3 | 0.2 | 93.5 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.8361 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9819 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.0066 | 65.7942 | 93.3 | 0.2 | 93.5 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9031 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 93.2 | 0.2 | 93.4 | 27.4 | 27.431 | 0.1 | 13.1 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 93.2 | 0.2 | 93.4 | 27.5 | 27.5 | 0.1 | 13 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 195.9 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 93.2 | 0.2 | 93.4 | 27.5 | 27.5 | 0.1 | 13 | 912 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.5 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4323 | 27.9323 | 0.2
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 93.2 | 0.2 | 93.4 | 27.5 | 27.5 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.1 | 59.864 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9893 | 0.2
2008-11-05 | 10:46 | 1.724 | 0.007 | 70 | 93.1 | 0.2 | 93.3 | 27.5 | 27.5 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 195.2 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9108 | 0.2
2008-11-05 | 10:46 | 1.73 | 0.007 | 70 | 93 | 0.2 | 93.2 | 27.5 | 27.5 | 0.1 | 12.8 | 915 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.6 | 59.864 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.729 | 0.007 | 70 | 93 | 0.2 | 93.2 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.2 | 5 | 195.9 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:46 | 1.721 | 0.007 | 70 | 93 | 0.2 | 93.2 | 27.5 | 27.5 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9244 | 0.2
2008-11-05 | 10:46 | 1.712 | 0.007 | 70 | 92.9 | 0.2 | 93.1 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4975 | 27.9975 | 0.2
2008-11-05 | 10:46 | 1.71 | 0.007 | 70 | 92.9 | 0.136 | 93 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4187 | 27.9187 | 0.1
2008-11-05 | 10:46 | 1.71 | 0.007 | 70 | 92.9 | 0.1 | 93 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1015 | 20.4598 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.71 | 0.007 | 70 | 92.9 | 0.1 | 93 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.8 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.711 | 0.0054 | 54.3131 | 92.9 | 0.1 | 93 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.9 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4833 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.719 | 0.0041 | 41.3445 | 92.8 | 0.1 | 92.9 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.5 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4048 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.008 | 80 | 92.9 | 0.1 | 93 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.1 | 59.9 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4736 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.008 | 80 | 92.8 | 0.1 | 92.9 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 195.2 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.008 | 80 | 92.8 | 0.1 | 92.9 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 195.6 | 59.836 | 0.0158 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.008 | 80 | 92.7 | 0.1 | 92.8 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.9 | 59.864 | 0.0157 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4911 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.008 | 80 | 92.7 | 0.1 | 92.8 | 27.5 | 27.5 | 0.1 | 12.8 | 916 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.36 | 20.4129 | 27.9 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.0077 | 77.4599 | 92.7 | 0.1 | 92.8 | 27.5 | 27.4693 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1054 | 1016 | 20.4 | 27.9657 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 92.7 | 0.1 | 92.8 | 27.5 | 27.5 | 0.1 | 12.8 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1016 | 20.4443 | 27.9557 | 0.1
2008-11-05 | 10:46 | 1.72 | 0.007 | 70 | 92.7 | 0.1 | 92.8 | 27.5 | 27.4667 | 0.1 | 12.9 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1054 | 1016 | 20.4773 | 27.9227 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.007 | 70 | 92.7 | 0.1 | 92.8 | 27.4 | 27.4316 | 0.1 | 13 | 912 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1016 | 20.4 | 27.9988 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.007 | 70 | 92.7 | 0.1 | 92.8 | 27.5 | 27.4565 | 0.1 | 13 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.921 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.0066 | 66.0982 | 92.7 | 0.1 | 92.8 | 27.4 | 27.4312 | 0.1 | 13.1 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.006 | 60 | 92.7 | 0.1 | 92.8 | 27.5 | 27.5 | 0.1 | 13.1 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.4 | 27.4301 | 0.1 | 13.1 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1017 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.72 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 13.1 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.3607 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.724 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 13 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4717 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.728 | 0.006 | 60 | 92.7 | 0.1 | 92.8 | 27.5 | 27.5 | 0.1 | 13.1 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4498 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.721 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 13 | 912 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9287 | 0.1
2008-11-05 | 10:47 | 1.729 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 12.8 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4071 | 27.9929 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 12.7 | 916 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.9 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4856 | 27.9144 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.64 | 20.3728 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5483 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.864 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5292 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3419 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.6 | 27.6 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4987 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.5 | 0.1 | 92.6 | 27.6 | 27.6 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.64 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.5 | 0.1 | 92.6 | 27.6 | 27.569 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.015636 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4438 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.5 | 0.1 | 92.6 | 27.5 | 27.5 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | 0.015664 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.73 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.61 | 0.006 | 60 | 92.6 | 0.1 | 92.7 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.8644 | 0.0157 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.225 | 0.0045 | 44.5462 | 91.2 | 0.1 | 91.3 | 27.5 | 27.5 | 0.1 | 12.7 | 917 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | 0.015444 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.47 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.148 | 0.0022 | 22.1898 | 73.2 | 0.1 | 73.3 | 27.1 | 27.1152 | 0.1 | 12.7 | 918 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.8 | 5 | 196 | 59.864 | 0.012528 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4518 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.132 | 0.0031 | 30.5299 | 66.1 | 0.1 | 66.2 | 22.7 | 22.6739 | 0.1 | 12.7 | 916 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.011408 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4265 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.0039 | 39.0769 | 63.8 | 0.1 | 63.9 | 19 | 19.0372 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.010936 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 63.6 | 0.1 | 63.7 | 18.4 | 18.4 | 0.1 | 12.7 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8644 | 0.0109 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 63.5 | 0.1 | 63.6 | 18.3 | 18.3 | 0.1 | 12.7 | 914 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | 0.0109 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.6387 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 63.5 | 0.1 | 63.6 | 18.3 | 18.3 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0109 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4187 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 63.4 | 0.1 | 63.5 | 18.3 | 18.3 | 0.1 | 12.7 | 913 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.864 | 0.0109 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33.315 | 60 | 1054 | 1016 | 20.338 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.138 | 0.004 | 40 | 63.1 | 0.1 | 63.2 | 18.3 | 18.3 | 0.1 | 12.9 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0108 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4951 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 63 | 0.1 | 63.1 | 18.3 | 18.3 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0108 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.424 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 63 | 0.1 | 63.1 | 18.3 | 18.3 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8359 | 0.0107 | 0.000136 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.138 | 0.004 | 40 | 62.9 | 0.1 | 63 | 18.3 | 18.3 | 0.1 | 13.1 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0107 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4327 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 62.8 | 0.1 | 62.9 | 18.3 | 18.2727 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0107 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4891 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 62.8 | 0.1 | 62.9 | 18.3 | 18.2696 | 0.1 | 13.1 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | 0.0107 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4103 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.13 | 0.004 | 40 | 62.7 | 0.1 | 62.8 | 18.2 | 18.2 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.9 | 59.8 | 0.0107 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4682 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.136 | 0.004 | 40 | 62.7 | 0.1 | 62.8 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.5 | 59.8 | 0.0107 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9467 | 0.1
2008-11-05 | 10:47 | 1.136 | 0.004 | 40 | 62.6 | 0.1 | 62.7 | 18.2 | 18.2 | 0.1 | 13 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.864 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4749 | 27.9749 | 0.1
2008-11-05 | 10:47 | 1.132 | 0.004 | 40 | 62.6 | 0.1 | 62.7 | 18.1 | 18.1318 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 195 | 59.836 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.36 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.9 | 912 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.8639 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195 | 59.864 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1015 | 20.4 | 27.9604 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40.497 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190.2 | 5 | 195.3 | 59.9 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4388 | 27.9612 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.0049 | 49.066 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.7 | 59.8361 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.5 | 27.9172 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.005 | 50 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.836 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1014.36 | 20.5 | 27.9953 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.005 | 50 | 62.5 | 0.1 | 62.6 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8361 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1014.6393 | 20.4259 | 28 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.005 | 50 | 62.4 | 0.1 | 62.5 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.864 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4527 | 28 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.005 | 50 | 62.4 | 0.1 | 62.5 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4689 | 27.9689 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.0049 | 49.1798 | 62.4 | 0.1 | 62.5 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.8645 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.0041 | 40.8948 | 62.4 | 0.1 | 62.5 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.3 | 0.1 | 62.4 | 18.1 | 18.1283 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 196 | 59.9 | 0.0106 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3336 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.3 | 0.1 | 62.4 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.8 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3449 | 27.9 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.3 | 0.1 | 62.4 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0105 | 0.0001 | -3.6e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9233 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.3 | 0.1 | 62.4 | 18.1 | 18.1 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0106 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4017 | 27.9983 | 0.1
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.3 | 0.16429 | 62.5 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8643 | 0.010536 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4799 | 27.9201 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.9 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33.315 | 60 | 1054 | 1016 | 20.5 | 27.9585 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 18.1 | 18.1 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 195.6 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 28 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 18.1 | 18.1 | 0.1 | 12.9 | 915 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.2 | 59.9 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9845 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 18 | 18.0482 | 0.1 | 12.9 | 913 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.1 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.906 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 18 | 18 | 0.1 | 13.1 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 195.5 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4281 | 27.9 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 18 | 18 | 0.1 | 13 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 195.9 | 59.9 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:47 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 18 | 18 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.9 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.4293 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 18 | 17.9701 | 0.1 | 13 | 910 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.4 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 18 | 17.9507 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.1 | 59.8 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 34 | 60 | 1054 | 1015.36 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 17.9 | 17.9 | 0.1 | 13.1 | 910 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.3 | 59.864 | 0.0105 | 0.0001 | 0 | -0.0001 |  | n/a |  | 33.685 | 60 | 1054 | 1015 | 20.5 | 27.9645 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.0045 | 44.6677 | 62.1 | 0.2 | 62.3 | 17.9 | 17.9 | 0.1 | 13.1 | 910 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 195.7 | 59.8642 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.5 | 28 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.006 | 59.5353 | 62.1 | 0.2 | 62.3 | 17.9 | 17.9 | 0.1 | 13.1 | 910 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33.315 | 60 | 1054 | 1015 | 20.5 | 28 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.0051 | 50.9294 | 62.1 | 0.2 | 62.3 | 17.9 | 17.9 | 0.1 | 12.9 | 912 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 34 | 60 | 1054 | 1015 | 20.5 | 28 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.005 | 50 | 62.1 | 0.2 | 62.3 | 17.9 | 17.8686 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33.685 | 60 | 1054 | 1015 | 20.4216 | 27.9216 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.005 | 50 | 62.1 | 0.2 | 62.3 | 17.8 | 17.8 | 0.1 | 12.7 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190.2 | 5 | 195.9 | 59.864 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4565 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.005 | 50 | 62.1 | 0.2 | 62.3 | 17.8 | 17.7689 | 0.1 | 12.8 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.5 | 59.836 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4651 | 27.9349 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.005 | 50 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 916 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 195.1 | 59.836 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4137 | 27.9863 | 0.2
2008-11-05 | 10:48 | 1.147 | 0.005 | 49.5517 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 190 | 5 | 195.2 | 59.864 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4922 | 27.9078 | 0.2
2008-11-05 | 10:48 | 1.15 | 0.0041 | 40.931 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 195.6 | 59.9 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.15 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8639 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.2
2008-11-05 | 10:48 | 1.148 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 195.7 | 59.9 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4724 | 27.9276 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 898 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 195.3 | 59.9 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.406 | 27.994 | 0.2
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 195.1 | 59.8639 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4845 | 27.9155 | 0.2
2008-11-05 | 10:48 | 1.147 | 0.004 | 40 | 62 | 0.136 | 62.1 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190.8 | 5 | 195.4 | 59.836 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4371 | 27.9 | 0.1
2008-11-05 | 10:48 | 1.15 | 0.004 | 40 | 62 | 0.1 | 62.1 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.8 | 59.864 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4413 | 27.9 | 0.1
2008-11-05 | 10:48 | 1.146 | 0.004 | 40 | 62 | 0.1 | 62.1 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.9 | 59.9 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 1.14 | 0.004 | 40 | 62 | 0.1 | 62.1 | 17.7 | 17.7 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 195.5 | 59.8 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015.6452 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 1.1 | 0.0033 | 32.5188 | 62 | 0.1 | 62.1 | 17.6 | 17.6 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 195.1 | 59.8 | 0.0105 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.676 | 0.0024 | 24.3523 | 61.8 | 0.1 | 61.9 | 17.6 | 17.6 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.2 | 59.9 | 0.010436 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.609 | 0.0038 | 37.6539 | 42.3 | 0.1 | 42.4 | 17.4 | 17.3574 | 0.1 | 12.7 | 917 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 195.6 | 59.9 | 0.007204 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4665 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.595 | 0.0013 | 12.6955 | 34.8 | 0.1 | 34.9 | 12.4 | 12.3953 | 0.1 | 12.9 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.005944 | 0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.59 | 0.001 | 10 | 33 | 0.1 | 33.1 | 9.1 | 9.05 | 0.1 | 12.9 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.005636 | 3.6e-05 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.588 | 0.001 | 10 | 32.5 | 0.1 | 32.6 | 8.5 | 8.5 | 0.1 | 13 | 910 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.8644 | 0.0055 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.001 | 10 | 32.5 | 0.1 | 32.6 | 8.4 | 8.4 | 0.1 | 13 | 909 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.9 | 0.0055 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4473 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.001 | 10 | 32.5 | 0.1 | 32.6 | 8.4 | 8.4 | 0.1 | 13 | 910 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0055 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4743 | 27.9257 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.0011 | 10.7304 | 32.4 | 0.1 | 32.5 | 8.4 | 8.4 | 0.1 | 13 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.8361 | 0.005436 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9958 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.0028 | 28.1217 | 32.3 | 0.1 | 32.4 | 8.3 | 8.3 | 0.1 | 13.1 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8361 | 0.0054 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9174 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.003 | 30 | 32.2 | 0.1 | 32.3 | 8.3 | 8.3316 | 0.1 | 13.1 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0054 | 0 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.461 | 27.961 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.003 | 30 | 32.1 | 0.1 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.005336 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4605 | 27.9605 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.003 | 30 | 32.1 | 0.1 | 32.2 | 8.3 | 8.3 | 0.1 | 12.9 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0053 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.0028 | 27.9276 | 32.1 | 0.1 | 32.2 | 8.3 | 8.3 | 0.1 | 12.8 | 913 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0053 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 32 | 0.1 | 32.1 | 8.3 | 8.3 | 0.1 | 12.8 | 913 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0053 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 32 | 0.1 | 32.1 | 8.3 | 8.3 | 0.1 | 12.7 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0053 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.3467 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.3 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.2 | 5 | 196 | 59.864 | 0.0053 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.3635 | 27.9318 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.3 | 0.1 | 12.7 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.005236 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4898 | 27.9898 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.3 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.64 | 20.4114 | 27.9114 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.3 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 190 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.3 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4455 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.2686 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | -6.4e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.4761 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.2 | 8.2 | 0.1 | 12.7 | 915 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4024 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.9 | 0.1 | 32 | 8.3 | 8.2697 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4808 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.8 | 0.1 | 31.9 | 8.2 | 8.2 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4407 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.8 | 0.1 | 31.9 | 8.2 | 8.2 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.8639 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.8 | 0.1 | 31.9 | 8.2 | 8.2 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 27.9161 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.2 | 0.1 | 12.9 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8641 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1015.6407 | 20.4 | 27.9945 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.2 | 0.1 | 12.9 | 912 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.927 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.2 | 0.1 | 13 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4515 | 27.9515 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.2 | 0.1 | 13 | 911 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4701 | 27.9701 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.2 | 0.1 | 13 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.2 | 8.1686 | 0.1 | 13 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.1 | 8.1 | 0.1 | 13 | 911 | 905 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9652 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.7 | 0.1 | 31.8 | 8.1 | 8.1 | 0.1 | 13 | 911 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4437 | 27.9563 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.6 | 0.1 | 31.7 | 8.1 | 8.1 | 0.1 | 13 | 911 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | -3.6e-05 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.6 | 0.1 | 31.7 | 8.1 | 8.1 | 0.1 | 13 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 191 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.6 | 0.1 | 31.7 | 8.1 | 8.1 | 0.1 | 12.8 | 913 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8.1 | 8.1 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8359 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32.685 | 60 | 1054 | 1015 | 20.4427 | 27.9573 | 0.1
2008-11-05 | 10:48 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8.1 | 8.0596 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.2 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32 | 60 | 1054 | 1015 | 20.4717 | 27.9642 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8 | 8 | 0.1 | 12.8 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32 | 60 | 1054 | 1015.64 | 20.5857 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8.1 | 8.0801 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32 | 60 | 1054 | 1016 | 20.5073 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8 | 8.0283 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32 | 60 | 1054 | 1016 | 20.429 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8.1 | 8.0686 | 0.1 | 12.7 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8361 | 0.0052 | -0.0001 | 0 | -0.0001 |  | n/a |  | 32 | 60 | 1054 | 1016 | 20.4 | 27.9495 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.5 | 0.1 | 31.6 | 8.1 | 8.1 | 0.1 | 12.8 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 32.315 | 60 | 1054 | 1016 | 20.4 | 27.9719 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.1 | 31.5 | 8.1 | 8.0697 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4066 | 27.9066 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.1 | 31.5 | 8 | 8 | 0.1 | 12.8 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015.3607 | 20.485 | 27.985 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.1 | 31.5 | 8 | 8 | 0.1 | 12.8 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3732 | 27.9366 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.1 | 31.5 | 8 | 8 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015.36 | 20.3418 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.0018 | 17.9167 | 31.4 | 0.1 | 31.5 | 8 | 8 | 0.1 | 12.8 | 914 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4203 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.58 | 0.001 | 10 | 31.4 | 0.2 | 31.6 | 8 | 8 | 0.1 | 12.7 | 915 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.864 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4987 | 27.9 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.001 | 10 | 31.4 | 0.2 | 31.6 | 8 | 7.9707 | 0.1 | 12.8 | 914 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4228 | 27.9 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.001 | 10 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 12.7 | 915 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4554 | 27.9554 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.001 | 10 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4661 | 28 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.0014 | 13.6717 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 12.9 | 914 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1015 | 20.4 | 28 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 912 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 28 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 912 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.836 | 0.0052 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4 | 27.9306 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 912 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0052 | -3.6e-05 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4478 | 27.9 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 911 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0.0052 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4737 | 27.9263 | 0.2
2008-11-05 | 10:49 | 0.58 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 911 | 906 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8639 | 0.0052 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4047 | 27.9953 | 0.2
2008-11-05 | 10:49 | 0.567 | 0.002 | 20 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 911 | 905 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0.0052 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053.9142 | 1016 | 20.4831 | 27.9169 | 0.2
2008-11-05 | 10:49 | 0.176 | 0.0014 | 13.7 | 31.4 | 0.2 | 31.6 | 7.9 | 7.9 | 0.1 | 13 | 911 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 191 | 5 | 196 | 59.9 | 0.0053 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9616 | 0.2
2008-11-05 | 10:49 | 0.038 | -0.001 | -10.4483 | 17.3 | 0.2 | 17.5 | 8 | 7.9684 | 0.1 | 13 | 911 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0.002874 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4601 | 27.9601 | 0.2
2008-11-05 | 10:49 | 0.005 | -0.0019 | -19.069 | 4.5 | 0.2 | 4.7 | 4.2 | 4.1694 | 0.1 | 12.8 | 912 | 907 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 4.2 | 196 | 59.836 | 0.00078 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4182 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 2 | 0.2 | 2.2 | -0.1 | 0 | 0.1 | 12.8 | 912 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 4 | 196 | 59.836 | 0.0003 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4969 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.8 | 0.136 | 2 | -1.2 | 0 | 0.1 | 12.7 | 912 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 4.8 | 196 | 59.836 | 0.0003 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4246 | 27.9 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.8 | 0.1 | 1.9 | -1.3 | 0 | 0.1 | 12.7 | 912 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.864 | 0.0003 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4 | 27.9538 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.7 | 0.1 | 1.8 | -1.3 | 0 | 0.1 | 12.7 | 912 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.9 | 0.000236 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4321 | 27.9679 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.6 | 0.1 | 1.7 | -1.4 | 0 | 0.1 | 12.7 | 911 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.8 | 0.0002 | 0 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4898 | 27.9 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.5 | 0.1 | 1.6 | -1.4 | 0 | 0.1 | 12.7 | 911 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 190 | 5 | 196 | 59.8 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.411 | 27.9 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.4 | 0.1 | 1.5 | -1.4 | 0 | 0.1 | 12.8 | 911 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.864 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4675 | 27.9 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.3 | 0.1 | 1.4 | -1.4 | 0 | 0.1 | 12.8 | 911 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190 | 5 | 196 | 59.836 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.946 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.3 | 0.1 | 1.4 | -1.5 | 0 | 0.1 | 12.7 | 911 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 190.8 | 5 | 196 | 59.836 | 0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4757 | 27.9757 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.2 | 0.1 | 1.3 | -1.5 | 0 | 0.1 | 12.8 | 911 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.000136 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1.2 | 0.2 | 1.4 | -1.5 | 0 | 0.1 | 12.7 | 911 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | 0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 1 | 0.2 | 1.2 | -1.5 | 0 | 0.1 | 12.8 | 911 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.5 | 0.2 | 0.7 | -1.6 | 0 | 0.1 | 12.7 | 911 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8639 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4382 | 27.9382 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.5 | 0.2 | 0.7 | -1.6 | 0 | 0.1 | 12.8 | 911 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4669 | 28 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.5 | 0.2 | 0.7 | -1.6 | 0 | 0.1 | 12.7 | 911 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.3109 | 28 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.5 | 0.2 | 0.7 | -1.7 | 0 | 0.1 | 12.9 | 909 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8355 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4467 | 27.9267 | 0.2
2008-11-05 | 10:49 | 0.007 | -0.002 | -20 | 0.5 | 0.2 | 0.7 | -1.7 | 0 | 0.1 | 13 | 908 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015.36 | 20.4481 | 27.9 | 0.2
2008-11-05 | 10:49 | 0.01 | -0.002 | -20 | 0.4 | 0.136 | 0.6 | -1.7 | 0 | 0.1 | 13 | 908 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.836 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015.64 | 20.4 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.01 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.7 | 0 | 0.1 | 13 | 908 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4088 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.01 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.7 | 0 | 0.1 | 13 | 909 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4871 | 27.9 | 0.1
2008-11-05 | 10:49 | 0.01 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.7 | 0 | 0.1 | 13 | 909 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 191 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9656 | 0.1
2008-11-05 | 10:49 | 0.001 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.7 | 0 | 0.1 | 13 | 909 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9558 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.7 | 0 | 0.1 | 13 | 909 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.5 | 27.9226 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.6 | 0 | 0.1 | 13 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.864 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015.36 | 20.5 | 27.999 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.4 | 0.1 | 0.5 | -1.6 | 0 | 0.1 | 12.8 | 911 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.836 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.921 | 0.1
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.164 | 0.5 | -1.6 | 0 | 0.1 | 12.7 | 913 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4424 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.5 | 0 | 0.1 | 12.7 | 914 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 191 | 5 | 196 | 59.9 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.5 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190.2 | 5 | 196 | 59.8 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.5 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8361 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.4 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.864 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9716 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.4 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.8361 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 28 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.4 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.8 | 190 | 5 | 196 | 59.8361 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4 | 27.9715 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.3 | 0.2 | 0.5 | -1.4 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190 | 5 | 196 | 59.864 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.407 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.3 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 190.8 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015.64 | 20.4855 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.4 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.8 | 191 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4364 | 27.9 | 0.2
2008-11-05 | 10:49 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.3 | 0 | 0.1 | 12.8 | 913 | 908 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1016 | 20.4 | 27.942 | 0.2
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.3 | 0 | 0.1 | 12.7 | 914 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.4207 | 27.9793 | 0.2
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.3 | 0 | 0.1 | 12.8 | 913 | 907 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.36 | 20.4991 | 27.9009 | 0.2
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.3 | 0 | 0.1 | 12.7 | 913 | 907 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9776 | 0.2
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.2 | 0.1 | 0.3 | -1.4 | 0 | 0.1 | 12.9 | 911 | 906 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4444 | 27.9444 | 0.1
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.1 | 0.1 | 0.2 | -1.3 | 0 | 0.1 | 13 | 909 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.434 | 27.9 | 0.1
2008-11-05 | 10:50 | 0 | -0.002 | -20 | 0.1 | 0.1 | 0.2 | -1.4 | 0 | 0.1 | 13.1 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:50 | 0 | -0.0013 | -12.8608 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13.1 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.864 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13.1 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 192 | 191 | 5 | 196 | 59.836 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9698 | 0.1
2008-11-05 | 10:50 | 0 | -0.0014 | -13.661 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13.1 | 908 | 904 | 893 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191.2 | 191 | 5 | 196 | 59.9 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.5 | 27.9519 | 0.1
2008-11-05 | 10:50 | -0.009 | -0.0018 | -17.8369 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4471 | 27.9264 | 0.1
2008-11-05 | 10:50 | -0.01 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13.1 | 908 | 905 | 894 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 191 | 191 | 5 | 196 | 59.8639 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.3102 | 27.9949 | 0.1
2008-11-05 | 10:50 | -0.01 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 13 | 909 | 905 | 895 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190.2 | 191 | 5 | 196 | 59.9 | -0.0002 | -0.000164 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4671 | 27.9165 | 0.1
2008-11-05 | 10:50 | -0.006 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.8 | 911 | 906 | 896 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016.64 | 20.5 | 27.962 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 191 | 5 | 196 | 59.864 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4599 | 27.9599 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.4 | 0 | 0.1 | 12.7 | 914 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190.2 | 5 | 196 | 59.8 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4183 | 27.9183 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.2 | 190 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4973 | 27.9973 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.864 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1054 | 1016 | 20.4242 | 27.9242 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.836 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.4542 | 27.9 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.9326 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.9891 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | 0 | 0.1 | 0.1 | -1.3 | 0 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189 | 190 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.9105 | 0.1
2008-11-05 | 10:50 | 0 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -1.2 | 0 | 0.1 | 12.7 | 915 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 189.8 | 190 | 5 | 196 | 59.864 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.9 | 0.1
2008-11-05 | 10:50 | 0 | -0.0009 | -9.4558 | -0.1 | 0.1 | 0.1 | -1.2 | 0 | 0.1 | 12.7 | 916 | 908 | 897 | 0 | 0 | 0 | 994 | -40 | 0 | 30 | 190 | 190 | 5 | 196 | 59.9 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 33 | 60 | 1053 | 1015 | 20.5 | 27.9 | 0.1
Summary Information:
Post Processor DLL Version | 6.01
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-11-05
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 50
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 28261
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6.1
C6H14 Span(ppm) | 2005
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 131
Version | 1.611
NO Span(ppm) | 301
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.30
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | Backhoe J Deere MN 700212 310G
License Plate
Engine Displacement | 4.5
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 10:29
Test End Time | 10:50
Test Duration (s) | 1262
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 994
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iAMBII_RPM
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
Faults:
0X0000 - 11/05/2008 10:28:51.564 - None Found
Warnings:
0Xe501 - 11/05/2008 08:35:29.997 - FID flame went out
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 122
Brake Specific Dropout Limit Count | 1262
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-11-05 | 10:29 | [RELEASE_VER=2.018 BUILD=50 BDATE=07/18/2008 IP=10.10.1.55]
No Information Available

## THC Lin
Test Date:
Test Time:
11/5/2008
10:28:52
FID Serial No:
177
SEMTECH Serial Number:
G05-SDS02

## CO2 Lin
Test Date:
Test Time:
11/5/2008
10:28:52
AMBII Serial No:
41304
SEMTECH Serial Number:
G05-SDS02

## CO Lin
Test Date:
Test Time:
11/5/2008
10:28:52
AMBII Serial No:
41304
SEMTECH Serial Number:
G05-SDS02

## NO Lin
Test Date:
Test Time:
11/5/2008
10:28:52
NDUV Serial No:
217
SEMTECH Serial Number:
G05-SDS02

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
11/5/2008
10:28:52
FID Serial No:
177
SEMTECH Serial Number:
G05-SDS02

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
11/5/2008
10:28:52
NDUV Serial No:
217
AMBII Serial No:
41304
SEMTECH Serial Number:
G05-SDS02

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6.07 | Rm = | 0.99269 | 0.99269 | 0.99269 | 0.99269 | 0.99269 | 0.99269 | 0.99269 | 0.99269 | 0.99269
| C.F.= | 0.099341 | 0.19883 | 0.29846 | 0.39824 | 0.49816 | 0.59824 | 0.69846 | 0.79882 | 0.89934
| gas = | 0.603 | 1.2069 | 1.8117 | 2.4173 | 3.0239 | 3.6313 | 4.2396 | 4.8489 | 5.459
Ro for CO2= | 0.8795