## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 22
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 37249 |  |  | CO - AMBII Serial Number: |  |  | 37249 |  | HC - AMBII Serial Number: |  |  | 37249 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-10-20 |  | 6 | X | 0.1 | 3 |  |  | 201 |  | 50 | 3 |  | 31 |  | 4 | 3 |  | 302 |  | 15 | 3 |  | 93 | 100 | 5 | 2
Test Time: | 16:54
Test Software: | 2.018 22
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 |  |  | 0 | 1004.5 | 1004.5 |  |  | 0 | -1.9 | -1.9 |  |  | 0 | -0.44 | -0.44
90 | 105 | 1 | 0.6 | 0.596 | 0.58 | -0.016 | -2.68 |  | 20.1 | 20 | 0 | 0 |  | 3.1 | 1004.9 | 1001.8 | 32316.13 | Fail | 30.2 | 28.7 | -1.5 | -4.97 |  | 9.3 | 8.83 | -0.47 | -5.05
150 | 165 | 2 | 1.2 | 1.193 | 1.15 | -0.043 | -3.6 |  | 40.2 | 40 | 0 | 0 |  | 6.2 | 1004.7 | 998.5 | 16104.84 | Fail | 60.4 | 60.3 | -0.1 | -0.17 |  | 18.6 | 18 | -0.6 | -3.23
210 | 225 | 3 | 1.8 | 1.791 | 1.726 | -0.065 | -3.63 |  | 60.3 | 67 | 7 | 11.61 |  | 9.3 | 1004.9 | 995.6 | 10705.38 | Fail | 90.6 | 90.2 | -0.4 | -0.44 |  | 27.9 | 27.19 | -0.71 | -2.54
270 | 285 | 4 | 2.4 | 2.39 | 2.32 | -0.07 | -2.93 |  | 80.4 | 80 | 0 | 0 |  | 12.4 | 1005 | 992.6 | 8004.84 | Fail | 120.8 | 119.3 | -1.5 | -1.24 |  | 37.2 | 36.45 | -0.75 | -2.02
330 | 345 | 5 | 3 | 2.989 | 2.92 | -0.069 | -2.31 |  | 100.5 | 100 | -1 | -1 |  | 15.5 | 1004.8 | 989.3 | 6382.58 | Fail | 151 | 150.3 | -0.7 | -0.46 |  | 46.5 | 45.56 | -0.94 | -2.02
390 | 405 | 6 | 3.6 | 3.59 | 3.53 | -0.06 | -1.67 |  | 120.6 | 120 | -1 | -0.83 |  | 18.6 | 1004.9 | 986.3 | 5302.69 | Fail | 181.2 | 182.1 | 0.9 | 0.5 |  | 55.8 | 55.13 | -0.67 | -1.2
450 | 465 | 7 | 4.2 | 4.191 | 4.15 | -0.041 | -0.98 |  | 140.7 | 140 | -1 | -0.71 |  | 21.7 | 1005 | 983.3 | 4531.34 | Fail | 211.4 | 212.7 | 1.3 | 0.61 |  | 65.1 | 64.02 | -1.08 | -1.66
510 | 525 | 8 | 4.8 | 4.793 | 4.76 | -0.033 | -0.69 |  | 160.8 | 158 | -3 | -1.87 |  | 24.8 | 1005 | 980.2 | 3952.42 | Fail | 241.6 | 243.3 | 1.7 | 0.7 |  | 74.4 | 73.47 | -0.93 | -1.25
570 | 585 | 9 | 5.4 | 5.396 | 5.386 | -0.01 | -0.19 |  | 180.9 | 175 | -6 | -3.32 |  | 27.9 | 1004.9 | 977 | 3501.79 | Fail | 271.8 | 274 | 2.2 | 0.81 |  | 83.7 | 82.78 | -0.92 | -1.1
630 | 645 | 10 | 6 | 6 | 5.988 | -0.012 | -0.2 |  | 201 | 194 | -7 | -3.48 |  | 31 | 1004.6 | 973.6 | 3140.65 | Fail | 302 | 302.8 | 0.8 | 0.26 |  | 93 | 92.26 | -0.74 | -0.8
690 | 705 | 9 | 5.4 | 5.396 | 5.382 | -0.014 | -0.26 |  | 180.9 | 170 | -11 | -6.08 |  | 27.9 | 1004.9 | 977 | 3501.79 | Fail | 271.8 | 274.5 | 2.7 | 0.99 |  | 83.7 | 82.77 | -0.93 | -1.11
750 | 765 | 8 | 4.8 | 4.793 | 4.766 | -0.027 | -0.56 |  | 160.8 | 159 | -2 | -1.24 |  | 24.8 | 1004.8 | 980 | 3951.61 | Fail | 241.6 | 243.5 | 1.9 | 0.79 |  | 74.4 | 73.32 | -1.08 | -1.45
810 | 825 | 7 | 4.2 | 4.191 | 4.16 | -0.031 | -0.74 |  | 140.7 | 138 | -3 | -2.13 |  | 21.7 | 1004.8 | 983.1 | 4530.41 | Fail | 211.4 | 213.7 | 2.3 | 1.09 |  | 65.1 | 64.02 | -1.08 | -1.66
870 | 885 | 6 | 3.6 | 3.59 | 3.54 | -0.05 | -1.39 |  | 120.6 | 120 | -1 | -0.83 |  | 18.6 | 1004.5 | 985.9 | 5300.54 | Fail | 181.2 | 181.7 | 0.5 | 0.28 |  | 55.8 | 54.84 | -0.96 | -1.72
930 | 945 | 5 | 3 | 2.989 | 2.922 | -0.067 | -2.24 |  | 100.5 | 100 | -1 | -1 |  | 15.5 | 1004.3 | 988.8 | 6379.35 | Fail | 151 | 151.3 | 0.3 | 0.2 |  | 46.5 | 45.48 | -1.02 | -2.19
990 | 1005 | 4 | 2.4 | 2.39 | 2.319 | -0.071 | -2.97 |  | 80.4 | 80 | 0 | 0 |  | 12.4 | 1004.4 | 992 | 8000 | Fail | 120.8 | 121.9 | 1.1 | 0.91 |  | 37.2 | 36.05 | -1.15 | -3.09
1050 | 1065 | 3 | 1.8 | 1.791 | 1.714 | -0.077 | -4.3 |  | 60.3 | 60 | 0 | 0 |  | 9.3 | 1004.8 | 995.5 | 10704.3 | Fail | 90.6 | 92.2 | 1.6 | 1.77 |  | 27.9 | 26.85 | -1.05 | -3.76
1110 | 1125 | 2 | 1.2 | 1.193 | 1.15 | -0.043 | -3.6 |  | 40.2 | 40 | 0 | 0 |  | 6.2 | 1004.6 | 998.4 | 16103.23 | Fail | 60.4 | 61.7 | 1.3 | 2.15 |  | 18.6 | 17.74 | -0.86 | -4.62
1170 | 1185 | 1 | 0.6 | 0.596 | 0.577 | -0.019 | -3.19 |  | 20.1 | 19 | -1 | -4.98 |  | 3.1 | 1004.7 | 1001.6 | 32309.68 | Fail | 30.2 | 31.6 | 1.4 | 4.64 |  | 9.3 | 8.32 | -0.98 | -10.54
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -9 | -9 |  |  | 0 | 1004.8 | 1004.8 |  |  | 0 | 0.8 | 0.8 |  |  | 0 | -1.11 | -1.11
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.04 | 0.06 |  |  |  | Intercept | 0 | 2 |  |  | Intercept | 1005 | 0 | Fail |  | Intercept | -0.4 | 3 |  |  | Intercept | -0.78 | 0.93
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.98 | 0.99 - 1.01 | Fail |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1.01 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0259 | 0.06 |  |  |  | SEE | 3.4944 | 2.01 | Fail |  | SEE | 0.2017 | 0.31 |  |  | SEE | 1.0916 | 3.02 |  |  | SEE | 0.203 | 0.93
|  |  | r2 | 0.9998 | 0.998 |  |  |  | r2 | 0.9969 | 0.998 | Fail |  | r2 | 0.0484 | 0.998 | Fail |  | r2 | 0.9999 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.9 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.9 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0009 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.9 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2024 | 0.0005 | 0.000898 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.8 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.8 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2 | 0.0005 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1039.8911 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.8 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2022 | 0.000498 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.9 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2024 | 0.0004 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.9 | 0.6 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.41 | 189 | 27 | 5 | 195 | 59.2972 | 0.0004 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.8 | 0.5976 | 0.6 | -0.1 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0008 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:54 | 0.01 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.1 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.59 | 189 | 27 | 5 | 195 | 59.2976 | 0.0004 | 0.000798 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2188 | 27.3 | 0.5
2008-10-20 | 16:55 | 0.01 | 0 | 0 | -1.7 | 0.5 | 0.5 | -0.2 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2976 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1364 | 27.3 | 0.5
2008-10-20 | 16:55 | 0.01 | 0 | 0 | -1.7 | 0.5 | 0.5 | -0.2 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2024 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.2907 | 27.3 | 0.5
2008-10-20 | 16:55 | 0.007 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.2 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 189 | 27 | 5 | 195 | 59.2976 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039.0905 | 1004.024 | 10.3 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.41 | 189 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.2006 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.1537 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2976 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.976 | 10.292 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.59 | 189 | 27 | 5 | 195 | 59.297 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.0298 | 10.1377 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.41 | 189 | 27 | 5 | 195 | 59.2 | 0.0004 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.9721 | 10.2168 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.59 | 189 | 27 | 5 | 195 | 59.3 | 0.000397 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.0259 | 10.3356 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.8 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.41 | 189 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1005 | 10.3873 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.59 | 189.4 | 27 | 5 | 195 | 59.2024 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.976 | 10.3101 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2022 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.978 | 10.2329 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.41 | 190 | 27 | 5 | 195 | 59.2976 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.2442 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004.024 | 10.3 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.5906 | 190 | 27 | 5 | 195 | 59.3 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1005 | 10.3 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -2 | 0.5 | 0.5 | -0.3 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2976 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2243 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -2 | 0.5 | 0.5 | -0.4 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2024 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.2529 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.4 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.33 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.4 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2026 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.3856 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.4 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2312 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2976 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2616 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2225 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1318 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.3 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2861 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 4.6 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.5 | 0.5 | -0.5 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 4 | 195 | 59.2976 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.2055 | 27.3 | 0.5
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.4 | 0.4 | -0.4 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190 | 27 | 4.4 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1244 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.4 | 0.4 | -0.4 | 0 | 0.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2015 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.4 | 0.4 | -0.4 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 190 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0007 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0278 | 10.2785 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.4 | 0.4 | -0.4 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 190.4 | 27 | 5 | 195 | 59.2 | 0.0003 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3555 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -1.9 | 0.4 | 0.4 | -0.4 | 0 | 0.2 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2976 | 0.0003 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3328 | 27.3672 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2024 | 0.000298 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3688 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1373 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2295 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2581 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2191 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2024 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2962 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.4 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.4 | 0.4 | -0.5 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.3 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.4
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.5 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.5 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2976 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2024 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2728 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1185 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 191.4 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2358 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3451 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 952 | 942 | 18.6 | 0.25 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3778 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 953 | 942 | 18 | 0.24 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3006 | 27.3 | 0.3
2008-10-20 | 16:55 | 0 | 0 | 0 | -2.1 | 0.3 | 0.3 | -0.6 | 0 | 0.1 | 13 | 955 | 953 | 942 | 18.4 | 0.25 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0.065 | 0.001 | 9.6308 | -2.1 | 0.2 | 0.2 | -0.5 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.525 | 0.0017 | 17.0536 | -2 | 0.2 | 0.2 | 0.3 | 0.2717 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.56 | 0.001 | 10 | 11.8 | 0.2 | 12 | 4.7 | 4.6702 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.002936 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.56 | 0.001 | 10 | 25.6 | 0.3 | 25.9 | 8.7 | 8.7173 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.005707 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0.569 | 0.001 | 10 | 27.5 | 0.29759 | 27.8 | 9.3 | 9.3 | 0.1 | 13.1 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.006002 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:55 | 0.58 | 0.001 | 10 | 28.3 | 0.2 | 28.5 | 9.2 | 9.2181 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2212 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.58 | 0.001 | 10 | 28.3 | 0.2 | 28.5 | 9.3 | 9.274 | 0.1 | 13.1 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1328 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.58 | 0.001 | 10 | 28.4 | 0.2 | 28.6 | 9.3 | 9.3 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2878 | 27.3 | 0.2
2008-10-20 | 16:55 | 0.58 | 0.001 | 10 | 28.5 | 0.2 | 28.7 | 9.3 | 9.3 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.203 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:56 | 0.58 | 0.001 | 10 | 28.5 | 0.2 | 28.7 | 9.3 | 9.3 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2026 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:56 | 0.58 | 0.001 | 10 | 28.6 | 0.2 | 28.8 | 9.2 | 9.2205 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:56 | 0.58 | 0.001 | 10 | 28.6 | 0.2 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:56 | 0.58 | 0.0014 | 14.3036 | 28.6 | 0.20239 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.2
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2432 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.234 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.3 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13.1 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2976 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0006 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3198 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.3 | 28.8 | 9.2 | 9.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.41 | 192 | 27 | 5 | 195 | 59.2026 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.397 | 27.3 | 0.3
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.4 | 28.9 | 9.2 | 9.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.59 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1775 | 27.3 | 0.4
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.5 | 0.4024 | 28.9 | 9.2 | 9.1937 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.41 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2027 | 27.3 | 0.4
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.5024 | 29.1 | 9.1 | 9.1 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.3 | 27.3 | 0.5
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.6024 | 29.2 | 9.1 | 9.1 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2943 | 27.3 | 0.6
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.7 | 29.3 | 9.1 | 9.1 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.59 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2171 | 27.3 | 0.7
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.7024 | 29.3 | 9 | 9.02 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.4096 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2601 | 27.3 | 0.7
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.8 | 29.4 | 9 | 9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26.5906 | 192 | 27 | 5 | 195 | 59.2978 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.9 | 29.5 | 9 | 9 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 26 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2711 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.6 | 0.9 | 29.5 | 8.9 | 8.9183 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 26.41 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1168 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1687 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2459 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2769 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1998 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1226 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2091 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2828 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.006198 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1289 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2261 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2028 | 0.0062 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3404 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.7 | 1 | 29.7 | 8.8 | 8.7999 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2976 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3478 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.7 | 8.7 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1166 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.7 | 8.7319 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.1718 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.7 | 8.7 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2491 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.764 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2213 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.0068 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1611 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191.6 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2578 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2 | 0.0061 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.7225 | 1005 | 10.2878 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 1 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2974 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2106 | 27.3 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9976 | 29.8 | 8.8 | 8.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2665 | 27.3665 | 1
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.8 | 8.8 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2978 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3563 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.8644 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.3 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2024 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1497 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2569 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3705 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.58 | 0.002 | 20 | 28.8 | 0.9 | 29.7 | 8.9 | 8.9 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2022 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:56 | 0.625 | 0.002 | 20 | 28.8 | 0.8 | 29.6 | 8.9 | 8.9 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0061 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:56 | 1.034 | 0.0023 | 22.5992 | 28.8 | 0.7 | 29.5 | 9 | 8.9736 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190.6 | 27 | 5 | 195 | 59.2 | 0.006108 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.116 | 0.003 | 30 | 39.9 | 0.7 | 40.6 | 11.8 | 11.775 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2026 | 0.008234 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 57 | 0.7 | 57.7 | 15.3 | 15.3144 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.011505 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3154 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 58.5 | 0.7 | 59.2 | 17.2 | 17.2087 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.011805 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3925 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 59.7 | 0.7 | 60.4 | 18 | 17.9933 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.012002 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.191 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.2 | 18.203 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.3 | 18.3 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1239 | 27.3 | 0.7
2008-10-20 | 16:56 | 1.13 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.3 | 18.3 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2009 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.142 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.3 | 18.3 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2779 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.3 | 18.3 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2451 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.2 | 18.2194 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2322 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.003 | 30 | 60.3 | 0.7 | 61 | 18.2 | 18.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2026 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.0033 | 32.6359 | 60.3 | 0.7 | 61 | 18.2 | 18.1687 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.6 | 60.9 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.6024 | 60.9 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.69759 | 61 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.7977 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.6 | 60.9 | 18.1 | 18.0693 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.6 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1998 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.6 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1272 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.2044 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2816 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2026 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.1826 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2022 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1261 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2805 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3674 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3554 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2028 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.2 | 0.7 | 60.9 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188.6 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2024 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2024 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2837 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2068 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2976 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2703 | 27.3703 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2527 | 27.3527 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2243 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.004 | 40 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.15 | 0.0041 | 40.8319 | 60.3 | 0.7 | 61 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.156 | 0.0049 | 49.1514 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2976 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.4 | 0.7 | 61.1 | 18 | 18 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.4 | 0.69759 | 61.1 | 18 | 18 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.4 | 0.6 | 61 | 18 | 18 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.1 | 18.0987 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3195 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3967 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.1 | 18.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9721 | 10.3262 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.2 | 18.161 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2026 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.2 | 18.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6 | 60.9 | 18.2 | 18.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.6024 | 60.9 | 18.2 | 18.2 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2976 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.6
2008-10-20 | 16:57 | 1.16 | 0.005 | 50 | 60.3 | 0.7 | 61 | 18.2 | 18.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2024 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.168 | 0.005 | 50 | 60.3 | 0.7 | 61 | 18.2 | 18.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 27 | 5 | 195 | 59.2 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3368 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.245 | 0.0059 | 58.6364 | 60.3 | 0.7 | 61 | 18.2 | 18.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.372 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.635 | 0.0067 | 67.362 | 67.2 | 0.7 | 67.9 | 18.5 | 18.5331 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0134 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2176 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.686 | 0.006 | 60 | 72.3 | 0.7 | 73 | 22.5 | 22.4542 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.014336 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2684 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.696 | 0.006 | 60 | 86.8 | 0.7 | 87.5 | 25.6 | 25.6346 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.017105 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.209 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.7 | 0.006 | 60 | 88.6 | 0.7 | 89.3 | 27 | 27.0336 | 0.1 | 13 | 959 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0174 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1681 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.7 | 0.006 | 60 | 88.9 | 0.7 | 89.6 | 27.5 | 27.5221 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2026 | 0.0175 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3995 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.7 | 0.006 | 60 | 89 | 0.7 | 89.7 | 27.6 | 27.6444 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2976 | 0.0175 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.246 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.7 | 0.006 | 60 | 89 | 0.7 | 89.7 | 27.6 | 27.6 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0175 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3083 | 27.3 | 0.7
2008-10-20 | 16:57 | 1.7 | 0.006 | 60 | 89 | 0.7 | 89.7 | 27.6 | 27.5682 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0175 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3687 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.7 | 0.006 | 60 | 89.6 | 0.7 | 90.3 | 27.5 | 27.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.017502 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.704 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.4 | 27.437 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.4 | 27.3634 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2206 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.3 | 27.3 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1334 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.3 | 27.3 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.2875 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2974 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2976 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.71 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.711 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.3 | 27.3 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.294 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.719 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.3 | 27.3314 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1397 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.72 | 0.006 | 60 | 89.9 | 0.7 | 90.6 | 27.4 | 27.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0176 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2147 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.72 | 0.006 | 60 | 90 | 0.7 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.72 | 0.006 | 60 | 90 | 0.7 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2022 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2883 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.722 | 0.006 | 60 | 90 | 0.7 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2112 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90 | 0.7 | 90.7 | 27.4 | 27.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2972 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.266 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.1 | 0.7 | 90.8 | 27.3 | 27.3184 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2136 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.1 | 0.7 | 90.8 | 27.3 | 27.3 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1407 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.1 | 0.7024 | 90.8 | 27.3 | 27.2684 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2976 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2951 | 27.3 | 0.7
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.1 | 0.8 | 90.9 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.1 | 0.8 | 90.9 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189.4 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.006 | 60 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.0061 | 60.7256 | 90.2 | 0.8 | 91 | 27.1 | 27.1383 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.0069 | 69.0659 | 90.2 | 0.8 | 91 | 27.2 | 27.1815 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.728 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.2972 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.9721 | 10.2853 | 27.3147 | 0.8
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.1 | 27.1195 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.208 | 27.392 | 0.8
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.1685 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3383 | 27.4 | 0.8
2008-10-20 | 16:58 | 1.727 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3537 | 27.3537 | 0.8
2008-10-20 | 16:58 | 1.73 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.726 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.8 | 91 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.9 | 91.1 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.9 | 91.1 | 27.2 | 27.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3091 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.9 | 91.1 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2972 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3861 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.9 | 91.1 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2974 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.3368 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.2 | 0.9 | 91.1 | 27.2 | 27.2 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.282 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.2 | 27.1688 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2049 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2723 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2011 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190.4 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1003.978 | 10.1532 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1003.024 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2976 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.1 | 27.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2972 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.2 | 27.1636 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2974 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.337 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.007 | 70 | 90.3 | 0.9 | 91.2 | 27.2 | 27.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5685 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.72 | 0.0064 | 64.0468 | 90.3 | 0.9 | 91.2 | 27.2 | 27.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2976 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3998 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.724 | 0.006 | 60 | 90.3 | 0.9 | 91.2 | 27.2 | 27.2312 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2122 | 27.3 | 0.9
2008-10-20 | 16:58 | 1.774 | 0.0063 | 62.5772 | 90.2 | 0.9 | 91.1 | 27.3 | 27.3 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2 | 0.0177 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.7254 | 1005 | 10.1632 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.14 | 0.007 | 70 | 90.4 | 0.9 | 91.3 | 27.8 | 27.8212 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2 | 0.017734 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3947 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.27 | 0.007 | 70 | 105.1 | 0.9 | 106 | 30.2 | 30.195 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2 | 0.020434 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1738 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.27 | 0.0077 | 77.3846 | 115.7 | 0.9 | 116.6 | 34 | 33.9568 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2976 | 0.022402 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2575 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.275 | 0.008 | 80 | 117.3 | 0.9 | 118.2 | 35.7 | 35.6797 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.3 | 0.022602 | 0.0003 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3407 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.29 | 0.008 | 80 | 118.2 | 0.9 | 119.1 | 36.3 | 36.2616 | 0.1 | 13.1 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2 | 0.0228 | 0.000303 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2136 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.29 | 0.008 | 80 | 118.2 | 0.9 | 119.1 | 36.5 | 36.5 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2028 | 0.0228 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3679 | 27.3 | 0.9
2008-10-20 | 16:58 | 2.29 | 0.008 | 80 | 118.3 | 0.9 | 119.2 | 36.6 | 36.6 | 0.1 | 13.1 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0228 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3389 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.296 | 0.008 | 80 | 118.3 | 0.9 | 119.2 | 36.6 | 36.6 | 0.1 | 13 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.4 | 28 | 5 | 195 | 59.2 | 0.0228 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3383 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.3 | 0.9 | 119.2 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0228 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.9759 | 10.3846 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.3 | 0.9 | 119.2 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.022802 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3073 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.4 | 0.9 | 119.3 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2976 | 0.0229 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.1603 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.5 | 0.9 | 119.4 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2 | 0.0229 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.147 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.5 | 0.9 | 119.4 | 36.6 | 36.6 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0229 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.248 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 118.8 | 0.8976 | 119.7 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2 | 0.0229 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3991 | 27.3 | 0.9
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 119.1 | 0.8 | 119.9 | 36.6 | 36.6 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3222 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 119.2 | 0.8 | 120 | 36.6 | 36.6 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2972 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3553 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 119.2 | 0.8 | 120 | 36.6 | 36.6 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3673 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.3 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3097 | 0.8
2008-10-20 | 16:59 | 2.308 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.6 | 36.5607 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3867 | 0.8
2008-10-20 | 16:59 | 2.31 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3359 | 0.8
2008-10-20 | 16:59 | 2.314 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4395 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3186 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3957 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3272 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.35 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.3 | 36.3302 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2028 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.3 | 36.3 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2974 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3957 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.8 | 120.2 | 36.3 | 36.3 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3185 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.3 | 36.3 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3586 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.8 | 120.2 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.9759 | 10.387 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.8 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.976 | 10.3099 | 27.3 | 0.8
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3444 | 27.3 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3216 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.4809 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3988 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3241 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.5 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.353 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.6 | 36.5819 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3698 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.6 | 36.6 | 0.1 | 13.1 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3073 | 27.3927 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.6 | 36.6 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3845 | 27.3155 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.6 | 36.6 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3383 | 27.3617 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.6 | 36.5684 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3388 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.384 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2976 | 0.023 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3069 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.3 | 0.7 | 120 | 36.4 | 36.4189 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2974 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3943 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4516 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.4 | 36.4 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2988 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.314 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2217 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.314 | 0.008 | 80 | 119.4 | 0.7 | 120.1 | 36.4 | 36.4 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2976 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3116 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.5 | 0.7 | 120.2 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.2024 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3668 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.5 | 0.7 | 120.2 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3103 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.6 | 0.7 | 120.3 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0231 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3875 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 119.6 | 0.7 | 120.3 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2024 | 0.023102 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 120 | 0.7 | 120.7 | 36.4 | 36.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.0232 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3582 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 120 | 0.7 | 120.7 | 36.5 | 36.4627 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2976 | 0.0232 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.32 | 0.008 | 80 | 120.1 | 0.7 | 120.8 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 27 | 5 | 195 | 59.2 | 0.0232 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.331 | 0.008 | 80 | 120.5 | 0.7 | 121.2 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2974 | 0.0233 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.453 | 0.0083 | 82.5469 | 120.5 | 0.7 | 121.2 | 36.5 | 36.5 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2 | 0.0233 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3495 | 0.7
2008-10-20 | 16:59 | 2.787 | 0.0091 | 90.7108 | 120.9 | 0.7 | 121.6 | 36.7 | 36.6981 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2 | 0.023348 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3276 | 27.3276 | 0.7
2008-10-20 | 16:59 | 2.91 | 0.0099 | 98.8807 | 139.5 | 0.7 | 140.2 | 39.8 | 39.8429 | 0.1 | 13.1 | 958 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.026626 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3952 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.91 | 0.01 | 100 | 147.6 | 0.7 | 148.3 | 43.3 | 43.2594 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2978 | 0.028102 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3181 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.91 | 0.01 | 100 | 149 | 0.7 | 149.7 | 44.9 | 44.9231 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0283 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.911 | 0.01 | 100 | 149.3 | 0.7 | 150 | 45.4 | 45.3944 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.919 | 0.01 | 100 | 149.3 | 0.7 | 150 | 45.6 | 45.6 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2024 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 16:59 | 2.92 | 0.01 | 100 | 149.3 | 0.7 | 150 | 45.7 | 45.6713 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.4 | 0.7 | 150.1 | 45.7 | 45.7 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2972 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.4 | 0.7 | 150.1 | 45.8 | 45.7796 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.5 | 0.7 | 150.2 | 45.7 | 45.7 | 0.1 | 13 | 959 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2024 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.6 | 0.7 | 150.3 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2024 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.6 | 0.7 | 150.3 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.6 | 0.7 | 150.3 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 27.4 | 5 | 195 | 59.2 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.7 | 0.7 | 150.4 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.7 | 0.8 | 150.5 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.7 | 0.8 | 150.5 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.3 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.7 | 0.8 | 150.5 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.362 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.8 | 0.8 | 150.6 | 45.7 | 45.7 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3608 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.8 | 0.8 | 150.6 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3489 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.8 | 0.8 | 150.6 | 45.8 | 45.7597 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5795 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.8 | 0.8 | 150.6 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0284 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4589 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 149.9 | 0.8 | 150.7 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.8
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150 | 0.7 | 150.7 | 45.8 | 45.8443 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2976 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.1 | 0.7 | 150.8 | 45.9 | 45.9 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2976 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3982 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.1 | 0.7 | 150.8 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3211 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.1 | 0.7 | 150.8 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 27.4 | 5 | 195 | 59.3 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.976 | 10.3 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.2 | 0.7 | 150.9 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3336 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.2 | 0.7 | 150.9 | 45.7 | 45.7191 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.3893 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.7 | 151 | 45.6 | 45.6276 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3122 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.7 | 151 | 45.6 | 45.6 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.365 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.7 | 151 | 45.6 | 45.6 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.7
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.6 | 150.9 | 45.6 | 45.6 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.6
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.6 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2024 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.6
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.6 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 959 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3264 | 27.4 | 0.6
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2978 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3508 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3721 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.3 | 0.5 | 150.8 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2024 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.305 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3822 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.3 | 0.0285 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.0286 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.4 | 0.5 | 150.9 | 45.5 | 45.5 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0286 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.5 | 0.5 | 151 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0286 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.9 | 0.5 | 151.4 | 45.5 | 45.5 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2 | 0.0286 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3549 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 150.9 | 0.5 | 151.4 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2 | 0.0286 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151 | 0.5 | 151.5 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2976 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.1 | 0.5 | 151.6 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3765 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.2 | 0.5 | 151.7 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.3 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4536 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.2 | 0.5 | 151.7 | 45.4 | 45.4191 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4385 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.2 | 0.5 | 151.7 | 45.4 | 45.4 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2972 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.2 | 0.5 | 151.7 | 45.4 | 45.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2974 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.4 | 45.4 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2976 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.4 | 45.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3393 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.5 | 45.5 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3837 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.6 | 45.562 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.2 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3066 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.6 | 45.6 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2296 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.2 | 0.5 | 151.7 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27.6 | 5 | 195 | 59.2976 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2949 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.92 | 0.01 | 100 | 151.3 | 0.5 | 151.8 | 45.7 | 45.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 27.4 | 5 | 195 | 59.3 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.5
2008-10-20 | 17:00 | 2.959 | 0.0104 | 103.7542 | 151.3 | 0.4 | 151.7 | 45.8 | 45.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.0287 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3984 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.41 | 0.011 | 110 | 151.9 | 0.4 | 152.3 | 45.9 | 45.922 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0288 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3207 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.47 | 0.011 | 110 | 159.6 | 0.4 | 160 | 49.1 | 49.0635 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2028 | 0.030136 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.487 | 0.011 | 110 | 176.7 | 0.4 | 177.1 | 52.3 | 52.2684 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2026 | 0.033105 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.5 | 0.011 | 110 | 178.9 | 0.4 | 179.3 | 54.1 | 54.0538 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0335 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.5 | 0.011 | 110 | 179.3 | 0.4 | 179.7 | 54.7 | 54.6887 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.033502 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.5 | 0.011 | 110 | 179.8 | 0.4 | 180.2 | 54.9 | 54.8628 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0336 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.5 | 0.011 | 110 | 179.9 | 0.4 | 180.3 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0336 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:00 | 3.507 | 0.011 | 110 | 180 | 0.4 | 180.4 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0337 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3193 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.51 | 0.011 | 110 | 180.6 | 0.4 | 181 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2024 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3965 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.51 | 0.0117 | 117.1324 | 180.6 | 0.4 | 181 | 54.8 | 54.8367 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.3 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.512 | 0.012 | 120 | 180.7 | 0.4 | 181.1 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.2845 | 1004 | 10.4508 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 180.8 | 0.4 | 181.2 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3883 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.1 | 0.4 | 181.5 | 54.8 | 54.8 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1153 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.2 | 0.4 | 181.6 | 54.8 | 54.7848 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0338 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3466 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.2 | 0.4 | 181.6 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.3 | 0.4 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3635 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.3 | 0.4 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.3 | 0.4 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.52 | 0.012 | 120 | 181.4 | 0.4 | 181.8 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2024 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3678 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.527 | 0.012 | 120 | 181.4 | 0.4 | 181.8 | 54.8 | 54.7612 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.355 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 181.4 | 0.4 | 181.8 | 54.8 | 54.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2976 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3221 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 181.6 | 0.4 | 182 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3993 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 181.6 | 0.4 | 182 | 55 | 54.98 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2974 | 0.0339 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3235 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 181.7 | 0.4 | 182.1 | 55.1 | 55.131 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.033902 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.4 | 182.4 | 55.2 | 55.2 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.6429 | 1005 | 10.3614 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.4 | 182.4 | 55.3 | 55.3 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4921 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.39779 | 182.4 | 55.3 | 55.3307 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2978 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.415 | 27.4 | 0.4
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55.4 | 55.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.8036 | 1005 | 10.4 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.4 | 55.3685 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3608 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.3 | 55.3 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3163 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.3 | 55.3 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3932 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.3 | 55.3 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2976 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4702 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.2 | 55.22 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.4 | 27.4 | 5 | 195 | 59.2972 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4057 | 27.4471 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.1 | 55.1302 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2974 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3244 | 27.4756 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 55.1 | 55.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3979 | 27.4021 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55.1 | 55.1 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3209 | 27.4791 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55.1 | 55.0681 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3563 | 27.4437 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55 | 55 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3665 | 27.4 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55 | 55 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2976 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4106 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 55 | 54.9689 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4877 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2024 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 54.9 | 54.9 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 54.9 | 54.9 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.2 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3191 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182.1 | 0.3 | 182.4 | 54.9 | 54.8636 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3962 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.53 | 0.012 | 120 | 182 | 0.3 | 182.3 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4266 | 0.3
2008-10-20 | 17:01 | 3.534 | 0.012 | 120 | 182 | 0.3 | 182.3 | 54.9 | 54.8684 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3494 | 27.4506 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.9 | 0.3 | 182.2 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3277 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.9 | 0.3 | 182.2 | 54.8 | 54.8 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3952 | 27.4952 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.8 | 0.3 | 182.1 | 54.8 | 54.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3181 | 27.4181 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.8 | 0.3 | 182.1 | 54.8 | 54.8 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.3 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4591 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.8 | 0.3 | 182.1 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2024 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3362 | 27.4638 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.8 | 0.3 | 182.1 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.2 | 0.034 | 0.0004 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3867 | 27.4133 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.7 | 0.3 | 182 | 54.8 | 54.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3096 | 27.4904 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.6 | 0.3 | 181.9 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.4 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3675 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.5 | 0.3 | 181.8 | 54.8 | 54.8 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3106 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.5 | 0.3 | 181.8 | 54.8 | 54.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2976 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2873 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.5 | 0.3 | 181.8 | 54.8 | 54.7691 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.3 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.976 | 10.596 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.5 | 0.3 | 181.8 | 54.7 | 54.7 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4477 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.012 | 120 | 181.4 | 0.3 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.6 | 5 | 195 | 59.2978 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.0124 | 123.619 | 181.4 | 0.3 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2024 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4608 | 27.4696 | 0.3
2008-10-20 | 17:01 | 3.54 | 0.013 | 130 | 181.4 | 0.3 | 181.7 | 54.7 | 54.7 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.034 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5775 | 27.4075 | 0.3
2008-10-20 | 17:01 | 3.94 | 0.013 | 129.7001 | 181.5 | 0.3 | 181.8 | 54.7 | 54.7 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2976 | 0.034007 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3461 | 27.4846 | 0.3
2008-10-20 | 17:01 | 4.115 | 0.0121 | 121.1311 | 191.8 | 0.3 | 192.1 | 56.7 | 56.7378 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2024 | 0.035755 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3618 | 27.5 | 0.3
2008-10-20 | 17:01 | 4.12 | 0.012 | 120 | 206.5 | 0.3 | 206.8 | 60.2 | 60.2382 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27 | 5 | 195 | 59.2 | 0.038205 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3611 | 27.5 | 0.3
2008-10-20 | 17:01 | 4.122 | 0.012 | 120 | 209.6 | 0.3 | 209.9 | 62.7 | 62.693 | 0.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 27.4 | 5 | 195 | 59.2028 | 0.0387 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.316 | 27.5 | 0.3
2008-10-20 | 17:01 | 4.13 | 0.012 | 120 | 210.2 | 0.3 | 210.5 | 63.6 | 63.6178 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2974 | 0.0388 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3931 | 27.5 | 0.3
2008-10-20 | 17:01 | 4.13 | 0.012 | 120 | 210.3 | 0.3 | 210.6 | 63.9 | 63.8711 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0388 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3299 | 27.5 | 0.3
2008-10-20 | 17:01 | 4.13 | 0.012 | 120.275 | 210.3 | 0.3 | 210.6 | 63.9 | 63.9319 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0388 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.347 | 27.5 | 0.3
2008-10-20 | 17:02 | 4.13 | 0.0129 | 128.6083 | 210.3 | 0.3 | 210.6 | 64 | 64 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0388 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.3
2008-10-20 | 17:02 | 4.13 | 0.013 | 130 | 210.7 | 0.3 | 211 | 64 | 64 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0389 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.399 | 27.5 | 0.3
2008-10-20 | 17:02 | 4.13 | 0.013 | 130 | 211.1 | 0.3 | 211.4 | 64 | 64 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.4 | 27.6 | 5 | 195 | 59.2976 | 0.0389 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3214 | 27.5 | 0.3
2008-10-20 | 17:02 | 4.13 | 0.013 | 130 | 211.2 | 0.4 | 211.6 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27 | 5 | 195 | 59.2024 | 0.038902 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3559 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.13 | 0.013 | 130 | 211.3 | 0.4 | 211.7 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.3 | 0.039 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.144 | 0.0131 | 130.654 | 211.3 | 0.4 | 211.7 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.039 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.0139 | 138.9321 | 211.4 | 0.4 | 211.8 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 27.6 | 5 | 195 | 59.2026 | 0.039 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 211.7 | 0.4 | 212.1 | 64 | 64 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 27.4 | 5 | 195 | 59.3 | 0.039 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3356 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 211.9 | 0.4 | 212.3 | 64 | 64 | 0.1 | 13.1 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4245 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.2 | 0.4 | 212.6 | 64 | 64 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.544 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.3 | 0.4 | 212.7 | 64 | 64 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3127 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.3 | 0.4 | 212.7 | 64 | 64 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3729 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.3 | 0.4 | 212.7 | 64 | 64 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.4 | 0.4 | 212.8 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3728 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.4 | 0.4 | 212.8 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.4 | 0.4024 | 212.8 | 64 | 64 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.4
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.4 | 0.5 | 212.9 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4756 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.157 | 0.014 | 140 | 212.4 | 0.5 | 212.9 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2974 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4931 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.16 | 0.014 | 140 | 212.4 | 0.5 | 212.9 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0391 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.156 | 0.014 | 140 | 212.5 | 0.5 | 213 | 64 | 64 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.5 | 213.3 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.5 | 213.3 | 64 | 64 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3882 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.5 | 213.3 | 64 | 64 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2022 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4788 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.5022 | 213.2 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4015 | 27.5 | 0.5
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.6 | 213.3 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.6
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.6022 | 213.3 | 64 | 64 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.6
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64 | 64 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3931 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.7 | 213.5 | 64 | 64 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.316 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64 | 64 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3611 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64 | 64 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3618 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.7 | 213.5 | 64 | 64 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3153 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3924 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2598 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.157 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1716 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.155 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.203 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.152 | 0.014 | 140 | 212.8 | 0.7 | 213.5 | 64.1 | 64.1 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 0.7
2008-10-20 | 17:02 | 4.16 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3683 | 27.5317 | 0.7
2008-10-20 | 17:02 | 4.151 | 0.014 | 140 | 212.7 | 0.7 | 213.4 | 64.1 | 64.1 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5912 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.7 | 0.7022 | 213.4 | 64.1 | 64.0692 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5142 | 0.7
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.8 | 213.6 | 64 | 64 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.8 | 0.8 | 213.6 | 64 | 64 | 0.1 | 13.1 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2978 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 64 | 64 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2822 | 27.5 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 64 | 64 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.976 | 10.205 | 27.5 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 64 | 64 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3442 | 27.5721 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 63.9 | 63.9191 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2022 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.6 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 63.9 | 63.9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2028 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3472 | 27.5736 | 0.8
2008-10-20 | 17:02 | 4.15 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 63.9 | 63.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2974 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2035 | 27.5035 | 0.8
2008-10-20 | 17:02 | 4.148 | 0.014 | 140 | 212.9 | 0.8 | 213.7 | 63.9 | 63.8619 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2806 | 27.5806 | 0.8
2008-10-20 | 17:02 | 4.14 | 0.014 | 140 | 212.8 | 0.8 | 213.6 | 63.8 | 63.8 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:02 | 4.14 | 0.0142 | 141.8265 | 212.8 | 0.8 | 213.6 | 63.8 | 63.8 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0392 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:02 | 4.538 | 0.015 | 150 | 212.9 | 0.8 | 213.7 | 63.9 | 63.8849 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.039217 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.0241 | 10.3 | 27.588 | 0.8
2008-10-20 | 17:02 | 4.72 | 0.015 | 150 | 224.9 | 0.8 | 225.7 | 66 | 65.9672 | 0.1 | 13.1 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.041248 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5109 | 0.8
2008-10-20 | 17:02 | 4.724 | 0.015 | 150 | 237.5 | 0.8 | 238.3 | 70 | 70.0326 | 0.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.043207 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3662 | 27.5 | 0.8
2008-10-20 | 17:02 | 4.73 | 0.015 | 150 | 240 | 0.8 | 240.8 | 72.2 | 72.1926 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 28 | 5 | 195 | 59.2 | 0.0437 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3567 | 27.5433 | 0.8
2008-10-20 | 17:02 | 4.73 | 0.015 | 150 | 241.7 | 0.8 | 242.5 | 73 | 73.0151 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0439 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:02 | 4.73 | 0.015 | 150 | 241.8 | 0.8 | 242.6 | 73.2 | 73.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0439 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:02 | 4.737 | 0.015 | 150 | 241.8 | 0.8 | 242.6 | 73.2 | 73.2 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2026 | 0.0439 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5253 | 0.8
2008-10-20 | 17:03 | 4.74 | 0.015 | 150 | 241.9 | 0.8 | 242.7 | 73.2 | 73.2317 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.043902 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1964 | 27.5 | 0.8
2008-10-20 | 17:03 | 4.74 | 0.015 | 150 | 241.9 | 0.8 | 242.7 | 73.3 | 73.3 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1578 | 27.5289 | 0.8
2008-10-20 | 17:03 | 4.74 | 0.015 | 150 | 241.9 | 0.8 | 242.7 | 73.3 | 73.3 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.74 | 0.015 | 150 | 242 | 0.8 | 242.8 | 73.3 | 73.3314 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.748 | 0.015 | 150 | 242.1 | 0.8 | 242.9 | 73.4 | 73.4 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1796 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.75 | 0.015 | 150.1493 | 242.1 | 0.8 | 242.9 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2978 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1746 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.755 | 0.0158 | 158.4411 | 242.2 | 0.8 | 243 | 73.4 | 73.4 | 0.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2712 | 27.5856 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.0153 | 153.2333 | 242.2 | 0.8 | 243 | 73.4 | 73.4 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.044 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.117 | 27.5085 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.015 | 150 | 242.4 | 0.8 | 243.2 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 28 | 5 | 195 | 59.3 | 0.044002 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2372 | 27.5686 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.015 | 150 | 242.8 | 0.8 | 243.6 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2543 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.015 | 150 | 242.8 | 0.8 | 243.6 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2026 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.015 | 150.1039 | 242.8 | 0.8 | 243.6 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.0159 | 158.7619 | 242.9 | 0.8024 | 243.7 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 0.8
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 242.9 | 0.9 | 243.8 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2543 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 242.9 | 0.9 | 243.8 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 242.9 | 0.9 | 243.8 | 73.4 | 73.4 | 0.1 | 13 | 958 | 952 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 242.9 | 0.9 | 243.8 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243 | 0.9 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2978 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2374 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243 | 0.9 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.297 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2396 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243 | 0.9 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0441 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.1 | 0.9 | 244 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.2 | 0.9 | 244.1 | 73.4 | 73.4148 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 28 | 5 | 195 | 59.2024 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.3 | 0.9 | 244.2 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.3 | 0.9 | 244.2 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.3 | 0.9 | 244.2 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.4 | 0.9 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.4 | 0.9024 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.3 | 1 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.016 | 160 | 243.3 | 1 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.016 | 159.8822 | 243.3 | 1 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.022 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.0151 | 151.4646 | 243.4 | 1 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.0157 | 156.7984 | 243.4 | 1 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.0155 | 154.9083 | 243.4 | 1 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:03 | 4.76 | 0.0153 | 153.4279 | 243.4 | 0.9 | 244.3 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.76 | 0.0158 | 158.223 | 243.4 | 0.9 | 244.3 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1528 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.781 | 0.015 | 150 | 243.4 | 0.9 | 244.3 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.785 | 0.015 | 150 | 243.4 | 0.9 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.4 | 0.9 | 244.3 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1099 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2641 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2638 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 0.9026 | 244.4 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 1 | 244.5 | 73.5 | 73.5 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 1
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 1 | 244.5 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2 | 27.6 | 1
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 1 | 244.5 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2446 | 27.6 | 1
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 1 | 244.5 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2782 | 27.6 | 1
2008-10-20 | 17:03 | 4.788 | 0.015 | 150 | 243.5 | 0.9976 | 244.5 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.2011 | 27.6 | 1
2008-10-20 | 17:03 | 4.79 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.2 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.78 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2978 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1469 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.77 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1604 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.77 | 0.015 | 150 | 243.4 | 0.9 | 244.3 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 4.775 | 0.0153 | 153.4108 | 243.4 | 0.9 | 244.3 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.1 | 0.0158 | 158.2587 | 243.3 | 0.9 | 244.2 | 73.4 | 73.4 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2026 | 0.0442 | 0.0005 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.302 | 0.015 | 150.3774 | 243.4 | 0.9 | 244.3 | 74.5 | 74.5378 | 0.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.044217 | 0.000502 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.2227 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.0167 | 166.7842 | 263 | 0.9 | 263.9 | 78.1 | 78.1398 | 0.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.047317 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.017 | 170 | 270.4 | 0.9 | 271.3 | 80.9 | 80.9099 | 0.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.048502 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.017 | 170 | 270.8 | 0.9 | 271.7 | 82.2 | 82.1504 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0486 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.017 | 170 | 271.3 | 0.9 | 272.2 | 82.5 | 82.5428 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.048602 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.194 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.017 | 170 | 271.4 | 0.9 | 272.3 | 82.7 | 82.6598 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0487 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:03 | 5.37 | 0.017 | 170 | 271.8 | 0.9 | 272.7 | 82.7 | 82.7 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0487 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.37 | 0.017 | 170 | 271.9 | 0.9 | 272.8 | 82.7 | 82.7 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.048702 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.373 | 0.017 | 170 | 272.1 | 0.9 | 273 | 82.7 | 82.7315 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2972 | 0.0488 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 272.5 | 0.9 | 273.4 | 82.8 | 82.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2026 | 0.0488 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 272.9 | 0.9 | 273.8 | 82.8 | 82.831 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0489 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 273.3 | 0.9022 | 274.2 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2022 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 273.3 | 1 | 274.3 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:04 | 5.385 | 0.017 | 170 | 273.3 | 1 | 274.3 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.3 | 1 | 274.3 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.4 | 1 | 274.4 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.4 | 1 | 274.4 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005.976 | 10.1555 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.5 | 1 | 274.5 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.5 | 1 | 274.5 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.153 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.5 | 1 | 274.5 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2964 | 27.6 | 1
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.5 | 1 | 274.5 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2193 | 27.6 | 1
2008-10-20 | 17:04 | 5.388 | 0.017 | 170 | 273.7 | 1 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2578 | 27.6 | 1
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 273.8 | 1 | 274.8 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2301 | 27.6 | 1
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 273.8 | 0.9 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.017 | 170 | 273.8 | 0.9 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.385 | 0.017 | 170 | 273.8 | 0.9 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.8 | 0.9 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.9 | 0.9 | 274.8 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.9 | 0.9 | 274.8 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 273.9 | 0.9 | 274.8 | 82.8 | 82.826 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2494 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 274 | 0.9 | 274.9 | 82.8 | 82.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 274 | 0.9 | 274.9 | 82.8 | 82.8 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 274 | 0.9 | 274.9 | 82.8 | 82.8 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.288 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.017 | 170 | 274 | 0.9 | 274.9 | 82.8 | 82.8 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1338 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.0173 | 173.3855 | 274 | 0.9 | 274.9 | 82.7 | 82.7 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.7 | 82.7317 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1746 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.383 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.8 | 82.7622 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2712 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.7 | 82.7 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.117 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.7 | 82.7 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.2 | 0.9 | 275.1 | 82.7 | 82.7 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.6 | 5 | 195 | 59.2028 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.7 | 82.7 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.2 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.8 | 82.7806 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.8 | 82.8 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.254 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.385 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.8 | 82.8 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.387 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.8 | 82.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.4 | 28 | 5 | 195 | 59.2976 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274 | 0.9 | 274.9 | 82.8 | 82.8 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2833 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1291 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.38 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.387 | 0.018 | 180 | 274.1 | 0.9 | 275 | 82.9 | 82.9 | 0.1 | 13 | 957 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 274 | 0.9 | 274.9 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1334 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 274 | 0.9 | 274.9 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2028 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2875 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 274 | 0.9 | 274.9 | 82.9 | 82.9 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0491 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1584 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 273.8 | 0.9 | 274.7 | 82.9 | 82.9 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.049098 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 273.5 | 0.9 | 274.4 | 82.9 | 82.9 | 0.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 273.5 | 0.9 | 274.4 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 273.4 | 0.9 | 274.3 | 82.9 | 82.9 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.39 | 0.018 | 180 | 273.4 | 0.9 | 274.3 | 82.9 | 82.9 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.474 | 0.018 | 180 | 273.2 | 0.9 | 274.1 | 83 | 82.9638 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.049 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.9 | 0.018 | 180 | 273.3 | 0.9 | 274.2 | 83.1 | 83.082 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.04901 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1213 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.926 | 0.018 | 180 | 288.8 | 0.9 | 289.7 | 85.6 | 85.5759 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2022 | 0.051411 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2755 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.951 | 0.018 | 180 | 291.9 | 0.9 | 292.8 | 89.5 | 89.534 | 0.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.051912 | 0.0006 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1703 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.96 | 0.018 | 180 | 298.8 | 0.9 | 299.7 | 91.5 | 91.5061 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.053006 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.96 | 0.0187 | 186.6612 | 300.6 | 0.9 | 301.5 | 92.2 | 92.2203 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.96 | 0.019 | 190 | 300.7 | 0.9 | 301.6 | 92.4 | 92.3824 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.96 | 0.019 | 190 | 300.7 | 0.9 | 301.6 | 92.4 | 92.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2466 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.968 | 0.019 | 190 | 300.7 | 0.9 | 301.6 | 92.4 | 92.4 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:04 | 5.97 | 0.019 | 190 | 300.8 | 0.9 | 301.7 | 92.4 | 92.4 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2725 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 300.8 | 0.9 | 301.7 | 92.4 | 92.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1955 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 300.9 | 0.9 | 301.8 | 92.3 | 92.338 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.053302 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1183 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 300.9 | 0.9 | 301.8 | 92.3 | 92.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0534 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1588 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 301 | 0.9 | 301.9 | 92.3 | 92.3 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0534 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 301.4 | 0.9 | 302.3 | 92.2 | 92.2205 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.053403 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.213 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 301.5 | 0.9 | 302.4 | 92.2 | 92.2 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0535 | 0.0007 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2901 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 301.5 | 0.9 | 302.4 | 92.1 | 92.1276 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0535 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 302.2 | 0.9 | 303.1 | 92.1 | 92.1 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0536 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 302.5 | 0.8 | 303.3 | 92.1 | 92.1 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2022 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 302.6 | 0.8 | 303.4 | 92.1 | 92.1 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 302.6 | 0.8 | 303.4 | 92.1 | 92.1 | 0.1 | 13.1 | 952 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3756 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.97 | 0.019 | 190 | 302.6 | 0.8 | 303.4 | 92 | 92 | 0.1 | 13 | 952 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2947 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.975 | 0.019 | 190 | 302.6 | 0.8 | 303.4 | 92 | 92 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2298 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.6 | 0.8 | 303.4 | 92 | 92 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.6 | 0.80279 | 303.4 | 92.1 | 92.0674 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191.4 | 28 | 5 | 195 | 59.2972 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.5 | 0.9 | 303.4 | 92.1 | 92.1 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.5 | 0.9 | 303.4 | 92.1 | 92.1 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2237 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.6 | 0.9 | 303.5 | 92.1 | 92.1 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.6 | 0.9 | 303.5 | 92 | 92.0311 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.1 | 92.1 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2388 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.1 | 92.1 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2069 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.98 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.1 | 92.1 | 0.1 | 13.1 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.988 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.1 | 92.1353 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.2 | 92.2 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.2 | 92.2 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1549 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.019 | 190 | 302.7 | 0.9 | 303.6 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2028 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2319 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.019 | 190 | 302.8 | 0.9 | 303.7 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0537 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.019 | 190 | 302.9 | 0.9 | 303.8 | 92.3 | 92.3 | 0.1 | 13 | 951 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.053702 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.0192 | 191.559 | 302.9 | 0.9 | 303.8 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.237 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303 | 0.9 | 303.9 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0241 | 10.1599 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 302.9 | 0.9 | 303.8 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1342 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 302.9 | 0.9 | 303.8 | 92.3 | 92.33 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2883 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 302.9 | 0.9 | 303.8 | 92.4 | 92.4 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.157 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 302.9 | 0.9 | 303.8 | 92.4 | 92.4 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303 | 0.9 | 303.9 | 92.5 | 92.4812 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1518 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.1 | 0.9 | 304 | 92.6 | 92.5622 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.294 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.1 | 0.9 | 304 | 92.7 | 92.6684 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0538 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1398 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.7 | 0.9 | 304.6 | 92.7 | 92.7 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1572 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.7 | 92.7 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2343 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.7 | 92.6703 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.6 | 92.6 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.5 | 92.5314 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.5 | 92.5 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2147 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.8 | 0.9 | 304.7 | 92.4 | 92.4189 | 0.1 | 13.1 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1394 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.4 | 92.4 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2935 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.4 | 92.4 | 0.1 | 13 | 950 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2261 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.99 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.4 | 92.4 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.149 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.997 | 0.02 | 200 | 304 | 0.9 | 304.9 | 92.3 | 92.3334 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 6 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 6 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 6 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.4 | 92.3689 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 6 | 0.02 | 200 | 304 | 0.9 | 304.9 | 92.3 | 92.3 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 6 | 0.02 | 200 | 303.9 | 0.9 | 304.8 | 92.3 | 92.3 | 0.1 | 13 | 950 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1271 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.628 | 0.02 | 199.7099 | 304 | 0.9 | 304.9 | 92 | 91.9797 | 0.1 | 13 | 950 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0539 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2813 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.43 | 0.0184 | 183.5939 | 303.3 | 0.9 | 304.2 | 90.3 | 90.2594 | 0.1 | 13.1 | 949 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.053847 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.43 | 0.018 | 180 | 280.7 | 0.9 | 281.6 | 86.3 | 86.2542 | 0.1 | 13 | 949 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.05039 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.43 | 0.018 | 180 | 277.4 | 0.9 | 278.3 | 84.1 | 84.0578 | 0.1 | 13 | 949 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.049798 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2562 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.43 | 0.018 | 180 | 276 | 0.9 | 276.9 | 83.3 | 83.3326 | 0.1 | 13 | 950 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0496 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1019 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.43 | 0.018 | 180 | 275.6 | 0.9 | 276.5 | 82.9 | 82.931 | 0.1 | 13 | 950 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.049598 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:05 | 5.423 | 0.018 | 180 | 275.2 | 0.9 | 276.1 | 82.8 | 82.8 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.42 | 0.018 | 180 | 275.1 | 0.9 | 276 | 82.7 | 82.7356 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1607 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.42 | 0.018 | 180 | 275.1 | 0.9 | 276 | 82.8 | 82.7748 | 0.1 | 13 | 951 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2852 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.415 | 0.018 | 180 | 275 | 1 | 276 | 82.8 | 82.8 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.131 | 27.6 | 1
2008-10-20 | 17:06 | 5.39 | 0.018 | 180 | 275 | 0.9 | 275.9 | 82.7 | 82.7 | 0.1 | 13 | 951 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2232 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.0178 | 178.4974 | 274.9 | 0.9 | 275.8 | 82.7 | 82.7 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 28 | 5 | 195 | 59.2 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2226 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170.1287 | 274.8 | 0.9 | 275.7 | 82.7 | 82.7 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0495 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1316 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.0178 | 178.1738 | 274.8 | 0.9 | 275.7 | 82.7 | 82.7 | 0.1 | 13 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.297 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2858 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.018 | 180 | 274.7 | 0.9 | 275.6 | 82.8 | 82.7817 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.0175 | 174.8799 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.347 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3518 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1988 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1218 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.385 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.9 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1646 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.9 | 82.9 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2814 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.9 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.978 | 10.1273 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.9 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1634 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1594 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.9 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.8 | 82.8187 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.381 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8321 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2444 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2013 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.382 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1264 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.8 | 82.8 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1965 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 0.9 | 275.4 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1194 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1577 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2348 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2763 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2028 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1221 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:06 | 5.385 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.7 | 82.719 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1864 | 27.6 | 1
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.5 | 1 | 275.5 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 28 | 5 | 195 | 59.2978 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2797 | 27.6 | 1
2008-10-20 | 17:06 | 5.389 | 0.017 | 170 | 274.6 | 1 | 275.6 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2026 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 1 | 275.6 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2745 | 27.6 | 1
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.387 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.7617 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2426 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.8 | 82.8 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2972 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.765 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.7 | 82.7 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.7 | 82.7314 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2032 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.8 | 82.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2076 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.39 | 0.017 | 170 | 274.6 | 0.9 | 275.5 | 82.9 | 82.8608 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.387 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.9 | 82.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2841 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 82.9 | 82.9163 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1299 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.38 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 83 | 83 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 5.164 | 0.017 | 170 | 274.7 | 0.9 | 275.6 | 83 | 83 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0494 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.889 | 0.017 | 170 | 274 | 0.9 | 274.9 | 82.6 | 82.6421 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0493 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.0178 | 178.025 | 260.4 | 0.9 | 261.3 | 79.5 | 79.5016 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.3 | 0.048354 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.018 | 180 | 250.7 | 0.9 | 251.6 | 76.1 | 76.0813 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.04569 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2407 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.0175 | 175.0464 | 247.2 | 0.9 | 248.1 | 74.7 | 74.6929 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.045098 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2052 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.017 | 170 | 246.3 | 0.9 | 247.2 | 74.1 | 74.1046 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.017 | 170 | 245.9 | 0.9 | 246.8 | 73.9 | 73.8983 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:06 | 4.77 | 0.017 | 170 | 245.6 | 0.9 | 246.5 | 73.7 | 73.7375 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2972 | 0.044897 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9721 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.017 | 170 | 245.5 | 0.9 | 246.4 | 73.7 | 73.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2974 | 0.0448 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.0164 | 163.5149 | 245.1 | 0.9 | 246 | 73.6 | 73.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1652 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 245.1 | 0.9 | 246 | 73.6 | 73.568 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0448 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 245 | 0.9 | 245.9 | 73.5 | 73.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0448 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.9 | 0.9 | 245.8 | 73.5 | 73.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.044798 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2356 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.5 | 73.4705 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2414 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.263 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1088 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.8 | 0.9 | 245.7 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.7 | 0.9 | 245.6 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1706 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.7 | 0.9 | 245.6 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2876 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.6 | 0.9 | 245.5 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0447 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.2106 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 244.2 | 0.9 | 245.1 | 73.4 | 73.3626 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1335 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.9 | 0.9 | 244.8 | 73.4 | 73.3604 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.9 | 0.9 | 244.8 | 73.4 | 73.4 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.9 | 0.9 | 244.8 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2972 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0278 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.8 | 0.9 | 244.7 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2974 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.7 | 0.9 | 244.6 | 73.3 | 73.3182 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.6 | 0.9 | 244.5 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0446 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.765 | 0.016 | 160 | 243.6 | 0.9 | 244.5 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0446 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.016 | 160 | 243.5 | 0.9 | 244.4 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0446 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005.0241 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.0155 | 155.038 | 243.4 | 0.9 | 244.3 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0446 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1006 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.015 | 150 | 243.5 | 0.9 | 244.4 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0446 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1749 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.0151 | 151.4204 | 243.5 | 0.9 | 244.4 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2709 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.016 | 159.6305 | 243.5 | 0.9 | 244.4 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.203 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1167 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.76 | 0.016 | 160 | 243.5 | 0.9 | 244.4 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2375 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.763 | 0.016 | 160 | 243.4 | 0.9 | 244.3 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2082 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.3 | 0.8976 | 244.2 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.123 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.3 | 0.8 | 244.1 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1999 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1228 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1628 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2915 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.771 | 0.016 | 160 | 243.2 | 0.8022 | 244 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2145 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.78 | 0.016 | 160 | 243.2 | 0.8976 | 244.1 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2626 | 27.6 | 0.9
2008-10-20 | 17:07 | 4.78 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2208 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.78 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.78 | 0.016 | 160 | 243 | 0.8 | 243.8 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.4 | 28 | 5 | 195 | 59.2026 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1006 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.78 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.779 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.4 | 73.4 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1037 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3378 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2586 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.4 | 73.3643 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1871 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.1 | 0.8 | 243.9 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2028 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.77 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.767 | 0.016 | 160 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.69 | 0.0155 | 155.1846 | 243.2 | 0.8 | 244 | 73.3 | 73.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0445 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.224 | 0.0153 | 153.0388 | 238.3 | 0.8 | 239.1 | 73 | 72.9909 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.044452 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.176 | 0.016 | 160 | 225.7 | 0.8 | 226.5 | 69.7 | 69.7335 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.041681 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1549 | 27.6 | 0.8
2008-10-20 | 17:07 | 4.17 | 0.016 | 160 | 219.5 | 0.8 | 220.3 | 66.3 | 66.3251 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0407 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2909 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.161 | 0.0144 | 144.2298 | 218.3 | 0.8 | 219.1 | 64.9 | 64.851 | 0.1 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.040493 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1368 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 216 | 0.8 | 216.8 | 64.3 | 64.2789 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0401 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2173 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 215.7 | 0.8 | 216.5 | 64.1 | 64.1306 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2972 | 0.04 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2285 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.9 | 0.8 | 215.7 | 64.1 | 64.1 | 0.1 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.2974 | 0.0399 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.9 | 0.8 | 215.7 | 64 | 64.0187 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0399 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.5 | 0.8 | 215.3 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.4 | 0.8 | 215.2 | 64 | 64 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.4 | 0.8 | 215.2 | 64 | 64.0014 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.3 | 0.8 | 215.1 | 64.1 | 64.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.2 | 0.8 | 215 | 64.1 | 64.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.2 | 0.8 | 215 | 64.2 | 64.2 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.1 | 0.8 | 214.9 | 64.1 | 64.1375 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214.1 | 0.8 | 214.9 | 64.2 | 64.1818 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2028 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 214 | 0.8 | 214.8 | 64.1 | 64.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 213.9 | 0.8 | 214.7 | 64.1 | 64.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0398 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.0391 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 213.9 | 0.9 | 214.8 | 64.1 | 64.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.076 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64.1 | 64.0689 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2151 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64 | 64 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2921 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.15 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64 | 64 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1615 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.153 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64 | 64 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1927 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.8 | 0.8 | 214.6 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.4 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0009 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.253 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.8 | 0.8024 | 214.6 | 64 | 64 | 0.1 | 13.1 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.9 | 0.9 | 214.8 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2096 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.8 | 0.9 | 214.7 | 64 | 64 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2364 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.7 | 0.9 | 214.6 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1177 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.0135 | 135.2315 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2718 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.6 | 0.9 | 214.5 | 64.1 | 64.0644 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.6 | 0.9 | 214.5 | 64.1 | 64.0708 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.18 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.6 | 0.9 | 214.5 | 64.1 | 64.0953 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.017 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.0138 | 137.8268 | 213.6 | 0.9 | 214.5 | 64 | 64.0406 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.0941 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2422 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13.1 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2032 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.9 | 214.5 | 64 | 64 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.6 | 0.8024 | 214.4 | 64 | 64 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.5 | 0.9 | 214.4 | 64 | 64 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.4 | 0.9 | 214.3 | 64 | 63.97 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.4 | 0.9 | 214.3 | 63.9 | 63.9 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.4 | 0.9 | 214.3 | 63.9 | 63.9 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0397 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.4 | 0.9 | 214.3 | 63.9 | 63.9 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.4 | 0.9 | 214.3 | 63.9 | 63.9 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.139 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.3 | 0.9 | 214.2 | 63.9 | 63.9 | 0.1 | 13.1 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2934 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.3 | 0.9 | 214.2 | 64 | 63.9815 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1524 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.014 | 140 | 213.3 | 0.9 | 214.2 | 64 | 64 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.0139 | 138.8676 | 213.3 | 0.9 | 214.2 | 64 | 64 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192.4 | 28 | 5 | 195 | 59.2024 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1558 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.0131 | 130.5412 | 213.2 | 0.9 | 214.1 | 64.1 | 64.1 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192.6 | 28 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.29 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.3 | 0.9 | 214.2 | 64.1 | 64.0682 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1358 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.3 | 0.9 | 214.2 | 64 | 64 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.2184 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.2 | 0.9 | 214.1 | 64 | 64 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:08 | 4.16 | 0.013 | 130 | 213.2 | 0.9 | 214.1 | 64 | 64 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0396 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.778 | 0.0131 | 131.1167 | 213 | 0.9 | 213.9 | 63.9 | 63.931 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.03959 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.588 | 0.0139 | 139.45 | 204.2 | 0.9 | 205.1 | 62.6 | 62.6051 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0381 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.0124 | 124.4592 | 199.2 | 0.9 | 200.1 | 58.8 | 58.7948 | 0.1 | 13.1 | 952 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2972 | 0.03735 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2106 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.012 | 120 | 185.6 | 0.9 | 186.5 | 56.2 | 56.1796 | 0.1 | 13 | 952 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034995 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.012 | 120 | 184.2 | 0.9 | 185.1 | 55.1 | 55.1452 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.012 | 120 | 184.1 | 0.9 | 185 | 54.7 | 54.7261 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0348 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.012 | 120 | 184 | 0.9 | 184.9 | 54.7 | 54.6692 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2022 | 0.0347 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.0469 | 27.6 | 0.9
2008-10-20 | 17:08 | 3.54 | 0.012 | 120 | 183.9 | 0.8 | 184.7 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0347 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.0301 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 183.9 | 0.8 | 184.7 | 54.6 | 54.6 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0347 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1145 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 183.5 | 0.8 | 184.3 | 54.6 | 54.6 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2022 | 0.034698 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2687 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 183 | 0.8 | 183.8 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0346 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1771 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.9 | 0.8 | 183.7 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0346 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1386 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.8 | 0.8 | 183.6 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1844 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.7 | 0.8 | 183.5 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1073 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.7 | 0.8 | 183.5 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.6 | 0.8 | 183.4 | 54.6 | 54.6 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1939 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.6 | 0.8 | 183.4 | 54.6 | 54.6 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.5 | 0.8 | 183.3 | 54.6 | 54.6 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2978 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.4 | 0.8 | 183.2 | 54.5 | 54.5406 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1438 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.4 | 0.8 | 183.2 | 54.6 | 54.5703 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0345 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2103 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182.3 | 0.8 | 183.1 | 54.5 | 54.5 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034498 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2356 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 182 | 0.8 | 182.8 | 54.6 | 54.5797 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.9 | 0.8 | 182.7 | 54.5 | 54.5283 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0298 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.9 | 0.8 | 182.7 | 54.6 | 54.563 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2266 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.6 | 54.6 | 0.1 | 13.1 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.9 | 0.8 | 182.7 | 54.5 | 54.5266 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.5 | 54.5 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.5 | 54.5 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2279 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.6 | 54.5704 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1508 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.5 | 54.5 | 0.1 | 13 | 953 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1526 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.8 | 0.8 | 182.6 | 54.6 | 54.6 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.7 | 54.6634 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.7 | 54.7 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1848 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.8 | 54.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2022 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1347 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.9 | 54.8804 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.1883 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.9 | 54.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2974 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1112 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.9 | 54.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2022 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 54.9 | 54.9319 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 28 | 5 | 195 | 59.2 | 0.0344 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 55 | 55 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.034398 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 55.1 | 55.0738 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 55.1 | 55.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 55.1 | 55.1 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2022 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.532 | 0.012 | 120 | 181.7 | 0.8 | 182.5 | 55.1 | 55.1 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.53 | 0.012 | 120 | 181.6 | 0.8 | 182.4 | 55 | 55.0182 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.53 | 0.012 | 120 | 181.6 | 0.8 | 182.4 | 54.9 | 54.9183 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.533 | 0.012 | 120 | 181.6 | 0.8 | 182.4 | 54.9 | 54.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2197 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.6 | 0.8 | 182.4 | 54.8 | 54.8178 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2262 | 27.6 | 0.8
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.6 | 0.9 | 182.5 | 54.8 | 54.8 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.5 | 0.9 | 182.4 | 54.7 | 54.7379 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.2 | 0.9 | 182.1 | 54.7 | 54.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 28 | 5 | 195 | 59.2 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.2 | 0.9 | 182.1 | 54.7 | 54.7 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.537 | 0.012 | 120 | 181.2 | 0.9 | 182.1 | 54.6 | 54.6 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.53 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.53 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.538 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1537 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2974 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1693 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0343 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1157 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2699 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181 | 0.9 | 181.9 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0342 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1759 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.54 | 0.012 | 120 | 181 | 0.9 | 181.9 | 54.6 | 54.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0342 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.1 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.319 | 0.012 | 120 | 181.1 | 0.9 | 182 | 54.6 | 54.6 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2978 | 0.0342 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1325 | 27.6 | 0.9
2008-10-20 | 17:09 | 3.072 | 0.012 | 120 | 180.6 | 0.9 | 181.5 | 54.4 | 54.3701 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2978 | 0.034171 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2865 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.0112 | 112.1048 | 167.7 | 0.9 | 168.6 | 50.6 | 50.5594 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.031966 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.011 | 110 | 156.2 | 0.9 | 157.1 | 47.5 | 47.4562 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029986 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2054 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.0106 | 105.7333 | 152.4 | 0.9 | 153.3 | 46.1 | 46.0949 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.203 | 0.0293 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1243 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.01 | 100 | 151.5 | 0.9 | 152.4 | 45.6 | 45.6396 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0292 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2014 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.01 | 100 | 151.4 | 0.9 | 152.3 | 45.5 | 45.5184 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0292 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.2784 | 27.6 | 0.9
2008-10-20 | 17:09 | 2.95 | 0.01 | 100 | 151.4 | 0.9 | 152.3 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0292 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2445 | 27.6 | 0.9
2008-10-20 | 17:10 | 2.946 | 0.01 | 100 | 151.4 | 0.9024 | 152.3 | 45.4 | 45.4362 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1675 | 27.6 | 0.9
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1192 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2743 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1714 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.3682 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.4 | 1 | 152.4 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2974 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1003 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1003.024 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.4 | 45.4 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1248 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.1 | 1 | 152.1 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2789 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.4819 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2126 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2972 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.175 | 27.6 | 1
2008-10-20 | 17:10 | 2.94 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2521 | 27.6 | 1
2008-10-20 | 17:10 | 2.932 | 0.01 | 100 | 151.2 | 1 | 152.2 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2708 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2 | 27.6 | 1
2008-10-20 | 17:10 | 2.917 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.5 | 45.4687 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2 | 27.6 | 1
2008-10-20 | 17:10 | 2.9 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | 0 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1395 | 27.6 | 1
2008-10-20 | 17:10 | 2.901 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.4 | 45.4017 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | -2e-06 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.909 | 0.01 | 100 | 151.4 | 1 | 152.4 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1293 | 27.6 | 1
2008-10-20 | 17:10 | 2.918 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2835 | 27.6 | 1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1623 | 27.6 | 1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.3 | 1 | 152.3 | 45.5 | 45.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2978 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.3 | 1.1 | 152.4 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.4 | 1.1 | 152.5 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.4 | 1.1 | 152.5 | 45.5 | 45.5 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.5 | 1.1 | 152.6 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.5 | 1.1 | 152.6 | 45.4 | 45.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.5 | 1.1 | 152.6 | 45.4 | 45.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.5 | 1.1 | 152.6 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.5 | 1.1 | 152.6 | 45.4 | 45.4 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1625 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.92 | 0.01 | 100 | 151.7 | 1.1 | 152.8 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2972 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2395 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.928 | 0.01 | 100 | 151.7 | 1.1 | 152.8 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0259 | 10.2669 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1.1 | 152.8 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.000798 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1128 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1.0976 | 152.8 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1954 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.01 | 100 | 151.7 | 1 | 152.7 | 45.3 | 45.2635 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.203 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2326 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.0106 | 106.25 | 151.7 | 1 | 152.7 | 45.2 | 45.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2972 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1108 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.011 | 110 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.1879 | 27.6 | 1
2008-10-20 | 17:10 | 2.93 | 0.011 | 110 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.265 | 27.6 | 1
2008-10-20 | 17:10 | 2.597 | 0.0109 | 109.0985 | 151.7 | 1 | 152.7 | 45.3 | 45.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0291 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.389 | 0.0101 | 100.7513 | 145 | 1 | 146 | 44.7 | 44.6856 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.027976 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.0092 | 92.4356 | 139.6 | 1 | 140.6 | 41 | 40.9944 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.026969 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.009 | 90 | 126.6 | 1 | 127.6 | 38.1 | 38.0856 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.024692 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.009 | 90 | 125 | 1 | 126 | 36.7 | 36.7091 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0244 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.009 | 90 | 124.1 | 1 | 125.1 | 36.4 | 36.3759 | 0.1 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1039.7747 | 1005 | 10.2724 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.009 | 90 | 123.1 | 1 | 124.1 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.4 | 28 | 5 | 195 | 59.2024 | 0.024 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1953 | 27.6 | 1
2008-10-20 | 17:10 | 2.32 | 0.009 | 90 | 122.7 | 1 | 123.7 | 36.3 | 36.3 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2022 | 0.024 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1182 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.009 | 90 | 122.7 | 1 | 123.7 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.023997 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1589 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.009 | 90 | 122.7 | 1 | 123.7 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.164 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.0086 | 85.5498 | 122.6 | 1 | 123.6 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1262 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.7 | 1 | 123.7 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2803 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.6 | 0.9976 | 123.6 | 36.3 | 36.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.5 | 0.9 | 123.4 | 36.2 | 36.2388 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.5 | 0.9 | 123.4 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2785 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.4 | 0.9 | 123.3 | 36.2 | 36.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2014 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.3 | 0.9 | 123.2 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0239 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2757 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.2 | 0.9 | 123.1 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.023898 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.2 | 0.9 | 123.1 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2403 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.1 | 0.9 | 123 | 36.2 | 36.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1139 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122.1 | 0.9 | 123 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2681 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 0.9 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 0.9 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2618 | 27.6 | 0.9
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1 | 123 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1847 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1 | 123 | 36.2 | 36.2 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1076 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1 | 123 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.4 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1933 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1.0024 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2525 | 27.6 | 1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.8 | 1.1 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1008 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.8 | 1.1 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2028 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1778 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.8 | 1.1 | 122.9 | 36.1 | 36.1385 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2026 | 0.0238 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2548 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.8 | 1.1 | 122.9 | 36.1 | 36.1 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.023798 | 0.000698 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2365 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1.1 | 123 | 36.1 | 36.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1088 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1.1 | 123 | 36.1 | 36.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1858 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1.1 | 123 | 36 | 36.0039 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2629 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 121.9 | 1.1 | 123 | 36 | 36 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2202 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 36 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1339 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 36 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.288 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 36 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1567 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 36 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.319 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 36 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1521 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.311 | 0.008 | 80 | 122 | 1.1 | 123.1 | 35.9 | 35.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 122 | 1.1 | 123.1 | 35.9 | 35.9 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.316 | 0.008 | 80 | 122 | 1.1 | 123.1 | 35.9 | 35.9 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.4 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0241 | 10.2427 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 35.9612 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2344 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.32 | 0.008 | 80 | 122 | 1.1 | 123.1 | 36 | 35.9571 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2885 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.319 | 0.008 | 80 | 122 | 1.1 | 123.1 | 35.9 | 35.9 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2114 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.311 | 0.008 | 80 | 122 | 1.1 | 123.1 | 35.9 | 35.9 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1.1 | 123 | 35.9 | 35.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1572 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1.1 | 123 | 35.9 | 35.9 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.314 | 0.008 | 80 | 121.9 | 1.0972 | 123 | 35.9 | 35.9 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:11 | 2.317 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36 | 35.9636 | 0.1 | 13.1 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 122 | 1 | 123 | 36 | 36 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2024 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 122 | 1 | 123 | 36 | 36 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2434 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36 | 36 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36 | 36.0317 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2978 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36.1 | 36.1 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36.2 | 36.1598 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36.2 | 36.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1138 | 27.6 | 1
2008-10-20 | 17:11 | 2.31 | 0.008 | 80 | 121.9 | 1 | 122.9 | 36.2 | 36.2297 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1909 | 27.6 | 1
2008-10-20 | 17:11 | 2.012 | 0.0078 | 77.6376 | 121.9 | 1 | 122.9 | 36.3 | 36.3 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2972 | 0.0237 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.132 | 27.6 | 1
2008-10-20 | 17:11 | 1.816 | 0.0052 | 52.2354 | 113.1 | 1 | 114.1 | 35.7 | 35.6717 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2974 | 0.022079 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 1.75 | 0.005 | 50 | 108.6 | 1 | 109.6 | 32 | 31.9909 | 0.1 | 13.1 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.021269 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1444 | 27.6 | 1
2008-10-20 | 17:11 | 1.747 | 0.005 | 50 | 96.8 | 1 | 97.8 | 29.1 | 29.1455 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.019093 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2986 | 27.6 | 1
2008-10-20 | 17:11 | 1.74 | 0.005 | 50 | 94 | 1 | 95 | 27.9 | 27.9101 | 0.1 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0186 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1472 | 27.6 | 1
2008-10-20 | 17:11 | 1.74 | 0.0053 | 52.6131 | 93.2 | 1 | 94.2 | 27.5 | 27.5418 | 0.1 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.018398 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:11 | 1.733 | 0.006 | 60 | 92.3 | 1 | 93.3 | 27.4 | 27.4 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0183 | 0.000598 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1612 | 27.6 | 1
2008-10-20 | 17:11 | 1.73 | 0.006 | 60 | 92.2 | 1 | 93.2 | 27.3 | 27.3 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.018298 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2846 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.9 | 1.1 | 93 | 27.3 | 27.3 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1305 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.9 | 1.1 | 93 | 27.3 | 27.3 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.9 | 1.1 | 93 | 27.2 | 27.2 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1779 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.8 | 1.1 | 92.9 | 27.2 | 27.17 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2679 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.8 | 1.1 | 92.9 | 27.1 | 27.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2974 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1137 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.7 | 1.1 | 92.8 | 27.1 | 27.1 | 0.1 | 13.1 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.0299 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.7 | 1.1 | 92.8 | 27.1 | 27.1 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.0471 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.7 | 1.1 | 92.8 | 27.1 | 27.097 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0182 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.124 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.7 | 1.0976 | 92.8 | 27 | 27 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.202 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 91.7 | 1 | 92.7 | 27 | 27 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3561 | 27.6 | 1
2008-10-20 | 17:12 | 1.723 | 0.006 | 60 | 91.8 | 1 | 92.8 | 27 | 27 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3449 | 27.6 | 1
2008-10-20 | 17:12 | 1.726 | 0.006 | 60 | 91.8 | 1 | 92.8 | 27 | 27 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2679 | 27.6 | 1
2008-10-20 | 17:12 | 1.721 | 0.006 | 60 | 91.8 | 1 | 92.8 | 26.9 | 26.9199 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.203 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 91.8 | 1.0028 | 92.8 | 27 | 26.972 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2972 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 91.8 | 1.1 | 92.9 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2974 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1362 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 91.9 | 1.1 | 93 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1412 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 91.9 | 1.1 | 93 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2183 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.018198 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2955 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2503 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2268 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3039 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.381 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.8 | 26.8377 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3419 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.8 | 26.8 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 28 | 5 | 195 | 59.2972 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2297 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.713 | 0.006 | 60 | 92 | 1.1 | 93.1 | 26.8 | 26.8 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1245 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.72 | 0.006 | 60 | 92.1 | 1.1 | 93.2 | 26.8 | 26.8 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2787 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.72 | 0.006 | 60 | 92.2 | 1.1 | 93.3 | 26.8 | 26.8 | 0.1 | 13 | 954 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.72 | 0.006 | 60 | 92.2 | 1.1 | 93.3 | 26.8 | 26.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.714 | 0.006 | 60 | 92.3 | 1.1 | 93.4 | 26.8 | 26.8 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.714 | 0.006 | 60 | 92.3 | 1.1 | 93.4 | 26.8 | 26.8 | 0.1 | 13 | 954 | 952 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.718 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3749 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.71 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3709 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.725 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2877 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2026 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1335 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.4 | 1.1 | 93.5 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.281 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.5 | 1.1 | 93.6 | 26.9 | 26.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3626 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.5 | 1.1 | 93.6 | 26.9 | 26.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 28 | 5 | 195 | 59.2978 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3145 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.5 | 1.1 | 93.6 | 26.9 | 26.9 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2024 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3916 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1.1 | 93.7 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1039 | 1004 | 10.4 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1.1 | 93.7 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3541 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1.1 | 93.7 | 26.9 | 26.9 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1.0976 | 93.7 | 27 | 26.9819 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3001 | 27.6 | 1.1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2972 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9721 | 10.3772 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3458 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.6 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 1
2008-10-20 | 17:12 | 1.724 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 4.6 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:12 | 1.724 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27 | 4.4 | 195 | 59.2024 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 27.4 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 27.6 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3394 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3836 | 27.6 | 1
2008-10-20 | 17:12 | 1.73 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 27.4 | 5 | 195 | 59.3 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.306 | 27.6 | 1
2008-10-20 | 17:12 | 1.491 | 0.006 | 60 | 92.6 | 1 | 93.6 | 27 | 27 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0181 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:12 | 1.251 | 0.006 | 60 | 92.3 | 1 | 93.3 | 26.7 | 26.6771 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2022 | 0.018098 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2039 | 27.6 | 1
2008-10-20 | 17:12 | 1.16 | 0.0066 | 65.8499 | 77.3 | 1 | 78.3 | 23.7 | 23.7249 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015261 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 1
2008-10-20 | 17:12 | 1.159 | 0.0062 | 61.6425 | 64.7 | 1 | 65.7 | 20 | 20.0292 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.012797 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1043 | 27.6 | 1
2008-10-20 | 17:12 | 1.151 | 0.005 | 50 | 63.9 | 1 | 64.9 | 18.5 | 18.5489 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0127 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2588 | 27.6 | 1
2008-10-20 | 17:12 | 1.15 | 0.005 | 50 | 63.6 | 1 | 64.6 | 18 | 17.96 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.012695 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2435 | 27.6 | 1
2008-10-20 | 17:12 | 1.15 | 0.005 | 50 | 62.7 | 1 | 63.7 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.012498 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2336 | 27.6 | 1
2008-10-20 | 17:13 | 1.15 | 0.0043 | 42.6115 | 62.7 | 1 | 63.7 | 17.8 | 17.7634 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.6 | 1 | 63.6 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 1
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.6 | 1 | 63.6 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3649 | 27.6 | 1
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.6 | 0.9976 | 63.6 | 17.7 | 17.7 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3581 | 27.6 | 1
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.7 | 17.7397 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3191 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3962 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.8 | 17.7787 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3496 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3275 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.5 | 0.9 | 63.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3954 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.4 | 0.9 | 63.3 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3183 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.4 | 0.9 | 63.3 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3588 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.4 | 0.9 | 63.3 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3641 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.4 | 0.9 | 63.3 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.012398 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.313 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.4 | 0.9 | 63.3 | 17.6 | 17.619 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3901 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.3 | 0.9 | 63.2 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.6 | 28 | 5 | 195 | 59.2 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.4 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.3 | 0.9 | 63.2 | 17.6 | 17.6194 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2972 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3557 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.3 | 0.89741 | 63.2 | 17.7 | 17.6785 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.9
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.3 | 0.8 | 63.1 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.2 | 0.8 | 63 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.2 | 0.8 | 63 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 62.2 | 0.8 | 63 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2024 | 0.0123 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.9 | 0.8 | 62.7 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.7809 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27.6 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3382 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3847 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 27.4 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3076 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.4 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5084 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.7688 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4603 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7315 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3474 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.8 | 17.8 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4983 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3443 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3091 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3862 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2735 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.6 | 0.8 | 62.4 | 17.6 | 17.6186 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2806 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.004 | 40 | 61.7 | 0.8 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.0041 | 40.7398 | 61.7 | 0.8 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.0049 | 49.2432 | 61.7 | 0.7976 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.6 | 0.8
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.7 | 0.7 | 62.4 | 17.6 | 17.6 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0221 | 10.3511 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2969 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2198 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3146 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.15 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.6 | 17.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.154 | 0.005 | 50 | 61.8 | 0.7 | 62.5 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2026 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.6 | 0.7
2008-10-20 | 17:13 | 1.16 | 0.005 | 50 | 61.8 | 0.6 | 62.4 | 17.7 | 17.7 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.6 | 0.6
2008-10-20 | 17:13 | 1.148 | 0.005 | 50 | 61.9 | 0.6 | 62.5 | 17.7 | 17.7 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0122 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1039 | 1004.976 | 10.3344 | 27.6 | 0.6
2008-10-20 | 17:13 | 1.009 | 0.0045 | 45.4059 | 61.9 | 0.6 | 62.5 | 17.7 | 17.7 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | 0.0122 | 0.000198 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1039.665 | 1005 | 10.2573 | 27.6 | 0.6
2008-10-20 | 17:13 | 0.716 | 0.003 | 30 | 60.6 | 0.6 | 61.2 | 16.9 | 16.9376 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.011964 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1801 | 27.6 | 0.6
2008-10-20 | 17:13 | 0.61 | 0.003 | 30 | 41.5 | 0.6 | 42.1 | 14.5 | 14.5245 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.008288 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.103 | 27.6 | 0.6
2008-10-20 | 17:13 | 0.598 | 0.0015 | 15.3683 | 39.1 | 0.6 | 39.7 | 10.8 | 10.7863 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.007781 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1741 | 27.6 | 0.6
2008-10-20 | 17:13 | 0.58 | 0.001 | 10 | 34.1 | 0.6 | 34.7 | 9.2 | 9.1522 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.006798 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2512 | 27.6 | 0.6
2008-10-20 | 17:13 | 0.58 | 0.001 | 10 | 33.2 | 0.5 | 33.7 | 8.7 | 8.6898 | 0.1 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0067 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3282 | 27.6 | 0.5
2008-10-20 | 17:13 | 0.58 | 0.001 | 10 | 33.1 | 0.5 | 33.6 | 8.5 | 8.5188 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | 0.0067 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3947 | 27.6 | 0.5
2008-10-20 | 17:13 | 0.58 | 0.001 | 10 | 33 | 0.5 | 33.5 | 8.5 | 8.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2028 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3176 | 27.6 | 0.5
2008-10-20 | 17:13 | 0.58 | 0.001 | 10 | 33 | 0.5 | 33.5 | 8.6 | 8.5734 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.1026 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 33 | 0.5 | 33.5 | 8.6 | 8.6 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 33 | 0.5 | 33.5 | 8.6 | 8.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.4 | 28 | 5 | 195 | 59.2 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 32.9 | 0.5 | 33.4 | 8.6 | 8.6 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.0013 | 12.5103 | 32.9 | 0.5 | 33.4 | 8.6 | 8.6 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0066 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1641 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 32.6 | 0.5 | 33.1 | 8.6 | 8.5987 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 32.5 | 0.5 | 33 | 8.5 | 8.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1443 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 32.5 | 0.5 | 33 | 8.5 | 8.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2985 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 32.5 | 0.5 | 33 | 8.5 | 8.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.0024 | 23.8301 | 32.4 | 0.5 | 32.9 | 8.5 | 8.5 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1931 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 32.4 | 0.5 | 32.9 | 8.4 | 8.4194 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2028 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1611 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 32.3 | 0.5 | 32.8 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 32.3 | 0.5 | 32.8 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 32.2 | 0.5 | 32.7 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0065 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 32.2 | 0.5 | 32.7 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.006498 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.2611 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 31.9 | 0.5 | 32.4 | 8.4 | 8.3694 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.184 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 31.8 | 0.5 | 32.3 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1069 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 31.7 | 0.5 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2405 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.003 | 30 | 31.7 | 0.5 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2978 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.022 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.0023 | 22.7492 | 31.7 | 0.5 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2513 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 988 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2972 | 0.0064 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1028 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.7 | 0.5 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2974 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2569 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.7 | 0.5 | 32.2 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4976 | 32.1 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.4 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1721 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1823 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.575 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.3 | 8.3319 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2028 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.155 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1992 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.4024 | 32 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.4 | 8.4 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.57 | 0.0018 | 17.8101 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2022 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.57 | 0.001 | 10 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9759 | 10.358 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.575 | 0.001 | 10 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3649 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 31.6 | 0.4 | 32 | 8.3 | 8.3 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.001 | 10 | 31.6 | 0.4 | 32 | 8.2 | 8.2181 | 0.1 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.0012 | 12.1765 | 31.6 | 0.4 | 32 | 8.2 | 8.2 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.4 | 28 | 5 | 195 | 59.2 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2131 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.58 | 0.002 | 20 | 31.6 | 0.4 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.1412 | 27.6 | 0.4
2008-10-20 | 17:14 | 0.574 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9741 | 10.2954 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13.1 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2878 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.6 | 0.5 | 32.1 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1336 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.5 | 0.5 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0063 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2205 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.5 | 0.5 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 9.8e-05 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.57 | 0.002 | 20 | 31.5 | 0.5 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.022 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.545 | 0.002 | 20 | 31.5 | 0.5 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2028 | 0.006297 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.9722 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.257 | 0.0011 | 11.4089 | 31.5 | 0.5 | 32 | 8.2 | 8.2 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0062 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1627 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.068 | -0.0019 | -19.2681 | 20.4 | 0.5 | 20.9 | 7 | 6.9636 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.004069 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1916 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.01 | -0.0011 | -10.7574 | 6.1 | 0.5 | 6.6 | 3.2 | 3.2203 | 0.1 | 13 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.001293 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.01 | -0.001 | -10 | 4.8 | 0.5 | 5.3 | 0.4 | 0.42 | 0.1 | 13.1 | 953 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.000995 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.01 | -0.001 | -10 | 2.5 | 0.5 | 3 | -0.7 | 0 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.000498 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.01 | -0.001 | -10 | 1.5 | 0.5 | 2 | -1 | 0 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1916 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.01 | -0.001 | -10 | 1.4 | 0.5 | 1.9 | -1 | 0 | 0.1 | 13 | 954 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.5
2008-10-20 | 17:14 | 0.003 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1.1 | 0 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.1168 | 27.6 | 0.5
2008-10-20 | 17:14 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.271 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.175 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2972 | 0.0003 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.179 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.5 | 1.8 | -1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1587 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.50222 | 1.8 | -1 | 0 | 0.1 | 13.1 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1954 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.6 | 1.9 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.3 | 0.6 | 1.9 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2964 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.2 | 0.6 | 1.8 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2978 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1414 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.2 | 0.6 | 1.8 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.2 | 0.6 | 1.8 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1336 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.2 | 0.6 | 1.8 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1893 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.2 | 0.6 | 1.8 | -1.2 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1122 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 1.1 | 0.6 | 1.7 | -1.2 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.0006 | -6.1333 | 1.1 | 0.6 | 1.7 | -1.2 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.184 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | 0 | 0 | 1.1 | 0.6 | 1.7 | -1.2 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | 0 | 0 | 1 | 0.6 | 1.6 | -1.2 | 0 | 0.1 | 13.1 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | 0 | 0 | 1 | 0.6 | 1.6 | -1.2 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | 0 | 0 | 1 | 0.6 | 1.6 | -1.2 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3504 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | 0 | 0 | 1 | 0.6 | 1.6 | -1.2 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0002 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3725 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.0004 | -3.85 | 1 | 0.6 | 1.6 | -1.2 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.000198 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3046 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.9 | 0.6 | 1.5 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3817 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.9 | 0.6 | 1.5 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3412 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.9 | 0.6 | 1.5 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2026 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2281 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.6 | 1.4 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2976 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.6 | 1.4 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.6 | 1.4 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2024 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2344 | 27.6 | 0.6
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.5 | 1.3 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2113 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.5 | 1.3 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1214 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.5 | 1.3 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.6 | 28 | 5 | 195 | 59.3 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1985 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.8 | 0.5 | 1.3 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1244 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.7 | 0.5 | 1.2 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.7 | 0.5 | 1.2 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1298 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2026 | 9.7e-05 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2069 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.284 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.6 | 0.4976 | 1.1 | -1.1 | 0 | 0.1 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2389 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.7 | 0.4 | 1.1 | -1 | 0 | 0.1 | 13 | 955 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2382 | 27.6 | 0.4
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.7 | 0.4 | 1.1 | -1.1 | 0 | 0.1 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2976 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0241 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:15 | -0.001 | -0.001 | -10 | 0.7 | 0.4 | 1.1 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:15 | -0.009 | -0.001 | -10 | 0.7 | 0.4 | 1.1 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:15 | -0.003 | -0.001 | -10 | 0.6 | 0.4 | 1 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2067 | 27.6 | 0.4
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.6 | 0.4 | 1 | -1.1 | 0 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1475 | 27.6 | 0.4
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.6 | 0.4 | 1 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2983 | 27.6 | 0.4
2008-10-20 | 17:15 | -0.002 | -0.001 | -10 | 0.6 | 0.40259 | 1 | -1.1 | 0 | 0.1 | 13 | 956 | 952 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.6 | 28 | 5 | 195 | 59.2026 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1443 | 27.6 | 0.4
2008-10-20 | 17:15 | -0.01 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.1 | 0 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.5
2008-10-20 | 17:15 | -0.01 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.024 | 10.1634 | 27.6 | 0.5
2008-10-20 | 17:15 | -0.01 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.2 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:15 | -0.004 | -0.001 | -10 | 0.6 | 0.5 | 1.1 | -1.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.0003 | -2.7666 | 0.6 | 0.5 | 1.1 | -1.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2976 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1744 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.0005 | -5.4538 | 0.6 | 0.5 | 1.1 | -1.2 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.5 | 0.5 | 1 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1341 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.5 | 0.5 | 1 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2896 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | 0 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1561 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | -3e-06 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.0279 | 10.1981 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2477 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1032 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1803 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.4 | 0.5 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.6 | 28 | 5 | 195 | 59.3 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2 | 27.6 | 0.5
2008-10-20 | 17:15 | 0 | -0.001 | -10 | 0.5 | 0.5 | 1 | -1.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.976 | 10.2233 | 27.6 | 0.5
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3775 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2025 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.297 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2972 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2485 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.5 | 0.4 | 0.9 | -1.3 | 0 | 0.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.4 | 0.4 | 0.8 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2715 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.4 | 0.4 | 0.8 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2976 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1944 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.3 | 0.4 | 0.7 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.1173 | 27.6 | 0.4
2008-10-20 | 17:16 | 0 | -0.001 | -10 | 0.3 | 0.4 | 0.7 | -1.3 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2024 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2195 | 27.6 | 0.4
2008-10-20 | 17:16 | -0.005 | -0.001 | -10 | 0.3 | 0.4 | 0.7 | -1.3 | 0 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2263 | 27.6 | 0.4
2008-10-20 | 17:16 | -0.006 | -0.001 | -10 | 0.3 | 0.4 | 0.7 | -1.3 | 0 | 0.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2978 | -0.0001 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.114 | 27.6 | 0.4
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-10-20
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 302
NO2 Span(ppm) | 255
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.00
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 16:54
Test End Time | 17:16
Test Duration (s) | 1283
Average Ambient Temperature (deg C) | -35
Average Ambient Pressure (mbar) | 988.86
Average Relative Humidity (%) | 18.998
Average Absolute Humidity (grains/lb dry air) | 0.257
Average Kh Factor | 0.833
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 10/20/2008 16:54:20.812 - None Found
Warnings:
0X0000 - 10/20/2008 16:54:20.812 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 144
Brake Specific Dropout Limit Count | 1283
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-10-20 | 16:54 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
Cerion
light w/generator

## THC Lin
Test Date:
Test Time:
10/20/2008
16:54:21
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
10/20/2008
16:54:21
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
10/20/2008
16:54:21
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
10/20/2008
16:54:21
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
10/20/2008
16:54:21
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
10/20/2008
16:54:21
NDUV Serial No:
208
AMBII Serial No:
37249
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6 | Rm = | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277
| C.F.= | 0.099349 | 0.19884 | 0.29848 | 0.39826 | 0.49819 | 0.59826 | 0.69847 | 0.79884 | 0.89935
| gas = | 0.59609 | 1.193 | 1.7909 | 2.3896 | 2.9891 | 3.5895 | 4.1908 | 4.793 | 5.3961
Ro for CO2= | 0.8795