## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-710C | GDU-002 | 2008-01-01 | 2009-01-01
0 | 0 | 0 | 0
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 50568
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.02 | Scott gas | ALM014749 | ±1% | 2010-01-06 | 93.06
200 ppm CO | 200.7
6% CO2 | 5.99
300 ppm NO | 302.6
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300.5
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | CO2 - AMBII Serial Number: |  |  | 50568 |  |  | CO - AMBII Serial Number: |  |  | 50568 |  | HC - AMBII Serial Number: |  |  | 50568 |  | NO - NDUV Serial Number: |  |  | 208 |  | THC - FID Serial Number: |  |  | 123
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-05-07 |  | 12 | X | 0.1 | 3 |  |  | 1201 |  | 50 | 3 |  | 319 |  | 4 | 3 |  | 1495 |  | 15 | 3 |  | 957 | 1000 | 5 | 2
Test Time: | 9:23
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -1 | -1 |  |  | 0 | 996.4 | 996.4 |  |  | 0 | 0.9 | 0.9 |  |  | 0 | -0.4 | -0.4
90 | 105 | 1 | 1.2 | 1.184 | 1.157 | -0.027 | -2.28 |  | 120.1 | 130 | 10 | 8.33 |  | 31.9 | 997 | 965.1 | 3025.39 | Fail | 149.5 | 148.3 | -1.2 | -0.8 |  | 95.7 | 94.6 | -1.1 | -1.15
150 | 165 | 2 | 2.4 | 2.372 | 2.33 | -0.042 | -1.77 |  | 240.2 | 240 | 0 | 0 |  | 63.8 | 996.7 | 932.9 | 1462.23 | Fail | 299 | 296.2 | -2.8 | -0.94 |  | 191.4 | 189.5 | -1.9 | -0.99
210 | 225 | 3 | 3.6 | 3.563 | 3.541 | -0.022 | -0.62 |  | 360.3 | 356 | -4 | -1.11 |  | 95.7 | 996.5 | 900.8 | 941.27 | Fail | 448.5 | 445.9 | -2.6 | -0.58 |  | 287.1 | 285 | -2.1 | -0.73
270 | 285 | 4 | 4.8 | 4.758 | 4.74 | -0.018 | -0.38 |  | 480.4 | 483 | 3 | 0.62 |  | 127.6 | 997.6 | 870 | 681.82 | Fail | 598 | 597.4 | -0.6 | -0.1 |  | 382.8 | 379.7 | -3.1 | -0.81
330 | 345 | 5 | 6 | 5.956 | 5.957 | 0.001 | 0.02 |  | 600.5 | 599 | -2 | -0.33 |  | 159.5 | 996.6 | 837.1 | 524.83 | Fail | 747.5 | 745.1 | -2.4 | -0.32 |  | 478.5 | 474.5 | -4 | -0.84
390 | 405 | 6 | 7.2 | 7.158 | 7.183 | 0.025 | 0.35 |  | 720.6 | 714 | -7 | -0.97 |  | 191.4 | 997.1 | 805.7 | 420.95 | Fail | 897 | 896.6 | -0.4 | -0.04 |  | 574.2 | 570.4 | -3.8 | -0.66
450 | 465 | 7 | 8.4 | 8.363 | 8.398 | 0.035 | 0.42 |  | 840.7 | 829 | -12 | -1.43 |  | 223.3 | 996.2 | 772.9 | 346.13 | Fail | 1046.5 | 1044.7 | -1.8 | -0.17 |  | 669.9 | 666.5 | -3.4 | -0.51
510 | 525 | 8 | 9.6 | 9.572 | 9.605 | 0.033 | 0.34 |  | 960.8 | 950 | -11 | -1.14 |  | 255.2 | 996.6 | 741.4 | 290.52 | Fail | 1196 | 1194.5 | -1.5 | -0.13 |  | 765.6 | 763.7 | -1.9 | -0.25
570 | 585 | 9 | 10.8 | 10.784 | 10.799 | 0.015 | 0.14 |  | 1080.9 | 1074 | -7 | -0.65 |  | 287.1 | 996.8 | 709.7 | 247.2 | Fail | 1345.5 | 1349.6 | 4.1 | 0.3 |  | 861.3 | 860.3 | -1 | -0.12
630 | 645 | 10 | 12 | 12 | 11.98 | -0.02 | -0.17 |  | 1201 | 1189 | -12 | -1 |  | 319 | 997.3 | 678.3 | 212.63 | Fail | 1495 | 1499.8 | 4.8 | 0.32 |  | 957 | 957.3 | 0.3 | 0.03
690 | 705 | 9 | 10.8 | 10.784 | 10.812 | 0.028 | 0.26 |  | 1080.9 | 1065 | -16 | -1.48 |  | 287.1 | 997.8 | 710.7 | 247.54 | Fail | 1345.5 | 1348.5 | 3 | 0.22 |  | 861.3 | 859.4 | -1.9 | -0.22
750 | 765 | 8 | 9.6 | 9.572 | 9.619 | 0.047 | 0.49 |  | 960.8 | 949 | -12 | -1.25 |  | 255.2 | 997.1 | 741.9 | 290.71 | Fail | 1196 | 1200 | 4 | 0.33 |  | 765.6 | 762.5 | -3.1 | -0.4
810 | 825 | 7 | 8.4 | 8.363 | 8.408 | 0.045 | 0.54 |  | 840.7 | 831 | -10 | -1.19 |  | 223.3 | 997.1 | 773.8 | 346.53 | Fail | 1046.5 | 1046.4 | -0.1 | -0.01 |  | 669.9 | 665.7 | -4.2 | -0.63
870 | 885 | 6 | 7.2 | 7.158 | 7.176 | 0.018 | 0.25 |  | 720.6 | 712 | -9 | -1.25 |  | 191.4 | 996.4 | 805 | 420.59 | Fail | 897 | 898.8 | 1.8 | 0.2 |  | 574.2 | 570.1 | -4.1 | -0.71
930 | 945 | 5 | 6 | 5.956 | 5.957 | 0.001 | 0.02 |  | 600.5 | 593 | -8 | -1.33 |  | 159.5 | 997 | 837.5 | 525.08 | Fail | 747.5 | 748.7 | 1.2 | 0.16 |  | 478.5 | 473.9 | -4.6 | -0.96
990 | 1005 | 4 | 4.8 | 4.758 | 4.743 | -0.015 | -0.32 |  | 480.4 | 471 | -9 | -1.87 |  | 127.6 | 997 | 869.4 | 681.35 | Fail | 598 | 599 | 1 | 0.17 |  | 382.8 | 379.1 | -3.7 | -0.97
1050 | 1065 | 3 | 3.6 | 3.563 | 3.54 | -0.023 | -0.65 |  | 360.3 | 360 | 0 | 0 |  | 95.7 | 996.7 | 901 | 941.48 | Fail | 448.5 | 448.5 | 0 | 0 |  | 287.1 | 284.9 | -2.2 | -0.77
1110 | 1125 | 2 | 2.4 | 2.372 | 2.33 | -0.042 | -1.77 |  | 240.2 | 240 | 0 | 0 |  | 63.8 | 996.9 | 933.1 | 1462.54 | Fail | 299 | 299.9 | 0.9 | 0.3 |  | 191.4 | 189.2 | -2.2 | -1.15
1170 | 1185 | 1 | 1.2 | 1.184 | 1.16 | -0.024 | -2.03 |  | 120.1 | 120 | 0 | 0 |  | 31.9 | 996.9 | 965 | 3025.08 | Fail | 149.5 | 151 | 1.5 | 1 |  | 95.7 | 93.8 | -1.9 | -1.99
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -19 | -19 |  |  | 0 | 996.3 | 996.3 |  |  | 0 | 0.9 | 0.9 |  |  | 0 | -1.3 | -1.3
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.026 | 0.12 |  |  |  | Intercept | 0 | 12 |  |  | Intercept | 997 | 3 | Fail |  | Intercept | -1 | 15 |  |  | Intercept | -2.2 | 9.6
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.99 | 0.99 - 1.01 |  |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0219 | 0.12 |  |  |  | SEE | 5.9251 | 12.01 |  |  | SEE | 0.3898 | 3.19 |  |  | SEE | 2.0363 | 14.95 |  |  | SEE | 1.3695 | 9.573
|  |  | r2 | 1 | 0.998 |  |  |  | r2 | 0.9998 | 0.998 |  |  | r2 | 0.1385 | 0.998 | Fail |  | r2 | 1 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-07 | 9:23 | 0.01 | 0 | 0 | 1 | 0.5 | 1.5 | 0.2 | 0.1737 | 0.1 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 |  |  | 0.5
2008-05-07 | 9:23 | 0.01 | -0.0006 | -5.97 | 1 | 0.5 | 1.5 | 0.1 | 0.1 | 0.1 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.4 | 26.1 | 0.5
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1 | 0.5 | 1.5 | 0.1 | 0.0685 | 0.1 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.4 | 26.163 | 0.5
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.5 | 1.6 | 0 | 0.0259 | 0.1 | 13 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.3577 | 26.2 | 0.5
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0.1 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 997 | 28.3 | 26.2 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0.1 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.000583 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.2983 | 26.2 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.1404 | 26.2 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.000784 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996 | 28.2184 | 26.2 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0 | 13 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 996.164 | 28.2614 | 26.2386 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.4 | 1.5 | 0 | 0 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1031 | 997 | 28.1822 | 26.2822 | 0.4
2008-05-07 | 9:23 | 0.01 | -0.001 | -10 | 1.1 | 0.3 | 1.4 | -0.1 | 0 | 0 | 13 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.103 | 26.203 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1.1 | 0.3 | 1.4 | -0.1 | 0 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2164 | 0.0007 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2525 | 26.1238 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1.1 | 0.3 | 1.4 | -0.1 | 0 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2836 | 0.0007 | 0.0005 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.3552 | 26.1552 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1.1 | 0.3 | 1.4 | -0.2 | 0 | 0 | 12.9 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2164 | 0.0007 | 0.000484 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.4 | 26.2 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1 | 0.3 | 1.3 | -0.1 | 0 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.000683 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.4 | 26.2 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1 | 0.3 | 1.3 | -0.2 | 0 | 0 | 12.9 | 950 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.4 | 26.2 | 0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1 | 0.2 | 1.2 | -0.1 | 0 | 0 | 12.9 | 950 | 956 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.3279 | 26.1279 | 0.2
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1 | 0.2 | 1.2 | -0.2 | 0 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1971 | 26.1515 | 0.2
2008-05-07 | 9:24 | 0.01 | -0.0002 | -2.317 | 0.9 | 0.2 | 1.1 | -0.2 | 0 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.162 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.01 | 0 | 0 | 0.8 | 0.2 | 1 | -0.2 | 0 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.000584 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2796 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.01 | 0 | 0 | 0.8 | 0.2 | 1 | -0.2 | 0 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1211 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.01 | 0 | 0 | 0.7 | 0.2 | 0.9 | -0.3 | 0 | 0 | 13 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1686 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.01 | 0 | 0 | 0.6 | 0.2 | 0.8 | -0.2 | 0 | 0 | 12.9 | 950 | 956 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.2957 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.01 | 0 | 0 | 0.6 | 0.2 | 0.8 | -0.3 | 0 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0004 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.2915 | 26.2 | 0.2
2008-05-07 | 9:24 | 0.006 | 0 | 0 | 0.6 | 0.2 | 0.8 | -0.3 | 0 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28 | 26.2 | 0.2
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.6 | 0.1 | 0.7 | -0.3 | 0 | 0 | 13 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0005 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.6 | 0.1 | 0.7 | -0.4 | 0 | 0 | 12.9 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.000484 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1298 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.6 | 0.1 | 0.7 | -0.3 | 0 | 0 | 12.9 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1117 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.7 | 0.1 | 0.8 | -0.4 | 0 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.0935 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.7 | 0.1 | 0.8 | -0.3 | 0 | 0 | 12.9 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.1623 | 28.3948 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.7 | 0.1 | 0.8 | -0.5 | 0 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2363 | 26.2 | 0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.8 | 0 | 0.8 | -0.3 | 0 | 0 | 12.9 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.0778 | 26.2 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.8 | 0 | 0.8 | -0.4 | 0 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.0805 | 26.2 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.8 | 0 | 0.8 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.8363 | 28.2 | 26.2 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.9 | 0 | 0.9 | -0.4 | 0 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.2 | 26.2 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.9 | 0 | 0.9 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.8377 | 28.0439 | 26.2781 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.9 | -0.0164 | 0.9 | -0.4 | 0 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2164 | 0.000384 | 0.0002 | 0.0001 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.1146 | 26.2427 | 0
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.9 | -0.1 | 0.9 | -0.4 | 0 | 0 | 12.9 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.1269 | 26.2 | -0.1
2008-05-07 | 9:24 | 0 | 0 | 0 | 0.9 | -0.1164 | 0.9 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.836 | 28.0314 | 26.2 | -0.1
2008-05-07 | 9:24 | 0 | -0.0003 | -2.5458 | 1 | -0.2 | 1 | -0.4 | 0 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.1894 | 26.2 | -0.2
2008-05-07 | 9:24 | 0 | -0.0009 | -9.1201 | 1 | -0.2 | 1 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.0517 | 26.2742 | -0.2
2008-05-07 | 9:24 | 0 | -0.0001 | -0.6599 | 1 | -0.2 | 1 | -0.4 | 0 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.164 | 28.0535 | 26.3 | -0.2
2008-05-07 | 9:24 | 0 | 0 | 0 | 1 | -0.2 | 1 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.0674 | 26.2674 | -0.2
2008-05-07 | 9:24 | 0 | 0 | 0 | 1 | -0.2 | 1 | -0.5 | 0 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.836 | 28.0346 | 26.2115 | -0.2
2008-05-07 | 9:24 | 0 | -0.0004 | -4.2083 | 1 | -0.2 | 1 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.2735 | 26.2912 | -0.2
2008-05-07 | 9:24 | 0 | -0.001 | -10 | 1 | -0.2 | 1 | -0.3 | 0 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.4409 | 26.3 | -0.2
2008-05-07 | 9:24 | 0 | -0.001 | -10 | 1 | -0.2 | 1 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2836 | 0.0003 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.351 | 26.3 | -0.2
2008-05-07 | 9:24 | 0 | -0.001 | -10 | 1 | -0.2 | 1 | -0.3 | 0 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996 | 28.2 | 26.2711 | -0.2
2008-05-07 | 9:24 | 0 | -0.0002 | -2.4813 | 1 | -0.21623 | 1 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.1623 | 28.2 | 26.2 | -0.2
2008-05-07 | 9:24 | 0 | -0.0006 | -6.0413 | 1.1 | -0.3 | 1.1 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0003 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2 | 26.2 | -0.3
2008-05-07 | 9:24 | 0 | -0.001 | -10 | 1.1 | -0.3 | 1.1 | -0.4 | 0 | 0 | 13 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2664 | 26.2 | -0.3
2008-05-07 | 9:24 | 0 | -0.0013 | -12.5167 | 1.1 | -0.3 | 1.1 | -0.3 | 0 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.2083 | 26.2458 | -0.3
2008-05-07 | 9:24 | 0 | -0.0019 | -19.1319 | 1.1 | -0.3 | 1.1 | -0.4 | 0 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.1754 | 26.2749 | -0.3
2008-05-07 | 9:24 | 0 | -0.0011 | -10.6213 | 1.1 | -0.3 | 1.1 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.3956 | 26.2 | -0.3
2008-05-07 | 9:24 | 0 | -0.001 | -10 | 1.1 | -0.3 | 1.1 | -0.2 | 0 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28.3164 | 26.2 | -0.3
2008-05-07 | 9:24 | 0.003 | -0.001 | -10 | 1.1 | -0.3 | 1.1 | -0.3 | 0 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 996.836 | 28.1116 | 26.2628 | -0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1.1 | -0.3 | 1.1 | -0.3 | 0 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 31.2707 | 53 | 1032 | 997 | 28.168 | 26.3 | -0.3
2008-05-07 | 9:24 | 0.01 | -0.001 | -10 | 1.1 | -0.3 | 1.1 | -0.4 | 0 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2164 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.836 | 28.3151 | 26.3 | -0.3
2008-05-07 | 9:24 | 0.01 | -0.0009 | -8.8732 | 1.1 | -0.3 | 1.1 | -0.3 | 0 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 31.73 | 53 | 1032 | 996.164 | 28 | 26.3 | -0.3
2008-05-07 | 9:24 | 0.499 | -0.0001 | -0.58823 | 1 | -0.3 | 1 | -0.4 | 0 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0002 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 53 | 1032 | 997 | 28 | 26.3 | -0.3
2008-05-07 | 9:24 | 0.971 | 0.0125 | 124.8027 | 33.6 | -0.3 | 33.6 | 5.5 | 5.5425 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.006651 | 0 | 0 | 0.0001 |  | n/a |  | 31.27 | 53 | 1032 | 997 | 28.2359 | 26.241 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.0142 | 142.1595 | 87.3 | -0.3 | 87.3 | 59.1 | 59.0946 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.017 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.4 | 26.2382 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 133.4 | -0.3 | 133.4 | 87.1 | 87.1107 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.025512 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.3826 | 26.3 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 147.3 | -0.3 | 147.3 | 92.2 | 92.2444 | 0 | 13 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.028 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.3033 | 26.3 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 147.5 | -0.3 | 147.5 | 93.4 | 93.4239 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.028016 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1482 | 26.3 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 148.3 | -0.3 | 148.3 | 94 | 93.9748 | 0 | 13 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0281 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.2653 | 26.2449 | -0.3
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 148.3 | -0.1836 | 148.3 | 94.2 | 94.1748 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.028116 | 0 | 0 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.836 | 28.297 | 26.2343 | -0.2
2008-05-07 | 9:24 | 1.14 | 0.013 | 130 | 148.4 | -0.1 | 148.4 | 94.2 | 94.2 | 0 | 13 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 1.6e-05 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1 | 26.3 | -0.1
2008-05-07 | 9:24 | 1.141 | 0.013 | 130 | 148.5 | -0.1 | 148.5 | 94.2 | 94.2 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1 | 26.3 | -0.1
2008-05-07 | 9:24 | 1.149 | 0.0133 | 132.582 | 148.6 | 0 | 148.6 | 94.5 | 94.5438 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1719 | 26.2281 | 0
2008-05-07 | 9:24 | 1.15 | 0.014 | 140 | 148.6 | 0.1 | 148.7 | 94.4 | 94.4228 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.2 | 26.2512 | 0.1
2008-05-07 | 9:25 | 1.15 | 0.014 | 140 | 148.7 | 0.2 | 148.9 | 94.6 | 94.6 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.2 | 26.3 | 0.2
2008-05-07 | 9:25 | 1.15 | 0.014 | 140 | 148.7 | 0.2164 | 148.9 | 94.6 | 94.5686 | 0 | 12.9 | 952 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.2 | 26.3 | 0.2
2008-05-07 | 9:25 | 1.15 | 0.014 | 140 | 148.7 | 0.3 | 149 | 94.7 | 94.675 | 0 | 12.9 | 953 | 957 | 940 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.2 | 26.3 | 0.3
2008-05-07 | 9:25 | 1.15 | 0.014 | 140 | 148.7 | 0.4 | 149.1 | 94.6 | 94.5514 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.2 | 26.3 | 0.4
2008-05-07 | 9:25 | 1.15 | 0.0138 | 137.5333 | 148.7 | 0.4164 | 149.1 | 94.6 | 94.6 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1053 | 26.2526 | 0.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.7 | 0.6 | 149.3 | 94.6 | 94.6233 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0796 | 26.2265 | 0.6
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.7 | 0.6 | 149.3 | 94.6 | 94.6463 | 0 | 12.9 | 953 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.3 | 26.3 | 0.6
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.6 | 0.6164 | 149.2 | 94.7 | 94.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.3 | 26.3 | 0.6
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.6 | 0.7 | 149.3 | 94.6 | 94.6187 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1073 | 26.3 | 0.7
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.6 | 0.7 | 149.3 | 94.8 | 94.7638 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 0.7
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.6 | 0.81637 | 149.4 | 94.6 | 94.6082 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.2836 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.8363 | 28.0687 | 26.3 | 0.8
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.7 | 1 | 149.7 | 94.8 | 94.7705 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.2936 | 26.3 | 1
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.7 | 2.1922 | 150.9 | 94.6 | 94.6261 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.000216 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995.8377 | 28.0559 | 26.3 | 2.2
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.8 | 4.911 | 153.7 | 94.7 | 94.6557 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.000316 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.0606 | 26.3 | 4.9
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.7 | 7.3 | 156 | 94.8 | 94.8 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0604 | 26.3 | 7.3
2008-05-07 | 9:25 | 1.151 | 0.013 | 130 | 148.7 | 7.4164 | 156.1 | 94.7 | 94.6511 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0559 | 26.3 | 7.4
2008-05-07 | 9:25 | 1.159 | 0.013 | 130 | 148.6 | 7.5 | 156.1 | 94.6 | 94.62 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.2947 | 26.3 | 7.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.6 | 7.5 | 156.1 | 94.4 | 94.4278 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.2225 | 26.3 | 7.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.4 | 7.5 | 155.9 | 94.6 | 94.6229 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997.164 | 28.0297 | 26.3 | 7.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.4 | 7.4 | 155.8 | 94.6 | 94.6 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28.0801 | 26.3 | 7.4
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.4 | 7.4 | 155.8 | 94.6 | 94.626 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.2169 | 0.0282 | 0 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28.3391 | 26.3 | 7.4
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 7.4 | 155.7 | 94.7 | 94.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997.836 | 28.0222 | 26.3 | 7.4
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 6.5 | 154.8 | 94.7 | 94.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.2211 | 26.3 | 6.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.7 | 154 | 94.5 | 94.5 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.3529 | 26.3 | 5.7
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.7 | 154 | 94.7 | 94.651 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.2723 | 26.3 | 5.7
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.5836 | 153.9 | 94.6 | 94.6291 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 5.6
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.5 | 153.8 | 94.6 | 94.5715 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 5.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.5 | 153.8 | 94.5 | 94.5407 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0699 | 26.3 | 5.5
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.3 | 5.5 | 153.8 | 94.6 | 94.6 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.8377 | 28.1 | 26.3 | 5.5
2008-05-07 | 9:25 | 1.153 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.6 | 94.6188 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.185 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.6 | 94.573 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 28.3699 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.7 | 94.6733 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0529 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.5 | 94.505 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.7 | 94.7458 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9548 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.6 | 94.6297 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9244 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.7 | 94.6757 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0036 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.7 | 94.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.0829 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.3 | 5.4 | 153.7 | 94.7 | 94.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 27.9758 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.2 | 5.4 | 153.6 | 94.6 | 94.6445 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 27.9414 | 26.3 | 5.4
2008-05-07 | 9:25 | 1.151 | 0.0126 | 125.9417 | 148.2 | 5.3 | 153.5 | 94.5 | 94.5274 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995.164 | 28 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.159 | 0.0122 | 122.3937 | 148.2 | 5.3 | 153.5 | 94.5 | 94.4951 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.2 | 5.3 | 153.5 | 94.5 | 94.5084 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.16 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.8 | 94.7741 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1746 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.156 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.6 | 94.6 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.225 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.8 | 94.7508 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28.0833 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.7 | 94.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0042 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.7 | 94.6784 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.15 | 0.013 | 130 | 148.1 | 5.3 | 153.4 | 94.7 | 94.6731 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0544 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.411 | 0.013 | 130 | 148 | 5.3164 | 153.3 | 94.7 | 94.6746 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0282 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0663 | 26.3 | 5.3
2008-05-07 | 9:25 | 1.872 | 0.0234 | 234.1519 | 158.7 | 5.3164 | 164 | 94.9 | 94.8622 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.030108 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0258 | 26.3 | 5.3
2008-05-07 | 9:25 | 2.272 | 0.0247 | 247.2523 | 199.1 | 5.4 | 204.5 | 140.8 | 140.8091 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.037 | 0.0001 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1842 | 26.3 | 5.4
2008-05-07 | 9:25 | 2.282 | 0.023 | 230 | 262 | 8 | 270 | 178.7 | 178.7333 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.047376 | 0.0001 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.129 | 26.3 | 8
2008-05-07 | 9:25 | 2.299 | 0.023 | 230 | 290.9 | 8.9164 | 299.8 | 186.7 | 186.6509 | 0 | 12.9 | 953 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.051916 | 0.0002 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0497 | 26.3 | 8.9
2008-05-07 | 9:25 | 2.3 | 0.0231 | 231.0042 | 292.5 | 9.8 | 302.3 | 188.5 | 188.4931 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.052132 | 0.0002 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.317 | 0.0239 | 239.3724 | 292.9 | 9.8 | 302.7 | 188.8 | 188.8097 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0523 | 0.0002 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9911 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.326 | 0.024 | 240 | 293 | 9.8 | 302.8 | 189.2 | 189.1949 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.052316 | 0.0002 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9119 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.32 | 0.024 | 240 | 293.3 | 9.8 | 303.1 | 189.3 | 189.3 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0524 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9674 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.32 | 0.024 | 240 | 293.4 | 9.8 | 303.2 | 189.3 | 189.3251 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0524 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0933 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.32 | 0.024 | 240 | 294.1 | 9.8 | 303.9 | 189.5 | 189.4698 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0526 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1482 | 26.3 | 9.8
2008-05-07 | 9:25 | 2.32 | 0.024 | 240 | 294.2 | 9.8 | 304 | 189.3 | 189.3253 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0526 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.8377 | 28.0051 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.024 | 240 | 294.3 | 9.8 | 304.1 | 189.6 | 189.6206 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0526 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0843 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.024 | 240 | 294.3 | 9.8 | 304.1 | 189.6 | 189.6 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0526 | 0.0003 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0367 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.024 | 240 | 294.2 | 9.8 | 304 | 189.9 | 189.8835 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0526 | 0.0004 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.024 | 240 | 294.1 | 9.8 | 303.9 | 189.5 | 189.5251 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0004 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9561 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.024 | 240 | 294.2 | 9.8 | 304 | 189.8 | 189.7887 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0004 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.8024 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.0241 | 240.7573 | 294.1 | 9.8 | 303.9 | 189.8 | 189.7868 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0004 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 27.961 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.0249 | 249.3632 | 294.1 | 9.8 | 303.9 | 189.6 | 189.6 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0004 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0597 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.025 | 250 | 294.1 | 9.8 | 303.9 | 189.7 | 189.6727 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1389 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.025 | 250 | 294.1 | 9.8 | 303.9 | 189.7 | 189.6579 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0527 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1819 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.025 | 250 | 294.1 | 9.8 | 303.9 | 189.7 | 189.7265 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.052716 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 30.2703 | 53 | 1032 | 997 | 28.1027 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.025 | 250 | 294.1 | 9.8 | 303.9 | 189.6 | 189.6267 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0528 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 30.73 | 53 | 1032 | 997 | 28.1765 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.32 | 0.0249 | 249.2721 | 294.1 | 9.8 | 303.9 | 189.4 | 189.4192 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0528 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 30.27 | 53 | 1032 | 997 | 28.1445 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.326 | 0.0241 | 241.0008 | 294.2 | 9.8 | 304 | 189.5 | 189.5278 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0528 | 0.0005 | 0.0006 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 28.0306 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 294.2 | 9.8 | 304 | 189.7 | 189.7 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0528 | 0.000516 | 0.0006 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.9142 | 26.3 | 9.8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 294.2 | 9.7 | 303.9 | 189.6 | 189.6 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0528 | 0.0006 | 0.0006 | 0.0001 |  | n/a |  | 30.73 | 53 | 1032 | 997 | 27.9934 | 26.3 | 9.7
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 294.2 | 9.7 | 303.9 | 189.7 | 189.6851 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0529 | 0.0006 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0727 | 26.3 | 9.7
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 294.3 | 9.6838 | 304 | 189.6 | 189.5935 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.2838 | 0.0529 | 0.0006 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0481 | 26.3 | 9.7
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 294.7 | 8.8 | 303.5 | 189.5 | 189.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.053 | 0.0006 | 0.0006 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 8.8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 295.1 | 7.9836 | 303.1 | 189.6 | 189.5722 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.053016 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0309 | 26.3 | 8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 295.4 | 7.8 | 303.2 | 189.6 | 189.6436 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0531 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.8377 | 28.2686 | 26.3 | 7.8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 295.6 | 6.9 | 302.5 | 189.6 | 189.553 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0531 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0937 | 26.3 | 6.9
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 295.6 | 6.8836 | 302.5 | 189.6 | 189.6 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0531 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 6.9
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296 | 6.6836 | 302.7 | 189.7 | 189.664 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0532 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 28 | 26.3 | 6.7
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.1 | 6.5836 | 302.7 | 189.2 | 189.2311 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0532 | 0.0006 | 0.0005 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 6.6
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.3 | 6.4 | 302.7 | 189.6 | 189.6103 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.0532 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 6.4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.4 | 6.1688 | 302.6 | 189.5 | 189.4711 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.053216 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 28 | 26.3 | 6.2
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.5 | 5.5 | 302 | 189.7 | 189.6737 | 0 | 12.9 | 953 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9118 | 26.3 | 5.5
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.5 | 5.4 | 301.9 | 189.5 | 189.4793 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.8466 | 26.3 | 5.4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.6 | 5.3836 | 302 | 189.4 | 189.3699 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0025 | 26.3 | 5.4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.5 | 5.2836 | 301.8 | 189.2 | 189.1662 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0817 | 26.3 | 5.3
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.5 | 5.2 | 301.7 | 188.9 | 188.8563 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0391 | 26.3 | 5.2
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.6 | 5.0836 | 301.7 | 189.2 | 189.2461 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0006 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.04 | 26.3 | 5.1
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.6 | 5 | 301.6 | 189.3 | 189.3031 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.062 | 26.3 | 5
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.7 | 4.9 | 301.6 | 189.6 | 189.6 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0.0004 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.903 | 26.3 | 4.9
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.8 | 4.8 | 301.6 | 189.5 | 189.4715 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.2164 | 0.0533 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9779 | 26.3 | 4.8
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.8 | 4.6838 | 301.5 | 189.7 | 189.6818 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0533 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995.8377 | 28 | 26.3 | 4.7
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.9 | 4.5836 | 301.5 | 189.7 | 189.7 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.053316 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28 | 26.3 | 4.6
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 297 | 4.4836 | 301.5 | 189.5 | 189.5 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995.164 | 28.0155 | 26.3 | 4.5
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.9 | 4.4 | 301.3 | 189.6 | 189.5906 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0946 | 26.3 | 4.4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.9 | 4.2836 | 301.2 | 189.7 | 189.7196 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0261 | 26.3 | 4.3
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.8 | 4.1836 | 301 | 189.7 | 189.7298 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 4.2
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.8 | 4.0836 | 300.9 | 189.8 | 189.8 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0645 | 26.3 | 4.1
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.8 | 4 | 300.8 | 189.9 | 189.8698 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.1777 | 26.3 | 4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 296.9 | 4 | 300.9 | 189.7 | 189.6589 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28.0186 | 26.3 | 4
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 297 | 3.9 | 300.9 | 189.9 | 189.9391 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.07 | 26.3 | 3.9
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 297.1 | 3.8836 | 301 | 189.9 | 189.8725 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 28.0507 | 26.3 | 3.9
2008-05-07 | 9:26 | 2.33 | 0.024 | 240 | 297.1 | 3.8 | 300.9 | 189.7 | 189.7229 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 27.9715 | 26.3 | 3.8
2008-05-07 | 9:26 | 2.33 | 0.0241 | 240.7301 | 297.1 | 3.7 | 300.8 | 189.8 | 189.8 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 995 | 27.9077 | 26.3 | 3.7
2008-05-07 | 9:26 | 2.397 | 0.0249 | 248.9336 | 297.1 | 3.6 | 300.7 | 189.7 | 189.7282 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0534 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9867 | 26.3 | 3.6
2008-05-07 | 9:26 | 2.851 | 0.0351 | 351.3311 | 300.5 | 3.5836 | 304.1 | 190 | 190.006 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.053924 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.934 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.489 | 0.0367 | 366.9257 | 347.1 | 3.5 | 350.6 | 227.2 | 227.1946 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.060954 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.764 | 26.3 | 3.5
2008-05-07 | 9:26 | 3.52 | 0.035 | 350 | 406 | 3.5 | 409.5 | 271.8 | 271.7598 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.069487 | 0.0007 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7222 | 26.3 | 3.5
2008-05-07 | 9:26 | 3.52 | 0.035 | 350 | 437.9 | 3.5164 | 441.4 | 281.2 | 281.1877 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.073947 | 0.0008 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0933 | 26.3 | 3.5
2008-05-07 | 9:26 | 3.52 | 0.035 | 350 | 443.6 | 3.6 | 447.2 | 283.5 | 283.5387 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0748 | 0.0008 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9341 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.531 | 0.035 | 350 | 444.5 | 3.6 | 448.1 | 284 | 283.9727 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3164 | 0.074916 | 0.0008 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0246 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.54 | 0.035 | 350 | 444.6 | 3.6 | 448.2 | 284.4 | 284.3946 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.075 | 0.000816 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.54 | 0.035 | 350 | 444.8 | 3.6 | 448.4 | 284.4 | 284.4319 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.075016 | 0.0009 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0586 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.541 | 0.0354 | 353.8935 | 444.7 | 3.6 | 448.3 | 285 | 284.9765 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0751 | 0.0009 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9002 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.549 | 0.036 | 360 | 444.8 | 3.6 | 448.4 | 284.9 | 284.9448 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.2836 | 0.0751 | 0.001 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.836 | 27.979 | 26.3 | 3.6
2008-05-07 | 9:26 | 3.55 | 0.036 | 360 | 444.8 | 3.6 | 448.4 | 284.8 | 284.8299 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.075116 | 0.001 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28 | 26.3 | 3.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 444.7 | 3.6 | 448.3 | 285 | 284.9769 | 0 | 13 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.2836 | 0.0752 | 0.001016 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0751 | 26.3 | 3.6
2008-05-07 | 9:27 | 3.546 | 0.036 | 360 | 444.7 | 3.6 | 448.3 | 284.7 | 284.7302 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.0752 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.15 | 26.3 | 3.6
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 444.7 | 3.6 | 448.3 | 284.9 | 284.874 | 0 | 13 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0752 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9133 | 26.3 | 3.6
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 444.8 | 3.5836 | 448.4 | 284.9 | 284.8762 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0752 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9751 | 26.3 | 3.6
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.2 | 3.5 | 448.7 | 285 | 285 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.075316 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.0544 | 26.3 | 3.5
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.7 | 3.5 | 449.2 | 284.8 | 284.7853 | 0 | 13 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0754 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1337 | 26.3 | 3.5
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.8 | 3.4 | 449.2 | 284.9 | 284.8602 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0754 | 0.0011 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 28.1226 | 26.3 | 3.4
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.8 | 3.3831 | 449.2 | 284.9 | 284.9261 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0754 | 0.001117 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6475 | 26.3 | 3.4
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.8 | 3.3 | 449.1 | 285.1 | 285.1 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0754 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.8841 | 26.3 | 3.3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.3 | 449.2 | 285.2 | 285.1547 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2838 | 0.0754 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8994 | 26.3 | 3.3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.3 | 449.2 | 285 | 285.02 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8594 | 26.3 | 3.3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.2836 | 449.2 | 285 | 285 | 0 | 13 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9911 | 26.3 | 3.3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.2 | 449.1 | 285 | 285 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9122 | 26.3 | 3.2
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.2 | 449.1 | 285.2 | 285.1642 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0346 | 26.3 | 3.2
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.2 | 449.1 | 285.1 | 285.0776 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0534 | 26.3 | 3.2
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 3.2 | 449.1 | 285.2 | 285.1568 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0012 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 3.2
2008-05-07 | 9:27 | 3.54 | 0.0364 | 364.4501 | 445.9 | 3.2 | 449.1 | 285 | 285 | 0 | 13 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.001216 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 3.2
2008-05-07 | 9:27 | 3.54 | 0.0378 | 377.8484 | 445.9 | 3.1 | 449 | 285.2 | 285.233 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.54 | 0.0353 | 353.3985 | 445.8 | 3.1 | 448.9 | 285 | 285 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0755 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9368 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.54 | 0.035 | 350 | 445.7 | 3.1 | 448.8 | 285 | 284.9819 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9424 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.54 | 0.035 | 350 | 445.8 | 3.1 | 448.9 | 285 | 284.9984 | 0 | 13 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0219 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.541 | 0.035 | 350 | 445.8 | 3.1 | 448.9 | 284.9 | 284.8713 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0965 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.549 | 0.035 | 350 | 445.9 | 3.1 | 449 | 285.1 | 285.05 | 0 | 13 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8599 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.543 | 0.0351 | 350.7373 | 446 | 3.0838 | 449.1 | 284.8 | 284.8258 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.000284 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9189 | 26.3 | 3.1
2008-05-07 | 9:27 | 3.54 | 0.0359 | 359.2119 | 446 | 3 | 449 | 285 | 284.9739 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9223 | 26.3 | 3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.9836 | 449 | 284.8 | 284.8433 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 27.8181 | 26.3 | 3
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 2.9 | 448.8 | 284.9 | 284.9 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8972 | 26.3 | 2.9
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.9 | 448.9 | 284.9 | 284.9 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9 | 26.3 | 2.9
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.9 | 448.9 | 285.1 | 285.1425 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0013 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9 | 26.3 | 2.9
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.8 | 448.8 | 284.9 | 284.8836 | 0 | 13 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9345 | 26.3 | 2.8
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.8 | 448.8 | 285.2 | 285.154 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0756 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9435 | 26.3 | 2.8
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 445.9 | 2.7 | 448.6 | 285.1 | 285.1 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6263 | 26.3 | 2.7
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.7 | 448.7 | 285 | 285.0273 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6726 | 26.3 | 2.7
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446 | 2.7 | 448.7 | 285.1 | 285.0885 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8036 | 26.3 | 2.7
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446.1 | 2.6834 | 448.8 | 285.1 | 285.1478 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8379 | 26.3 | 2.7
2008-05-07 | 9:27 | 3.54 | 0.036 | 360 | 446.1 | 2.6 | 448.7 | 285.3 | 285.2602 | 0 | 13 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7102 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.546 | 0.036 | 360 | 446.2 | 2.6 | 448.8 | 285 | 284.9532 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7895 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.3 | 2.6 | 448.9 | 285 | 284.9545 | 0 | 12.9 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7313 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.3 | 2.6 | 448.9 | 285 | 284.9859 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7952 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.4 | 2.6 | 449 | 285.1 | 285.0781 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0757 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.82 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.4 | 2.6 | 449 | 285 | 285 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0758 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.1637 | 27.6128 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.4 | 2.6 | 449 | 285 | 285 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0758 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7712 | 26.3 | 2.6
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.5 | 2.5 | 449 | 284.9 | 284.9285 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0758 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8 | 26.3 | 2.5
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 446.7 | 2.5 | 449.2 | 284.9 | 284.8746 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0758 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.756 | 26.3 | 2.5
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 447 | 2.5 | 449.5 | 285.2 | 285.1754 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0758 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7232 | 26.3 | 2.5
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 447.1 | 2.5 | 449.6 | 284.9 | 284.9258 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0759 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7951 | 26.3 | 2.5
2008-05-07 | 9:27 | 3.55 | 0.036 | 360 | 447.2 | 2.5 | 449.7 | 284.9 | 284.9279 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0759 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032.31 | 996.836 | 27.6366 | 26.3 | 2.5
2008-05-07 | 9:27 | 3.55 | 0.0371 | 370.8553 | 447.3 | 2.5 | 449.8 | 285 | 284.9792 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0759 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 27.7218 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.132 | 0.0494 | 494.2105 | 450.5 | 2.5 | 453 | 285.1 | 285.1259 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.076294 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.836 | 27.8796 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.658 | 0.0481 | 481.3629 | 524 | 2.5 | 526.5 | 320.5 | 320.5267 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.082696 | 0.0014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9812 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.0465 | 464.5088 | 571.1 | 2.5 | 573.6 | 362.6 | 362.5739 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.092 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9014 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.046 | 460 | 584.9 | 2.5 | 587.4 | 375.5 | 375.4903 | 0 | 13 | 952 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.093684 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.8222 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.046 | 460 | 591.7 | 2.5 | 594.2 | 378 | 378.0192 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.094549 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032.7123 | 996 | 27.6859 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.0461 | 461.4239 | 593.8 | 2.5 | 596.3 | 378.7 | 378.7359 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.094816 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 27.6725 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.0478 | 477.6052 | 595.2 | 2.5 | 597.7 | 378.9 | 378.8853 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.095016 | 0.0015 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8154 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.048 | 480 | 596.7 | 2.5162 | 599.2 | 379 | 379.0257 | 0 | 13 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0953 | 0.0016 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8945 | 26.3 | 2.5
2008-05-07 | 9:27 | 4.72 | 0.048 | 480 | 596.8 | 2.6 | 599.4 | 379.3 | 379.3 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0953 | 0.0016 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9739 | 26.3 | 2.6
2008-05-07 | 9:27 | 4.721 | 0.048 | 480 | 596.9 | 2.6 | 599.5 | 379.4 | 379.4 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0953 | 0.0016 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8937 | 26.3 | 2.6
2008-05-07 | 9:27 | 4.729 | 0.048 | 480 | 596.8 | 2.6 | 599.4 | 379.5 | 379.4766 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.095316 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8963 | 26.3 | 2.6
2008-05-07 | 9:27 | 4.73 | 0.048 | 480 | 596.9 | 2.7 | 599.6 | 379.6 | 379.5731 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.8363 | 28.0563 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597 | 2.7 | 599.7 | 379.5 | 379.5452 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.737 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597.1 | 2.7 | 599.8 | 379.6 | 379.6 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9101 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597.1 | 2.7 | 599.8 | 379.6 | 379.6281 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8521 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597.1 | 2.7 | 599.8 | 379.7 | 379.7485 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6715 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597.1 | 2.8 | 599.9 | 379.7 | 379.7263 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0954 | 0.0017 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.623 | 26.3 | 2.8
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 597 | 2.7 | 599.7 | 379.7 | 379.7265 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.095416 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 27.8606 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.73 | 0.048 | 480 | 596.8 | 2.7 | 599.5 | 380 | 379.9708 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9661 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.734 | 0.048 | 480 | 596.8 | 2.7 | 599.5 | 379.7 | 379.7 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.6 | 2.7 | 599.3 | 379.8 | 379.8383 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032.2802 | 997 | 27.9268 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.0484 | 483.6781 | 596.6 | 2.7 | 599.3 | 379.8 | 379.7517 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997.164 | 27.6969 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.0488 | 487.9029 | 596.5 | 2.7 | 599.2 | 379.9 | 379.8739 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0018 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.6171 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.5 | 2.7 | 599.2 | 379.6 | 379.5968 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0955 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.8491 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.5 | 2.7 | 599.2 | 379.6 | 379.5692 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.095516 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.8339 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.6 | 2.7 | 599.3 | 379.7 | 379.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.5 | 2.7 | 599.2 | 379.8 | 379.7865 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.5 | 2.7 | 599.2 | 379.9 | 379.9278 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.6 | 2.7 | 599.3 | 380 | 379.9744 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.8 | 2.7 | 599.5 | 379.9 | 379.8785 | 0 | 12.9 | 951 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7128 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597 | 2.7 | 599.7 | 379.7 | 379.7 | 0 | 12.9 | 951 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.0019 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.165 | 27.8662 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597 | 2.7 | 599.7 | 379.7 | 379.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.2836 | 0.0956 | 0.001916 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7078 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597 | 2.6836 | 599.7 | 379.8 | 379.8 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0956 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7 | 26.3 | 2.7
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.9 | 2.6 | 599.5 | 379.8 | 379.7733 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.095616 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6455 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.9 | 2.6 | 599.5 | 379.6 | 379.6 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0957 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.5663 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.7 | 2.6 | 599.3 | 379.7 | 379.6788 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0957 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.5646 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 596.8 | 2.6 | 599.4 | 379.7 | 379.727 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.2836 | 0.0957 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.9608 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597.6 | 2.6 | 600.2 | 379.7 | 379.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28.1425 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597.7 | 2.6 | 600.3 | 379.6 | 379.6261 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.8971 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597.7 | 2.6 | 600.3 | 379.5 | 379.4699 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.7193 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.048 | 480 | 597.8 | 2.6 | 600.4 | 379.5 | 379.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.0481 | 481.4578 | 597.7 | 2.6 | 600.3 | 379.4 | 379.4274 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 28 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.0498 | 497.8378 | 597.7 | 2.6 | 600.3 | 379.6 | 379.6287 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0958 | 0.002 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.9325 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.0493 | 492.856 | 597.8 | 2.6 | 600.4 | 379.6 | 379.5737 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.002016 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.9 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.049 | 490 | 597.8 | 2.6 | 600.4 | 379.8 | 379.8416 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.874 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.049 | 490 | 597.8 | 2.6 | 600.4 | 379.7 | 379.7279 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.8 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.049 | 490 | 597.8 | 2.6 | 600.4 | 379.7 | 379.7 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.8 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.74 | 0.0489 | 489.3218 | 597.9 | 2.6 | 600.5 | 379.5 | 379.5269 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7365 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.745 | 0.0481 | 481.0505 | 598 | 2.6 | 600.6 | 379.5 | 379.4523 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.8283 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.3 | 379.312 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 28 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.2 | 379.1606 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9946 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.2 | 2.6 | 600.8 | 379.4 | 379.3804 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.6777 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.3 | 379.2707 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.7196 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.3 | 379.2679 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.164 | 27.839 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598 | 2.6 | 600.6 | 379.3 | 379.2737 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.9182 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.1 | 379.1204 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9975 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.2 | 379.171 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.9234 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.2 | 2.6 | 600.8 | 379.5 | 379.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.7329 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.048 | 480 | 598.1 | 2.6 | 600.7 | 379.5 | 379.5 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.6 | 26.3 | 2.6
2008-05-07 | 9:28 | 4.75 | 0.0481 | 480.6656 | 598.1 | 2.6 | 600.7 | 379.4 | 379.434 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 996.1637 | 27.6144 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.112 | 0.0489 | 488.7825 | 598 | 2.6 | 600.6 | 379.5 | 379.5261 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0959 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6936 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.606 | 0.0561 | 561.0702 | 608.7 | 2.6 | 611.3 | 380 | 380.0458 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.097209 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6272 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.93 | 0.0585 | 584.5042 | 683.4 | 2.6 | 686 | 428.8 | 428.8384 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.10609 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6521 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.93 | 0.0591 | 591.3079 | 709.8 | 2.6 | 712.4 | 464.3 | 464.2813 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.10908 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6062 | 26.3313 | 2.6
2008-05-07 | 9:28 | 5.93 | 0.061 | 610 | 737.2 | 2.6 | 739.8 | 472 | 472.0091 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11208 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.4314 | 26.3895 | 2.6
2008-05-07 | 9:28 | 5.937 | 0.061 | 610 | 741 | 2.6 | 743.6 | 473.2 | 473.1571 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1126 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6691 | 26.3103 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.061 | 610 | 742 | 2.6 | 744.6 | 473.7 | 473.6713 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.11272 | 0.0023 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.061 | 610 | 742.7 | 2.6 | 745.3 | 473.8 | 473.8435 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.1129 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.7 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.061 | 610 | 743 | 2.6 | 745.6 | 474 | 474.0357 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11292 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1032.6536 | 997 | 27.7814 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.0606 | 606.2276 | 743.3 | 2.6 | 745.9 | 474.2 | 474.2 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.113 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9867 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.06 | 600 | 743.3 | 2.6 | 745.9 | 474.5 | 474.4518 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11302 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8283 | 26.3 | 2.6
2008-05-07 | 9:28 | 5.94 | 0.06 | 600 | 743.5 | 2.6 | 746.1 | 474.4 | 474.4362 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1131 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9301 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 743.5 | 2.6 | 746.1 | 473.8 | 473.791 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1131 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9114 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 743.8 | 2.6 | 746.4 | 474.1 | 474.1056 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1132 | 0.002516 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 744 | 2.6 | 746.6 | 474 | 473.9725 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1132 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7974 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 744 | 2.6 | 746.6 | 474 | 474.0368 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1132 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7181 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.0598 | 597.9792 | 744.3 | 2.6 | 746.9 | 474.1 | 474.0721 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11322 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6388 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.059 | 590 | 744.6 | 2.6 | 747.2 | 474.2 | 474.1747 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1133 | 0.002616 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6403 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.059 | 590 | 744.6 | 2.6 | 747.2 | 474.1 | 474.1461 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1133 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.7 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.059 | 590 | 744.7 | 2.6 | 747.3 | 474.2 | 474.2 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1134 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.7 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.0595 | 595.4958 | 744.6 | 2.6 | 747.2 | 474.2 | 474.1545 | 0 | 12.9 | 952 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1134 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8563 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 744.6 | 2.6 | 747.2 | 474.2 | 474.1843 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1134 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.94 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.3 | 474.2574 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8634 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.947 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.4 | 474.3586 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.8471 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.95 | 0.06 | 600 | 744.8 | 2.6 | 747.4 | 474.3 | 474.3133 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.0839 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.95 | 0.06 | 600 | 744.8 | 2.6 | 747.4 | 474.3 | 474.3258 | 0 | 12.9 | 952 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.002816 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 28.0259 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.952 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.3 | 474.3407 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3164 | 0.1135 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8929 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.3 | 474.3248 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.1135 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.3 | 474.3204 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8229 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 744.7 | 2.6 | 747.3 | 474.2 | 474.2284 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1135 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9823 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 744.8 | 2.6 | 747.4 | 474.4 | 474.3599 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 744.9 | 2.6 | 747.5 | 474.4 | 474.356 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.8363 | 27.9006 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.958 | 0.0595 | 594.5042 | 745.1 | 2.6 | 747.7 | 474.1 | 474.1186 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.7711 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.95 | 0.059 | 590 | 745.2 | 2.6 | 747.8 | 474.5 | 474.4725 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.7241 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.95 | 0.059 | 590 | 745.2 | 2.6 | 747.8 | 474.3 | 474.2774 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9605 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.957 | 0.059 | 590.2332 | 745.1 | 2.6 | 747.7 | 474.6 | 474.5889 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9336 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.0599 | 598.8687 | 745.1 | 2.6 | 747.7 | 474.4 | 474.4297 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.9 | 26.3 | 2.6
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.2 | 2.5 | 747.7 | 474.5 | 474.5474 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.925 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.1 | 2.5 | 747.6 | 474.4 | 474.3672 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1136 | 0.002916 | 0.0002 | 0.0001 |  | n/a |  | 30.2707 | 53 | 1033 | 997 | 27.9914 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745 | 2.5 | 747.5 | 474.3 | 474.271 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11362 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.833 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.2 | 2.5 | 747.7 | 474.5 | 474.5271 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6745 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.1 | 2.5 | 747.6 | 474.5 | 474.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7259 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.955 | 0.06 | 600 | 745.2 | 2.5 | 747.7 | 474.8 | 474.775 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9212 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.95 | 0.0607 | 607.0503 | 745.2 | 2.5 | 747.7 | 474.5 | 474.5307 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9988 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.952 | 0.061 | 610 | 745.3 | 2.5 | 747.8 | 474.7 | 474.6733 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7613 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.0606 | 606.2571 | 745.5 | 2.5 | 748 | 474.7 | 474.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1122 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.6 | 2.5 | 748.1 | 474.6 | 474.5922 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.209 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.6 | 2.5 | 748.1 | 474.6 | 474.6436 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.5 | 2.4 | 747.9 | 475 | 474.9805 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3836 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9 | 26.3 | 2.4
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.5 | 2.4 | 747.9 | 475.1 | 475.0756 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1137 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7483 | 26.3 | 2.4
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.4 | 2.4 | 747.8 | 475 | 475.027 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.11372 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7 | 26.3 | 2.4
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.4 | 2.4 | 747.8 | 474.9 | 474.8545 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8369 | 26.3 | 2.4
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.5 | 2.5 | 748 | 474.9 | 474.8794 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0731 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.7 | 2.5 | 748.2 | 474.9 | 474.8933 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9146 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.7 | 2.5 | 748.2 | 474.7 | 474.7 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6847 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.06 | 600 | 745.7 | 2.5 | 748.2 | 474.9 | 474.8758 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.7021 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.0595 | 594.5042 | 745.8 | 2.5 | 748.3 | 474.7 | 474.731 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8609 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.059 | 590 | 745.7 | 2.5 | 748.2 | 475 | 474.9891 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.059 | 590 | 745.6 | 2.5 | 748.1 | 474.9 | 474.9 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.059 | 590 | 745.7 | 2.5 | 748.2 | 474.8 | 474.8 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1138 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 5.96 | 0.059 | 590 | 745.9 | 2.5 | 748.4 | 474.6 | 474.647 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.11382 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 6.366 | 0.0676 | 675.7426 | 748.4 | 2.5 | 750.9 | 474.8 | 474.7713 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1139 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9736 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.061 | 0.0715 | 715.4958 | 782.6 | 2.5 | 785.1 | 489.4 | 489.3642 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3 | 0.11776 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8942 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.11 | 0.072 | 720 | 809.5 | 2.5 | 812 | 546.1 | 546.1095 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.12062 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.815 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.118 | 0.072 | 720 | 881.8 | 2.5 | 884.3 | 564.5 | 564.491 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3836 | 0.12801 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8641 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.12 | 0.072 | 720 | 892.4 | 2.5 | 894.9 | 568.4 | 568.4471 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1291 | 0.003116 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9434 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.136 | 0.072 | 720 | 892.5 | 2.5 | 895 | 570.1 | 570.1263 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.12912 | 0.0032 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.157 | 0.072 | 720 | 893.1 | 2.5 | 895.6 | 570.5 | 570.5 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.2164 | 0.12922 | 0.0032 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.17 | 0.072 | 720 | 893.4 | 2.5 | 895.9 | 570.5 | 570.5082 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.12932 | 0.003216 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.17 | 0.072 | 720 | 894.2 | 2.5 | 896.7 | 570.5 | 570.5301 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1294 | 0.0033 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8791 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.179 | 0.072 | 720 | 894.5 | 2.5 | 897 | 570.7 | 570.7 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1295 | 0.0033 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8395 | 26.3 | 2.5
2008-05-07 | 9:29 | 7.18 | 0.072 | 720 | 894.4 | 2.4 | 896.8 | 570.5 | 570.5262 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1295 | 0.0033 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9371 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.6 | 2.4 | 897 | 570.5 | 570.52 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1295 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0963 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.5 | 2.4 | 896.9 | 570.5 | 570.5312 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.12952 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9452 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.5 | 2.4 | 896.9 | 570.5 | 570.4795 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1296 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 27.9 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.4 | 2.4 | 896.8 | 570.6 | 570.6249 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1296 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9355 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.4 | 2.4 | 896.8 | 570.6 | 570.6 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1296 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.9 | 2.4 | 897.3 | 570.6 | 570.6277 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1297 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 894.8 | 2.4 | 897.2 | 570.3 | 570.3365 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1297 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7792 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 895.3 | 2.4 | 897.7 | 570.3 | 570.3188 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1298 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8586 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 895.6 | 2.4 | 898 | 570.4 | 570.3534 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1298 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 895.9 | 2.4 | 898.3 | 570.4 | 570.3741 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.1298 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0318 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896 | 2.4 | 898.4 | 570.4 | 570.366 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.12982 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.2713 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.1 | 2.4 | 898.5 | 570.3 | 570.2683 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1299 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0209 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.3 | 2.4 | 898.7 | 570.3 | 570.3 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1299 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.6 | 2.4 | 899 | 570.3 | 570.2812 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1299 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.0715 | 714.5455 | 896.6 | 2.4 | 899 | 570.2 | 570.2245 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.12992 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 27.8774 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.2 | 570.2 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6397 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.5 | 570.5218 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 27.7979 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.3 | 570.3315 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9452 | 26.2548 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.8 | 2.4 | 899.2 | 570.3 | 570.3269 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 28.0732 | 26.2244 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.4 | 570.4283 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.2893 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.6 | 570.6194 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2164 | 0.13 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0514 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.5 | 570.5481 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.003616 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2378 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.6 | 570.6076 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2414 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.6 | 570.5746 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.8 | 2.4 | 899.2 | 570.4 | 570.3807 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.5 | 570.514 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.0715 | 715.4545 | 896.6 | 2.4 | 899 | 570.5 | 570.4503 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.6 | 2.4 | 899 | 570.6 | 570.6364 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.5 | 2.4 | 898.9 | 570.4 | 570.3727 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.3 | 2.4 | 898.7 | 570.3 | 570.3082 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.18 | 0.072 | 720 | 896.3 | 2.4 | 898.7 | 570.3 | 570.329 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 27.9247 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.187 | 0.072 | 720 | 896.6 | 2.4 | 899 | 570.3 | 570.3 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192.2 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.2 | 0.0715 | 714.5455 | 896.7 | 2.4 | 899.1 | 570.2 | 570.1731 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192.8 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8662 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.2 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.2 | 570.1681 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8261 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.2 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.3 | 570.3394 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9846 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.2 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.3 | 570.3 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0711 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.195 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.5 | 570.5 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0037 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 28.0495 | 26.2495 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.5 | 2.4 | 898.9 | 570.2 | 570.2206 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.003716 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.23 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.6 | 2.4 | 899 | 570.5 | 570.5371 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0092 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.7 | 2.4 | 899.1 | 570.3 | 570.3266 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0884 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.0712 | 711.725 | 896.7 | 2.4 | 899.1 | 570.5 | 570.4526 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0323 | 26.2323 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.072 | 720 | 896.7 | 2.4 | 899.1 | 570.5 | 570.4746 | 0 | 13 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.247 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.072 | 720 | 896.5 | 2.4 | 898.9 | 570.5 | 570.516 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.072 | 720 | 896.6 | 2.4 | 899 | 570.5 | 570.5 | 0 | 13 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1301 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.0715 | 714.7596 | 896.5 | 2.4 | 898.9 | 570.4 | 570.4326 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13012 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.3 | 2.4 | 898.7 | 570.7 | 570.6755 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1302 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9361 | 26.3 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.2 | 2.4 | 898.6 | 570.5 | 570.5194 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1302 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9 | 26.2569 | 2.4
2008-05-07 | 9:30 | 7.19 | 0.071 | 710 | 896.5 | 2.4 | 898.9 | 570.6 | 570.6 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1302 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 27.9224 | 26.2224 | 2.4
2008-05-07 | 9:30 | 7.561 | 0.071 | 710 | 896.8 | 2.4 | 899.2 | 570.4 | 570.3866 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1302 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 8.056 | 0.0797 | 796.9544 | 905.4 | 2.4 | 907.8 | 571.3 | 571.2766 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13106 | 0.0038 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 28 | 26.3 | 2.4
2008-05-07 | 9:30 | 8.37 | 0.0825 | 824.7591 | 979.6 | 2.4169 | 982 | 625 | 625.0365 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.13809 | 0.0039 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0602 | 26.3 | 2.4
2008-05-07 | 9:30 | 8.37 | 0.081 | 810 | 1023.2 | 2.5 | 1025.7 | 657.3 | 657.3376 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14215 | 0.0039 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1394 | 26.2606 | 2.5
2008-05-07 | 9:30 | 8.37 | 0.081 | 810 | 1035.3 | 2.5 | 1037.8 | 663.6 | 663.621 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14333 | 0.0039 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.163 | 26.2 | 2.5
2008-05-07 | 9:30 | 8.377 | 0.081 | 810.0167 | 1038.3 | 2.6 | 1040.9 | 665.3 | 665.253 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14362 | 0.003916 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0043 | 26.2 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.0827 | 826.7112 | 1040.7 | 2.6 | 1043.3 | 665.7 | 665.6763 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14392 | 0.004 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1542 | 26.2 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.083 | 830 | 1043 | 2.6 | 1045.6 | 666 | 666.0221 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1441 | 0.004 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1438 | 26.2 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.0835 | 835.1472 | 1043.3 | 2.6 | 1045.9 | 666.1 | 666.1 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1442 | 0.0041 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0646 | 26.2354 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.0836 | 836.4213 | 1043.5 | 2.6 | 1046.1 | 666.1 | 666.1 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1442 | 0.0041 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9852 | 26.3 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.0832 | 831.942 | 1043.4 | 2.6 | 1046 | 666.2 | 666.1743 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14422 | 0.004116 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9059 | 26.3 | 2.6
2008-05-07 | 9:30 | 8.38 | 0.084 | 840 | 1043.7 | 2.6 | 1046.3 | 666 | 666.0083 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1443 | 0.0042 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9 | 26.3 | 2.6
2008-05-07 | 9:31 | 8.38 | 0.0832 | 831.6736 | 1043.7 | 2.6 | 1046.3 | 666 | 666.0478 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1443 | 0.0042 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8474 | 26.3 | 2.6
2008-05-07 | 9:31 | 8.38 | 0.083 | 830 | 1043.6 | 2.7 | 1046.3 | 666.2 | 666.2 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1443 | 0.0042 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8634 | 26.3 | 2.7
2008-05-07 | 9:31 | 8.38 | 0.083 | 830 | 1043.5 | 2.7 | 1046.2 | 666.2 | 666.2 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1443 | 0.0042 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0213 | 26.2893 | 2.7
2008-05-07 | 9:31 | 8.38 | 0.083 | 830 | 1043.3 | 2.7 | 1046 | 666.2 | 666.2 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.1443 | 0.0042 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1806 | 26.2097 | 2.7
2008-05-07 | 9:31 | 8.38 | 0.083 | 830 | 1043.3 | 2.7 | 1046 | 666.2 | 666.2302 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.3391 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.38 | 0.083 | 830 | 1043.5 | 2.7 | 1046.2 | 666.3 | 666.3 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.2536 | 26.1512 | 2.7
2008-05-07 | 9:31 | 8.385 | 0.083 | 830 | 1043.7 | 2.7 | 1046.4 | 666.3 | 666.328 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0719 | 26.1 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1043.7 | 2.7 | 1046.4 | 666.6 | 666.6438 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0073 | 26.1073 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1043.7 | 2.7 | 1046.4 | 666.6 | 666.6416 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 28.0866 | 26.1866 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.1 | 2.7 | 1046.8 | 666.8 | 666.7737 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1444 | 0.0043 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9026 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.398 | 0.083 | 830 | 1044.2 | 2.7 | 1046.9 | 666.9 | 666.8917 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 27.89 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.083 | 830 | 1044.2 | 2.7 | 1046.9 | 666.8 | 666.8234 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.083 | 830 | 1044.2 | 2.7 | 1046.9 | 666.9 | 666.9 | 0 | 13 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.0823 | 823.3583 | 1044.5 | 2.7 | 1047.2 | 666.9 | 666.914 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.402 | 0.082 | 820 | 1044.7 | 2.7 | 1047.4 | 666.7 | 666.7053 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995 | 28.1239 | 26.138 | 2.7
2008-05-07 | 9:31 | 8.41 | 0.082 | 820 | 1044.7 | 2.7 | 1047.4 | 666.7 | 666.7 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3164 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995 | 28.2 | 26.1412 | 2.7
2008-05-07 | 9:31 | 8.41 | 0.082 | 820 | 1044.7 | 2.7 | 1047.4 | 666.5 | 666.5282 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995.1623 | 28.1591 | 26.1796 | 2.7
2008-05-07 | 9:31 | 8.41 | 0.082 | 820 | 1044.4 | 2.7 | 1047.1 | 666.5 | 666.5487 | 0 | 12.9 | 953 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3165 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.0006 | 26.1003 | 2.7
2008-05-07 | 9:31 | 8.41 | 0.082 | 820 | 1044.4 | 2.7 | 1047.1 | 666.7 | 666.6741 | 0 | 12.9 | 953 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.0789 | 26.1789 | 2.7
2008-05-07 | 9:31 | 8.41 | 0.0827 | 826.6389 | 1044.5 | 2.7 | 1047.2 | 666.8 | 666.7786 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1445 | 0.0044 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.1 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.408 | 0.083 | 830 | 1044.5 | 2.7 | 1047.2 | 666.7 | 666.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.14452 | 0.004416 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.1 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.083 | 830 | 1044.6 | 2.7 | 1047.3 | 666.6 | 666.6 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.0833 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.083 | 830 | 1044.8 | 2.7 | 1047.5 | 666.5 | 666.5395 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995 | 28.0043 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.407 | 0.083 | 830 | 1045 | 2.7 | 1047.7 | 666.3 | 666.3271 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995 | 27.9249 | 26.2751 | 2.7
2008-05-07 | 9:31 | 8.405 | 0.083 | 830 | 1045.1 | 2.7 | 1047.8 | 666.4 | 666.4048 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.1178 | 26.2456 | 2.7
2008-05-07 | 9:31 | 8.4 | 0.083 | 830 | 1044.8 | 2.7 | 1047.5 | 666.6 | 666.5718 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28.199 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.398 | 0.083 | 830 | 1044.6 | 2.7 | 1047.3 | 666.2 | 666.2 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.5 | 2.7 | 1047.2 | 666.4 | 666.3762 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.5 | 2.7 | 1047.2 | 666.4 | 666.4 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.8577 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.6 | 2.7 | 1047.3 | 666.6 | 666.6 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1446 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.8 | 26.2505 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.6 | 2.7 | 1047.3 | 666.5 | 666.4709 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0045 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996.8363 | 27.8299 | 26.2701 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.8 | 2.7 | 1047.5 | 666.7 | 666.6714 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.8 | 2.7 | 1047.5 | 666.6 | 666.6132 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.8 | 2.7 | 1047.5 | 666.5 | 666.5448 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.9 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.7 | 2.7 | 1047.4 | 666.6 | 666.5541 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30.73 | 53 | 1032 | 996.164 | 27.9 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.5 | 2.7 | 1047.2 | 666.5 | 666.5192 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30.27 | 53 | 1032 | 997 | 27.8739 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.083 | 830 | 1044.4 | 2.7 | 1047.1 | 666.7 | 666.7 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.7947 | 26.2053 | 2.7
2008-05-07 | 9:31 | 8.39 | 0.0837 | 836.6694 | 1044.3 | 2.7 | 1047 | 666.8 | 666.8284 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.7154 | 26.2846 | 2.7
2008-05-07 | 9:31 | 8.395 | 0.084 | 840 | 1044.3 | 2.7 | 1047 | 666.9 | 666.8781 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.8916 | 26.2361 | 2.7
2008-05-07 | 9:31 | 8.42 | 0.0837 | 836.5587 | 1044.3 | 2.7 | 1047 | 667.1 | 667.0737 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.42 | 0.0832 | 831.5833 | 1044.5 | 2.7 | 1047.2 | 667 | 666.9685 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.42 | 0.084 | 839.9167 | 1044.9 | 2.7 | 1047.6 | 667 | 667.0191 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.2 | 2.7
2008-05-07 | 9:31 | 8.42 | 0.0832 | 831.75 | 1045.1 | 2.6 | 1047.7 | 667.2 | 667.1733 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 8.42 | 0.083 | 830 | 1045.3 | 2.6 | 1047.9 | 666.6 | 666.6276 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995.836 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 8.42 | 0.083 | 830 | 1045.5 | 2.6 | 1048.1 | 666.6 | 666.647 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1447 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 995 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 8.42 | 0.083 | 830 | 1045.4 | 2.6 | 1048 | 667 | 666.987 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1448 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 8.42 | 0.083 | 830 | 1045.2 | 2.6 | 1047.8 | 666.4 | 666.4177 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.2164 | 0.1448 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 8.42 | 0.083 | 830 | 1045 | 2.6 | 1047.6 | 666.5 | 666.5254 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1448 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032 | 997 | 27.8458 | 26.2771 | 2.6
2008-05-07 | 9:31 | 8.445 | 0.083 | 830 | 1045 | 2.6 | 1047.6 | 666.5 | 666.4706 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.1448 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1032.4814 | 997 | 28.0808 | 26.2438 | 2.6
2008-05-07 | 9:31 | 8.822 | 0.0955 | 954.8256 | 1046 | 2.6 | 1048.6 | 666.7 | 666.6531 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.14488 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1939 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.56 | 0.0986 | 985.7868 | 1104.2 | 2.6 | 1106.8 | 694.8 | 694.8176 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.14994 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.56 | 0.095 | 950 | 1173.5 | 2.6 | 1176.1 | 746.2 | 746.1667 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1558 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.576 | 0.095 | 950 | 1179.3 | 2.6 | 1181.9 | 758.7 | 758.6632 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15628 | 0.0046 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.58 | 0.095 | 950 | 1188 | 2.6 | 1190.6 | 761.9 | 761.934 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15692 | 0.0047 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.1051 | 26.1475 | 2.6
2008-05-07 | 9:31 | 9.585 | 0.095 | 950 | 1187.9 | 2.6 | 1190.5 | 762.6 | 762.5594 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.157 | 0.0047 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1368 | 26.1316 | 2.6
2008-05-07 | 9:31 | 9.59 | 0.095 | 950 | 1188.1 | 2.6 | 1190.7 | 762.9 | 762.8706 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15702 | 0.004716 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 28.0104 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.59 | 0.095 | 950 | 1189.6 | 2.6 | 1192.2 | 763 | 763 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1572 | 0.0048 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0902 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.59 | 0.095 | 950 | 1191 | 2.6 | 1193.6 | 763.1 | 763.1049 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1573 | 0.0048 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.1695 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.59 | 0.095 | 950 | 1190.9 | 2.6 | 1193.5 | 763.1 | 763.1297 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15732 | 0.0048 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0061 | 26.2 | 2.6
2008-05-07 | 9:31 | 9.597 | 0.095 | 950 | 1191.2 | 2.6 | 1193.8 | 763.2 | 763.1713 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1574 | 0.0049 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8551 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1191.6 | 2.6 | 1194.2 | 763.3 | 763.3 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15742 | 0.0049 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.3 | 2.6 | 1194.9 | 763.4 | 763.3723 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1575 | 0.0049 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.5 | 2.6 | 1195.1 | 763.6 | 763.5692 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15752 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.1314 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.7 | 2.6 | 1195.3 | 763.5 | 763.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1576 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.1101 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.8 | 2.6 | 1195.4 | 763.4 | 763.3731 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1576 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0725 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.7 | 2.6 | 1195.3 | 763.3 | 763.3158 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15762 | 0.0051 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.2898 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.3 | 2.6 | 1194.9 | 763.4 | 763.4151 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1577 | 0.0051 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0519 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.2 | 2.6 | 1194.8 | 763.4 | 763.4311 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1577 | 0.0051 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.3 | 2.6 | 1194.9 | 763.5 | 763.4907 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15772 | 0.0052 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.1 | 2.6 | 1194.7 | 763.3 | 763.3328 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1578 | 0.0052 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1192.8 | 2.6 | 1195.4 | 763.5 | 763.5471 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.15782 | 0.0052 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.607 | 0.095 | 950 | 1194 | 2.6 | 1196.6 | 763.7 | 763.678 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.158 | 0.0052 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 27.7634 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.605 | 0.095 | 950 | 1194 | 2.6 | 1196.6 | 763.8 | 763.7994 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.158 | 0.005216 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997.8357 | 27.8742 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.2 | 2.6 | 1196.8 | 763.8 | 763.8442 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3164 | 0.158 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.112 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.2 | 2.6 | 1196.8 | 763.7 | 763.6518 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.158 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.2173 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.4 | 2.6 | 1197 | 763.7 | 763.6723 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.158 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8218 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.2 | 2.6 | 1196.8 | 763.9 | 763.8567 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.158 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9501 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.7 | 763.7 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2544 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.7 | 763.7 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0053 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9664 | 26.2664 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.6 | 763.5552 | 0 | 12.9 | 952 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.005316 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9128 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.5 | 763.5249 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.9921 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.5 | 763.4706 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.1 | 2.6 | 1196.7 | 763.9 | 763.8524 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0504 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.607 | 0.095 | 950 | 1194.1 | 2.6 | 1196.7 | 763.9 | 763.9 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.1 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.1 | 2.6 | 1196.7 | 763.8 | 763.8 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.091 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.3 | 2.6 | 1196.9 | 763.7 | 763.7205 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 28.0121 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.5 | 2.6 | 1197.1 | 763.8 | 763.7741 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.6 | 2.6 | 1197.2 | 763.8 | 763.8 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0935 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.6 | 2.6 | 1197.2 | 763.7 | 763.725 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.226 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.8 | 2.6 | 1197.4 | 763.7 | 763.6732 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1581 | 0.0054 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.2948 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.605 | 0.095 | 950 | 1194.9 | 2.6 | 1197.5 | 763.7 | 763.705 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.005416 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.2155 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195 | 2.6 | 1197.6 | 763.7 | 763.7269 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 30.7293 | 53 | 1033 | 996 | 28.2 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195.1 | 2.6 | 1197.7 | 763.5 | 763.5 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 28.1149 | 26.2425 | 2.6
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1194.8 | 2.6 | 1197.4 | 763.6 | 763.5758 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0666 | 26.2778 | 2.6
2008-05-07 | 9:32 | 9.608 | 0.095 | 950 | 1194.8 | 2.5836 | 1197.4 | 763.4 | 763.4 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.2955 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.603 | 0.095 | 950 | 1194.9 | 2.5 | 1197.4 | 763.7 | 763.6633 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0578 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195.2 | 2.5 | 1197.7 | 763.6 | 763.5538 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.1199 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195.2 | 2.5 | 1197.7 | 763.7 | 763.6728 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.1609 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195.4 | 2.5 | 1197.9 | 763.5 | 763.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.1184 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.6 | 0.095 | 950 | 1195.6 | 2.6 | 1198.2 | 763.5 | 763.526 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.1976 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.608 | 0.095 | 950 | 1195.4 | 2.5164 | 1197.9 | 763.5 | 763.5258 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.0463 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1195.5 | 2.5836 | 1198.1 | 763.3 | 763.3165 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2 | 2.6
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1195.2 | 2.5 | 1197.7 | 763.5 | 763.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1582 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.1623 | 27.8942 | 26.2353 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.9 | 2.5 | 1197.4 | 763.3 | 763.3443 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.15822 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7293 | 26.2853 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.9 | 2.5 | 1197.4 | 763.4 | 763.3562 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.1583 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8878 | 26.2061 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.7 | 2.5 | 1197.2 | 762.8 | 762.7533 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1583 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 28.1194 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.7 | 2.5 | 1197.2 | 763.5 | 763.5194 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1583 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 28.1476 | 26.2 | 2.5
2008-05-07 | 9:32 | 9.61 | 0.095 | 950 | 1194.7 | 2.5 | 1197.2 | 763.4 | 763.3686 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1583 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.9738 | 26.2315 | 2.5
2008-05-07 | 9:32 | 9.936 | 0.095 | 950 | 1195 | 2.5 | 1197.5 | 763.6 | 763.5727 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1583 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7319 | 26.2894 | 2.5
2008-05-07 | 9:32 | 10.44 | 0.103 | 1030.0648 | 1206.3 | 2.5 | 1208.8 | 764.2 | 764.2179 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1589 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.97 | 26.21 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.1065 | 1064.9787 | 1281.2 | 2.5 | 1283.7 | 813.3 | 813.2542 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.16506 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.2078 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1326 | 2.5 | 1328.5 | 847.3 | 847.2611 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.16837 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0583 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1339.1 | 2.5 | 1341.6 | 856.8 | 856.7855 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1694 | 0.0055 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8551 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.793 | 0.106 | 1060 | 1341.4 | 2.5 | 1343.9 | 858.8 | 858.8367 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.16962 | 0.005516 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.9788 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.795 | 0.106 | 1060 | 1343.7 | 2.5 | 1346.2 | 859.3 | 859.2541 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.16972 | 0.0056 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7411 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1344.4 | 2.5 | 1346.9 | 859.5 | 859.5 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1698 | 0.0056 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8966 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1344.7 | 2.5 | 1347.2 | 859.5 | 859.5276 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1698 | 0.0056 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0448 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1344.7 | 2.5 | 1347.2 | 859.8 | 859.7741 | 0 | 12.9 | 952 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1699 | 0.0056 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0761 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1344.7 | 2.5 | 1347.2 | 859.6 | 859.6384 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1699 | 0.0057 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.5
2008-05-07 | 9:32 | 10.8 | 0.106 | 1060 | 1345.2 | 2.5 | 1347.7 | 859.6 | 859.6 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1699 | 0.0057 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.106 | 1060 | 1345.4 | 2.5 | 1347.9 | 859.7 | 859.6673 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.16992 | 0.0057 | 0.0002 | 0.0001 |  | n/a |  | 30.27 | 53 | 1033 | 996 | 28.1233 | 26.1384 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.1067 | 1066.5228 | 1345.7 | 2.5 | 1348.2 | 859.6 | 859.5707 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.17 | 0.0057 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.1182 | 26.1 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.6 | 2.5 | 1348.1 | 859.8 | 859.8494 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.17 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 28.0202 | 26.1202 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.9 | 2.5 | 1348.4 | 859.9 | 859.8888 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.17 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 30.7293 | 53 | 1033 | 996 | 28.0994 | 26.1994 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.8 | 2.5 | 1348.3 | 859.8 | 859.8 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.17 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28.0213 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.7 | 2.5 | 1348.2 | 859.7 | 859.7 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3836 | 0.17 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 28.1158 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.6 | 2.5 | 1348.1 | 859.9 | 859.9496 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3 | 0.17 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.2743 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.6 | 2.5 | 1348.1 | 859.9 | 859.9 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.3672 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.4 | 2.5 | 1347.9 | 859.8 | 859.8 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.2089 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.5 | 2.5 | 1348 | 859.8 | 859.8 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.2748 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.6 | 2.4836 | 1348.1 | 859.8 | 859.8 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.0058 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.1919 | 26.2 | 2.5
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.9 | 2.4 | 1348.3 | 859.9 | 859.9 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.005816 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0667 | 26.2 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.8 | 2.4 | 1348.2 | 859.8 | 859.8483 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1701 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1345.8 | 2.4 | 1348.2 | 860 | 860 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3833 | 0.1701 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1346.1 | 2.4 | 1348.5 | 860.1 | 860.1332 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.17012 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.1419 | 26.2 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1346.1 | 2.4 | 1348.5 | 860.6 | 860.6212 | 0 | 12.9 | 952 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1702 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0996 | 26.2 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1346.1 | 2.3836 | 1348.5 | 860.5 | 860.5299 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1702 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9705 | 26.2295 | 2.4
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1346.4 | 2.3 | 1348.7 | 860.5 | 860.4762 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1702 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9261 | 26.2826 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.1078 | 1078.2847 | 1347.2 | 2.3 | 1349.5 | 860.4 | 860.4 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.17022 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.1634 | 26.1244 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.108 | 1080 | 1348.8 | 2.3 | 1351.1 | 860.6 | 860.5713 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1704 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0657 | 26.1671 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.1075 | 1075 | 1349.4 | 2.3 | 1351.7 | 860.6 | 860.6207 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1704 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.4 | 2.3 | 1351.7 | 860.6 | 860.6292 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1704 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.4 | 2.3 | 1351.7 | 860.9 | 860.8825 | 0 | 13 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1704 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0197 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.6 | 2.3 | 1351.9 | 860.8 | 860.8253 | 0 | 13 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1704 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.3366 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.6 | 2.3 | 1351.9 | 860.8 | 860.8267 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.17042 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.1474 | 26.2631 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.7 | 2.3 | 1352 | 860.3 | 860.3123 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0424 | 26.2576 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.107 | 1070 | 1349.9 | 2.3 | 1352.2 | 860.3 | 860.3299 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 28.0781 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.1073 | 1073.0283 | 1350.2 | 2.3 | 1352.5 | 859.9 | 859.9277 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2011 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.108 | 1080 | 1350.2 | 2.3 | 1352.5 | 859.9 | 859.8751 | 0 | 12.9 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.28 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.108 | 1080 | 1350.2 | 2.3 | 1352.5 | 860.1 | 860.1439 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.0059 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2406 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.108 | 1080 | 1350 | 2.3 | 1352.3 | 860 | 860 | 0 | 12.9 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.8 | 0.1073 | 1073.4445 | 1349.8 | 2.2836 | 1352.1 | 859.8 | 859.82 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.3
2008-05-07 | 9:33 | 10.797 | 0.107 | 1070 | 1349.9 | 2.2 | 1352.1 | 859.9 | 859.9394 | 0 | 12.9 | 951 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.2836 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2 | 2.2
2008-05-07 | 9:33 | 10.781 | 0.107 | 1070 | 1349.6 | 2.2 | 1351.8 | 860 | 860.008 | 0 | 12.9 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.2765 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.107 | 1070 | 1349.6 | 2.2 | 1351.8 | 859.8 | 859.8 | 0 | 13 | 951 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 28.0554 | 26.2446 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.107 | 1070 | 1349.5 | 2.2 | 1351.7 | 859.8 | 859.8267 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 28.1 | 26.2346 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.107 | 1070 | 1349.4 | 2.2 | 1351.6 | 859.5 | 859.5095 | 0 | 13 | 951 | 956 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 28.0858 | 26.2858 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.1063 | 1063.4734 | 1349.6 | 2.2 | 1351.8 | 859.8 | 859.7501 | 0 | 12.9 | 951 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28.0066 | 26.2066 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.5 | 2.2 | 1351.7 | 859.3 | 859.3303 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 28.1453 | 26.2 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.3 | 2.2 | 1351.5 | 859.6 | 859.6324 | 0 | 12.9 | 950 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 28.0962 | 26.2519 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.3 | 2.2 | 1351.5 | 859.7 | 859.7 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 28 | 26.2377 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.3 | 2.2 | 1351.5 | 859.6 | 859.6 | 0 | 12.9 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28 | 26.1104 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.3 | 2.2 | 1351.5 | 859.5 | 859.518 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28 | 26.1896 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.1 | 2.2 | 1351.3 | 859.5 | 859.4696 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.0689 | 26.2 | 2.2
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.1 | 2.1 | 1351.2 | 859.6 | 859.6 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.0523 | 26.2 | 2.1
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.3 | 2.1 | 1351.4 | 859.4 | 859.4261 | 0 | 13 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.9733 | 26.2 | 2.1
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.5 | 2.1 | 1351.6 | 859.7 | 859.7415 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9065 | 26.2 | 2.1
2008-05-07 | 9:33 | 10.78 | 0.106 | 1060 | 1349.7 | 2.1 | 1351.8 | 859.6 | 859.591 | 0 | 13 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1705 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9857 | 26.2 | 2.1
2008-05-07 | 9:33 | 10.947 | 0.106 | 1060 | 1349.6 | 2.1 | 1351.7 | 859.6 | 859.5715 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.17052 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.265 | 2.1
2008-05-07 | 9:33 | 11.368 | 0.1157 | 1157.2246 | 1350.7 | 2.1 | 1352.8 | 859.4 | 859.4 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.17065 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 28.1326 | 26.2558 | 2.1
2008-05-07 | 9:33 | 11.931 | 0.119 | 1190.4813 | 1383.7 | 2.1 | 1385.8 | 896.1 | 896.0953 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3164 | 0.17296 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.2297 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.94 | 0.117 | 1170 | 1471.8 | 2.1 | 1473.9 | 941.6 | 941.585 | 0 | 13 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1791 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0026 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.94 | 0.117 | 1170 | 1492.9 | 2.1 | 1495 | 952.9 | 952.8536 | 0 | 12.9 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.18053 | 0.006 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996.164 | 28.0819 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.94 | 0.117 | 1170 | 1496.6 | 2.1 | 1498.7 | 955.6 | 955.5528 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1808 | 0.006016 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.1611 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.945 | 0.117 | 1170 | 1496.6 | 2.1 | 1498.7 | 956.5 | 956.4558 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.18082 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.079 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.953 | 0.117 | 1170 | 1497.4 | 2.1 | 1499.5 | 956.7 | 956.71 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3164 | 0.1809 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9587 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.96 | 0.117 | 1170 | 1497.5 | 2.1 | 1499.6 | 956.8 | 956.7727 | 0 | 13 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.18092 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.1637 | 28.1964 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.96 | 0.117 | 1170 | 1497.4 | 2.1 | 1499.5 | 957 | 957 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.181 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.044 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.96 | 0.117 | 1170 | 1497.2 | 2.1 | 1499.3 | 957.1 | 957.0584 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.181 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28 | 26.2 | 2.1
2008-05-07 | 9:33 | 11.967 | 0.117 | 1170 | 1497.1 | 2.1 | 1499.2 | 957.3 | 957.2745 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.181 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28 | 26.2365 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.117 | 1170 | 1497.4 | 2.1164 | 1499.5 | 957.5 | 957.472 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.181 | 0.006216 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.8363 | 28 | 26.2843 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.117 | 1170 | 1497.6 | 2.2 | 1499.8 | 957.7 | 957.6732 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.18102 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2053 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.117 | 1170 | 1497.8 | 2.2 | 1500 | 957.5 | 957.4564 | 0 | 13 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.1811 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.1173 | 1172.8908 | 1497.9 | 2.2 | 1500.1 | 957.4 | 957.3674 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1811 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.118 | 1180 | 1497.9 | 2.2 | 1500.1 | 957.6 | 957.5605 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.1811 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8696 | 26.2326 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.118 | 1180 | 1498.1 | 2.2 | 1500.3 | 957.2 | 957.1562 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.18112 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6469 | 26.2883 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.118 | 1180 | 1498 | 2.2 | 1500.2 | 957.1 | 957.1 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1812 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.9642 | 26.209 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.1193 | 1192.9331 | 1497.9 | 2.2 | 1500.1 | 957.4 | 957.3741 | 0 | 13 | 949 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1812 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8594 | 26.2 | 2.2
2008-05-07 | 9:34 | 11.97 | 0.12 | 1200 | 1498.1 | 2.1 | 1500.2 | 957.3 | 957.3315 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.1812 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8496 | 26.2496 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.12 | 1200 | 1497.9 | 2.1 | 1500 | 957.1 | 957.1062 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.1812 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9576 | 26.2712 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.12 | 1200 | 1497.7 | 2.1 | 1499.8 | 957.4 | 957.4049 | 0 | 12.9 | 950 | 957 | 939 | 19.2 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.18122 | 0.006416 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.092 | 26.208 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.12 | 1200 | 1497.7 | 2.1 | 1499.8 | 957.4 | 957.3522 | 0 | 13 | 949 | 956 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1813 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.0131 | 26.2869 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.1192 | 1192.1177 | 1498.6 | 2.1 | 1500.7 | 957.3 | 957.2758 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1813 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9337 | 26.1673 | 2.1
2008-05-07 | 9:34 | 11.973 | 0.119 | 1190 | 1498.7 | 2.1 | 1500.8 | 957.3 | 957.3282 | 0 | 13 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.18132 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9 | 26.1457 | 2.1
2008-05-07 | 9:34 | 11.979 | 0.119 | 1190 | 1498.7 | 2.1 | 1500.8 | 957.1 | 957.1255 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8752 | 26.2248 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.119 | 1190 | 1498.8 | 2.1 | 1500.9 | 957.2 | 957.1807 | 0 | 12.9 | 949 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8079 | 26.296 | 2.1
2008-05-07 | 9:34 | 11.978 | 0.119 | 1190 | 1498.8 | 2.1 | 1500.9 | 957.3 | 957.2736 | 0 | 12.9 | 949 | 956 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9667 | 26.2167 | 2.1
2008-05-07 | 9:34 | 11.974 | 0.119 | 1190 | 1499.1 | 2.1 | 1501.2 | 957.9 | 957.8788 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7496 | 26.2626 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.119 | 1190 | 1499.4 | 2.1 | 1501.5 | 957.7 | 957.6721 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7674 | 26.2582 | 2.1
2008-05-07 | 9:34 | 11.97 | 0.119 | 1190 | 1499.6 | 2.1 | 1501.7 | 957.4 | 957.3811 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9789 | 26.2211 | 2.1
2008-05-07 | 9:34 | 11.971 | 0.119 | 1190 | 1499.8 | 2.1 | 1501.9 | 957.1 | 957.1136 | 0 | 13 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9003 | 26.2997 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500 | 2.1 | 1502.1 | 957.3 | 957.3278 | 0 | 12.9 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9792 | 26.2208 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500 | 2.1 | 1502.1 | 957.4 | 957.3692 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2586 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500 | 2.1 | 1502.1 | 957.4 | 957.4 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0759 | 26.262 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500 | 2.1 | 1502.1 | 957.4 | 957.4 | 0 | 13 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1814 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.1656 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1499.8 | 2.1 | 1501.9 | 957.4 | 957.4 | 0 | 12.9 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0071 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.1185 | 1185.067 | 1500 | 2.1 | 1502.1 | 957.3 | 957.2658 | 0 | 12.9 | 950 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2756 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.118 | 1180 | 1499.8 | 2.1 | 1501.9 | 957.1 | 957.0745 | 0 | 12.9 | 950 | 957 | 939 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2451 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.1181 | 1181.2941 | 1499.8 | 2.1 | 1501.9 | 957.2 | 957.1747 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9659 | 26.2341 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1189.6975 | 1499.8 | 2.1 | 1501.9 | 957.1 | 957.1 | 0 | 12.9 | 950 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.8601 | 26.2867 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1499.8 | 2.1 | 1501.9 | 957.4 | 957.4 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.1637 | 27.6223 | 26.2074 | 2.1
2008-05-07 | 9:34 | 11.983 | 0.119 | 1190 | 1499.6 | 2.1 | 1501.7 | 957.1 | 957.0909 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3164 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8142 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1499.6 | 2.1 | 1501.7 | 957 | 957 | 0 | 12.9 | 950 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9507 | 26.2507 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500 | 2.1 | 1502.1 | 957.1 | 957.0865 | 0 | 12.9 | 950 | 957 | 939 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2698 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.1 | 2.1 | 1502.2 | 957.1 | 957.1 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1815 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9906 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.4 | 2.1 | 1502.5 | 957.2 | 957.2 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9113 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.7 | 2.1 | 1502.8 | 957.1 | 957.1297 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0358 | 26.2679 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.8 | 2.1 | 1502.9 | 957 | 956.9686 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0529 | 26.2529 | 2.1
2008-05-07 | 9:34 | 11.988 | 0.119 | 1190 | 1500.7 | 2.1 | 1502.8 | 956.9 | 956.9 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9208 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500.5 | 2.1 | 1502.6 | 957 | 956.9783 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7169 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500.5 | 2.1 | 1502.6 | 956.8 | 956.8 | 0 | 12.9 | 950 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9544 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500.6 | 2.1 | 1502.7 | 957 | 957.0496 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 28 | 26.2636 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500.4 | 2.1 | 1502.5 | 957.1 | 957.0673 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.0429 | 26.2571 | 2.1
2008-05-07 | 9:34 | 11.983 | 0.119 | 1190 | 1500.2 | 2.1 | 1502.3 | 957 | 957.021 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28.0775 | 26.2 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.4 | 2.1 | 1502.5 | 957.2 | 957.2318 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3164 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9983 | 26.2017 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.7 | 2.1 | 1502.8 | 956.9 | 956.9266 | 0 | 12.9 | 950 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.9191 | 26.2809 | 2.1
2008-05-07 | 9:34 | 11.99 | 0.119 | 1190 | 1500.5 | 2.1 | 1502.6 | 957 | 956.9709 | 0 | 12.9 | 950 | 957 | 940 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 28.0203 | 26.3 | 2.1
2008-05-07 | 9:34 | 11.983 | 0.119 | 1190 | 1500.8 | 2.1 | 1502.9 | 957 | 956.9539 | 0 | 12.9 | 950 | 956 | 939 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0606 | 26.3 | 2.1
2008-05-07 | 9:34 | 11.98 | 0.119 | 1190 | 1500.8 | 2.1 | 1502.9 | 957 | 957 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1816 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.6316 | 997 | 27.9441 | 26.3 | 2.1
2008-05-07 | 9:34 | 11.634 | 0.1094 | 1093.7489 | 1500.5 | 2.1 | 1502.6 | 956.9 | 956.9141 | 0 | 12.9 | 949 | 957 | 939 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1817 | 0.006716 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.7064 | 26.3 | 2.1
2008-05-07 | 9:34 | 10.857 | 0.1059 | 1059.3909 | 1447.8 | 2.1 | 1449.9 | 947 | 947.0007 | 0 | 12.9 | 950 | 957 | 939 | 19.8 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.17814 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.777 | 26.3 | 2.1
2008-05-07 | 9:34 | 10.841 | 0.1077 | 1076.8571 | 1386.8 | 2.1 | 1388.9 | 883.2 | 883.2355 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.17387 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.9118 | 26.3 | 2.1
2008-05-07 | 9:34 | 10.832 | 0.107 | 1070 | 1363.1 | 2.1 | 1365.2 | 865.1 | 865.1111 | 0 | 12.9 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.17215 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9299 | 26.3 | 2.1
2008-05-07 | 9:34 | 10.83 | 0.107 | 1070 | 1354.8 | 2.1 | 1356.9 | 861.3 | 861.2769 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.17158 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.8377 | 27.8147 | 26.2853 | 2.1
2008-05-07 | 9:34 | 10.825 | 0.107 | 1070 | 1353 | 2.1 | 1355.1 | 860.2 | 860.1609 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8939 | 26.2061 | 2.1
2008-05-07 | 9:34 | 10.82 | 0.107 | 1070 | 1352.9 | 2.1 | 1355 | 859.9 | 859.9276 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9 | 26.2 | 2.1
2008-05-07 | 9:34 | 10.82 | 0.107 | 1070 | 1352.8 | 2 | 1354.8 | 859.9 | 859.8777 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9524 | 26.2 | 2
2008-05-07 | 9:34 | 10.82 | 0.107 | 1070 | 1352.6 | 2 | 1354.6 | 859.9 | 859.9 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9684 | 26.2316 | 2
2008-05-07 | 9:34 | 10.812 | 0.107 | 1070 | 1352.8 | 2 | 1354.8 | 859.7 | 859.7289 | 0 | 12.9 | 948 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9108 | 26.2892 | 2
2008-05-07 | 9:34 | 10.81 | 0.107 | 1070 | 1352.8 | 2 | 1354.8 | 859.9 | 859.9419 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9901 | 26.2099 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1352.8 | 2 | 1354.8 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.2693 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1352.6 | 2 | 1354.6 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0963 | 26.2518 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1352.6 | 2 | 1354.6 | 859.7 | 859.6569 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.1728 | 26.2 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1352.5 | 2 | 1354.5 | 859.6 | 859.6459 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0793 | 26.2069 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1352.4 | 2 | 1354.4 | 859.5 | 859.4889 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26.2 | 5 | 195 | 59.3 | 0.1714 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.8416 | 26.2861 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1351.4 | 2 | 1353.4 | 859.5 | 859.5451 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 27 | 5 | 195 | 59.3 | 0.1713 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.9308 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1351.4 | 2 | 1353.4 | 859.6 | 859.6177 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26.8 | 5 | 195 | 59.3162 | 0.1713 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.866 | 26.2553 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1351.6 | 2 | 1353.6 | 859.6 | 859.5793 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1713 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7477 | 26.2238 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1351.6 | 2 | 1353.6 | 859.6 | 859.6313 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1713 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.6307 | 997 | 27.897 | 26.297 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1351.1 | 2 | 1353.1 | 859.5 | 859.5 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.17128 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8178 | 26.2177 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1350.5 | 2 | 1352.5 | 859.4 | 859.4 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1712 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8 | 26.2615 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1349.7 | 2 | 1351.7 | 859.3 | 859.3 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3171 | 0.1712 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8809 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1349.2 | 2 | 1351.2 | 859.3 | 859.3 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9225 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.1064 | 1063.634 | 1349 | 2 | 1351 | 859.2 | 859.1537 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.6035 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.2 | 859.1729 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.9137 | 26.2216 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.9 | 2 | 1350.9 | 859.2 | 859.1901 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3164 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.8271 | 26.2576 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.9 | 2 | 1350.9 | 859.3 | 859.3094 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3835 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.6632 | 26.2632 | 2
2008-05-07 | 9:35 | 10.816 | 0.106 | 1060 | 1348.9 | 2 | 1350.9 | 859.3 | 859.3276 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.2164 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.6482 | 26.2161 | 2
2008-05-07 | 9:35 | 10.82 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.2 | 859.2239 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.3658 | 998 | 27.8857 | 26.2952 | 2
2008-05-07 | 9:35 | 10.82 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859 | 859.0287 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9745 | 26.2255 | 2
2008-05-07 | 9:35 | 10.819 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.1 | 859.1 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8388 | 26.2537 | 2
2008-05-07 | 9:35 | 10.811 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.3 | 859.2586 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7327 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.1 | 859.0716 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7884 | 26.2884 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.2 | 859.2 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7087 | 26.2087 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.7 | 2 | 1350.7 | 859.1 | 859.1259 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.8412 | 26.2 | 2
2008-05-07 | 9:35 | 10.81 | 0.106 | 1060 | 1348.8 | 2 | 1350.8 | 859.2 | 859.2 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8003 | 26.2 | 2
2008-05-07 | 9:35 | 10.81 | 0.1065 | 1064.579 | 1348.8 | 2 | 1350.8 | 859.3 | 859.3124 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7582 | 26.2291 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.6 | 2 | 1350.6 | 859.6 | 859.6135 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9083 | 26.2917 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.6 | 2 | 1350.6 | 859.9 | 859.9452 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.9874 | 26.2126 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.6 | 2 | 1350.6 | 859.9 | 859.854 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7332 | 26.2667 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.3 | 2 | 1350.3 | 859.9 | 859.8723 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6918 | 26.2541 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7749 | 26.2251 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.8 | 859.8 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.7131 | 26.2956 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.1 | 2 | 1350.1 | 859.7 | 859.7 | 0 | 12.9 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9508 | 26.2164 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.1711 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8115 | 26.2628 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.1 | 2 | 1350.1 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7421 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.3 | 2 | 1350.3 | 859.8 | 859.8 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3164 | 0.171 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8426 | 26.3 | 2
2008-05-07 | 9:35 | 10.81 | 0.107 | 1070 | 1348.3 | 2 | 1350.3 | 859.7 | 859.7 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9981 | 26.3 | 2
2008-05-07 | 9:35 | 10.809 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.9 | 859.9353 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 31.73 | 53 | 1034 | 997 | 27.6814 | 26.3 | 2
2008-05-07 | 9:35 | 10.8 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.8 | 859.7815 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0068 | 0.0002 | 0.0001 |  | n/a |  | 32 | 53 | 1034 | 997 | 27.7768 | 26.2411 | 2
2008-05-07 | 9:35 | 10.808 | 0.107 | 1070 | 1348 | 2 | 1350 | 859.8 | 859.7754 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.9 | 26.2382 | 2
2008-05-07 | 9:35 | 10.791 | 0.107 | 1070 | 1347.8 | 2 | 1349.8 | 860.1 | 860.0508 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31.27 | 53 | 1033 | 997.1623 | 27.9174 | 26.3 | 2
2008-05-07 | 9:35 | 10.78 | 0.1062 | 1062.1817 | 1347.6 | 2 | 1349.6 | 859.8 | 859.827 | 0 | 12.9 | 949 | 957 | 938 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.9966 | 26.3 | 2
2008-05-07 | 9:35 | 10.783 | 0.106 | 1060 | 1347.9 | 2 | 1349.9 | 859.7 | 859.6816 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 28 | 26.3 | 2
2008-05-07 | 9:35 | 10.79 | 0.106 | 1060 | 1347.9 | 2 | 1349.9 | 859.5 | 859.5303 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3167 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8348 | 26.3 | 2
2008-05-07 | 9:35 | 10.79 | 0.106 | 1060 | 1347.9 | 2 | 1349.9 | 859.2 | 859.214 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8028 | 26.3 | 2
2008-05-07 | 9:35 | 10.79 | 0.106 | 1060 | 1347.8 | 2 | 1349.8 | 859.1 | 859.1 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.3 | 2
2008-05-07 | 9:35 | 10.79 | 0.106 | 1060 | 1348 | 2.1 | 1350.1 | 859.1 | 859.073 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28 | 26.3 | 2.1
2008-05-07 | 9:35 | 10.79 | 0.1076 | 1075.6215 | 1348 | 2.1 | 1350.1 | 859.3 | 859.3 | 0 | 12.9 | 949 | 957 | 939 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7845 | 26.3 | 2.1
2008-05-07 | 9:35 | 10.49 | 0.1007 | 1007.3605 | 1348.2 | 2.1 | 1350.3 | 859.1 | 859.1 | 0 | 12.9 | 949 | 956 | 939 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.171 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 28.0064 | 26.3 | 2.1
2008-05-07 | 9:35 | 9.708 | 0.0947 | 946.5243 | 1305.7 | 2.1 | 1307.8 | 846 | 846.0172 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.16786 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 28.1484 | 26.3 | 2.1
2008-05-07 | 9:35 | 9.633 | 0.0933 | 932.9559 | 1225.8 | 2.1 | 1227.9 | 793.4 | 793.3574 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.16316 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996.164 | 27.8 | 26.3 | 2.1
2008-05-07 | 9:35 | 9.622 | 0.094 | 940 | 1214.6 | 2.1 | 1216.7 | 769.3 | 769.3008 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.16083 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.6393 | 996.836 | 27.8 | 26.3 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1205.1 | 2.1 | 1207.2 | 764.6 | 764.6191 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.16008 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.868 | 26.232 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1203.6 | 2.1 | 1205.7 | 763.9 | 763.8522 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.15998 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.7584 | 26.2472 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1203.3 | 2.1 | 1205.4 | 763.4 | 763.4489 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.1599 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.1637 | 27.7055 | 26.3 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1203.1 | 2.1 | 1205.2 | 763.3 | 763.3 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1599 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2944 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1202.9 | 2.1 | 1205 | 763.2 | 763.2275 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1599 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 28 | 26.2151 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1202.2 | 2.1 | 1204.3 | 763 | 763.0286 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.1598 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.744 | 26.264 | 2.1
2008-05-07 | 9:35 | 9.62 | 0.094 | 940 | 1201.9 | 2.1 | 1204 | 762.9 | 762.888 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1598 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7297 | 26.3 | 2.1
2008-05-07 | 9:36 | 9.616 | 0.094 | 940 | 1201.6 | 2.0838 | 1203.7 | 762.8 | 762.8 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1598 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9225 | 26.3 | 2.1
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1201.4 | 2 | 1203.4 | 762.9 | 762.9295 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1598 | 0.0067 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9983 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1201.3 | 2 | 1203.3 | 762.8 | 762.7758 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.919 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1200.8 | 2 | 1202.8 | 762.6 | 762.6253 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.9102 | 998 | 27.9602 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.0941 | 941.2941 | 1200.7 | 2 | 1202.7 | 762.7 | 762.675 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3836 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 28 | 26.2606 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 949.6975 | 1200.8 | 2 | 1202.8 | 762.6 | 762.6204 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.9629 | 26.2186 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.8 | 2 | 1202.8 | 762.5 | 762.4752 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8043 | 26.2979 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.6 | 762.6 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1597 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.6459 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.4 | 2 | 1202.4 | 762.4 | 762.4265 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.2836 | 0.15968 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.6562 | 26.2438 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.5 | 762.548 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7 | 26.2354 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.6 | 762.5903 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1200.4 | 2 | 1202.4 | 762.3 | 762.3264 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7 | 26.3 | 2
2008-05-07 | 9:36 | 9.615 | 0.095 | 950 | 1200.7 | 2 | 1202.7 | 762.5 | 762.4742 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7731 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.5 | 762.5 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8524 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.6 | 762.6442 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9315 | 26.2685 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.6 | 2 | 1202.6 | 762.5 | 762.5267 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9681 | 26.2106 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.3 | 2 | 1202.3 | 762.7 | 762.6784 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0066 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.73 | 26.29 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.3 | 2 | 1202.3 | 762.3 | 762.3287 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.006584 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.9078 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1199.9 | 2 | 1201.9 | 762.3 | 762.2861 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1596 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.9515 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1199.9 | 2 | 1201.9 | 762.4 | 762.3529 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3836 | 0.1596 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.817 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200 | 2 | 1202 | 762.3 | 762.3 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.15958 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6138 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1199.9 | 2 | 1201.9 | 762.5 | 762.4806 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7723 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1199.9 | 2 | 1201.9 | 762.4 | 762.3739 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6693 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1199.9 | 2 | 1201.9 | 762.5 | 762.5 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5554 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.1 | 2 | 1202.1 | 762.3 | 762.3 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5711 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.1 | 2 | 1202.1 | 762.5 | 762.5217 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200 | 2 | 1202 | 762.3 | 762.3463 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8 | 26.3 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200 | 2 | 1202 | 762.4 | 762.3735 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6772 | 26.3 | 2
2008-05-07 | 9:36 | 9.612 | 0.095 | 950 | 1200 | 2 | 1202 | 762.3 | 762.3162 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0065 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.354 | 997 | 27.6406 | 26.3 | 2
2008-05-07 | 9:36 | 9.616 | 0.095 | 950 | 1200.2 | 2 | 1202.2 | 762.2 | 762.2142 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7198 | 26.2802 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.2 | 2 | 1202.2 | 762.6 | 762.5721 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.799 | 26.201 | 2
2008-05-07 | 9:36 | 9.62 | 0.0944 | 944.0216 | 1200.1 | 2 | 1202.1 | 762.7 | 762.6727 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6434 | 26.2783 | 2
2008-05-07 | 9:36 | 9.62 | 0.0944 | 944.484 | 1200.1 | 2 | 1202.1 | 762.9 | 762.8951 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8301 | 26.2425 | 2
2008-05-07 | 9:36 | 9.62 | 0.095 | 950 | 1200.1 | 2 | 1202.1 | 762.9 | 762.8756 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9265 | 26.2367 | 2
2008-05-07 | 9:36 | 9.62 | 0.0949 | 948.7156 | 1200.1 | 2 | 1202.1 | 763.2 | 763.2194 | 0 | 12.9 | 948 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3164 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8 | 26.2841 | 2
2008-05-07 | 9:36 | 9.614 | 0.094 | 940.3753 | 1200.1 | 2 | 1202.1 | 763.2 | 763.1755 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8 | 26.2051 | 2
2008-05-07 | 9:36 | 9.61 | 0.0948 | 947.7016 | 1200 | 2 | 1202 | 763.1 | 763.1 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7257 | 26.2743 | 2
2008-05-07 | 9:36 | 9.61 | 0.0944 | 944.0417 | 1200 | 2 | 1202 | 762.9 | 762.9304 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1595 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7536 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.9 | 2 | 1201.9 | 763.1 | 763.1283 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.15948 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.7344 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.6 | 2 | 1201.6 | 763.1 | 763.1 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.612 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.4 | 2 | 1201.4 | 762.7 | 762.7239 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6913 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.5 | 2 | 1201.5 | 762.7 | 762.656 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.7705 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.4 | 2 | 1201.4 | 762.5 | 762.5259 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7006 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.3 | 2 | 1201.3 | 762.6 | 762.6 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.094 | 940 | 1199.6 | 2 | 1201.6 | 762.6 | 762.6 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6163 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.0943 | 942.7006 | 1199.5 | 2 | 1201.5 | 762.6 | 762.6422 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7742 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1199.2 | 2 | 1201.2 | 762.7 | 762.7 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.667 | 26.3 | 2
2008-05-07 | 9:36 | 9.61 | 0.095 | 950 | 1199.2 | 2 | 1201.2 | 762.9 | 762.8537 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7831 | 26.3 | 2
2008-05-07 | 9:36 | 9.381 | 0.095 | 950 | 1199.1 | 2 | 1201.1 | 762.6 | 762.5713 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1594 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.9002 | 26.3 | 2
2008-05-07 | 9:36 | 8.915 | 0.0873 | 873.0149 | 1195.4 | 2 | 1197.4 | 762.7 | 762.6801 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.15912 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6042 | 26.3 | 2
2008-05-07 | 9:36 | 8.44 | 0.0833 | 833.1073 | 1119.6 | 2 | 1121.6 | 724.5 | 724.4672 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.15279 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.8377 | 27.6834 | 26.3 | 2
2008-05-07 | 9:36 | 8.44 | 0.085 | 850 | 1074.4 | 2 | 1076.4 | 683.4 | 683.379 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14887 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.888 | 26.3 | 2
2008-05-07 | 9:36 | 8.44 | 0.0848 | 847.5281 | 1059.7 | 2 | 1061.7 | 670 | 669.9634 | 0 | 12.9 | 948 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3164 | 0.14757 | 0.0064 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.8325 | 26.3 | 2
2008-05-07 | 9:36 | 8.434 | 0.0831 | 831.4767 | 1055.5 | 2 | 1057.5 | 667.2 | 667.1763 | 0 | 12.9 | 948 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14718 | 0.006384 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5578 | 26.3 | 2
2008-05-07 | 9:36 | 8.417 | 0.083 | 830 | 1052.7 | 2 | 1054.7 | 666.4 | 666.3966 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1469 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.4016 | 26.3 | 2
2008-05-07 | 9:36 | 8.4 | 0.083 | 830 | 1051.5 | 2 | 1053.5 | 666.2 | 666.2257 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14678 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996.164 | 27.7965 | 26.3 | 2
2008-05-07 | 9:36 | 8.4 | 0.0834 | 834.4184 | 1050.1 | 1.9 | 1052 | 666.6 | 666.5511 | 0 | 12.9 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1467 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7828 | 26.3 | 1.9
2008-05-07 | 9:36 | 8.4 | 0.084 | 840 | 1050 | 1.9 | 1051.9 | 666.6 | 666.5756 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14668 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6621 | 26.3 | 1.9
2008-05-07 | 9:36 | 8.4 | 0.084 | 840 | 1050 | 1.9 | 1051.9 | 666.4 | 666.4 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0063 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6343 | 26.3 | 1.9
2008-05-07 | 9:36 | 8.4 | 0.084 | 840 | 1050 | 1.9 | 1051.9 | 665.9 | 665.9 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3164 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7929 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.8 | 1.9 | 1051.7 | 665.9 | 665.9 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6488 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.7 | 1.9 | 1051.6 | 666 | 666.0294 | 0 | 13 | 948 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.8377 | 27.6 | 26.2452 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.8 | 1.9 | 1051.7 | 665.9 | 665.9 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.1637 | 27.6 | 26.234 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.5 | 1.9 | 1051.4 | 665.9 | 665.9293 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.4 | 1.9 | 1051.3 | 665.8 | 665.8 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1466 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.2 | 1.9 | 1051.1 | 666 | 665.9619 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.084 | 840 | 1049.3 | 1.9 | 1051.2 | 665.7 | 665.6737 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.0062 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.5492 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.0834 | 834.1049 | 1049.4 | 1.9 | 1051.3 | 665.8 | 665.8 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.006184 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.4 | 1.9 | 1051.3 | 665.6 | 665.6 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5187 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.7 | 1.9 | 1051.6 | 665.8 | 665.773 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6772 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.6 | 1.9 | 1051.5 | 665.6 | 665.6273 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.4 | 0.1465 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6322 | 26.3 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.7 | 1.9 | 1051.6 | 665.8 | 665.7607 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1465 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.647 | 26.253 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.6 | 1.8836 | 1051.5 | 665.6 | 665.5546 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1464 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.6475 | 26.2262 | 1.9
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1049.3 | 1.8 | 1051.1 | 665.5 | 665.5 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1464 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.4945 | 26.3 | 1.8
2008-05-07 | 9:37 | 8.401 | 0.083 | 830 | 1049.1 | 1.8 | 1050.9 | 665.7 | 665.6693 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1464 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033.6524 | 997 | 27.4154 | 26.3 | 1.8
2008-05-07 | 9:37 | 8.409 | 0.083 | 830 | 1048.1 | 1.8 | 1049.9 | 665.6 | 665.6261 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14638 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5275 | 26.3 | 1.8
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1047.4 | 1.8 | 1049.2 | 665.7 | 665.6749 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1463 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.643 | 26.3 | 1.8
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1047.3 | 1.7836 | 1049.1 | 665.4 | 665.3853 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1463 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6331 | 26.3 | 1.8
2008-05-07 | 9:37 | 8.41 | 0.0838 | 837.7534 | 1047.2 | 1.7 | 1048.9 | 665.6 | 665.628 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4032 | 26.2984 | 1.7
2008-05-07 | 9:37 | 8.407 | 0.0834 | 834.0325 | 1047.4 | 1.7 | 1049.1 | 665.4 | 665.426 | 0 | 13 | 948 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5615 | 26.2192 | 1.7
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1047.2 | 1.7 | 1048.9 | 665.5 | 665.5261 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.26 | 1.7
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1047.1 | 1.7 | 1048.8 | 665.5 | 665.4562 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6784 | 26.2608 | 1.7
2008-05-07 | 9:37 | 8.408 | 0.083 | 830 | 1047 | 1.7 | 1048.7 | 665.5 | 665.4765 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.7448 | 26.2184 | 1.7
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046.8 | 1.7 | 1048.5 | 665.7 | 665.6733 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1462 | 0.0061 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.8377 | 27.5071 | 26.2976 | 1.7
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046.5 | 1.6836 | 1048.2 | 665.7 | 665.7 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.14618 | 0.0061 | 0.000184 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.164 | 27.5768 | 26.3 | 1.7
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.8 | 1.6 | 1047.4 | 665.6 | 665.6289 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.0061 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997.836 | 27.4881 | 26.2441 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.9 | 1.6 | 1047.5 | 665.6 | 665.6279 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.4703 | 26.2351 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046 | 1.6 | 1047.6 | 665.8 | 665.7816 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.7 | 1.6 | 1047.3 | 665.7 | 665.6743 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.6 | 1.6 | 1047.2 | 665.5 | 665.5 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 998 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.7 | 1.6 | 1047.3 | 665.6 | 665.6 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.4435 | 26.3 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1045.9 | 1.5836 | 1047.5 | 665.8 | 665.8134 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.4255 | 26.3 | 1.6
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 666.1 | 666.1 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6894 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046.2 | 1.5 | 1047.7 | 665.9 | 665.9433 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.6102 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046.2 | 1.5 | 1047.7 | 665.7 | 665.6516 | 0 | 12.9 | 948 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.7 | 665.7068 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3836 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.41 | 0.083 | 830 | 1046 | 1.4 | 1047.4 | 665.7 | 665.6525 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3162 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.4
2008-05-07 | 9:37 | 8.407 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.9 | 665.873 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5933 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.9 | 665.9254 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5141 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046.1 | 1.5 | 1047.6 | 665.9 | 665.8804 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1461 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.5651 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046.1 | 1.5 | 1047.6 | 665.7 | 665.7427 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.6443 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046.1 | 1.5 | 1047.6 | 665.7 | 665.6778 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996 | 27.6295 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046.1 | 1.5 | 1047.6 | 665.5 | 665.5261 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 996.164 | 27.4052 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1045.8 | 1.5 | 1047.3 | 665.7 | 665.7 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5638 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1045.8 | 1.5 | 1047.3 | 665.8 | 665.762 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 666 | 666.0373 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6807 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.9 | 665.9196 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.761 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.4 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.9 | 665.9285 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6022 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.396 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 665.8 | 665.8 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.006 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.318 | 0.083 | 830 | 1046 | 1.5 | 1047.5 | 666 | 665.9906 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.146 | 0.0059 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6573 | 26.3 | 1.5
2008-05-07 | 9:37 | 8.029 | 0.0772 | 772.2776 | 1045.5 | 1.5 | 1047 | 665.9 | 665.9 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.14597 | 0.0059 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.6635 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.286 | 0.0697 | 697.3311 | 1012.1 | 1.5 | 1013.6 | 644.7 | 644.6534 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3836 | 0.14295 | 0.0059 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1034 | 997 | 27.5687 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.222 | 0.072 | 720 | 936.9 | 1.5 | 938.4 | 593.5 | 593.5256 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1359 | 0.0059 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033.7632 | 997 | 27.4104 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.22 | 0.072 | 720 | 907 | 1.5 | 908.5 | 575 | 574.9582 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.13377 | 0.0059 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7703 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.22 | 0.072 | 720 | 900.7 | 1.5 | 902.2 | 571.8 | 571.807 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13237 | 0.0058 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7933 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.22 | 0.0714 | 714.1965 | 899.9 | 1.5 | 901.4 | 570.9 | 570.9084 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1322 | 0.0058 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7325 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.219 | 0.071 | 710 | 899.8 | 1.5 | 901.3 | 570.4 | 570.3765 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.1322 | 0.0058 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7883 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.211 | 0.071 | 710 | 899.5 | 1.5 | 901 | 570.6 | 570.5919 | 0 | 12.9 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13218 | 0.005784 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.709 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.21 | 0.071 | 710 | 899.4 | 1.5 | 900.9 | 570.4 | 570.4 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1321 | 0.0057 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4191 | 26.3 | 1.5
2008-05-07 | 9:37 | 7.198 | 0.071 | 710 | 899.3 | 1.5 | 900.8 | 570.2 | 570.1533 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1321 | 0.0057 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4483 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.19 | 0.071 | 710 | 899.2 | 1.5 | 900.7 | 570.1 | 570.1 | 0 | 12.9 | 950 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1321 | 0.0057 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5713 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.19 | 0.071 | 710 | 898.9 | 1.5 | 900.4 | 570.1 | 570.1 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13208 | 0.0057 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.5235 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.188 | 0.0714 | 714.2179 | 898.7 | 1.5 | 900.2 | 569.9 | 569.9273 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.132 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.7606 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.171 | 0.072 | 720 | 898.7 | 1.5 | 900.2 | 569.8 | 569.7764 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.132 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.7338 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.8 | 569.8 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.132 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6545 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.8 | 569.7787 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.132 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5753 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.7 | 1.5 | 900.2 | 569.9 | 569.8659 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13198 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5117 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.173 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.8 | 569.8118 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1319 | 0.0056 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7496 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.9 | 1.5 | 900.4 | 569.8 | 569.8 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1319 | 0.005584 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7376 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 570 | 569.9947 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1319 | 0.0055 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5751 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 570 | 570 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1319 | 0.0055 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4417 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 899 | 1.5 | 900.5 | 569.9 | 569.9 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13188 | 0.0055 | 0.0001 | 0.0001 |  | n/a |  | 30.27 | 53 | 1033 | 997 | 27.6002 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.9 | 569.9 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1318 | 0.005484 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7581 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 570.3 | 570.2846 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1318 | 0.0054 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.9 | 569.9351 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3836 | 0.1318 | 0.0054 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.8 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.18 | 0.072 | 720 | 898.9 | 1.5 | 900.4 | 570 | 569.9503 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1318 | 0.0054 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7831 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.172 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 569.9 | 569.9264 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1318 | 0.0054 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7039 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.8 | 1.5 | 900.3 | 570 | 569.9616 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.005384 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5493 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.9 | 1.5164 | 900.4 | 570 | 569.9729 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6638 | 26.3 | 1.5
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 899.1 | 1.6 | 900.7 | 570.1 | 570.1157 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.7324 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.17 | 0.072 | 720 | 898.9 | 1.6 | 900.5 | 570.3 | 570.3 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.17 | 0.0717 | 717.3188 | 898.6 | 1.6 | 900.2 | 570.2 | 570.2474 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.8377 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.177 | 0.071 | 710 | 898.8 | 1.6 | 900.4 | 570.4 | 570.3722 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1317 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6712 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.174 | 0.071 | 710 | 898.8 | 1.6 | 900.4 | 570.2 | 570.2447 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6495 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.174 | 0.071 | 710 | 898.9 | 1.6 | 900.5 | 570 | 570.008 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6299 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.8 | 1.6 | 900.4 | 570 | 569.9818 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0053 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6727 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.7 | 1.6 | 900.3 | 570 | 569.9576 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.005284 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.4352 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.6 | 1.6 | 900.2 | 569.9 | 569.9 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.535 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.8 | 1.6 | 900.4 | 569.8 | 569.8196 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6 | 26.2532 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.6 | 1.6 | 900.2 | 570.2 | 570.1703 | 0 | 12.9 | 950 | 956 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.226 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.7 | 1.6 | 900.3 | 570.2 | 570.189 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5895 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.8 | 1.6 | 900.4 | 570 | 570 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.4312 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.179 | 0.071 | 710 | 898.7 | 1.6 | 900.3 | 570.1 | 570.1172 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1316 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.527 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.171 | 0.071 | 710 | 898.5 | 1.6 | 900.1 | 570 | 570.0173 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.13158 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.2573 | 1.6
2008-05-07 | 9:38 | 7.178 | 0.071 | 710 | 898.2 | 1.6 | 899.8 | 570.2 | 570.228 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6221 | 26.2221 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.2 | 1.6 | 899.8 | 570.5 | 570.4718 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6987 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898 | 1.6 | 899.6 | 570.3 | 570.2518 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 27.6195 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898 | 1.6 | 899.6 | 570.1 | 570.1 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 998 | 27.6597 | 26.2403 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898 | 1.6 | 899.6 | 570.4 | 570.4 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997.836 | 27.5442 | 26.239 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.1 | 1.6 | 899.7 | 570.4 | 570.3819 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.3726 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.4 | 570.4 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6895 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.186 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.2 | 570.228 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6234 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.186 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.4 | 570.4393 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0052 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6557 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.4 | 1.6 | 900 | 570.2 | 570.2287 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1315 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6651 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.2 | 1.6 | 899.8 | 570.4 | 570.3949 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.1315 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.1 | 570.1 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13148 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.2 | 570.2492 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1314 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.18 | 0.071 | 710 | 898.3 | 1.6 | 899.9 | 570.3 | 570.3465 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1314 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5481 | 26.3 | 1.6
2008-05-07 | 9:38 | 7.112 | 0.071 | 710 | 898.5 | 1.6 | 900.1 | 570.2 | 570.1707 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1314 | 0.0051 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.836 | 27.5623 | 26.3 | 1.6
2008-05-07 | 9:38 | 6.784 | 0.0623 | 622.5301 | 896.2 | 1.6 | 897.8 | 570.5 | 570.451 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.13115 | 0.0051 | 0.000116 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6896 | 26.3 | 1.6
2008-05-07 | 9:38 | 6.043 | 0.0581 | 580.7865 | 849.6 | 1.7 | 851.3 | 546.3 | 546.2673 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.12701 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6105 | 26.3 | 1.7
2008-05-07 | 9:38 | 5.966 | 0.061 | 610 | 776.1 | 1.7 | 777.8 | 494.9 | 494.9302 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11878 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6 | 26.3 | 1.7
2008-05-07 | 9:38 | 5.96 | 0.061 | 610 | 758.3 | 1.7 | 760 | 479.6 | 479.5884 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1169 | 0.005 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.6958 | 26.3 | 1.7
2008-05-07 | 9:38 | 5.96 | 0.061 | 610 | 755.4 | 1.7 | 757.1 | 475.9 | 475.8895 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11657 | 0.004984 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7459 | 26.3 | 1.7
2008-05-07 | 9:38 | 5.96 | 0.0603 | 602.5544 | 753 | 1.7 | 754.7 | 475.1 | 475.0735 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11628 | 0.0049 | 0.0002 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5936 | 26.3 | 1.7
2008-05-07 | 9:38 | 5.96 | 0.06 | 600 | 752.6 | 1.6 | 754.2 | 474.9 | 474.9051 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11618 | 0.0049 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.164 | 27.5143 | 26.3 | 1.6
2008-05-07 | 9:38 | 5.96 | 0.06 | 600 | 752.2 | 1.6 | 753.8 | 474.9 | 474.8861 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11608 | 0.0048 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.4351 | 26.3 | 1.6
2008-05-07 | 9:38 | 5.96 | 0.0598 | 597.6446 | 750.5 | 1.6 | 752.1 | 474.8 | 474.7701 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11588 | 0.0048 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.4882 | 26.3 | 1.6
2008-05-07 | 9:38 | 5.96 | 0.059 | 590 | 750.1 | 1.6 | 751.7 | 474.5 | 474.5478 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.11578 | 0.0048 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.5534 | 26.3 | 1.6
2008-05-07 | 9:38 | 5.96 | 0.059 | 590 | 749.7 | 1.6 | 751.3 | 474.5 | 474.4733 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1157 | 0.0048 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.4127 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.956 | 0.059 | 590 | 749.6 | 1.6 | 751.2 | 474.6 | 474.6 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1157 | 0.004784 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.8086 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 749.4 | 1.6 | 751 | 474.4 | 474.4341 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1157 | 0.0047 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.5953 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.951 | 0.059 | 590 | 749.3 | 1.6 | 750.9 | 474.5 | 474.4743 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1157 | 0.0047 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996 | 27.4401 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.959 | 0.0592 | 592.349 | 749.2 | 1.6 | 750.8 | 474.3 | 474.3224 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1156 | 0.0047 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 996.1637 | 27.5192 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.953 | 0.06 | 600 | 749.2 | 1.6 | 750.8 | 474.2 | 474.2 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1156 | 0.004684 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5986 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.06 | 600 | 749.1 | 1.6 | 750.7 | 474.2 | 474.2 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.1156 | 0.0046 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7557 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.954 | 0.06 | 600 | 749 | 1.6 | 750.6 | 474.1 | 474.1147 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.1155 | 0.0046 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6287 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.06 | 600 | 748.8 | 1.6 | 750.4 | 474.5 | 474.474 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1155 | 0.0046 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997.164 | 27.5726 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.959 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 474.1 | 474.1482 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1155 | 0.0046 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997.836 | 27.7 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.951 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 474.2 | 474.1792 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1155 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.06 | 600 | 748.8 | 1.6 | 750.4 | 474.1 | 474.0871 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1155 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.626 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.956 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 474.1 | 474.1269 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3164 | 0.11548 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 473.9 | 473.9207 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5353 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.957 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 474.1 | 474.1367 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.4229 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.0596 | 596.0217 | 748.6 | 1.6 | 750.2 | 474.1 | 474.091 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3164 | 0.1154 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997.164 | 27.5816 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 474 | 474 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.67 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 474 | 474 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0045 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6507 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.956 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 473.9 | 473.9257 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.657 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 474 | 473.9793 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.7693 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.7 | 1.6 | 750.3 | 474 | 473.9712 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.4528 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.7 | 1.5836 | 750.3 | 474.2 | 474.1745 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.5321 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.8 | 1.5 | 750.3 | 474 | 473.9557 | 0 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1154 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.8 | 1.5 | 750.3 | 474 | 474.0259 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 30.73 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.8 | 1.5 | 750.3 | 473.9 | 473.9205 | 0 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.6 | 1.5 | 750.1 | 473.8 | 473.8 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.96 | 0.059 | 590 | 748.7 | 1.5 | 750.2 | 474 | 473.9863 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4755 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.959 | 0.059 | 590 | 748.7 | 1.5 | 750.2 | 474 | 474 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.3586 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.951 | 0.059 | 590 | 748.7 | 1.5 | 750.2 | 473.9 | 473.9272 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4241 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.95 | 0.0591 | 590.7521 | 748.6 | 1.5 | 750.1 | 473.9 | 473.8889 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.8995 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.95 | 0.0599 | 599.2991 | 748.5 | 1.5165 | 750 | 473.9 | 473.872 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.979 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.95 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 473.8 | 473.8139 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0044 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.7085 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.06 | 600 | 748.6 | 1.5838 | 750.2 | 473.7 | 473.7241 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.004384 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5376 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.952 | 0.06 | 600 | 748.6 | 1.5 | 750.1 | 473.7 | 473.7481 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5832 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.969 | 0.06 | 600 | 748.6 | 1.5 | 750.1 | 473.8 | 473.7754 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.5042 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.8 | 1.5 | 750.3 | 473.7 | 473.7198 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4248 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.9 | 1.5 | 750.4 | 473.8 | 473.7746 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1153 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4 | 26.3 | 1.5
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.9 | 1.6 | 750.5 | 474 | 474 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3836 | 0.11528 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 995.836 | 27.4673 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.9 | 1.6 | 750.5 | 473.9 | 473.8725 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.6 | 1.6 | 750.2 | 474 | 474.026 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.97 | 0.06 | 600 | 748.5 | 1.6 | 750.1 | 473.9 | 473.9049 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3869 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.955 | 0.06 | 600 | 748.4 | 1.6 | 750 | 474.1 | 474.0751 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4007 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.06 | 600 | 748.4 | 1.6 | 750 | 473.9 | 473.9 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4702 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.0593 | 592.6792 | 748.3 | 1.6 | 749.9 | 474 | 473.9731 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.2 | 1.6 | 749.8 | 473.8 | 473.8279 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.4 | 1.6 | 750 | 474 | 473.9733 | 0 | 12.9 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4674 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 473.8 | 473.8385 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 473.8 | 473.8 | 0 | 12.9 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0043 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5258 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.95 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 474 | 474.017 | 0 | 13 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.004284 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5899 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.761 | 0.059 | 590 | 748.6 | 1.6 | 750.2 | 473.7 | 473.7436 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1152 | 0.0042 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4315 | 26.3 | 1.6
2008-05-07 | 9:39 | 5.303 | 0.0517 | 517.4293 | 744.6 | 1.6 | 746.2 | 473.9 | 473.8847 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.11475 | 0.0042 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5264 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.769 | 0.0468 | 467.8102 | 688.8 | 1.6 | 690.4 | 449.1 | 449.0648 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.10851 | 0.0042 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.1623 | 27.515 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.761 | 0.048 | 480 | 619.1 | 1.6 | 620.7 | 394.7 | 394.716 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.1005 | 0.0042 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.378 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.76 | 0.048 | 480 | 605.2 | 1.6 | 606.8 | 382.9 | 382.9136 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.098667 | 0.004184 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.3036 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.76 | 0.048 | 480 | 604.3 | 1.6 | 605.9 | 380.6 | 380.6471 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.098484 | 0.0041 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5407 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.76 | 0.048 | 480 | 603.2 | 1.6 | 604.8 | 380 | 380.0268 | 0 | 13 | 948 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.098384 | 0.0041 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.758 | 0.048 | 480 | 602.2 | 1.6 | 603.8 | 379.8 | 379.8474 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0982 | 0.0041 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.6
2008-05-07 | 9:39 | 4.75 | 0.048 | 480 | 600.6 | 1.5 | 602.1 | 379.6 | 379.5751 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.098 | 0.004 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.582 | 26.3 | 1.5
2008-05-07 | 9:39 | 4.75 | 0.048 | 480 | 600.5 | 1.5 | 602 | 379.6 | 379.572 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0979 | 0.004 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5028 | 26.3 | 1.5
2008-05-07 | 9:39 | 4.75 | 0.048 | 480 | 599.7 | 1.5 | 601.2 | 379.5 | 379.5151 | 0 | 13 | 948 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0978 | 0.004 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5765 | 26.3 | 1.5
2008-05-07 | 9:39 | 4.75 | 0.048 | 480 | 599.5 | 1.5 | 601 | 379.6 | 379.6442 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0978 | 0.004 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4891 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.4 | 1.4 | 600.8 | 379.5 | 379.5308 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.097784 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4693 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.2 | 1.4 | 600.6 | 379.2 | 379.2196 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3164 | 0.0977 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.5858 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.1 | 1.4 | 600.5 | 379.3 | 379.3244 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0977 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.5067 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.1 | 1.4 | 600.5 | 379.3 | 379.3 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0977 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3548 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.2 | 1.4 | 600.6 | 379.3 | 379.3 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0977 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3518 | 26.2482 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.048 | 480 | 599.2 | 1.4 | 600.6 | 379.1 | 379.1274 | 0 | 13 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0977 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4 | 26.231 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.0473 | 472.7325 | 599.1 | 1.4 | 600.5 | 379.3 | 379.3 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4204 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379 | 379.0377 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3836 | 0.0976 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.8363 | 27.5789 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379.1 | 379.1451 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0039 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3939 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599 | 1.4 | 600.4 | 379.3 | 379.3 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3477 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.0468 | 467.8186 | 599.1 | 1.4 | 600.5 | 379.2 | 379.1749 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.4537 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.0443 | 443.3088 | 599.2 | 1.4 | 600.6 | 379.3 | 379.2741 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 27.5808 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.0447 | 447.2083 | 599 | 1.4164 | 600.4 | 379.1 | 379.1124 | 0 | 12.9 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3432 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.0462 | 461.5251 | 598.9 | 1.5 | 600.4 | 379.3 | 379.2721 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4295 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.1 | 1.5 | 600.6 | 379 | 379.0316 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.456 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.1 | 1.4 | 600.5 | 379.1 | 379.1 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0976 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4464 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.1 | 1.4 | 600.5 | 379 | 379.0292 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.097584 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 599.1 | 1.4 | 600.5 | 379.2 | 379.214 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 598.9 | 1.4 | 600.3 | 379.1 | 379.0596 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0038 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5391 | 26.2391 | 1.4
2008-05-07 | 9:40 | 4.75 | 0.047 | 470 | 598.8 | 1.4 | 600.2 | 379.1 | 379.1122 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.003784 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.3801 | 26.24 | 1.4
2008-05-07 | 9:40 | 4.749 | 0.047 | 470 | 598.8 | 1.4 | 600.2 | 379.3 | 379.3 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2758 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.741 | 0.047 | 470 | 598.8 | 1.4 | 600.2 | 379.1 | 379.0673 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5943 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.8 | 1.4 | 600.2 | 379.2 | 379.1538 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3667 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.8 | 1.4 | 600.2 | 379 | 379.0251 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.4
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.9 | 1.5 | 600.4 | 379.2 | 379.1545 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3724 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 599.1 | 1.5 | 600.6 | 379 | 379.0263 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4539 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379.2 | 379.1768 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0975 | 0.0037 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2162 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379.1 | 379.0537 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.003684 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2739 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379.1 | 379.0713 | 0 | 12.9 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3531 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 599.1 | 1.5 | 600.6 | 379.2 | 379.2 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4322 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.744 | 0.047 | 470 | 599.1 | 1.5 | 600.6 | 379.2 | 379.1564 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5111 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.748 | 0.047 | 470 | 599.2 | 1.5 | 600.7 | 379.1 | 379.1246 | 0 | 13 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5907 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.0474 | 474.0138 | 599.1 | 1.5 | 600.6 | 379.1 | 379.1461 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3201 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.048 | 480 | 599.1 | 1.5164 | 600.6 | 379.2 | 379.2 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2983 | 26.3 | 1.5
2008-05-07 | 9:40 | 4.74 | 0.048 | 480 | 599.1 | 1.6 | 600.7 | 379.3 | 379.3 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0974 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2865 | 26.3 | 1.6
2008-05-07 | 9:40 | 4.74 | 0.048 | 480 | 599.1 | 1.6 | 600.7 | 379.3 | 379.3 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.097384 | 0.0036 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.0228 | 26.3 | 1.6
2008-05-07 | 9:40 | 4.74 | 0.048 | 480 | 598.7 | 1.6 | 600.3 | 379.2 | 379.2272 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.8 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0036 | 0.000116 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2601 | 26.3 | 1.6
2008-05-07 | 9:40 | 4.74 | 0.0474 | 474.4854 | 598.6 | 1.6164 | 600.2 | 379.3 | 379.2741 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0036 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.6
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.5 | 1.7 | 600.2 | 379.4 | 379.4 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.6 | 1.7 | 600.3 | 379.2 | 379.1711 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.5 | 1.7 | 600.2 | 379.2 | 379.2187 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.4 | 1.7 | 600.1 | 379.1 | 379.1256 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.2 | 1.7 | 599.9 | 379.2 | 379.2152 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.2836 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.8363 | 27.2379 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.3 | 1.7 | 600 | 379.2 | 379.2281 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2414 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.3 | 1.7 | 600 | 379.2 | 379.2295 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3206 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.3 | 1.7 | 600 | 379.1 | 379.1099 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0973 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3998 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.4 | 1.7 | 600.1 | 379.3 | 379.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0972 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3211 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.3 | 1.7 | 600 | 379 | 378.9568 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0972 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3582 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.74 | 0.047 | 470 | 598.2 | 1.7 | 599.9 | 379.1 | 379.1328 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0972 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3625 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.681 | 0.047 | 470 | 598.1 | 1.7 | 599.8 | 379.1 | 379.0544 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0972 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 4.408 | 0.0424 | 424.414 | 596.2 | 1.7 | 597.9 | 378.9 | 378.9447 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.8 | 26 | 5 | 195 | 59.3 | 0.0972 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.712 | 0.0357 | 356.6031 | 561.8 | 1.7 | 563.5 | 364.9 | 364.9255 | 0 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.092726 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5247 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.577 | 0.0378 | 377.6057 | 479 | 1.7 | 480.7 | 309.4 | 309.4047 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.082056 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4373 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.56 | 0.0354 | 353.6632 | 464.4 | 1.7 | 466.1 | 290.5 | 290.5367 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.080136 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.3 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.56 | 0.035 | 350 | 453.7 | 1.7 | 455.4 | 286.8 | 286.8455 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.078668 | 0.0035 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3127 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.56 | 0.035 | 350 | 452 | 1.7 | 453.7 | 285.5 | 285.4636 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.078368 | 0.003484 | 0.000184 | 8.4e-05 |  | n/a |  | 31.73 | 53 | 1033 | 996 | 27.3919 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.56 | 0.0354 | 353.8347 | 451 | 1.7164 | 452.7 | 285 | 285.0216 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3836 | 0.078184 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.329 | 26.3 | 1.7
2008-05-07 | 9:40 | 3.56 | 0.036 | 360 | 450 | 1.8 | 451.8 | 285.2 | 285.2 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.078 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.3 | 26.3 | 1.8
2008-05-07 | 9:40 | 3.56 | 0.036 | 360 | 449.8 | 1.8 | 451.6 | 285.2 | 285.1708 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.078 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.2111 | 26.3 | 1.8
2008-05-07 | 9:40 | 3.56 | 0.036 | 360 | 449.7 | 1.8 | 451.5 | 285.1 | 285.125 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.078 | 0.003384 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.0353 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.6 | 1.8 | 451.4 | 285 | 285 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0779 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 996.164 | 27.3508 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.6 | 1.7836 | 451.4 | 284.9 | 284.9276 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0779 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.2657 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.4 | 1.7 | 451.1 | 284.9 | 284.9402 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.0779 | 0.003284 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997.164 | 27.293 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.2 | 1.7 | 450.9 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.077884 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 998 | 27.4 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.1 | 1.7 | 450.8 | 284.9 | 284.8818 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3164 | 0.0778 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 998 | 27.3951 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449 | 1.7 | 450.7 | 284.7 | 284.7488 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.0778 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 998 | 27.3159 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.56 | 0.036 | 360 | 449.1 | 1.7 | 450.8 | 284.8 | 284.7737 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0778 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.2369 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.556 | 0.036 | 360 | 449.1 | 1.7 | 450.8 | 284.6 | 284.6 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.077784 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3271 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 449 | 1.7 | 450.7 | 284.7 | 284.7496 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0777 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4346 | 26.3 | 1.7
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 449 | 1.8 | 450.8 | 284.7 | 284.7142 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0777 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.202 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.9 | 1.8 | 450.7 | 284.7 | 284.7241 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.077684 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3598 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.9 | 1.8 | 450.7 | 284.6 | 284.5523 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3407 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.7 | 284.7492 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3773 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.9 | 284.9 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4643 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.7 | 1.8 | 450.5 | 284.7 | 284.7372 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3059 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.9 | 284.8897 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.003 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4526 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.003 | 0.0001 | 0 |  | n/a |  | 31.27 | 53 | 1033 | 997 | 27.3893 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.8 | 284.8442 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0776 | 0.003 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.3 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.7 | 1.8 | 450.5 | 285 | 284.9876 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.328 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.8 | 284.8257 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.0002 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4864 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.8 | 1.8 | 450.6 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.0002 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.5 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.7 | 1.8 | 450.5 | 284.9 | 284.9 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.0002 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5514 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.6 | 1.8 | 450.4 | 285 | 284.9731 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.0002 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.6 | 1.8 | 450.4 | 284.9 | 284.9 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.003 | 0.000184 | 0 |  | n/a |  | 31 | 53 | 1033 | 997.164 | 27.6 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.8 | 450.2 | 285 | 285 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0775 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.6 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.8 | 450.2 | 284.8 | 284.8278 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.3944 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.8 | 450.2 | 284.9 | 284.9157 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.8164 | 450.2 | 284.9 | 284.8503 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.9 | 450.3 | 285 | 284.9762 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3836 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.3124 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.5 | 1.9 | 450.4 | 284.9 | 284.9 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4707 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.5 | 1.9 | 450.4 | 284.8 | 284.8246 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3712 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.4 | 1.9 | 450.3 | 284.9 | 284.9298 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.2 | 1.9 | 450.1 | 285 | 284.9778 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.323 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.1 | 1.9 | 450 | 285 | 285 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0774 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3979 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448 | 1.9 | 449.9 | 284.9 | 284.9 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3186 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.1 | 1.9 | 450 | 284.8 | 284.8184 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3606 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.2 | 1.9 | 450.1 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4793 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.1 | 1.9 | 450 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.5439 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.2 | 1.9 | 450.1 | 284.7 | 284.7263 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188.8 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0029 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3053 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.1 | 1.9 | 450 | 285 | 285.0097 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.002884 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 1.9
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448 | 1.8 | 449.8 | 285 | 284.9875 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4131 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448 | 1.8 | 449.8 | 285 | 285.0439 | 0 | 12.9 | 950 | 956 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 188.2 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.5357 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448 | 1.8 | 449.8 | 284.8 | 284.8092 | 0 | 12.9 | 950 | 956 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.5698 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 447.9 | 1.8 | 449.7 | 285 | 285.0226 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4114 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 447.8 | 1.8 | 449.6 | 284.8 | 284.8382 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4735 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448 | 1.8 | 449.8 | 284.8 | 284.805 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0773 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5528 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.1 | 1.8 | 449.9 | 284.8 | 284.7549 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0772 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.8363 | 27.5363 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.54 | 0.036 | 360 | 448.2 | 1.8 | 450 | 284.8 | 284.8 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0772 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4 | 26.3 | 1.8
2008-05-07 | 9:41 | 3.314 | 0.036 | 360 | 448.2 | 1.8 | 450 | 284.8 | 284.8139 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0772 | 0.0028 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4 | 26.3 | 1.8
2008-05-07 | 9:41 | 2.826 | 0.0272 | 272.0672 | 445.4 | 1.8 | 447.2 | 284.7 | 284.7 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.076823 | 0.002784 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3304 | 26.3 | 1.8
2008-05-07 | 9:41 | 2.376 | 0.0219 | 218.8465 | 373 | 1.8 | 374.8 | 249.4 | 249.3603 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.066649 | 0.0027 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | 1.8
2008-05-07 | 9:41 | 2.328 | 0.025 | 250 | 312.7 | 1.7 | 314.4 | 203.8 | 203.7779 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.059408 | 0.0027 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3561 | 26.3 | 1.7
2008-05-07 | 9:41 | 2.327 | 0.025 | 250 | 304.5 | 1.6836 | 306.2 | 193.2 | 193.2348 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.056418 | 0.0026 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.5144 | 26.3 | 1.7
2008-05-07 | 9:41 | 2.33 | 0.025 | 250 | 301.9 | 1.6 | 303.5 | 190.9 | 190.9426 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.055984 | 0.0026 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6729 | 26.3 | 1.6
2008-05-07 | 9:41 | 2.33 | 0.025 | 250 | 301.3 | 1.5836 | 302.8 | 190.3 | 190.2721 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.055884 | 0.002584 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.5686 | 26.3 | 1.6
2008-05-07 | 9:41 | 2.33 | 0.025 | 250 | 300.7 | 1.5 | 302.2 | 190 | 189.9801 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0557 | 0.0025 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4551 | 26.3 | 1.5
2008-05-07 | 9:41 | 2.33 | 0.0243 | 242.7727 | 300.6 | 1.5 | 302.1 | 189.8 | 189.8 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.055684 | 0.0025 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.448 | 26.3 | 1.5
2008-05-07 | 9:41 | 2.33 | 0.023 | 230 | 300.6 | 1.4 | 302 | 189.8 | 189.8 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3836 | 0.0556 | 0.0025 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.6 | 26.3 | 1.4
2008-05-07 | 9:41 | 2.33 | 0.023 | 230 | 300.5 | 1.4 | 301.9 | 189.6 | 189.6271 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0556 | 0.0024 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.6 | 26.3 | 1.4
2008-05-07 | 9:41 | 2.335 | 0.023 | 230 | 300.6 | 1.4 | 302 | 189.6 | 189.589 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0556 | 0.0024 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.4766 | 26.3 | 1.4
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.5 | 1.4 | 301.9 | 189.5 | 189.5283 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0556 | 0.0024 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.3591 | 26.3 | 1.4
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.4 | 1.4 | 301.8 | 189.6 | 189.6 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0556 | 0.0024 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.3603 | 26.3 | 1.4
2008-05-07 | 9:42 | 2.339 | 0.023 | 230 | 300.4 | 1.3 | 301.7 | 189.6 | 189.5891 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0555 | 0.0023 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.5979 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.331 | 0.023 | 230 | 300.4 | 1.3 | 301.7 | 189.5 | 189.5 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0555 | 0.0023 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.4429 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.33 | 0.023 | 230 | 300.4 | 1.3 | 301.7 | 189.4 | 189.426 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.055484 | 0.002284 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.5155 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.33 | 0.023 | 230 | 300.4 | 1.3 | 301.7 | 189.5 | 189.4522 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0554 | 0.0022 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.489 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.33 | 0.0237 | 237.3505 | 300.5 | 1.3 | 301.8 | 189.4 | 189.424 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0554 | 0.0022 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997.836 | 27.3324 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.332 | 0.0234 | 234.4043 | 300.5 | 1.3 | 301.8 | 189.4 | 189.3581 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0554 | 0.0022 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.4904 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.6 | 1.3 | 301.9 | 189.5 | 189.4802 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.4255 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.6 | 1.2 | 301.8 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.3462 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.5 | 1.2 | 301.7 | 189.3 | 189.3136 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189.2 | 26 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0.0001 | 0 |  | n/a |  | 31.7297 | 53 | 1033 | 997 | 27.333 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.4 | 1.2 | 301.6 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0553 | 0.0021 | 0.0001 | 0 |  | n/a |  | 31.27 | 53 | 1033 | 997 | 27.4 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.4 | 1.2 | 301.6 | 189.4 | 189.3728 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0552 | 0.0021 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.4 | 1.2 | 301.6 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0552 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.3 | 1.2 | 301.5 | 189.4 | 189.3737 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.2164 | 0.0552 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3501 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.3 | 1.2 | 301.5 | 189.2 | 189.2318 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0552 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3291 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.1 | 1.2 | 301.3 | 189.2 | 189.1743 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4166 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.34 | 0.023 | 230 | 300.1 | 1.2 | 301.3 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5747 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.336 | 0.0237 | 237.1262 | 300.1 | 1.2 | 301.3 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26.2 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5333 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300 | 1.2 | 301.2 | 189.2 | 189.2204 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4541 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300 | 1.2 | 301.2 | 189.2 | 189.2292 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26.8 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4251 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300.1 | 1.2 | 301.3 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4957 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300.1 | 1.2 | 301.3 | 189.1 | 189.0721 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.4164 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300.1 | 1.2 | 301.3 | 189.3 | 189.272 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.002 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3372 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300.1 | 1.2 | 301.3 | 189.2 | 189.2 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0551 | 0.001984 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.836 | 27.384 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300.1 | 1.2 | 301.3 | 189.3 | 189.2716 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055084 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.165 | 27.4788 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 300 | 1.2 | 301.2 | 189.3 | 189.3 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4008 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.9 | 1.2 | 301.1 | 189.3 | 189.2513 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5591 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.8 | 1.2 | 301 | 189.2 | 189.1562 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5412 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.8 | 1.2 | 301 | 189.1 | 189.0512 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.7 | 1.2 | 300.9 | 189.2 | 189.1836 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4828 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.6 | 1.2 | 300.8 | 189.2 | 189.1701 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4035 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.6 | 1.2 | 300.8 | 189.3 | 189.2939 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190.2 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.6 | 1.2 | 300.8 | 189.2 | 189.2 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3451 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.6 | 1.2 | 300.8 | 189.3 | 189.2728 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4024 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.5 | 1.2 | 300.7 | 189.1 | 189.0889 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.055 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.5 | 1.2 | 300.7 | 188.9 | 188.9288 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.054984 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.0245 | 245.4958 | 299.5 | 1.2 | 300.7 | 189.1 | 189.077 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3164 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5283 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.5 | 1.2 | 300.7 | 189.1 | 189.131 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5509 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.3 | 1.2 | 300.5 | 189.2 | 189.2278 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5097 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.3 | 1.2 | 300.5 | 189.1 | 189.0628 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997.1637 | 27.3 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.2 | 1.2 | 300.4 | 189.2 | 189.1792 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.3 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.2 | 1.2 | 300.4 | 189.1 | 189.1 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4355 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.025 | 250 | 299.2 | 1.2 | 300.4 | 189.1 | 189.1 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.5 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.0246 | 246.255 | 299.2 | 1.2164 | 300.4 | 189.2 | 189.2 | 0 | 13 | 949 | 956 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997.836 | 27.4476 | 26.3 | 1.2
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.2 | 1.3 | 300.5 | 189.3 | 189.2706 | 0 | 12.9 | 950 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3054 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.3 | 1.3 | 300.6 | 189.4 | 189.3743 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3836 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3846 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.33 | 0.024 | 240 | 299.3 | 1.3 | 300.6 | 189.2 | 189.1749 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4 | 26.3 | 1.3
2008-05-07 | 9:42 | 2.1 | 0.024 | 240 | 299.2 | 1.3 | 300.5 | 189.4 | 189.3705 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0549 | 0.0019 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4426 | 26.3 | 1.3
2008-05-07 | 9:42 | 1.623 | 0.016 | 160.3087 | 292.1 | 1.3 | 293.4 | 189.2 | 189.2229 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.2 | 26 | 5 | 195 | 59.3 | 0.05431 | 0.0018 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5222 | 26.3 | 1.3
2008-05-07 | 9:42 | 1.2 | 0.0104 | 104.1 | 245.5 | 1.3 | 246.8 | 151.1 | 151.1353 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.046382 | 0.0018 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5957 | 26.3 | 1.3
2008-05-07 | 9:42 | 1.2 | 0.013 | 130 | 186.2 | 1.2836 | 187.5 | 109.4 | 109.417 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.036531 | 0.0018 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.358 | 26.3 | 1.3
2008-05-07 | 9:42 | 1.186 | 0.013 | 130 | 156.1 | 1.2 | 157.3 | 97.6 | 97.6406 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.031268 | 0.0017 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3599 | 26.3 | 1.2
2008-05-07 | 9:42 | 1.18 | 0.013 | 130 | 154.7 | 1.2 | 155.9 | 95.2 | 95.2286 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.030984 | 0.0017 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4 | 26.3 | 1.2
2008-05-07 | 9:42 | 1.18 | 0.013 | 130 | 153.9 | 1.2 | 155.1 | 94.9 | 94.9 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.030751 | 0.001684 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3817 | 26.3 | 1.2
2008-05-07 | 9:42 | 1.18 | 0.013 | 130 | 152 | 1.1838 | 153.2 | 94.5 | 94.4504 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.030384 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3025 | 26.3 | 1.2
2008-05-07 | 9:42 | 1.18 | 0.013 | 130 | 151.5 | 1.1 | 152.6 | 94.4 | 94.4 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0303 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.4533 | 26.3 | 1.1
2008-05-07 | 9:42 | 1.18 | 0.0128 | 128.15 | 151.4 | 1.1 | 152.5 | 94.2 | 94.2 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0302 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1.1
2008-05-07 | 9:42 | 1.18 | 0.012 | 120 | 151.3 | 1.1 | 152.4 | 94.2 | 94.2 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0302 | 0.0016 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1.1
2008-05-07 | 9:42 | 1.175 | 0.012 | 120 | 151.3 | 1 | 152.3 | 94.1 | 94.1293 | 0 | 13 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0302 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1
2008-05-07 | 9:43 | 1.166 | 0.012 | 120 | 151.3 | 1 | 152.3 | 94.1 | 94.1 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0302 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5 | 26.3 | 1
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.9836 | 152.3 | 94.1 | 94.1492 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0302 | 0.0015 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3545 | 26.3 | 1
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.9 | 152.2 | 94 | 94 | 0 | 12.9 | 949 | 956 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0301 | 0.001484 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.9
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.8836 | 152.2 | 94.1 | 94.0815 | 0 | 12.9 | 950 | 956 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0301 | 0.0014 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3311 | 26.3 | 0.9
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.8 | 152.2 | 93.9 | 93.9259 | 0 | 13 | 949 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0301 | 0.0014 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3794 | 26.3 | 0.8
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.7 | 152.1 | 94.1 | 94.0719 | 0 | 12.9 | 949 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0301 | 0.0014 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2209 | 26.3 | 0.7
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.7 | 152.1 | 94 | 94 | 0 | 13 | 949 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.03 | 0.001384 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4749 | 26.3 | 0.7
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.7 | 152.1 | 93.9 | 93.9148 | 0 | 12.9 | 949 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 0.7
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.7 | 152.1 | 94 | 93.9892 | 0 | 13 | 950 | 957 | 940 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5463 | 26.3 | 0.7
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.8 | 93.8 | 0 | 12.9 | 949 | 957 | 940 | 19 | 0.29 | 0.01 | 979 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4128 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.9 | 0 | 13 | 950 | 957 | 940 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.5712 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.165 | 0.012 | 120 | 151.5 | 0.6 | 152.1 | 93.8 | 93.8228 | 0 | 12.9 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.03 | 0.0013 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.6 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.17 | 0.012 | 120 | 151.5 | 0.6 | 152.1 | 93.8 | 93.8263 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0299 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4681 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.17 | 0.012 | 120 | 151.5 | 0.6 | 152.1 | 93.9 | 93.8716 | 0 | 13 | 949 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0299 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.17 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.8798 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0299 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2952 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.17 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.8702 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0299 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1369 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.17 | 0.012 | 120 | 151.4 | 0.6 | 152 | 94 | 93.9864 | 0 | 12.9 | 949 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.029884 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2822 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.165 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.8 | 93.8 | 0 | 13 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0298 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3602 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.9264 | 0 | 12.9 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0298 | 0.0012 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.8859 | 0 | 13 | 949 | 956 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.4 | 0.6 | 152 | 93.9 | 93.8713 | 0 | 12.9 | 950 | 956 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.9 | 93.9067 | 0 | 13 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2432 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.7 | 93.7199 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.2 | 0.6 | 151.8 | 93.9 | 93.8737 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.7 | 93.7289 | 0 | 13 | 949 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.8 | 93.7804 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3472 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.8 | 93.8271 | 0 | 13 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.029784 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4528 | 26.3528 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.3 | 0.6 | 151.9 | 93.9 | 93.888 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.4041 | 26.368 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 151.1 | 0.6 | 151.7 | 93.8 | 93.8 | 0 | 13 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.029684 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.3 | 0.6 | 150.9 | 93.7 | 93.7381 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0296 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.3 | 0.6 | 150.9 | 93.9 | 93.8835 | 0 | 13 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3836 | 0.029584 | 0.0011 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1303 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.3 | 0.6 | 150.9 | 93.7 | 93.7 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3443 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.3 | 0.6 | 150.9 | 93.8 | 93.8 | 0 | 13 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.5439 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.3 | 0.6 | 150.9 | 93.7 | 93.7 | 0 | 13 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3927 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.2 | 0.6 | 150.8 | 93.8 | 93.7751 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 27 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.3135 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7277 | 0 | 13 | 949 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.8 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2344 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.8 | 93.7825 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2448 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7283 | 0 | 13 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2759 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.6 | 93.6 | 0 | 12.9 | 949 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 192 | 26.2 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.2032 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7 | 0 | 13 | 950 | 957 | 940 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 191.8 | 26.8 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.2825 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.6 | 93.6376 | 0 | 12.9 | 949 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26.2 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.3617 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.8 | 93.7804 | 0 | 13 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 998 | 27.4409 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150 | 0.6 | 150.6 | 93.7 | 93.6723 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26.8 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31 | 53 | 1033 | 997.8363 | 27.4598 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 31.73 | 53 | 1033 | 997 | 27.3014 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.5 | 93.549 | 0 | 13 | 949 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26.2 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.8 | 93.7739 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 27 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.2423 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7269 | 0 | 13 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26.8 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3108 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.8 | 93.7806 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.5 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.16 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.7 | 93.7 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.5 | 26.3 | 0.6
2008-05-07 | 9:43 | 1.059 | 0.012 | 120 | 150.1 | 0.6 | 150.7 | 93.8 | 93.7934 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.0295 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3512 | 26.3 | 0.6
2008-05-07 | 9:43 | 0.576 | 0.0062 | 61.573 | 148.8 | 0.6 | 149.4 | 93.8 | 93.8 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.029187 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.6
2008-05-07 | 9:43 | 0.135 | -0.0031 | -31.3062 | 95.7 | 0.6 | 96.3 | 73.1 | 73.0892 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.019567 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.333 | 26.3 | 0.6
2008-05-07 | 9:43 | 0.022 | 0.001 | 9.759 | 20.7 | 0.5 | 21.2 | 17.4 | 17.3596 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 191 | 26 | 5 | 195 | 59.3 | 0.005139 | 0.001 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4122 | 26.3 | 0.5
2008-05-07 | 9:43 | 0.01 | -0.0025 | -24.6644 | 5 | 0.4836 | 5.5 | 3 | 3.0474 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190.8 | 26.2 | 5 | 195 | 59.3 | 0.001984 | 0.000984 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4914 | 26.3 | 0.5
2008-05-07 | 9:43 | 0.01 | -0.003 | -30 | 3.9 | 0.4 | 4.3 | 0.3 | 0.2923 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 27 | 5 | 195 | 59.3 | 0.001767 | 0.0009 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.4295 | 26.3 | 0.4
2008-05-07 | 9:43 | 0.01 | -0.0025 | -24.5455 | 3 | 0.4 | 3.4 | -0.3 | 0 | 0 | 12.9 | 949 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26.8 | 5 | 195 | 59.3 | 0.001484 | 0.0009 | 0.0001 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3502 | 26.3 | 0.4
2008-05-07 | 9:43 | 0.01 | -0.002 | -20 | 2.5 | 0.3 | 2.8 | -0.6 | 0 | 0 | 12.9 | 950 | 958 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.001384 | 0.0009 | 0.0001 | 0 |  | n/a |  | 31.27 | 53 | 1033 | 997 | 27.3 | 26.3 | 0.3
2008-05-07 | 9:43 | 0.01 | -0.002 | -20 | 2.4 | 0.28363 | 2.7 | -0.8 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0013 | 0.0008 | 8.4e-05 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2836 | 26.3 | 0.3
2008-05-07 | 9:43 | 0.01 | -0.002 | -20 | 2.4 | 0.2 | 2.6 | -0.8 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0013 | 0.0008 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1263 | 26.3 | 0.2
2008-05-07 | 9:43 | 0.01 | -0.002 | -20 | 2.3 | 0.1 | 2.4 | -0.9 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0012 | 0.0008 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.2329 | 26.3 | 0.1
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.9 | 0.083633 | 2 | -0.9 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.001184 | 0.0008 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | 0.1
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.3 | -0.0164 | 1.3 | -1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.001 | 0.0008 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | 0
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.2 | -0.1 | 1.2 | -1.1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.001 | 0.0007 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | -0.1
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.3 | -0.1164 | 1.3 | -1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.000984 | 0.0007 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.3 | 26.3 | -0.1
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.3 | -0.2 | 1.3 | -1.1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0 | 0 |  | n/a |  | 31.73 | 53 | 1033 | 996.164 | 27.3 | 26.3 | -0.2
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.2 | -0.3 | 1.2 | -1.1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 190 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3419 | 26.3 | -0.3
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.2 | -0.3 | 1.2 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189.8 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3789 | 26.3 | -0.3
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1.1 | -0.3164 | 1.1 | -1.1 | 0 | 0 | 12.9 | 951 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0009 | 0.0007 | 0 | 0 |  | n/a |  | 31.2707 | 53 | 1033 | 997 | 27.3009 | 26.3 | -0.3
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 1 | -0.4 | 1 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 940 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.2164 | 0.0008 | 0.0006 | 0 | 0 |  | n/a |  | 31.7297 | 53 | 1033 | 997 | 27.5374 | 26.3 | -0.4
2008-05-07 | 9:44 | 0.01 | -0.002 | -20 | 0.9 | -0.4164 | 0.9 | -1.1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0 | 0 |  | n/a |  | 31.27 | 53 | 1033 | 997 | 27.483 | 26.3 | -0.4
2008-05-07 | 9:44 | 0.003 | -0.002 | -20 | 0.8 | -0.5 | 0.8 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2862 | 26.3 | -0.5
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.5 | 0.8 | -1.2 | 0 | 0 | 12.9 | 951 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1341 | 26.3 | -0.5
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.5164 | 0.8 | -1.1 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0008 | 0.0006 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2921 | 26.3 | -0.5
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.7 | -0.6 | 0.7 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.000784 | 0.0006 | 0 | 0 |  | n/a |  | 31.73 | 53 | 1033 | 997.1637 | 27.2246 | 26.3 | -0.6
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.7 | -0.6 | 0.7 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 997.836 | 27.2547 | 26.3 | -0.6
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.6 | 0.6 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 997 | 27.3 | 26.3 | -0.6
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.61637 | 0.6 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 996.8363 | 27.3 | 26.3 | -0.6
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.7 | 0.6 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.3 | 26.3 | -0.7
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.8 | 0.6 | -1.2 | 0 | 0 | 12.9 | 951 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.1574 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.8 | 0.6 | -1.3 | 0 | 0 | 12.9 | 950 | 956 | 940 | 20 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0006 | 0 | 0 |  | n/a |  | 32 | 53 | 1033 | 996 | 27.1 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.6 | -0.8 | 0.6 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.000584 | 0 | 0 |  | n/a |  | 31.27 | 53 | 1033 | 996.164 | 27.0401 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.7 | -0.8 | 0.7 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0007 | 0.0005 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 26.9367 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.7 | -0.8 | 0.7 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0005 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2532 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.8 | 0.8 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0005 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2325 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.8 | 0.8 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0005 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1532 | 26.3 | -0.8
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.9 | 0.8 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0006 | 0.0005 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996 | 27.178 | 26.3 | -0.9
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.8 | -0.9 | 0.8 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.4 | 26.3 | -0.9
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.9 | -0.9 | 0.9 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.4 | 26.3 | -0.9
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 0.9 | -0.9164 | 0.9 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 940 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.2164 | 0.0006 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.3365 | 26.3 | -0.9
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 1.3 | -1 | 1.3 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0006 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.3 | 26.3 | -1
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 1.3 | -1 | 1.3 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.212 | 26.3 | -1
2008-05-07 | 9:44 | 0 | -0.002 | -20 | 1.4 | -1 | 1.4 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 26.9047 | 26.3 | -1
2008-05-07 | 9:44 | 0 | -0.0016 | -16.3882 | 1.4 | -1 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 996 | 27.2192 | 26.3 | -1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 996.164 | 27.2407 | 26.3 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1612 | 26.3388 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 27 | 5 | 195 | 59.3 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.1361 | 26.3819 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 27 | 5 | 195 | 59.2164 | 0.0006 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 53 | 1033 | 997 | 27.2945 | 26.3028 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.4 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.3765 | 26.3 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.2 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.3444 | 26.3 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.1 | 1.5 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.3 | 26.3 | -1.1
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.2 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.3 | 26.3 | -1.2
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.2 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.8 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.3 | 26.3 | -1.2
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 19.2 | 0.29 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.2 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.012 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26.8 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 26.9 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 26.9308 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 998 | 27.0101 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.4 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.0893 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.4 | -1.3 | 1.4 | -1.3 | 0 | 0 | 12.9 | 951 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.237 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.3 | 1.5 | -1.4 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0005 | 0.0002 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.3 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.3 | 1.5 | -1.3 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0001 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.2734 | 26.3 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.3 | 1.5 | -1.4 | 0 | 0 | 12.9 | 950 | 957 | 941 | 20 | 0.3 | 0.01 | 980 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0001 | 0 | 0 |  | n/a |  | 31 | 53 | 1032 | 997 | 27.1878 | 26.3061 | -1.3
2008-05-07 | 9:44 | 0 | -0.001 | -10 | 1.5 | -1.3 | 1.5 | -1.4 | 0 | 0 | 12.9 | 950 | 957 | 940 | 19.8 | 0.3 | 0.01 | 979 | -34 | 0 | 27 | 189 | 26 | 5 | 195 | 59.3 | 0.0004 | 0.0001 | 0 | 0 |  | n/a |  | 31 | 53 | 1029.4476 | 996 | 27.0295 | 26.3853 | -1.3
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
No Processing Errors Detected
Test Date | 2008-05-07
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 50568
Version | 6.907
CO Span(%) | 0.12
CO2 Span(%) | 12
C6H14 Span(ppm) | 325
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 1495
NO2 Span(ppm) | 507.8
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 957.0
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 2
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 9:23
Test End Time | 9:44
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -34
Average Ambient Pressure (mbar) | 979.622
Average Relative Humidity (%) | 19.081
Average Absolute Humidity (grains/lb dry air) | 0.288
Average Kh Factor | 0.886
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Percent Time in NTE Zone | 0
Total Number of NTE Windows | 0
Total Number of non-consecutive NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/kW-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/07/2008 09:23:19.836 - None Found
Warnings:
0X0000 - 05/07/2008 09:23:19.836 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 121
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-05-07 | 9:23 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## THC Lin
Test Date:
Test Time:
5/7/2008
9:23:20
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## CO2 Lin
Test Date:
Test Time:
5/7/2008
9:23:20
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## CO Lin
Test Date:
Test Time:
5/7/2008
9:23:20
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## NO Lin
Test Date:
Test Time:
5/7/2008
9:23:20
NDUV Serial No:
208
SEMTECH Serial Number:
D06-SDS01

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
5/7/2008
9:23:20
FID Serial No:
123
SEMTECH Serial Number:
D06-SDS01

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
5/7/2008
9:23:20
NDUV Serial No:
208
AMBII Serial No:
50568
SEMTECH Serial Number:
D06-SDS01

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
12 | Rm = | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554
| C.F.= | 0.098697 | 0.19768 | 0.29695 | 0.39651 | 0.49636 | 0.5965 | 0.69693 | 0.79766 | 0.89868
| gas = | 1.1844 | 2.3722 | 3.5634 | 4.7581 | 5.9563 | 7.158 | 8.3632 | 9.5719 | 10.7842
Ro for CO2= | 0.8795