## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 22
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 37249
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | NO2 NDUV Serial Number: |  |  | 208 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-10-20 |  | 518.6 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 17:26
Test software: | 2.018 22
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0.2 | 0.2 |  |  | 0 | 0.17 | 0.17
90 | 105 | 1 | 51.86 | 51.3 | -0.6 | -1.16 |  | 2.09 | 2.1 | 0.01 | 0.48
150 | 165 | 2 | 103.72 | 104.5 | 0.8 | 0.77 |  | 4.18 | 4.1 | -0.08 | -1.91
210 | 225 | 3 | 155.58 | 156.7 | 1.1 | 0.71 |  | 6.27 | 6.1 | -0.17 | -2.71
270 | 285 | 4 | 207.44 | 208.5 | 1.1 | 0.53 |  | 8.36 | 8.1 | -0.26 | -3.11
330 | 345 | 5 | 259.3 | 261.6 | 2.3 | 0.89 |  | 10.45 | 10.1 | -0.35 | -3.35
390 | 405 | 6 | 311.16 | 312.9 | 1.7 | 0.55 |  | 12.54 | 12.2 | -0.34 | -2.71
450 | 465 | 7 | 363.02 | 364.8 | 1.8 | 0.5 |  | 14.63 | 14.2 | -0.43 | -2.94
510 | 525 | 8 | 414.88 | 416.3 | 1.4 | 0.34 |  | 16.72 | 16.3 | -0.42 | -2.51
570 | 585 | 9 | 466.74 | 467.4 | 0.7 | 0.15 |  | 18.81 | 18.3 | -0.51 | -2.71
630 | 645 | 10 | 518.6 | 520.1 | 1.5 | 0.29 |  | 20.9 | 20.5 | -0.4 | -1.91
690 | 705 | 9 | 466.74 | 470.6 | 3.9 | 0.84 |  | 18.81 | 18.4 | -0.41 | -2.18
750 | 765 | 8 | 414.88 | 419.1 | 4.2 | 1.01 |  | 16.72 | 16.4 | -0.32 | -1.91
810 | 825 | 7 | 363.02 | 367.4 | 4.4 | 1.21 |  | 14.63 | 14.3 | -0.33 | -2.26
870 | 885 | 6 | 311.16 | 315.2 | 4 | 1.29 |  | 12.54 | 12.3 | -0.24 | -1.91
930 | 945 | 5 | 259.3 | 264.1 | 4.8 | 1.85 |  | 10.45 | 10.2 | -0.25 | -2.39
990 | 1005 | 4 | 207.44 | 210.8 | 3.4 | 1.64 |  | 8.36 | 8.2 | -0.16 | -1.91
1050 | 1065 | 3 | 155.58 | 158.7 | 3.1 | 1.99 |  | 6.27 | 6.2 | -0.07 | -1.12
1110 | 1125 | 2 | 103.72 | 106.3 | 2.6 | 2.51 |  | 4.18 | 4.2 | 0.02 | 0.48
1170 | 1185 | 1 | 51.86 | 52.8 | 0.9 | 1.74 |  | 2.09 | 2.2 | 0.11 | 5.26
1230 | 1245 | 0 | 0 | 0.5 | 0.5 |  |  | 0 | 0.18 | 0.18
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | 1 | 5.2 |  | Intercept | 0.1 | 0.21
|  |  |  | Slope | 1 | 0.99 - 1.01 |  | Slope | 0.97 | 0.99 - 1.01 | Fail
|  |  |  | SEE | 1.427 | 5.186 |  | SEE | 0.0736 | 0.209
|  |  |  | r2 | 0.9999 | 0.998 |  | r2 | 0.9999 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.2 | 0.3 | 0.3 | -1 | 0 | 0.2 | 13 | 956 | 954 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0011 | 0.0012 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 954 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.0012 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2088 | 10.4347 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4555 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 957 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3352 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1039 | 27.5 | 0.2
2008-10-20 | 17:26 | 0 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 957 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.2517 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.002 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2211 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.353 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3699 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0011 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3072 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3842 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13.1 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2795 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.2165 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.3 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2207 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.1763 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13.1 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3467 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2209 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3758 | 27.5 | 0.2
2008-10-20 | 17:26 | -0.01 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3556 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.0009 | -8.5179 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3673 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | 0 | -0.19151 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13.1 | 956 | 952 | 942 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | 0 | 0 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.0007 | -6.5325 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3642 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.008 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3587 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.001 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3498 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.396 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2793 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3188 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.008 | -0.0009 | -8.5488 | -0.6 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13.1 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.0009 | 0.0011 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3583 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | 0 | -0.20851 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.001021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3646 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | 0 | 0 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.0007 | -6.6277 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.0015 | -14.9708 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.002 | -20 | -0.5 | 0.2 | 0.2 | -1.1 | 0 | 0.2 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.344 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.002 | -20 | -0.5 | 0.2 | 0.2 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3789 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.002 | -20 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3018 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.01 | -0.002 | -20 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.007 | -0.002 | -20 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.0015 | -15.1199 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2207 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.2
2008-10-20 | 17:27 | -0.006 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.361 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.362 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2209 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2209 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.01 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2212 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | -0.004 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.5 | 0.3
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13.1 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3 | 27.4449 | 0.3
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.4 | 0.3 | 0.3 | -1 | 0 | 0.1 | 13 | 956 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0009 | 0.001 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3323 | 27.4323 | 0.3
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13.1 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0009 | 0.001 | 2.1e-05 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.4 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2209 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1040.4577 | 1005 | 10.4 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.5 | 0.2 | 0.2 | -1.1 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0009 | 0.001 | 0 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3592 | 27.5 | 0.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | -0.3 | 8.684 | 8.7 | -1.1 | 0 | 0.14 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0009 | 0.001 | 0.000574 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.264 | 27.5 | 8.7
2008-10-20 | 17:27 | 0 | -0.001 | -10 | 0.3 | 39.3535 | 39.7 | -0.9 | 0 | 0.9 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.001079 | 0.0011 | 0.002167 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1097 | 27.5 | 39.4
2008-10-20 | 17:27 | 0 | -0.001 | -10 | 0.8 | 48.954 | 49.8 | -0.9 | 0 | 1.6 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.001753 | 0.001753 | 0.002858 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2445 | 27.5 | 49
2008-10-20 | 17:27 | 0 | -0.001 | -10 | 1.2 | 50.2139 | 51.4 | -0.9 | 0 | 1.89 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.2207 | 0.002459 | 0.002379 | 0.0029 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 50.2
2008-10-20 | 17:27 | 0 | -0.001 | -10 | 1.2 | 50.4 | 51.6 | -0.8 | 0 | 2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0025 | 0.0024 | 0.0029 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 50.4
2008-10-20 | 17:27 | 0 | -0.001 | -10 | 1.2 | 50.5 | 51.7 | -0.8 | 0 | 2 | 13.1 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0026 | 0.0025 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3036 | 27.5 | 50.5
2008-10-20 | 17:27 | 0.002 | -0.001 | -10 | 1.1 | 50.5 | 51.6 | -0.8 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0027 | 0.002679 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3807 | 27.5 | 50.5
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.1 | 50.6792 | 51.8 | -0.7 | 0 | 2.1 | 13.1 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2792 | 0.0028 | 0.002779 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3422 | 27.5 | 50.7
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1 | 50.779 | 51.8 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.002879 | 0.0028 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.335 | 27.5 | 50.8
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.5 | 50.8 | 52.3 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2209 | 0.003079 | 0.0029 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3879 | 27.5 | 50.8
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.5 | 50.8 | 52.3 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0031 | 0.002979 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3108 | 27.5 | 50.8
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.6 | 50.9 | 52.5 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.003179 | 0.003 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3664 | 27.5 | 50.9
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.6 | 50.9 | 52.5 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0032 | 0.003079 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 50.9
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.8 | 51 | 52.8 | -0.7 | 0 | 2.1 | 13 | 960 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0033 | 0.0031 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3793 | 27.5 | 51
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.8 | 51 | 52.8 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0034 | 0.0032 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3022 | 27.5 | 51
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.8 | 51.1 | 52.9 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0034 | 0.0032 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 51.1
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.7 | 51.1 | 52.8 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.003479 | 0.003279 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.352 | 27.5 | 51.1
2008-10-20 | 17:27 | 0.01 | -0.001 | -10 | 1.7 | 51.179 | 52.9 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0035 | 0.0033 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3708 | 27.5 | 51.2
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.7 | 51.2 | 52.9 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.2 | 0.0035 | 0.0034 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 51.2
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.6 | 51.2 | 52.8 | -0.7 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.003579 | 0.0034 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 51.2
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.6 | 51.2 | 52.8 | -0.8 | 0 | 2.1 | 13 | 960 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0036 | 0.0034 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3605 | 27.5 | 51.2
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.6 | 51.279 | 52.9 | -0.8 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0036 | 0.0034 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3623 | 27.4623 | 51.3
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.6 | 51.3 | 52.9 | -0.8 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2788 | 0.0036 | 0.003479 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3148 | 27.4148 | 51.3
2008-10-20 | 17:28 | 0.01 | -0.001 | -10 | 1.6 | 51.3 | 52.9 | -0.9 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2205 | 0.0036 | 0.0035 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3919 | 27.4919 | 51.3
2008-10-20 | 17:28 | 0.009 | -0.001 | -10 | 1.6 | 51.3 | 52.9 | -0.9 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0037 | 0.0035 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3312 | 27.5 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.3 | 52.9 | -0.9 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2208 | 0.0037 | 0.0035 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4373 | 27.4542 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.221 | 52.8 | -0.9 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0037 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.4227 | 51.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.2 | 52.8 | -0.9 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.003779 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.4998 | 51.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.2 | 52.8 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.221 | 0.0038 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2127 | 27.4225 | 51.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.2789 | 52.9 | -1 | 0 | 2.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2211 | 0.0038 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2094 | 27.4547 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.6 | 51.3 | 52.9 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0038 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3318 | 27.5 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.7 | 51.3 | 53 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0038 | 0.0036 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.491 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 1.9 | 51.3 | 53.2 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.003879 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4139 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3367 | 27.4 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4404 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4824 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.21 | 10.3 | 27.4053 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0039 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3718 | 27.4718 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.003979 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.351 | 27.5 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.3 | 53.3 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.004 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3261 | 27.5 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.379 | 53.4 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.004 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3968 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2 | 51.4 | 53.4 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.004 | 0.0037 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3197 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.4 | 53.5 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.004 | 0.003779 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.2425 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.4 | 53.5 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.004 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2346 | 27.4654 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.4 | 53.5 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.004 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4117 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.3207 | 53.5 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.004 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4888 | 51.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.3792 | 53.6 | -1 | 0 | 2.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2208 | 0.004 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.366 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.4 | 53.6 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.4 | 53.6 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4797 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.4 | 53.6 | -1 | 0 | 2.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4025 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.479 | 53.7 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.279 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3254 | 27.4 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.5 | 53.7 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4517 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.421 | 53.6 | -1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 51.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.5 | 53.7 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040.0706 | 1005 | 10.3059 | 27.5 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.521 | 53.7 | -1.1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.0041 | 0.0038 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.383 | 27.5 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.2 | 51.5 | 53.7 | -1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.2205 | 0.0041 | 0.003879 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.5 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.5 | 53.6 | -1.1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0041 | 0.0039 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.4627 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.579 | 53.7 | -1.1 | 0 | 2.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.221 | 0.0041 | 0.0039 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3856 | 27.4144 | 51.6
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.6 | 53.7 | -1.1 | 0 | 2.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0041 | 0.0039 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040.4348 | 1005 | 10.3083 | 27.4917 | 51.6
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.1 | 51.5 | 53.6 | -1.1 | 0 | 2.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0041 | 0.0039 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3688 | 27.5 | 51.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 2.8 | 65.724 | 68.6 | -1 | 0 | 2.19 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.004258 | 0.0039 | 0.003811 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3542 | 27.4542 | 65.7
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 5.8 | 90.9952 | 96.8 | -1 | 0 | 2.97 | 13 | 960 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.005175 | 0.004295 | 0.005312 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4227 | 91
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 5.9 | 99.0228 | 104.9 | -0.9 | 0 | 3.69 | 13 | 960 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.005937 | 0.005116 | 0.005779 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4995 | 99
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 5.5 | 102.1 | 107.6 | -0.9 | 0 | 3.9 | 13 | 960 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0064 | 0.005679 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4543 | 27.5 | 102.1
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 5.2 | 102.1 | 107.3 | -0.9 | 0 | 4 | 13 | 960 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.006716 | 0.006016 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4456 | 27.4456 | 102.1
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 5 | 102.179 | 107.2 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.006879 | 0.0062 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4314 | 27.4314 | 102.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.9 | 102.2 | 107.1 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.007 | 0.006379 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.483 | 27.4915 | 102.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.9 | 102.3 | 107.2 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0071 | 0.0065 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3284 | 27.4142 | 102.3
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.8 | 102.479 | 107.3 | -0.9 | 0 | 4.1 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.279 | 0.0072 | 0.0066 | 0.0059 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4258 | 27.4629 | 102.5
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.7 | 103.211 | 107.9 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.007279 | 0.006679 | 0.005979 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4599 | 27.5 | 103.2
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.7 | 103.4 | 108.1 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.007379 | 0.006779 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 103.4
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.5 | 103.579 | 108.1 | -0.9 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.221 | 0.0074 | 0.0069 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4 | 27.5 | 103.6
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.5 | 103.6 | 108.1 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0075 | 0.006979 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.7904 | 10.5431 | 27.5 | 103.6
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.4 | 103.6 | 108 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2212 | 0.007579 | 0.007079 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4513 | 103.6
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.2 | 103.7 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.007679 | 0.007179 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5485 | 27.4258 | 103.7
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4.1 | 103.7 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.007779 | 0.0073 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3971 | 27.5 | 103.7
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4 | 103.7 | 107.7 | -1 | 0 | 4.1 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0078 | 0.007379 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.32 | 27.5 | 103.7
2008-10-20 | 17:28 | 0 | -0.001 | -10 | 4 | 103.8792 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.007879 | 0.0074 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3571 | 27.5 | 103.9
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.9 | 104 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0079 | 0.0075 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.9 | 104.1 | 108 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.007979 | 0.0075 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.1
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.8 | 104.1 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.008 | 0.0076 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.1
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.7 | 104.1791 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.008079 | 0.007679 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.7 | 104.2 | 107.9 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0081 | 0.0077 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.4572 | 104.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.7 | 104.2 | 107.9 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0081 | 0.0077 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.4 | 104.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.6 | 104.2 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0081 | 0.007779 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 104.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.6 | 104.2 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.008179 | 0.0078 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4 | 27.4742 | 104.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.6 | 104.2795 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0082 | 0.0078 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2052 | 10.4 | 27.4486 | 104.3
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.5 | 104.3 | 107.8 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2793 | 0.0082 | 0.007879 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 104.3
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.5 | 104.3 | 107.8 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0082 | 0.0079 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3944 | 27.4056 | 104.3
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.5 | 104.3 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2788 | 0.0082 | 0.0079 | 0.006 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.3172 | 27.4828 | 104.3
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.5 | 104.379 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.008279 | 0.0079 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.21 | 10.3 | 27.5 | 104.4
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.5 | 104.4 | 107.9 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0083 | 0.0079 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.337 | 27.463 | 104.4
2008-10-20 | 17:29 | -0.008 | -0.001 | -10 | 3.4 | 104.4 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2209 | 0.0083 | 0.007979 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4142 | 104.4
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.4 | 104.479 | 107.9 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0083 | 0.008 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4913 | 104.5
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.3 | 104.5 | 107.8 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0083 | 0.008 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.537 | 27.5 | 104.5
2008-10-20 | 17:29 | -0.01 | -0.0004 | -3.6378 | 3.2 | 104.6 | 107.8 | -1 | 0 | 4.1 | 13 | 958 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.008379 | 0.0081 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5087 | 27.4544 | 104.6
2008-10-20 | 17:29 | -0.01 | 0 | 0 | 3.2 | 104.6 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2205 | 0.0084 | 0.0081 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4227 | 104.6
2008-10-20 | 17:29 | -0.01 | -0.0003 | -2.8774 | 3.1 | 104.6793 | 107.8 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0081 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4995 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3 | 104.7 | 107.7 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.008179 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3231 | 27.4231 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3 | 104.7 | 107.7 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0082 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4082 | 27.4 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3 | 104.7 | 107.7 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0082 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4377 | 27.4311 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.9 | 104.7 | 107.6 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0084 | 0.0082 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3163 | 27.5 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.9 | 104.7 | 107.6 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0082 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4704 | 27.5 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.8 | 104.7 | 107.5 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0084 | 0.0082 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5625 | 27.5 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.8 | 104.6 | 107.4 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.008479 | 0.008279 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5204 | 27.5 | 104.6
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.7 | 104.7 | 107.4 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.0085 | 0.0083 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.7
2008-10-20 | 17:29 | -0.005 | -0.001 | -10 | 2.7 | 104.7 | 107.4 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0085 | 0.0083 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.7
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 2.6 | 104.679 | 107.3 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0085 | 0.0083 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 104.7
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 2.5 | 104.7 | 107.2 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0085 | 0.0084 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4516 | 104.7
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 2.5 | 104.7 | 107.2 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0085 | 0.0084 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 104.7
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 2.4 | 104.8 | 107.2 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.008579 | 0.0084 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4026 | 27.4 | 104.8
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 2.3 | 104.8 | 107.1 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0086 | 0.008479 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4798 | 27.4 | 104.8
2008-10-20 | 17:29 | -0.005 | -0.001 | -10 | 2.3 | 104.8 | 107.1 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3293 | 27.4 | 104.8
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.3 | 104.8 | 107.1 | -1 | 0 | 4.1 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.2681 | 27.434 | 104.8
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.3 | 104.7209 | 107 | -1 | 0 | 4.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2791 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.7908 | 10.4 | 27.4888 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.3 | 104.7 | 107 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4117 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.3 | 104.7 | 107 | -1 | 0 | 4.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4654 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 2.3 | 104.7 | 107 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4574 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.0002 | -1.8228 | 2.3 | 104.7 | 107 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0086 | 0.0085 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3803 | 27.4197 | 104.7
2008-10-20 | 17:29 | -0.01 | 0 | 0 | 2.2 | 104.7 | 106.9 | -1 | 0 | 4.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0087 | 0.0086 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3032 | 27.4968 | 104.7
2008-10-20 | 17:29 | -0.01 | -0.0005 | -4.5046 | 2.5 | 117.452 | 119.9 | -1 | 0 | 4.28 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0087 | 0.0086 | 0.006795 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.21 | 10.374 | 27.5 | 117.5
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 4.8 | 146.767 | 151.5 | -0.9 | 0 | 5.07 | 13 | 960 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.009774 | 0.009316 | 0.008516 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.79 | 10.3489 | 27.5 | 146.8
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 4.5 | 152.965 | 157.5 | -0.9 | 0 | 5.73 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.010853 | 0.010511 | 0.008858 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3282 | 27.5 | 153
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 4 | 154.479 | 158.5 | -0.9 | 0 | 5.99 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.011379 | 0.011079 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 154.5
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.9 | 154.5 | 158.4 | -0.8 | 0 | 6 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2211 | 0.0114 | 0.011179 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 154.5
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.8 | 154.6 | 158.4 | -0.9 | 0 | 6.1 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.011579 | 0.011279 | 0.0089 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4404 | 154.6
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.6 | 154.6 | 158.2 | -0.8 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.221 | 0.011679 | 0.011479 | 0.008979 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 154.6
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.5 | 155.411 | 158.9 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.011779 | 0.011579 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4139 | 27.4 | 155.4
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.5 | 155.6 | 159.1 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0118 | 0.011679 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.491 | 27.4 | 155.6
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.4 | 155.7 | 159.1 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0119 | 0.011779 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3636 | 27.4 | 155.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.4 | 155.7 | 159.1 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.012 | 0.0118 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3907 | 27.4454 | 155.7
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.3 | 155.8 | 159.1 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.012 | 0.0119 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4775 | 27.4775 | 155.8
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.3 | 155.9 | 159.2 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0121 | 0.012 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4005 | 27.4005 | 155.9
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.3 | 155.9791 | 159.3 | -0.9 | 0 | 6.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.012 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4768 | 156
2008-10-20 | 17:29 | -0.01 | -0.001 | -10 | 3.3 | 156.0209 | 159.3 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2791 | 0.0122 | 0.0121 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4461 | 156
2008-10-20 | 17:29 | -0.004 | -0.001 | -10 | 3.7 | 156.179 | 159.9 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.012358 | 0.0121 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4309 | 27.4 | 156.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.8 | 156.2 | 160 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0124 | 0.012179 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4921 | 27.4 | 156.2
2008-10-20 | 17:29 | 0 | -0.001 | -10 | 3.8 | 156.3 | 160.1 | -1 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.012479 | 0.0122 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4151 | 27.4 | 156.3
2008-10-20 | 17:30 | 0 | -0.0005 | -5.3977 | 3.7 | 156.3 | 160 | -1 | 0 | 6.1 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0125 | 0.0123 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 156.3
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 156.3 | 160 | -1 | 0 | 6.1 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0125 | 0.0123 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3608 | 27.4 | 156.3
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 156.3208 | 160 | -1 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0126 | 0.0124 | 0.009 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4167 | 156.3
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.6 | 156.479 | 160.1 | -0.9 | 0 | 6.1 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.279 | 0.0126 | 0.0124 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4939 | 156.5
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 156.4 | 160.1 | -1 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0126 | 0.0124 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.513 | 27.429 | 156.4
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 156.579 | 160.3 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0126 | 0.0124 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5038 | 27.4 | 156.6
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 156.6 | 160.3 | -1 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.012679 | 0.012479 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4252 | 27.4252 | 156.6
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.6 | 156.6 | 160.2 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0127 | 0.0125 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4976 | 27.4976 | 156.6
2008-10-20 | 17:30 | -0.008 | 0 | 0 | 3.6 | 156.679 | 160.3 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0127 | 0.0125 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4205 | 27.4205 | 156.7
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 3.5 | 156.6 | 160.1 | -1 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0127 | 0.012579 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4566 | 156.6
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 3.5 | 156.6 | 160.1 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0127 | 0.0126 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 156.6
2008-10-20 | 17:30 | -0.007 | 0 | 0 | 3.4 | 156.6 | 160 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0128 | 0.0126 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4891 | 156.6
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.3 | 156.6 | 159.9 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0128 | 0.0127 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.412 | 156.6
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.2 | 156.6 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2792 | 0.0128 | 0.0127 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5303 | 27.4652 | 156.6
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.2 | 156.7 | 159.9 | -0.9 | 0 | 6.1 | 13.1 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.012779 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3884 | 27.5 | 156.7
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.1 | 156.7 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.0128 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.1585 | 27.5 | 156.7
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.1 | 156.7 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0129 | 0.0128 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3898 | 27.5 | 156.7
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.8 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0129 | 0.0128 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3262 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.8 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2789 | 0.0129 | 0.0128 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.8 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.012879 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.272 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.8 | 159.8 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2103 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.779 | 159.8 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0129 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3645 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3 | 156.8 | 159.8 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2207 | 0.0129 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.013 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4731 | 27.4635 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.013 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5863 | 27.4137 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.013 | 0.0129 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5092 | 27.4908 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.013 | 0.012979 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.432 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.013 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.9 | 156.8 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.013 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3777 | 27.5 | 156.8
2008-10-20 | 17:30 | -0.003 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.013 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3006 | 27.5 | 156.8
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 957 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.013 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3765 | 27.5 | 156.8
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.013079 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 156.8
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 2.8 | 156.8795 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 156.9
2008-10-20 | 17:30 | -0.004 | 0 | 0 | 2.8 | 156.8207 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3924 | 27.5 | 156.8
2008-10-20 | 17:30 | -0.004 | 0 | 0 | 2.8 | 156.8792 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.0131 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3153 | 27.5 | 156.9
2008-10-20 | 17:30 | -0.01 | 0 | 0 | 2.8 | 156.821 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3617 | 27.4383 | 156.8
2008-10-20 | 17:30 | -0.009 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.013 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4388 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.8 | 156.9 | 159.7 | -0.9 | 0 | 6.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.0131 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4159 | 27.5 | 156.9
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 957 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2791 | 0.0131 | 0.0131 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4935 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.0131 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3585 | 27.5 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.8 | 156.8 | 159.6 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0131 | 0.0131 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4521 | 156.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 2.8 | 156.8796 | 159.7 | -0.9 | 0 | 6.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2204 | 0.01318 | 0.0131 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.275 | 27.425 | 156.9
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.8 | 172.409 | 176.2 | -0.8 | 0 | 6.16 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.013279 | 0.0131 | 0.00999 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2043 | 27.5 | 172.4
2008-10-20 | 17:30 | 0 | -0.0003 | -2.7 | 5.2 | 198.5528 | 203.7 | -0.8 | 0 | 7.07 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.014214 | 0.013835 | 0.011514 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3585 | 27.5 | 198.6
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 4.7 | 205.0539 | 209.7 | -0.7 | 0 | 7.67 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2792 | 0.014937 | 0.014717 | 0.011858 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3436 | 27.5 | 205.1
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 4.5 | 205.879 | 210.4 | -0.7 | 0 | 7.99 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.015179 | 0.014979 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4342 | 27.5 | 205.9
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 4.3 | 206.453 | 210.7 | -0.7 | 0 | 8 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.015437 | 0.015237 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6787 | 27.5 | 206.5
2008-10-20 | 17:30 | 0 | -0.0004 | -3.9227 | 4.1 | 206.8 | 210.9 | -0.7 | 0 | 8.1 | 13.1 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0156 | 0.015479 | 0.0119 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5244 | 27.5 | 206.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 4.1 | 207.7 | 211.8 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2209 | 0.015779 | 0.015679 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3702 | 27.5 | 207.7
2008-10-20 | 17:30 | 0 | -0.0003 | -2.7607 | 4 | 207.7 | 211.7 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.015879 | 0.015779 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.342 | 27.458 | 207.7
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 4 | 207.7 | 211.7 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.016 | 0.0159 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 207.7
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 3.9 | 207.7 | 211.6 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.2 | 0.0161 | 0.016079 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 207.7
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 3.8 | 207.7 | 211.5 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0162 | 0.0161 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4735 | 207.7
2008-10-20 | 17:30 | 0 | -0.001 | -10 | 4 | 207.8 | 211.8 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0163 | 0.016279 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 207.8
2008-10-20 | 17:30 | 0 | -0.0006 | -5.6547 | 3.9 | 207.8 | 211.7 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2208 | 0.016479 | 0.016379 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4556 | 27.4722 | 207.8
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.8 | 207.979 | 211.8 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.016579 | 0.016479 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5852 | 27.4 | 208
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.8 | 208 | 211.8 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0166 | 0.016579 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3539 | 27.4 | 208
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 208.079 | 211.8 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0167 | 0.016679 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4591 | 208.1
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.7 | 208.1 | 211.8 | -0.7 | 0 | 8.1 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0167 | 0.0167 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3725 | 27.4637 | 208.1
2008-10-20 | 17:30 | 0 | 0 | 0 | 3.6 | 208.1 | 211.7 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2791 | 0.0168 | 0.016779 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4866 | 27.4134 | 208.1
2008-10-20 | 17:31 | 0 | 0 | 0 | 3.5 | 208.021 | 211.5 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.016879 | 0.016879 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.7904 | 10.4095 | 27.4905 | 208
2008-10-20 | 17:31 | 0 | 0 | 0 | 3.4 | 208.1 | 211.5 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0169 | 0.0169 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4 | 27.5 | 208.1
2008-10-20 | 17:31 | 0 | 0 | 0 | 3.4 | 208.021 | 211.4 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.017 | 0.017 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4897 | 27.4552 | 208
2008-10-20 | 17:31 | 0 | 0 | 0 | 3.4 | 207.979 | 211.4 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.017 | 0.017079 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.556 | 27.422 | 208
2008-10-20 | 17:31 | 0 | 0 | 0 | 3.4 | 208 | 211.4 | -0.6 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.017 | 0.0171 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4017 | 27.4992 | 208
2008-10-20 | 17:31 | 0 | -0.0006 | -6.1084 | 3.4 | 208.1792 | 211.6 | -0.6 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.017079 | 0.0171 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2084 | 10.4763 | 27.5 | 208.2
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.4 | 208.2 | 211.6 | -0.6 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.0171 | 0.0171 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5533 | 27.4467 | 208.2
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.4 | 208.3 | 211.7 | -0.6 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0171 | 0.0171 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5393 | 27.4303 | 208.3
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.4 | 211.9 | -0.6 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0171 | 0.017179 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4073 | 27.5 | 208.4
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.4 | 211.9 | -0.6 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4844 | 27.5 | 208.4
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.6 | 208.4 | 212 | -0.6 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4386 | 27.4386 | 208.4
2008-10-20 | 17:31 | -0.003 | -0.001 | -10 | 3.5 | 208.3202 | 211.8 | -0.7 | 0 | 8.1 | 13 | 959 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4384 | 27.4 | 208.3
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.5 | 208.4 | 211.9 | -0.7 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4844 | 27.4 | 208.4
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.5 | 208.4 | 211.9 | -0.7 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2795 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4068 | 27.4 | 208.4
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.5 | 208.4 | 211.9 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0172 | 0.0172 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2068 | 10.4 | 27.4705 | 208.4
2008-10-20 | 17:31 | -0.004 | -0.001 | -10 | 3.5 | 208.5 | 212 | -0.7 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.012 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3524 | 27.4524 | 208.5
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.421 | 211.9 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.012079 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3742 | 27.4247 | 208.4
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.579 | 212.1 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5963 | 27.5 | 208.6
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.5 | 212 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.442 | 27.5 | 208.5
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.6 | 212.1 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3439 | 27.5 | 208.6
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.6 | 212.1 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.3 | 0.017379 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3665 | 27.4667 | 208.6
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.5 | 208.6 | 212.1 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4792 | 27.4104 | 208.6
2008-10-20 | 17:31 | -0.004 | -0.001 | -10 | 3.5 | 208.6 | 212.1 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.325 | 27.4875 | 208.6
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.5 | 208.6 | 212.1 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3647 | 27.4353 | 208.6
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.5 | 208.6206 | 212.1 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0174 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 208.6
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.6788 | 212.1 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2788 | 0.0174 | 0.0174 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3811 | 27.4189 | 208.7
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.7792 | 212.2 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3039 | 27.4961 | 208.8
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.8 | 212.2 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3732 | 27.4268 | 208.8
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.9 | 212.3 | -0.8 | 0 | 8.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4503 | 27.4503 | 208.9
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.9 | 212.3 | -0.8 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4176 | 27.4725 | 208.9
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 208.979 | 212.4 | -0.8 | 0 | 8.1 | 13 | 958 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.21 | 10.2046 | 27.4 | 209
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.4 | 209 | 212.4 | -0.8 | 0 | 8.1 | 13 | 958 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2817 | 27.4 | 209
2008-10-20 | 17:31 | -0.002 | -0.001 | -10 | 3.4 | 209 | 212.4 | -0.7 | 0 | 8.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0175 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3589 | 27.4 | 209
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.4 | 208.9212 | 212.3 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.017579 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.436 | 27.436 | 208.9
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.4 | 208.9 | 212.3 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0175 | 0.0176 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4869 | 27.5 | 208.9
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 3.4 | 208.9 | 212.3 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0176 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4097 | 27.5 | 208.9
2008-10-20 | 17:31 | -0.001 | -0.001 | -10 | 3.3 | 208.9 | 212.2 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0176 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5348 | 27.4326 | 208.9
2008-10-20 | 17:31 | -0.009 | -0.001 | -10 | 3.3 | 208.921 | 212.2 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0176 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4664 | 27.4445 | 208.9
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.3 | 208.979 | 212.3 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0177 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3217 | 27.4783 | 209
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.2 | 208.9 | 212.1 | -0.7 | 0 | 8.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0176 | 0.0177 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3988 | 27.4012 | 208.9
2008-10-20 | 17:31 | -0.01 | -0.001 | -10 | 3.2 | 208.9 | 212.1 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.0177 | 0.0177 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3241 | 27.4759 | 208.9
2008-10-20 | 17:31 | -0.007 | -0.001 | -10 | 3.2 | 208.9788 | 212.2 | -0.7 | 0 | 8.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0177 | 0.017779 | 0.0121 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.247 | 27.447 | 209
2008-10-20 | 17:31 | 0 | -0.0006 | -5.7798 | 4.2 | 225.8162 | 230 | -0.7 | 0 | 8.27 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.017858 | 0.0178 | 0.01309 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.2603 | 27.4302 | 225.8
2008-10-20 | 17:31 | 0 | -0.0003 | -2.6557 | 6.4 | 252.15 | 258.5 | -0.6 | 0 | 9.1 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.279 | 0.018911 | 0.018532 | 0.014611 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3927 | 27.5 | 252.2
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 5.9 | 258.364 | 264.2 | -0.6 | 0 | 9.74 | 13 | 960 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.019716 | 0.019416 | 0.014979 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3158 | 27.5 | 258.4
2008-10-20 | 17:31 | 0 | -0.001 | -10 | 5.7 | 258.7795 | 264.5 | -0.6 | 0 | 9.98 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.019879 | 0.0196 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4835 | 27.5 | 258.8
2008-10-20 | 17:31 | 0 | -0.0002 | -2.164 | 5.4 | 259 | 264.4 | -0.6 | 0 | 10.03 | 13 | 960 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0202 | 0.02 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5619 | 27.5 | 259
2008-10-20 | 17:31 | 0 | 0 | 0 | 5.2 | 259.7333 | 265 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2208 | 0.020379 | 0.020179 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4848 | 27.4848 | 259.7
2008-10-20 | 17:31 | 0 | 0 | 0 | 5.2 | 260 | 265.2 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.279 | 0.020479 | 0.0203 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4078 | 27.4078 | 260
2008-10-20 | 17:31 | 0 | 0 | 0 | 5 | 260.1 | 265.1 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.020579 | 0.0204 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4699 | 260.1
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.9 | 260.179 | 265.1 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0207 | 0.0206 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 260.2
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.8 | 260.2 | 265 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.279 | 0.0208 | 0.0207 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4244 | 27.4756 | 260.2
2008-10-20 | 17:31 | 0 | 0 | 0 | 5.1 | 260.3 | 265.4 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0209 | 0.0208 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5015 | 27.4015 | 260.3
2008-10-20 | 17:31 | 0 | 0 | 0 | 5.1 | 260.3789 | 265.5 | -0.6 | 0 | 10.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0209 | 0.0208 | 0.015 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5786 | 27.4786 | 260.4
2008-10-20 | 17:31 | 0 | 0 | 0 | 5.1 | 260.479 | 265.6 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.021 | 0.0209 | 0.015079 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4884 | 27.5 | 260.5
2008-10-20 | 17:31 | 0 | 0 | 0 | 5 | 260.6795 | 265.7 | -0.6 | 0 | 10.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2795 | 0.021079 | 0.020979 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4671 | 260.7
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.9 | 260.7 | 265.6 | -0.6 | 0 | 10.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.021179 | 0.021079 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.42 | 27.41 | 260.7
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.9 | 260.7 | 265.6 | -0.6 | 0 | 10.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2792 | 0.0212 | 0.0211 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5742 | 27.4871 | 260.7
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.9 | 260.7 | 265.6 | -0.7 | 0 | 10.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0212 | 0.0211 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4072 | 27.5 | 260.7
2008-10-20 | 17:31 | 0 | 0 | 0 | 4.9 | 260.879 | 265.8 | -0.7 | 0 | 10.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.021279 | 0.0212 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3414 | 27.5 | 260.9
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.9 | 260.821 | 265.7 | -0.7 | 0 | 10.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.0213 | 0.0212 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4815 | 260.8
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 260.7788 | 265.6 | -0.7 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.021379 | 0.0213 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4042 | 260.8
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 260.9 | 265.7 | -0.7 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0214 | 0.0213 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 260.9
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 261 | 265.8 | -0.7 | 0 | 10.1 | 13 | 959 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.021479 | 0.0214 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.45 | 27.45 | 261
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 261.1 | 265.9 | -0.7 | 0 | 10.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0215 | 0.0214 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4729 | 27.4729 | 261.1
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 261.1 | 265.9 | -0.7 | 0 | 10.1 | 13 | 958 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0215 | 0.0214 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3958 | 27.4042 | 261.1
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.8 | 261.1 | 265.9 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.021579 | 0.0215 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3187 | 27.4813 | 261.1
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.7 | 261.2 | 265.9 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0216 | 0.0215 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3585 | 27.5 | 261.2
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.7 | 261.3 | 266 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2208 | 0.0216 | 0.0216 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4712 | 27.5 | 261.3
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.6 | 261.221 | 265.8 | -0.6 | 0 | 10.1 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0217 | 0.0216 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5618 | 27.5 | 261.2
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.6 | 261.279 | 265.9 | -0.6 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0217 | 0.021679 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3303 | 27.5 | 261.3
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.5 | 261.4 | 265.9 | -0.5 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0217 | 0.0217 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.367 | 27.433 | 261.4
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.5 | 261.5791 | 266.1 | -0.5 | 0 | 10.1 | 13 | 959 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0218 | 0.0217 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 261.6
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.4 | 261.5211 | 265.9 | -0.5 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0218 | 0.0218 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3788 | 27.4212 | 261.5
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.4 | 261.6 | 266 | -0.5 | 0 | 10.1 | 13 | 958 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0218 | 0.0218 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3015 | 27.4985 | 261.6
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.4 | 261.7 | 266.1 | -0.5 | 0 | 10.1 | 13 | 958 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0218 | 0.0218 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.5 | 261.7
2008-10-20 | 17:32 | 0 | -0.0008 | -7.853 | 4.4 | 261.7 | 266.1 | -0.5 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0218 | 0.0219 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3527 | 27.5 | 261.7
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.4 | 261.7 | 266.1 | -0.5 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0219 | 0.0219 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.7
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.4 | 261.7 | 266.1 | -0.5 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0219 | 0.0219 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4931 | 261.7
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.4 | 261.6 | 266 | -0.5 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0219 | 0.0219 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.416 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.4 | 261.6 | 266 | -0.5 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0219 | 0.0219 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4611 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.3 | 261.6 | 265.9 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2788 | 0.0219 | 0.021979 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.3 | 261.6 | 265.9 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0219 | 0.022 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.415 | 27.485 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.3 | 261.5 | 265.8 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2209 | 0.022 | 0.022 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4919 | 27.4081 | 261.5
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.2 | 261.6 | 265.8 | -0.3 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.022 | 0.022 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.431 | 27.469 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.2 | 261.6 | 265.8 | -0.3 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.022 | 0.022079 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4464 | 27.5 | 261.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.2 | 261.521 | 265.7 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.022 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4761 | 27.4761 | 261.5
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.2 | 261.5 | 265.7 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.022 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4009 | 261.5
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.1 | 261.6 | 265.7 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4782 | 261.6
2008-10-20 | 17:32 | 0 | -0.0004 | -4.0595 | 4.1 | 261.6 | 265.7 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.6
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.1 | 261.5 | 265.6 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.279 | 0.0221 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.5
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.1 | 261.4788 | 265.6 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.2788 | 0.0221 | 0.0221 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.212 | 10.4193 | 27.4904 | 261.5
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.1 | 261.4212 | 265.5 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.022179 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.7884 | 10.5736 | 27.4132 | 261.4
2008-10-20 | 17:32 | -0.006 | 0 | 0 | 4.1 | 261.4 | 265.5 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.5361 | 27.4639 | 261.4
2008-10-20 | 17:32 | -0.01 | 0 | 0 | 4.1 | 261.4 | 265.5 | -0.4 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.459 | 27.5 | 261.4
2008-10-20 | 17:32 | -0.01 | 0.0006 | 5.9372 | 4.1 | 261.4789 | 265.6 | -0.4 | 0 | 10.1 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2112 | 10.4 | 27.5 | 261.5
2008-10-20 | 17:32 | -0.01 | 0.0006 | 5.875 | 4.1 | 261.5 | 265.6 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0221 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.5
2008-10-20 | 17:32 | -0.01 | 0 | 0 | 4.2 | 261.6 | 265.8 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0221 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.5 | 261.6
2008-10-20 | 17:32 | -0.01 | 0 | 0 | 4.2 | 261.7793 | 266 | -0.4 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0222 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3503 | 27.5 | 261.8
2008-10-20 | 17:32 | -0.01 | 0 | 0 | 4.2 | 261.7208 | 265.9 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0222 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3803 | 27.5 | 261.7
2008-10-20 | 17:32 | -0.006 | -0.0008 | -7.8363 | 4.2 | 261.7 | 265.9 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0222 | 0.0222 | 0.0151 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5923 | 27.5 | 261.7
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.2 | 264.249 | 268.4 | -0.3 | 0 | 10.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0222 | 0.0222 | 0.015258 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.438 | 27.5 | 264.2
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.9 | 286.731 | 291.7 | -0.3 | 0 | 10.43 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.022358 | 0.022379 | 0.016569 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.5163 | 27.5 | 286.7
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 5.2 | 305.3344 | 310.6 | -0.2 | 0 | 11.33 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.023733 | 0.023733 | 0.017617 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 305.3
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 6.1 | 310.1653 | 316.3 | -0.2 | 0 | 11.83 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.024395 | 0.024216 | 0.0179 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5752 | 27.5 | 310.2
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 5.9 | 310.6 | 316.5 | -0.2 | 0 | 12 | 13 | 960 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.024758 | 0.024637 | 0.0179 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.421 | 27.5 | 310.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 5.7 | 310.6 | 316.3 | -0.3 | 0 | 12.1 | 13 | 960 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0249 | 0.0249 | 0.0179 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5333 | 27.4333 | 310.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 5.5 | 310.6 | 316.1 | -0.3 | 0 | 12.1 | 13 | 960 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025079 | 0.024979 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 310.6
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 5.1 | 310.7788 | 315.9 | -0.3 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025379 | 0.0254 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4209 | 310.8
2008-10-20 | 17:32 | 0 | -0.001 | -10 | 4.9 | 310.9 | 315.8 | -0.4 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0255 | 0.025579 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4981 | 310.9
2008-10-20 | 17:32 | -0.007 | -0.001 | -10 | 4.8 | 311.8792 | 316.7 | -0.4 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0256 | 0.0257 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 311.9
2008-10-20 | 17:32 | -0.01 | -0.001 | -10 | 4.7 | 312 | 316.7 | -0.4 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.3 | 0.0257 | 0.025879 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 312
2008-10-20 | 17:32 | -0.01 | -0.001 | -10 | 4.6 | 312 | 316.6 | -0.4 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0258 | 0.0259 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 312
2008-10-20 | 17:32 | -0.01 | -0.001 | -10 | 4.6 | 312.1789 | 316.8 | -0.4 | 0 | 12.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0258 | 0.0259 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5867 | 27.4934 | 312.2
2008-10-20 | 17:32 | -0.01 | -0.001 | -10 | 4.6 | 312.0209 | 316.6 | -0.5 | 0 | 12.1 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0258 | 0.026 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4324 | 27.4162 | 312
2008-10-20 | 17:32 | -0.01 | -0.001 | -10 | 4.5 | 312.1 | 316.6 | -0.5 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0259 | 0.026079 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7045 | 27.4609 | 312.1
2008-10-20 | 17:32 | -0.003 | -0.0009 | -8.879 | 4.5 | 312.1 | 316.6 | -0.5 | 0 | 12.1 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.025979 | 0.0261 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7861 | 27.5 | 312.1
2008-10-20 | 17:32 | 0 | -0.0001 | -0.81452 | 4.5 | 312.021 | 316.5 | -0.4 | 0 | 12.17 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.026 | 0.0262 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 312
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.5 | 312 | 316.5 | -0.4 | 0 | 12.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.026 | 0.0262 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.5 | 312
2008-10-20 | 17:32 | 0 | 0 | 0 | 4.5 | 312.1 | 316.6 | -0.4 | 0 | 12.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0261 | 0.0262 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4307 | 312.1
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.1208 | 316.6 | -0.4 | 0 | 12.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0261 | 0.0263 | 0.018 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5069 | 27.4 | 312.1
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.279 | 316.8 | -0.4 | 0 | 12.2 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.026179 | 0.0263 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4471 | 27.4236 | 312.3
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.379 | 316.9 | -0.4 | 0 | 12.2 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0262 | 0.0263 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5988 | 27.5 | 312.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.5 | 317 | -0.3 | 0 | 12.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0262 | 0.0264 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4447 | 27.5 | 312.5
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.5 | 317 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0263 | 0.0264 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.445 | 312.5
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 312.4 | 316.9 | -0.4 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0263 | 0.026479 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 312.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.4 | 312.5 | 316.9 | -0.3 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0263 | 0.0265 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4188 | 27.4094 | 312.5
2008-10-20 | 17:33 | 0 | -0.0002 | -2.4481 | 4.4 | 312.579 | 317 | -0.3 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0264 | 0.0265 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5731 | 27.4866 | 312.6
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 312.6 | 316.9 | -0.3 | 0 | 12.2 | 13 | 958 | 952 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.3 | 0.0264 | 0.0266 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4726 | 27.4363 | 312.6
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 312.6 | 316.9 | -0.3 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0264 | 0.0266 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4816 | 27.4408 | 312.6
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 312.7 | 317 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0264 | 0.0266 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5641 | 27.5 | 312.7
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 312.7 | 317 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.026479 | 0.0266 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4099 | 27.5 | 312.7
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 312.779 | 317.1 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0265 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4277 | 312.8
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.4 | 312.821 | 317.2 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0265 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4494 | 27.4 | 312.8
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.4 | 312.879 | 317.3 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0265 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.4 | 312.9
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 312.8791 | 317.4 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.026579 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5037 | 27.4 | 312.9
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 312.9 | 317.4 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5808 | 27.4 | 312.9
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 312.9 | 317.4 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0266 | 0.0267 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5421 | 27.4 | 312.9
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 312.9 | 317.4 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0266 | 0.026779 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5351 | 27.4 | 312.9
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313 | 317.5 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5756 | 27.4 | 313
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.0793 | 317.6 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4213 | 27.4 | 313.1
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.3 | 317.8 | -0.5 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4436 | 27.4 | 313.3
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5208 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5979 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0266 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.026679 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4955 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.2 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 313.2
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.3 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4194 | 27.4 | 313.3
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.5 | 313.3 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0268 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6509 | 27.4 | 313.3
2008-10-20 | 17:33 | 0 | -0.0006 | -5.9328 | 4.5 | 313.321 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5177 | 27.4 | 313.3
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 313.3 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 313.3
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 313.4 | 317.9 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.415 | 313.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 313.4 | 317.9 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4919 | 313.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 313.4 | 317.9 | -0.4 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5383 | 27.4308 | 313.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.5 | 313.4 | 317.9 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5537 | 27.4 | 313.4
2008-10-20 | 17:33 | 0 | -0.0004 | -4.0657 | 4.4 | 313.4 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.4233 | 313.4
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.4 | 313.421 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0267 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.4996 | 313.4
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 313.479 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192.3 | 28 | 5 | 195 | 59.3 | 0.026779 | 0.0269 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.4226 | 313.5
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 313.5212 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192.7 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.027 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5 | 27.4 | 313.5
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 313.5 | 317.8 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.027 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4678 | 27.4 | 313.5
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.3 | 313.5793 | 317.9 | -0.4 | 0 | 12.2 | 13 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.027 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4185 | 27.4093 | 313.6
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.2 | 313.5208 | 317.7 | -0.4 | 0 | 12.2 | 13 | 958 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.027 | 0.0181 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5728 | 27.4864 | 313.5
2008-10-20 | 17:33 | 0 | -0.001 | -10 | 4.2 | 315.0265 | 319.2 | -0.3 | 0 | 12.2 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0268 | 0.027 | 0.018179 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4729 | 27.4365 | 315
2008-10-20 | 17:33 | 0 | -0.0009 | -9.317 | 5.7 | 339.558 | 345.2 | -0.2 | 0 | 12.48 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.027116 | 0.027079 | 0.019648 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4813 | 27.4 | 339.6
2008-10-20 | 17:33 | 0 | -0.0001 | -0.88533 | 6 | 358.883 | 364.8 | -0.2 | 0 | 13.42 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.028474 | 0.028532 | 0.020795 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5466 | 27.4 | 358.9
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.6 | 362.3 | 367.9 | -0.1 | 0 | 13.92 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.029 | 0.0291 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3153 | 27.4 | 362.3
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.1 | 362.9536 | 368.1 | -0.1 | 0 | 14.1 | 13.1 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.029279 | 0.029479 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3721 | 27.4 | 363
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.9 | 363.3581 | 368.3 | -0.1 | 0 | 14.12 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.029479 | 0.029679 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4985 | 27.4 | 363.4
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.8 | 363.379 | 368.2 | -0.2 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.029579 | 0.029879 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 363.4
2008-10-20 | 17:33 | -0.007 | -0.0001 | -0.66102 | 4.7 | 363.421 | 368.1 | -0.2 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0297 | 0.029979 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5929 | 27.4 | 363.4
2008-10-20 | 17:33 | -0.004 | -0.0009 | -9.1356 | 5.5 | 363.5 | 369 | -0.2 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.029979 | 0.0301 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4386 | 27.4 | 363.5
2008-10-20 | 17:33 | 0 | -0.0003 | -2.6393 | 5.4 | 363.6 | 369 | -0.2 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.03 | 0.0302 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5157 | 27.4578 | 363.6
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.4 | 363.679 | 369.1 | -0.2 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0301 | 0.0302 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4952 | 27.4651 | 363.7
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.3 | 363.779 | 369.1 | -0.2 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.030179 | 0.0303 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3361 | 27.4 | 363.8
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.2 | 363.8788 | 369.1 | -0.2 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.030279 | 0.030479 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5676 | 27.4 | 363.9
2008-10-20 | 17:33 | 0 | 0 | 0 | 5.1 | 363.921 | 369 | -0.2 | 0 | 14.2 | 13.1 | 958 | 952 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0303 | 0.0305 | 0.021 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4673 | 27.4664 | 363.9
2008-10-20 | 17:33 | -0.006 | 0 | 0 | 5 | 364 | 369 | -0.2 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.030379 | 0.0306 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3565 | 27.4565 | 364
2008-10-20 | 17:33 | -0.006 | 0 | 0 | 4.9 | 364 | 368.9 | -0.3 | 0 | 14.2 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0304 | 0.0307 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3206 | 27.4 | 364
2008-10-20 | 17:33 | 0 | 0 | 0 | 4.8 | 364.1 | 368.9 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0305 | 0.030779 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3978 | 27.4 | 364.1
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.7 | 364.079 | 368.8 | -0.3 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0306 | 0.030879 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3251 | 27.4 | 364.1
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.7 | 364.2 | 368.9 | -0.3 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0306 | 0.0309 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4563 | 27.4 | 364.2
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.6 | 364.3586 | 369 | -0.3 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0307 | 0.030979 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5416 | 27.4 | 364.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.6 | 364.4 | 369 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.2208 | 0.0307 | 0.031 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3873 | 27.4 | 364.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.6 | 364.4 | 369 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0307 | 0.031 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.233 | 27.4 | 364.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.5 | 364.4 | 368.9 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.030779 | 0.0311 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3212 | 27.4 | 364.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.5 | 364.321 | 368.8 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0308 | 0.0311 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4 | 27.4 | 364.3
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.5 | 364.4 | 368.9 | -0.3 | 0 | 14.2 | 13.1 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0308 | 0.0312 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041.2487 | 1005 | 10.4297 | 27.4 | 364.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.4 | 364.5 | 368.9 | -0.3 | 0 | 14.2 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0309 | 0.0312 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5838 | 27.4 | 364.5
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.4 | 364.5 | 368.9 | -0.3 | 0 | 14.2 | 13.1 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0309 | 0.0312 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3931 | 27.4 | 364.5
2008-10-20 | 17:34 | 0 | -0.0008 | -7.5063 | 4.4 | 364.5 | 368.9 | -0.4 | 0 | 14.2 | 13 | 958 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0309 | 0.0312 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.346 | 27.4 | 364.5
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.6 | 369.1 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0309 | 0.031279 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.377 | 27.4 | 364.6
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.7 | 369.2 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.030979 | 0.0313 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3001 | 27.4001 | 364.7
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.7 | 369.2 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.031 | 0.0313 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3771 | 27.4771 | 364.7
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.7 | 369.2 | -0.4 | 0 | 14.2 | 13 | 959 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.031 | 0.031379 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4459 | 364.7
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.7 | 369.2 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.031 | 0.0314 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4314 | 27.4 | 364.7
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.8 | 369.3 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.0314 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5091 | 27.4 | 364.8
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.4 | 364.7 | 369.1 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.0314 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5863 | 27.4 | 364.7
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.4 | 364.8791 | 369.3 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.0314 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4097 | 27.4 | 364.9
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.4 | 364.821 | 369.2 | -0.5 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0311 | 0.0314 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3406 | 27.4 | 364.8
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.8 | 369.3 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.031479 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3823 | 27.4 | 364.8
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.5 | 364.879 | 369.4 | -0.5 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.0315 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3052 | 27.4 | 364.9
2008-10-20 | 17:34 | 0 | -0.0003 | -2.5332 | 4.5 | 365 | 369.5 | -0.4 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0311 | 0.0315 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3719 | 27.4 | 365
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.4 | 365 | 369.4 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.031179 | 0.0315 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.4 | 365 | 369.4 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0312 | 0.0315 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365
2008-10-20 | 17:34 | 0 | -0.0002 | -2.3268 | 4.4 | 365.021 | 369.4 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0312 | 0.0315 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.4 | 365 | 369.4 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0312 | 0.0316 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.0212 | 369.3 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2212 | 0.0312 | 0.0316 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.1 | 369.4 | -0.5 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0312 | 0.0316 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3652 | 27.4 | 365.1
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.1 | 369.4 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0312 | 0.0316 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.312 | 27.4 | 365.1
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.179 | 369.5 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0313 | 0.0316 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3891 | 27.4 | 365.2
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.2 | 369.5 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0313 | 0.031679 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 365.2
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.358 | 369.7 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.221 | 0.0313 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.21 | 10.4867 | 27.4 | 365.4
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.4 | 369.7 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0313 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.559 | 27.4 | 365.4
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.4207 | 369.7 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.031379 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4048 | 27.4 | 365.4
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.4 | 369.7 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4 | 27.4 | 365.4
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.242 | 369.5 | -0.4 | 0 | 14.2 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.79 | 10.4 | 27.452 | 365.2
2008-10-20 | 17:34 | 0 | -0.001 | -10 | 4.3 | 365.279 | 369.6 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4291 | 27.4709 | 365.3
2008-10-20 | 17:34 | 0 | -0.0009 | -9.4249 | 4.3 | 365.4 | 369.7 | -0.4 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0317 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4875 | 27.4 | 365.4
2008-10-20 | 17:34 | 0 | -0.0001 | -0.8412 | 4.3 | 365.2209 | 369.5 | -0.4 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0318 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3332 | 27.4 | 365.2
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.2 | 365.1 | 369.3 | -0.4 | 0 | 14.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0318 | 0.0211 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3605 | 27.4 | 365.1
2008-10-20 | 17:34 | 0 | 0 | 0 | 4.3 | 367.47 | 371.7 | -0.4 | 0 | 14.2 | 13 | 959 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0314 | 0.0318 | 0.021258 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 367.5
2008-10-20 | 17:34 | 0 | 0 | 0 | 6 | 394.311 | 400.3 | -0.4 | 0 | 14.78 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.031995 | 0.032116 | 0.022774 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4296 | 27.4 | 394.3
2008-10-20 | 17:34 | 0 | 0 | 0 | 6.8 | 409.9152 | 416.7 | -0.4 | 0 | 15.65 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.033294 | 0.033315 | 0.023737 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5838 | 27.4 | 409.9
2008-10-20 | 17:34 | 0 | 0 | 0 | 6.1 | 415.1793 | 421.3 | -0.3 | 0 | 16.02 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.033938 | 0.034117 | 0.023959 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.7932 | 10.4623 | 27.4 | 415.2
2008-10-20 | 17:34 | 0 | 0 | 0 | 6 | 415.3791 | 421.4 | -0.3 | 0 | 16.17 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0341 | 0.034279 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 415.4
2008-10-20 | 17:34 | 0 | 0 | 0 | 6 | 415.5 | 421.5 | -0.3 | 0 | 16.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0342 | 0.0344 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 415.5
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.9 | 415.579 | 421.5 | -0.3 | 0 | 16.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034279 | 0.0345 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 415.6
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.9 | 415.6 | 421.5 | -0.3 | 0 | 16.2 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034379 | 0.0346 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 415.6
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.8 | 415.7582 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2209 | 0.034479 | 0.034779 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4547 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.7 | 415.8 | 421.5 | -0.4 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0346 | 0.0348 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4682 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.7 | 416 | 421.7 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034679 | 0.0349 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4182 | 27.4 | 416
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.7 | 415.9 | 421.6 | -0.4 | 0 | 16.3 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0347 | 0.035 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5725 | 27.4 | 415.9
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.6 | 415.9 | 421.5 | -0.4 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034779 | 0.035 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4731 | 27.4 | 415.9
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.6 | 415.8 | 421.4 | -0.4 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0348 | 0.0351 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4811 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.5 | 415.8 | 421.3 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034879 | 0.035179 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.4 | 415.8 | 421.2 | -0.4 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.034979 | 0.035279 | 0.024 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.4 | 415.8 | 421.2 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.035 | 0.0353 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4561 | 27.4 | 415.8
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.4 | 416 | 421.4 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.035 | 0.0354 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4491 | 27.4 | 416
2008-10-20 | 17:34 | 0 | 0 | 0 | 5.4 | 415.9 | 421.3 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0351 | 0.0354 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4737 | 27.4 | 415.9
2008-10-20 | 17:35 | 0 | -0.0004 | -4 | 5.4 | 415.9791 | 421.4 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0351 | 0.0354 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4069 | 27.4 | 416
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 5.4 | 416.1 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0352 | 0.0355 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5609 | 27.4 | 416.1
2008-10-20 | 17:35 | 0 | -0.0009 | -9.4588 | 5.4 | 416.1 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0352 | 0.035579 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4847 | 27.4 | 416.1
2008-10-20 | 17:35 | 0 | -0.0001 | -0.8677 | 5.4 | 416.2 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0353 | 0.0356 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4349 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | -0.0007 | -7.2676 | 5.4 | 416.2 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0353 | 0.0356 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.488 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | -0.0004 | -4.4463 | 5.5 | 416.2 | 421.7 | -0.3 | 0 | 16.3 | 13 | 958 | 954 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0353 | 0.0356 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4109 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.4 | 416.2 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.035379 | 0.0357 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4663 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.4 | 416.379 | 421.8 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0354 | 0.0357 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041.742 | 1005 | 10.4566 | 27.4 | 416.4
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.221 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.035479 | 0.0358 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.2 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0355 | 0.0358 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.2 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0355 | 0.0359 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4749 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.121 | 421.4 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0355 | 0.0359 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5 | 27.4 | 416.1
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.1 | 421.4 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0356 | 0.0359 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5292 | 27.4 | 416.1
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.1 | 421.4 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0356 | 0.0359 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041.2557 | 1005 | 10.5873 | 27.4 | 416.1
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.2 | 416.258 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0356 | 0.035979 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.433 | 27.4 | 416.3
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.2 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0356 | 0.036 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4606 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.1792 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0356 | 0.036 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2084 | 10.4245 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.279 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.035679 | 0.036 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3446 | 27.4 | 416.3
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.4212 | 421.7 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0357 | 0.036 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5755 | 27.4 | 416.4
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.242 | 421.5 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0357 | 0.0361 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5309 | 27.4 | 416.2
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.258 | 421.6 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0357 | 0.0361 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.21 | 10.4538 | 27.4 | 416.3
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.6791 | 422 | -0.3 | 0 | 16.3 | 13 | 958 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0358 | 0.0361 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4465 | 27.4 | 416.7
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.3 | 416.6 | 421.9 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0358 | 0.036179 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5994 | 27.4 | 416.6
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.2 | 416.621 | 421.8 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0358 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4453 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.001 | 0 | 0 | 5.2 | 416.679 | 421.9 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0358 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 416.7
2008-10-20 | 17:35 | -0.009 | 0 | 0 | 5.2 | 416.7 | 421.9 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0358 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4642 | 27.4 | 416.7
2008-10-20 | 17:35 | -0.01 | -0.0001 | -0.5802 | 5.2 | 416.7 | 421.9 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0358 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5815 | 27.4 | 416.7
2008-10-20 | 17:35 | -0.01 | -0.0009 | -9.1126 | 5.2 | 416.7 | 421.9 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.035879 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4273 | 27.4 | 416.7
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5.2 | 416.6 | 421.8 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.321 | 0.0359 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5.1 | 416.6789 | 421.8 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2211 | 0.0359 | 0.0362 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3593 | 27.4 | 416.7
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5.1 | 416.779 | 421.9 | -0.2 | 0 | 16.3 | 13 | 959 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.036279 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 416.8
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5.1 | 416.8 | 421.9 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.4 | 416.8
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5.1 | 416.5 | 421.6 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3721 | 27.4 | 416.5
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5 | 416.4 | 421.4 | -0.2 | 0 | 16.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 416.4
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5 | 416.5577 | 421.6 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4529 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5 | 416.6 | 421.6 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 5 | 416.4 | 421.4 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.0359 | 0.0363 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.4 | 416.4
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 4.9 | 416.4 | 421.3 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.036 | 0.036379 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5421 | 27.4 | 416.4
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 4.9 | 416.558 | 421.5 | -0.2 | 0 | 16.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.036 | 0.0364 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.535 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.01 | -0.001 | -10 | 4.9 | 416.6 | 421.5 | -0.2 | 0 | 16.3 | 13 | 958 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.036 | 0.0364 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5637 | 27.4 | 416.6
2008-10-20 | 17:35 | -0.006 | -0.001 | -10 | 4.9 | 416.3 | 421.2 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.036 | 0.0364 | 0.0241 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.3322 | 27.4 | 416.3
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 4.9 | 418.7529 | 423.7 | -0.3 | 0 | 16.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.036 | 0.0364 | 0.024258 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4993 | 27.4 | 418.8
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.6 | 446.6155 | 454.2 | -0.2 | 0 | 16.85 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.036854 | 0.036816 | 0.025875 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.6 | 27.4 | 446.6
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.6 | 461.9581 | 469.5 | -0.2 | 0 | 17.71 | 13 | 959 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.038337 | 0.038337 | 0.0267 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005.2096 | 10.5585 | 27.4 | 462
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.4 | 464.432 | 471.8 | -0.1 | 0 | 18.11 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.038479 | 0.038579 | 0.0269 | 0 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4043 | 27.4 | 464.4
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.3 | 464.958 | 472.3 | -0.1 | 0 | 18.2 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0386 | 0.0387 | 0.0269 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.325 | 27.4 | 465
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.2 | 465.2 | 472.4 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0387 | 0.0388 | 0.0269 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3522 | 27.4 | 465.2
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 7.1 | 465.4 | 472.5 | -0.1 | 0 | 18.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0388 | 0.038979 | 0.0269 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4588 | 27.4 | 465.4
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.6 | 466.2208 | 472.8 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039179 | 0.0394 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5869 | 27.4 | 466.2
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.5 | 466.3 | 472.8 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.039279 | 0.0395 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4327 | 27.4 | 466.3
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.4 | 466.3212 | 472.7 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2788 | 0.0393 | 0.0396 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5216 | 27.4 | 466.3
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.3 | 466.379 | 472.7 | -0.1 | 0 | 18.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0394 | 0.0397 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5241 | 27.4 | 466.4
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.2 | 466.3 | 472.5 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0395 | 0.039879 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.415 | 27.4 | 466.3
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6.1 | 466.3787 | 472.5 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039579 | 0.0399 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4919 | 27.4 | 466.4
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 6 | 466.321 | 472.3 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.039679 | 0.04 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.431 | 27.4 | 466.3
2008-10-20 | 17:35 | 0 | -0.001 | -10 | 5.9 | 466.458 | 472.4 | -0.1 | 0 | 18.3 | 13 | 957 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0397 | 0.0401 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 466.5
2008-10-20 | 17:35 | 0 | -0.0009 | -9.2147 | 5.9 | 466.658 | 472.6 | -0.1 | 0 | 18.3 | 13 | 958 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039779 | 0.0401 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4232 | 27.4 | 466.7
2008-10-20 | 17:35 | 0 | -0.0001 | -0.86049 | 5.8 | 466.7207 | 472.5 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0398 | 0.0402 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4983 | 27.4 | 466.7
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.7 | 466.5792 | 472.3 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0403 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3439 | 27.4 | 466.6
2008-10-20 | 17:35 | 0 | 0 | 0 | 5.7 | 466.6 | 472.3 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0403 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4105 | 27.4 | 466.6
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.7 | 466.7 | 472.4 | -0.2 | 0 | 18.3 | 13 | 957 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0399 | 0.0404 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4676 | 27.4 | 466.7
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.6 | 466.779 | 472.4 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.04 | 0.0404 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 466.8
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.6 | 466.9 | 472.5 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.04 | 0.0404 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 466.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.6 | 467 | 472.6 | -0.2 | 0 | 18.3 | 13 | 958 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.04 | 0.0405 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.6 | 467.0419 | 472.6 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0405 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.359 | 27.4 | 467
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.5 | 467.079 | 472.6 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0401 | 0.0405 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3181 | 27.4 | 467.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.5 | 467.0212 | 472.5 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2788 | 0.0401 | 0.0406 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3953 | 27.4 | 467
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.5 | 467.079 | 472.6 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0401 | 0.0406 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3276 | 27.4 | 467.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.4 | 467.1 | 472.5 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0402 | 0.040679 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3495 | 27.4 | 467.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.2 | 467.1 | 472.3 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0402 | 0.0407 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3733 | 27.4 | 467.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 5.1 | 467.1 | 472.2 | -0.2 | 0 | 18.3 | 13.1 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0402 | 0.040779 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3039 | 27.4 | 467.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 5 | 467.179 | 472.2 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0402 | 0.0408 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3811 | 27.4 | 467.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.9 | 467.279 | 472.2 | -0.1 | 0 | 18.3 | 13.1 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0408 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5165 | 27.4 | 467.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.9 | 467.2 | 472.1 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0409 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4584 | 27.4 | 467.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.9 | 467.3 | 472.2 | -0.1 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0409 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.225 | 27.4 | 467.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.3 | 472.1 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0403 | 0.0409 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3793 | 27.4 | 467.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.3 | 472.1 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0409 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.3 | 472.1 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.379 | 0.0403 | 0.0409 | 0.027 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.5792 | 472.4 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0409 | 0.027079 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467.6
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.6 | 472.4 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0403 | 0.0409 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467.6
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.8 | 467.679 | 472.5 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.040379 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 467.7
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.858 | 472.6 | -0.2 | 0 | 18.3 | 13.1 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5053 | 27.4 | 467.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.8212 | 472.5 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5702 | 27.4 | 467.8
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.8 | 472.5 | -0.2 | 0 | 18.3 | 13.1 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4931 | 27.4 | 467.8
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.8 | 472.5 | -0.2 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4159 | 27.4 | 467.8
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.821 | 472.5 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4613 | 27.4 | 467.8
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 468 | 472.7 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5 | 27.4 | 468
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.9 | 472.6 | -0.2 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0404 | 0.041 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4846 | 27.4 | 467.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.9207 | 472.6 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2793 | 0.0404 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4076 | 27.4 | 467.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.7 | 467.9 | 472.6 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2792 | 0.0404 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5388 | 27.4 | 467.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 4.6 | 467.979 | 472.6 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.040479 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5535 | 27.4 | 468
2008-10-20 | 17:36 | -0.009 | -0.0002 | -2.3058 | 4.6 | 468 | 472.6 | -0.3 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0405 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4052 | 27.4 | 468
2008-10-20 | 17:36 | -0.003 | -0.001 | -10 | 4.6 | 468.158 | 472.8 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.1056 | 27.4 | 468.2
2008-10-20 | 17:36 | 0 | -0.001 | -10 | 4.6 | 468.2 | 472.8 | -0.3 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4142 | 27.4 | 468.2
2008-10-20 | 17:36 | 0 | -0.0003 | -2.903 | 4.6 | 468.3 | 472.9 | -0.3 | 0 | 18.3 | 13.1 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4443 | 27.4 | 468.3
2008-10-20 | 17:36 | -0.002 | 0 | 0 | 4.6 | 468.3 | 472.9 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0411 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 468.3
2008-10-20 | 17:36 | -0.01 | 0 | 0 | 4.6 | 468.2 | 472.8 | -0.3 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0405 | 0.041179 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 468.2
2008-10-20 | 17:36 | -0.01 | 0 | 0 | 4.6 | 468.2 | 472.8 | -0.4 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0412 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.4 | 468.2
2008-10-20 | 17:36 | -0.01 | 0 | 0 | 4.6 | 468.1 | 472.7 | -0.4 | 0 | 18.3 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0405 | 0.0412 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4 | 27.4 | 468.1
2008-10-20 | 17:36 | -0.005 | 0 | 0 | 4.6 | 468.0793 | 472.7 | -0.4 | 0 | 18.3 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2793 | 0.040579 | 0.0412 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005.2068 | 10.483 | 27.4 | 468.1
2008-10-20 | 17:36 | 0 | -0.0007 | -6.9858 | 4.6 | 468 | 472.6 | -0.4 | 0 | 18.3 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0406 | 0.0412 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1006 | 10.5627 | 27.4 | 468
2008-10-20 | 17:36 | 0 | -0.001 | -10 | 4.6 | 468.158 | 472.8 | -0.4 | 0 | 18.36 | 13 | 957 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0406 | 0.0412 | 0.0271 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005.21 | 10.4083 | 27.4 | 468.2
2008-10-20 | 17:36 | 0 | -0.0006 | -6.0338 | 5.7 | 497.369 | 503.1 | -0.3 | 0 | 18.94 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.041095 | 0.041616 | 0.028774 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005.79 | 10.4 | 27.4 | 497.4
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.3 | 515.108 | 521.4 | -0.3 | 0 | 19.83 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.042632 | 0.043132 | 0.029837 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4 | 27.4 | 515.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.7 | 516.4791 | 523.1 | -0.2 | 0 | 20.21 | 13 | 958 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.042979 | 0.043379 | 0.0299 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4547 | 27.3727 | 516.5
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.7 | 517.8908 | 524.6 | -0.2 | 0 | 20.3 | 13.1 | 957 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0431 | 0.0435 | 0.0299 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.5909 | 27.3046 | 517.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.6 | 518.721 | 525.3 | -0.2 | 0 | 20.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.043279 | 0.043679 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.4366 | 27.3817 | 518.7
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.6 | 518.858 | 525.5 | -0.2 | 0 | 20.4 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0433 | 0.0437 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1042 | 1005 | 10.3412 | 27.4 | 518.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 7.2 | 518.942 | 526.1 | -0.2 | 0 | 20.4 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.043579 | 0.043879 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041.8892 | 1005 | 10.4438 | 27.4 | 518.9
2008-10-20 | 17:36 | 0 | 0 | 0 | 7.1 | 519.021 | 526.1 | -0.2 | 0 | 20.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.043679 | 0.043979 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6738 | 27.4 | 519
2008-10-20 | 17:36 | 0 | 0 | 0 | 7.1 | 519.1 | 526.2 | -0.2 | 0 | 20.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0437 | 0.044079 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5194 | 27.4 | 519.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 7 | 519.1 | 526.1 | -0.1 | 0 | 20.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0438 | 0.044179 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4326 | 27.4 | 519.1
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.9 | 519.221 | 526.1 | -0.1 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.043979 | 0.044279 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3554 | 519.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.9 | 519.2 | 526.1 | -0.1 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.044 | 0.0443 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 519.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.8 | 519.2 | 526 | -0.1 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.044079 | 0.044479 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 519.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.7 | 519.2 | 525.9 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0441 | 0.0445 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5522 | 27.3761 | 519.2
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.6 | 519.2581 | 525.9 | -0.2 | 0 | 20.4 | 13.1 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0446 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4402 | 27.3467 | 519.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.5 | 519.3 | 525.8 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0443 | 0.0447 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3 | 27.3304 | 519.3
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.4 | 519.379 | 525.8 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.044379 | 0.0448 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3221 | 27.3926 | 519.4
2008-10-20 | 17:36 | 0 | 0 | 0 | 6.4 | 519.4 | 525.8 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0444 | 0.0448 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5543 | 27.3152 | 519.4
2008-10-20 | 17:37 | 0 | 0 | 0 | 6.3 | 519.558 | 525.9 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.279 | 0.0444 | 0.044879 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4763 | 27.3619 | 519.6
2008-10-20 | 17:37 | 0 | -0.0002 | -2.1647 | 6.2 | 519.8 | 526 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2792 | 0.0444 | 0.0449 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4 | 27.3611 | 519.8
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 6.1 | 519.8 | 525.9 | -0.2 | 0 | 20.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0445 | 0.045 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4319 | 27.3 | 519.8
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 6 | 519.842 | 525.9 | -0.2 | 0 | 20.4 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.044579 | 0.045079 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.586 | 27.3 | 519.8
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.9 | 519.879 | 525.8 | -0.2 | 0 | 20.4 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0446 | 0.0451 | 0.03 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.6 | 27.3705 | 519.9
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.9 | 519.9 | 525.8 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2791 | 0.0446 | 0.045179 | 0.030079 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4567 | 27.3522 | 519.9
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.9 | 519.6209 | 525.5 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0446 | 0.0452 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.3249 | 27.3249 | 519.6
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 519.679 | 525.5 | -0.1 | 0 | 20.5 | 13.1 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0452 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4021 | 27.3979 | 519.7
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 519.621 | 525.4 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0447 | 0.0453 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4793 | 27.3207 | 519.6
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 519.6 | 525.4 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2795 | 0.0447 | 0.0453 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5564 | 27.3 | 519.6
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 519.6793 | 525.5 | 0 | 0 | 20.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2207 | 0.0447 | 0.0453 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5328 | 27.3336 | 519.7
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 519.8583 | 525.7 | 0 | 0 | 20.5 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.044779 | 0.0453 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4216 | 27.4 | 519.9
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.8 | 520.0579 | 525.9 | 0 | 0 | 20.5 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.045379 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004.2104 | 10.5759 | 27.4 | 520.1
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.6 | 520.2 | 525.8 | 0 | 0 | 20.5 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.0454 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.4698 | 27.4 | 520.2
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.5 | 520.6 | 526.1 | 0 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0448 | 0.0454 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1003 | 10.4846 | 27.3577 | 520.6
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.5 | 520.4 | 525.9 | 0 | 0 | 20.5 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.044879 | 0.0455 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.5417 | 27.3 | 520.4
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.4 | 520.3 | 525.7 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2209 | 0.0449 | 0.0455 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1041 | 1004 | 10.3102 | 27.3 | 520.3
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.4 | 520.1419 | 525.5 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.0449 | 0.0455 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040.3952 | 1004 | 10.3738 | 27.3 | 520.1
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.4 | 520.1 | 525.5 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.045579 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5019 | 27.3 | 520.1
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 5.4 | 520.021 | 525.4 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.0449 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.5719 | 27.3281 | 520
2008-10-20 | 17:37 | -0.007 | -0.001 | -9.7587 | 5.4 | 519.8 | 525.2 | 0 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3947 | 519.8
2008-10-20 | 17:37 | -0.01 | -0.0001 | -1.4393 | 5.4 | 520.1586 | 525.6 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0449 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2068 | 10.5 | 27.3175 | 520.2
2008-10-20 | 17:37 | -0.01 | -0.0007 | -6.8974 | 5.4 | 520.0417 | 525.4 | 0 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.044979 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7916 | 10.5596 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.0005 | -4.5769 | 5.4 | 520 | 525.4 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.0004 | -3.862 | 5.4 | 519.9 | 525.3 | -0.1 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.979 | 525.4 | -0.1 | 0 | 20.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0456 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.5 | 519.9 | 525.4 | -0.1 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0457 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.9 | 525.3 | -0.1 | 0 | 20.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.045 | 0.0457 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5546 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.9 | 525.3 | -0.1 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0457 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4775 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.9 | 525.3 | -0.1 | 0 | 20.5 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0457 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4005 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.258 | 525.7 | -0.2 | 0 | 20.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0457 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.553 | 27.3 | 520.3
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.3 | 525.7 | -0.2 | 0 | 20.5 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 520.3
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.1417 | 525.5 | -0.2 | 0 | 20.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7916 | 10.5374 | 27.3 | 520.1
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.1 | 525.5 | -0.2 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4166 | 27.3 | 520.1
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.179 | 525.6 | -0.2 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5707 | 27.3 | 520.2
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520.042 | 525.4 | -0.2 | 0 | 20.5 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2865 | 27.3627 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.921 | 525.3 | -0.2 | 0 | 20.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.2199 | 27.36 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520 | 525.4 | -0.3 | 0 | 20.5 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0451 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4171 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520 | 525.4 | -0.2 | 0 | 20.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0452 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4943 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.9 | 525.3 | -0.2 | 0 | 20.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0452 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5715 | 27.3715 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520 | 525.4 | -0.2 | 0 | 20.5 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0452 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3514 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 520 | 525.4 | -0.2 | 0 | 20.5 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0452 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5485 | 27.3 | 520
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 5.4 | 519.9207 | 525.3 | -0.2 | 0 | 20.5 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2793 | 0.0452 | 0.0458 | 0.0301 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4059 | 27.3 | 519.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 4.5 | 494.01 | 498.5 | -0.2 | 0 | 20.18 | 13 | 951 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.045042 | 0.0458 | 0.02855 | 7.9e-05 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5602 | 27.3 | 494
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.1 | 477.23 | 480.3 | -0.3 | 0 | 19.28 | 13 | 950 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.044742 | 0.045721 | 0.027642 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 477.2
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.3 | 473.1212 | 476.4 | -0.3 | 0 | 18.78 | 13 | 951 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.044321 | 0.045221 | 0.0274 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 473.1
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.4 | 472.4 | 475.8 | -0.3 | 0 | 18.6 | 13 | 951 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0442 | 0.0451 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5767 | 27.3 | 472.4
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.4 | 472.321 | 475.7 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.044121 | 0.045021 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4224 | 27.3 | 472.3
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.2 | 472.1418 | 475.4 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043942 | 0.044821 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 472.1
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 2.7 | 472.1 | 474.8 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043642 | 0.044642 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3431 | 472.1
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 2.6 | 471.842 | 474.4 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043521 | 0.044521 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.4 | 27.3797 | 471.8
2008-10-20 | 17:37 | -0.005 | -0.001 | -10 | 2.8 | 471.821 | 474.6 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0434 | 0.044321 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3025 | 471.8
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 2.9 | 471.6207 | 474.5 | -0.3 | 0 | 18.5 | 13 | 952 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0433 | 0.044221 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 471.6
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 2.9 | 471.4417 | 474.3 | -0.3 | 0 | 18.5 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043221 | 0.0442 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 471.4
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 3.1 | 471.3 | 474.4 | -0.3 | 0 | 18.5 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043121 | 0.044121 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 471.3
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 3.2 | 471.279 | 474.5 | -0.3 | 0 | 18.44 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0431 | 0.044021 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4122 | 27.3 | 471.3
2008-10-20 | 17:37 | 0 | -0.001 | -10 | 3.2 | 471.142 | 474.3 | -0.3 | 0 | 18.4 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043 | 0.043921 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5666 | 27.3 | 471.1
2008-10-20 | 17:37 | -0.005 | -0.001 | -10 | 3.2 | 470.942 | 474.1 | -0.3 | 0 | 18.4 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.043 | 0.0439 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.479 | 27.3 | 470.9
2008-10-20 | 17:37 | -0.01 | -0.001 | -10 | 3.3 | 470.9 | 474.2 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2791 | 0.0429 | 0.0438 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.9
2008-10-20 | 17:37 | -0.01 | -0.0005 | -4.8376 | 3.4 | 470.979 | 474.4 | -0.4 | 0 | 18.4 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0429 | 0.0438 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4148 | 27.3 | 471
2008-10-20 | 17:38 | -0.01 | 0 | 0 | 3.4 | 471.2 | 474.6 | -0.4 | 0 | 18.4 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0428 | 0.0437 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4919 | 27.3 | 471.2
2008-10-20 | 17:38 | -0.01 | -0.0002 | -2.0227 | 3.6 | 471.2 | 474.8 | -0.4 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042721 | 0.043621 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4312 | 27.3 | 471.2
2008-10-20 | 17:38 | -0.01 | -0.001 | -10 | 3.7 | 471.3 | 475 | -0.4 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0427 | 0.043521 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4917 | 27.3 | 471.3
2008-10-20 | 17:38 | -0.005 | -0.001 | -10 | 3.7 | 471.2207 | 474.9 | -0.4 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0427 | 0.0435 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5542 | 27.3 | 471.2
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.8 | 471.4 | 475.2 | -0.4 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0426 | 0.0434 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4002 | 27.3 | 471.4
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.8 | 471.479 | 475.3 | -0.4 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.0426 | 0.0434 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 471.5
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.8 | 471.4 | 475.2 | -0.4 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.2 | 0.0426 | 0.0434 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5099 | 27.3 | 471.4
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.8 | 471.242 | 475 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0425 | 0.0433 | 0.0273 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 471.2
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.9 | 471.3209 | 475.2 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0425 | 0.0433 | 0.027221 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5907 | 27.3 | 471.3
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.9 | 471.1419 | 475 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0425 | 0.0433 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5135 | 27.3 | 471.1
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.9 | 470.9 | 474.8 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0425 | 0.043221 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3727 | 27.3 | 470.9
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.9 | 470.742 | 474.6 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042421 | 0.0432 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3816 | 27.3408 | 470.7
2008-10-20 | 17:38 | 0 | -0.0007 | -6.5333 | 4 | 470.442 | 474.4 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0424 | 0.0432 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.518 | 27.382 | 470.4
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.9 | 470.2 | 474.1 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0424 | 0.0432 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1006 | 10.5951 | 27.3049 | 470.2
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.6 | 470.2 | 473.8 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0423 | 0.043121 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040.0422 | 1006 | 10.6 | 27.3 | 470.2
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.4 | 473.7 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042221 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.4
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.3 | 473.6 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040.0637 | 1005 | 10.6 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.4 | 470.458 | 473.8 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5884 | 27.3 | 470.5
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.4 | 470.6583 | 474.1 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3569 | 27.3 | 470.7
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.8579 | 474.2 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.2104 | 10.5329 | 27.3582 | 470.9
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.8422 | 474.1 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6646 | 27.3646 | 470.8
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.4 | 470.8 | 474.2 | -0.3 | 0 | 18.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 470.8
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.8 | 474.1 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0422 | 0.043021 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 470.8
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.7207 | 474 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042121 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.7932 | 10.6 | 27.3 | 470.7
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.2 | 470.2627 | 473.5 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.2 | 470.2 | 473.4 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.2
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.2 | 470.2 | 473.4 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.2
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.2 | 473.5 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.2
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.1 | 473.4 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4946 | 27.3 | 470.1
2008-10-20 | 17:38 | 0 | -0.0005 | -5.0542 | 3.3 | 470.1 | 473.4 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.1
2008-10-20 | 17:38 | 0 | -0.0007 | -6.5538 | 3.3 | 470.2581 | 473.6 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | 0 | 0 | 3.3 | 470.3 | 473.6 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0421 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | 0 | -0.06617 | 3.2 | 470.1 | 473.3 | -0.3 | 0 | 18.4 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042021 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.1
2008-10-20 | 17:38 | 0 | -0.0008 | -8.3375 | 3.2 | 470.179 | 473.4 | -0.3 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042 | 0.042921 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 470.2
2008-10-20 | 17:38 | -0.003 | -0.001 | -10 | 3.2 | 470.3793 | 473.6 | -0.2 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3847 | 27.3 | 470.4
2008-10-20 | 17:38 | -0.008 | -0.001 | -10 | 3.2 | 470.4 | 473.6 | -0.2 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3077 | 27.3 | 470.4
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.3 | 470.3 | 473.6 | -0.2 | 0 | 18.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.508 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 3.3 | 470.3 | 473.6 | -0.2 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 470.3
2008-10-20 | 17:38 | 0 | -0.001 | -9.9833 | 3.3 | 470.1421 | 473.4 | -0.2 | 0 | 18.4 | 13 | 955 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 470.1
2008-10-20 | 17:38 | 0 | -0.0002 | -1.6361 | 3.2 | 469.842 | 473 | -0.2 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005.79 | 10.5954 | 27.3 | 469.8
2008-10-20 | 17:38 | -0.004 | 0 | 0 | 3.2 | 469.7209 | 472.9 | -0.2 | 0 | 18.4 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.3639 | 27.3 | 469.7
2008-10-20 | 17:38 | -0.01 | 0 | 0 | 3.2 | 469.779 | 473 | -0.2 | 0 | 18.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1006 | 10.4676 | 27.3 | 469.8
2008-10-20 | 17:38 | -0.01 | -0.0004 | -3.5274 | 3.2 | 469.9 | 473.1 | -0.2 | 0 | 18.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.042 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5339 | 27.3 | 469.9
2008-10-20 | 17:38 | -0.01 | -0.001 | -10 | 2.4 | 452.297 | 454.7 | -0.2 | 0 | 18.29 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 4.7 | 195 | 59.3 | 0.041842 | 0.042821 | 0.026189 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 452.3
2008-10-20 | 17:38 | -0.01 | -0.001 | -9.593 | 1.4 | 431.748 | 433.1 | -0.3 | 0 | 17.37 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 4.3 | 195 | 59.3 | 0.040984 | 0.042084 | 0.025005 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 431.7
2008-10-20 | 17:38 | -0.01 | -0.0002 | -1.6121 | 1.8 | 422.4207 | 424.2 | -0.3 | 0 | 16.8 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039821 | 0.0409 | 0.0244 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 422.4
2008-10-20 | 17:38 | -0.01 | 0 | 0 | 2.2 | 420.6 | 422.8 | -0.3 | 0 | 16.5 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039483 | 0.040483 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4 | 27.3 | 420.6
2008-10-20 | 17:38 | -0.01 | -0.0005 | -5.2191 | 2.3 | 420.521 | 422.8 | -0.3 | 0 | 16.45 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0394 | 0.040321 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.438 | 27.3 | 420.5
2008-10-20 | 17:38 | -0.01 | -0.001 | -10 | 2.4 | 420.421 | 422.8 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039221 | 0.0402 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5923 | 27.3 | 420.4
2008-10-20 | 17:38 | -0.01 | -0.001 | -10 | 2.4 | 420.242 | 422.6 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.039121 | 0.040021 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 420.2
2008-10-20 | 17:38 | -0.003 | -0.001 | -10 | 2.4 | 420.279 | 422.7 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0391 | 0.04 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 420.3
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 2.6 | 420.1 | 422.7 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.038921 | 0.039821 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 420.1
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 2.7 | 420.1212 | 422.8 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.038821 | 0.039721 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 420.1
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 2.8 | 420.1 | 422.9 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0388 | 0.039621 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 420.1
2008-10-20 | 17:38 | 0 | -0.0006 | -6.4961 | 2.8 | 419.821 | 422.6 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0387 | 0.039521 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6592 | 27.3 | 419.8
2008-10-20 | 17:38 | 0 | 0 | 0 | 2.9 | 419.821 | 422.7 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.038621 | 0.039421 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6637 | 27.3 | 419.8
2008-10-20 | 17:38 | 0 | 0 | -0.041356 | 2.9 | 419.6417 | 422.5 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.3 | 0.0386 | 0.0394 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5865 | 27.3 | 419.6
2008-10-20 | 17:38 | 0 | -0.0008 | -8.3127 | 2.9 | 419.7 | 422.6 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.038521 | 0.039321 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5093 | 27.3 | 419.7
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 2.9 | 419.5 | 422.4 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.038421 | 0.0393 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5678 | 27.3 | 419.5
2008-10-20 | 17:38 | 0 | -0.001 | -10 | 2.9 | 419.421 | 422.3 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.379 | 0.0384 | 0.0392 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.4
2008-10-20 | 17:39 | 0 | -0.001 | -10 | 2.9 | 419.3212 | 422.2 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0384 | 0.0392 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.3
2008-10-20 | 17:39 | 0 | -0.001 | -10 | 3 | 419.3 | 422.3 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0383 | 0.0391 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.3
2008-10-20 | 17:39 | -0.007 | -0.001 | -10 | 3 | 419.3 | 422.3 | -0.3 | 0 | 16.4 | 13 | 956 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0383 | 0.0391 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.3
2008-10-20 | 17:39 | -0.01 | -0.0002 | -1.6597 | 3.1 | 419.1 | 422.2 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0382 | 0.039 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 3.2 | 419.2 | 422.4 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0381 | 0.0389 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.2
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 3.2 | 419.2 | 422.4 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0381 | 0.0389 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5923 | 27.3 | 419.2
2008-10-20 | 17:39 | -0.01 | -0.0003 | -3.3361 | 3.2 | 419.3 | 422.5 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0381 | 0.0388 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5153 | 27.3 | 419.3
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.2 | 419.1 | 422.3 | -0.3 | 0 | 16.4 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.038021 | 0.0388 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5618 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.3 | 419.1 | 422.4 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.038 | 0.038721 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.3 | 419.2 | 422.5 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.038 | 0.0387 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5665 | 27.3 | 419.2
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.4 | 419.1 | 422.5 | -0.3 | 0 | 16.4 | 13.1 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.321 | 0.038 | 0.0387 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4122 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.4 | 419.1 | 422.5 | -0.3 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2209 | 0.038 | 0.0387 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5421 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.01 | -0.001 | -10 | 3.4 | 419.1 | 422.5 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.037921 | 0.0387 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.1
2008-10-20 | 17:39 | -0.002 | -0.001 | -10 | 3.4 | 419.021 | 422.4 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0379 | 0.0386 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6254 | 27.3 | 419
2008-10-20 | 17:39 | 0 | -0.001 | -10 | 3.5 | 419 | 422.5 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.3204 | 0.0379 | 0.0386 | 0.0243 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6974 | 27.3 | 419
2008-10-20 | 17:39 | 0 | -0.0018 | -18.3236 | 3.5 | 419.0795 | 422.6 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0379 | 0.0386 | 0.0243 | 2.1e-05 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6202 | 27.3 | 419.1
2008-10-20 | 17:39 | 0 | -0.0013 | -13.3361 | 3.4 | 419.1 | 422.5 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0379 | 0.0386 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6569 | 27.3 | 419.1
2008-10-20 | 17:39 | 0 | -0.0005 | -4.9958 | 3.4 | 419.1 | 422.5 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.037821 | 0.0386 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6659 | 27.3 | 419.1
2008-10-20 | 17:39 | 0 | -0.0003 | -3.4629 | 3.4 | 419.179 | 422.6 | -0.1 | 0 | 16.4 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0378 | 0.0385 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.2
2008-10-20 | 17:39 | 0 | -0.0008 | -8.0544 | 3.4 | 419.1 | 422.5 | -0.1 | 0 | 16.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0378 | 0.0385 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.1
2008-10-20 | 17:39 | 0 | 0 | -0.048038 | 3.4 | 419.1 | 422.5 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0378 | 0.0385 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 419.1
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.9209 | 422.3 | -0.1 | 0 | 16.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0378 | 0.0385 | 0.0243 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.9 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.037721 | 0.038421 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.32 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.9 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0377 | 0.0384 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3971 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.9 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2788 | 0.0377 | 0.0384 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3257 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.921 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0377 | 0.0384 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.8586 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0377 | 0.0384 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6286 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 418.8208 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0377 | 0.038321 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6941 | 27.3 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 418.958 | 422.6 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.037621 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.617 | 27.3 | 419
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 418.842 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0376 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4795 | 27.3 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 419 | 422.6 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0376 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4748 | 27.3 | 419
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 418.921 | 422.5 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0376 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5854 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.6 | 418.9 | 422.5 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0376 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5083 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.9 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0376 | 0.0383 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5689 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.5 | 418.9 | 422.4 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.9 | 422.3 | -0.1 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.623 | 27.3 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.821 | 422.2 | -0.1 | 0 | 16.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6999 | 27.3001 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.8207 | 422.2 | -0.2 | 0 | 16.4 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6229 | 27.3771 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.9 | 422.3 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3458 | 418.9
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.721 | 422.1 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.7
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.8 | 422.2 | -0.2 | 0 | 16.4 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 3.4 | 418.8 | 422.2 | -0.2 | 0 | 16.4 | 13 | 954 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0375 | 0.0382 | 0.0242 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 418.8
2008-10-20 | 17:39 | 0 | 0 | 0 | 1.7 | 394.7305 | 396.5 | -0.3 | 0 | 16 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.037142 | 0.0381 | 0.022826 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 394.7
2008-10-20 | 17:39 | -0.008 | 0 | 0 | 0.7 | 373.0004 | 373.7 | -0.4 | 0 | 15.04 | 13 | 952 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.035505 | 0.036753 | 0.021684 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 373
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 0.8 | 369.9096 | 370.7 | -0.4 | 0 | 14.6 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034863 | 0.035863 | 0.0214 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6178 | 27.3 | 369.9
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 0.9 | 368.989 | 369.9 | -0.4 | 0 | 14.45 | 13 | 952 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034621 | 0.035621 | 0.021321 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.695 | 27.3 | 369
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 0.9 | 368.7 | 369.6 | -0.4 | 0 | 14.4 | 13 | 952 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034521 | 0.0356 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5557 | 27.3 | 368.7
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1 | 368.521 | 369.5 | -0.4 | 0 | 14.4 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0345 | 0.0355 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5493 | 27.3 | 368.5
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.1 | 368.4414 | 369.5 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034321 | 0.035321 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5735 | 27.3 | 368.4
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.2 | 368.4 | 369.6 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0343 | 0.035221 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5037 | 27.3 | 368.4
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.3 | 368.3 | 369.6 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034121 | 0.0351 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5809 | 27.3 | 368.3
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.4 | 368.1 | 369.5 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0341 | 0.035 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.716 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.5 | 368.1 | 369.6 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0349 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.7648 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.5 | 368.1 | 369.6 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.034 | 0.0349 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6627 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.7 | 368.2 | 369.9 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0339 | 0.0348 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4313 | 27.3 | 368.2
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.8 | 368.079 | 369.9 | -0.4 | 0 | 14.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0338 | 0.0347 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4667 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.8 | 368.1 | 369.9 | -0.3 | 0 | 14.3 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0338 | 0.034621 | 0.0213 | 0 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5878 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.9 | 368.121 | 370 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0337 | 0.0346 | 0.0213 | 7.9e-05 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6368 | 27.3 | 368.1
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.9 | 368.1793 | 370.1 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0337 | 0.034521 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4053 | 27.3 | 368.2
2008-10-20 | 17:39 | -0.01 | 0 | 0 | 1.9 | 368.2 | 370.1 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0337 | 0.0345 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4754 | 27.3 | 368.2
2008-10-20 | 17:40 | -0.01 | 0 | 0 | 1.9 | 368.079 | 370 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0336 | 0.0345 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5526 | 27.3 | 368.1
2008-10-20 | 17:40 | -0.01 | 0 | 0 | 1.9 | 368 | 369.9 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0336 | 0.0344 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 368
2008-10-20 | 17:40 | -0.003 | 0 | 0 | 2.1 | 368 | 370.1 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.033521 | 0.034321 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5861 | 27.3 | 368
2008-10-20 | 17:40 | 0 | -0.0017 | -16.5523 | 2.1 | 368.0791 | 370.2 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0335 | 0.0343 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1004.2087 | 10.4318 | 27.3 | 368.1
2008-10-20 | 17:40 | 0 | -0.0013 | -13.2079 | 2.2 | 367.8418 | 370 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0335 | 0.0343 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5832 | 27.3 | 367.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.2 | 367.8 | 370 | -0.3 | 0 | 14.3 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0335 | 0.0343 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6618 | 27.3 | 367.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.3 | 367.721 | 370 | -0.3 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0334 | 0.0342 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.569 | 27.3 | 367.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.3 | 367.6 | 369.9 | -0.3 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.3 | 0.0334 | 0.0341 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.4149 | 27.3 | 367.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.5 | 369.9 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0333 | 0.0341 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.5391 | 27.3 | 367.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.5792 | 370 | -0.3 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0333 | 0.034021 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.679 | 370.2 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0333 | 0.034 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.621 | 370.1 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0333 | 0.034 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.6 | 370.1 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0333 | 0.034 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.521 | 370 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0332 | 0.034 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.5209 | 369.9 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0332 | 0.034 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.3213 | 369.8 | -0.4 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0332 | 0.0339 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.3 | 369.7 | -0.5 | 0 | 14.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0332 | 0.0339 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.3 | 369.7 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0332 | 0.0339 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1041 | 1005 | 10.6646 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.3 | 369.7 | -0.5 | 0 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0332 | 0.0339 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040.0251 | 1005 | 10.6583 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.2207 | 369.6 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.033121 | 0.0339 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 367.2
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.2792 | 369.8 | -0.5 | 0 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.033821 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.321 | 369.8 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.0338 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6732 | 27.3 | 367.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367.142 | 369.6 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.0338 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.7 | 27.3 | 367.1
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 367.1 | 369.5 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.0338 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6725 | 27.3 | 367.1
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 367 | 369.5 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.0338 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5906 | 27.3 | 367
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 366.921 | 369.4 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0331 | 0.0338 | 0.0213 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4361 | 27.3 | 366.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 366.9 | 369.4 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0338 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040.9644 | 1005 | 10.5182 | 27.3 | 366.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 366.8 | 369.2 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.033721 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4912 | 27.3 | 366.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 366.9 | 369.3 | -0.5 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3269 | 27.3 | 366.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 366.9 | 369.3 | -0.4 | 0 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4813 | 27.3 | 366.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.4 | 366.7208 | 369.1 | -0.4 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4322 | 27.3 | 366.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 366.6 | 369.1 | -0.3 | 0 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 366.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.5 | 366.521 | 369 | -0.2 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.033 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4443 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.6 | 366.5 | 369.1 | -0.1 | 0 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.3 | 0.032921 | 0.0337 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5986 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.6 | 366.5 | 369.1 | 0 | 0.0406 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4471 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.6 | 366.5 | 369.1 | 0.1 | 0.1 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.6 | 366.5 | 369.1 | 0.2 | 0.2402 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4616 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.6 | 366.579 | 369.2 | 0.3 | 0.3 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6079 | 27.3 | 366.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.7 | 366.5 | 369.2 | 0.4 | 0.4 | 14.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6849 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.7 | 366.4 | 369.1 | 0.4 | 0.4 | 14.3 | 13 | 955 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5142 | 27.3 | 366.4
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.7 | 366.5 | 369.2 | 0.5 | 0.4693 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.7 | 366.5 | 369.2 | 0.5 | 0.5 | 14.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4335 | 27.3 | 366.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 2.7 | 366.4 | 369.1 | 0.5 | 0.4796 | 14.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0329 | 0.0336 | 0.0212 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5878 | 27.3 | 366.4
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.8 | 348.095 | 349.9 | 0.4 | 0.4 | 14.1 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.032742 | 0.0335 | 0.020126 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3868 | 27.3 | 348.1
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 0.2 | 322.676 | 322.9 | 0.4 | 0.3816 | 13.22 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.031763 | 0.032784 | 0.018684 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4447 | 27.3 | 322.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 0.2 | 318.7902 | 319 | 0.3 | 0.3 | 12.65 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.031221 | 0.032221 | 0.018442 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5492 | 27.3 | 318.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 0.8 | 317.2882 | 318.1 | 0.2 | 0.2 | 12.4 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.030584 | 0.031484 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 317.3
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.2 | 317.2 | 318.4 | 0.2 | 0.2 | 12.3 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0302 | 0.031 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 317.2
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.3 | 317.121 | 318.4 | 0.2 | 0.1612 | 12.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.030121 | 0.030842 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 317.1
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.4 | 317.021 | 318.4 | 0.1 | 0.1 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029921 | 0.030721 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 317
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.5 | 316.9 | 318.4 | 0.1 | 0.1 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029821 | 0.0306 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4225 | 27.3 | 316.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.5 | 316.8 | 318.3 | 0.1 | 0.1 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0297 | 0.030421 | 0.0184 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5769 | 27.3 | 316.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.6 | 316.721 | 318.3 | 0.1 | 0.1 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0296 | 0.030321 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4031 | 27.3 | 316.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.7 | 316.621 | 318.3 | 0.1 | 0.1 | 12.3 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029521 | 0.0303 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4285 | 27.3 | 316.6
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.7 | 315.9 | 317.6 | 0.1 | 0.1 | 12.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0295 | 0.0302 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.6 | 27.3 | 315.9
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.8 | 315.721 | 317.5 | 0.1 | 0.0607 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0294 | 0.0301 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 315.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.7 | 315.7791 | 317.5 | 0 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029321 | 0.030021 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6744 | 27.3 | 315.8
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.8 | 315.721 | 317.5 | -0.1 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0293 | 0.03 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315.7
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.8 | 315.521 | 317.3 | -0.1 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0292 | 0.0299 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6713 | 27.3 | 315.5
2008-10-20 | 17:40 | 0 | -0.001 | -10 | 1.9 | 315.5 | 317.4 | -0.2 | 0 | 12.3 | 13 | 954 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.0292 | 0.0298 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6059 | 27.3 | 315.5
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 1.9 | 315.5 | 317.4 | -0.2 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.0291 | 0.0298 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.683 | 27.3 | 315.5
2008-10-20 | 17:41 | 0 | -0.0009 | -8.5262 | 2 | 315.5793 | 317.6 | -0.3 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2793 | 0.0291 | 0.029721 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6398 | 27.3 | 315.6
2008-10-20 | 17:41 | 0 | 0 | -0.19983 | 2 | 315.6 | 317.6 | -0.3 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029021 | 0.0297 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5627 | 27.3 | 315.6
2008-10-20 | 17:41 | 0 | 0 | 0 | 2 | 315.6 | 317.6 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029 | 0.029621 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4855 | 27.3 | 315.6
2008-10-20 | 17:41 | 0 | 0 | 0 | 2 | 315.6 | 317.6 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.029 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4083 | 27.3 | 315.6
2008-10-20 | 17:41 | 0 | 0 | 0 | 2 | 315.5 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.029 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 315.5
2008-10-20 | 17:41 | 0 | 0 | 0 | 2 | 315.4 | 317.4 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.028921 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.492 | 27.3 | 315.4
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.4 | 317.5 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0296 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 315.4
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.321 | 317.4 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.0289 | 0.0295 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6003 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | -0.0008 | -8.2572 | 2.1 | 315.3 | 317.4 | -0.5 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.028821 | 0.0295 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6774 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.1 | 315.4 | 317.5 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0295 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315.4
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.1 | 315.4 | 317.5 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0288 | 0.0295 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6684 | 27.3 | 315.4
2008-10-20 | 17:41 | 0 | -0.0007 | -6.7237 | 2.1 | 315.2208 | 317.3 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0294 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 315.2
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.121 | 317.2 | -0.5 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0288 | 0.0294 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.3 | 317.4 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.028721 | 0.0294 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.221 | 317.3 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0294 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5186 | 27.3 | 315.2
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.279 | 317.4 | -0.5 | 0 | 12.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.029321 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.2096 | 10.4358 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.3 | 317.4 | -0.5 | 0 | 12.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5901 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.479 | 317.6 | -0.5 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 315.5
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.321 | 317.4 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6494 | 27.3 | 315.3
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.2 | 317.3 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315.2
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.1 | 317.2 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.1 | 315.0417 | 317.1 | -0.5 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0287 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.2 | 315 | 317.2 | -0.5 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0286 | 0.0293 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 315
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.2 | 314.9 | 317.1 | -0.5 | 0 | 12.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0286 | 0.029221 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 314.9
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.2 | 314.979 | 317.2 | -0.5 | 0 | 12.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0286 | 0.0292 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6876 | 27.3 | 315
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.3 | 315.179 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0286 | 0.029121 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6104 | 27.3 | 315.2
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.4 | 315.1 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.028521 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4664 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.4 | 315.1 | 317.5 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.4 | 315.1 | 317.5 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4423 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.4 | 315.1 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5966 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.4 | 315.1208 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.8264 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.5 | 315.1 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2207 | 0.0285 | 0.0291 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7421 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.5 | 315.1 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.029 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6596 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.5 | 315.079 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.029 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.793 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.6 | 315.1 | 317.7 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0285 | 0.029 | 0.0183 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7158 | 27.3 | 315.1
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.6 | 315 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.028421 | 0.029 | 0.018221 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5165 | 27.3 | 315
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.7 | 315 | 317.7 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0289 | 0.0182 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4766 | 27.3 | 315
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.7 | 314.9211 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0289 | 0.0182 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 314.9
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.7 | 314.9 | 317.6 | -0.4 | 0 | 12.3 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0289 | 0.0182 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 314.9
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.7 | 314.8 | 317.5 | -0.4 | 0 | 12.3 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0289 | 0.0182 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6697 | 27.3 | 314.8
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.7 | 314.9 | 317.6 | -0.4 | 0 | 12.2 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0284 | 0.0289 | 0.0182 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.653 | 27.3 | 314.9
2008-10-20 | 17:41 | 0 | -0.0003 | -3.2607 | 2 | 283.5408 | 285.6 | -0.5 | 0 | 11.53 | 13 | 953 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.028003 | 0.028582 | 0.016467 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 283.5
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 1.5 | 267.6207 | 269.2 | -0.5 | 0 | 10.71 | 13 | 953 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.026783 | 0.027383 | 0.0155 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 267.6
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 1.4 | 266.7 | 268.1 | -0.5 | 0 | 10.41 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.2792 | 0.026221 | 0.026821 | 0.0155 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 266.7
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 1.7 | 265.821 | 267.5 | -0.6 | 0 | 10.3 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025863 | 0.026463 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.4881 | 27.3 | 265.8
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.2 | 265.721 | 267.9 | -0.6 | 0 | 10.3 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025484 | 0.025984 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.4993 | 27.3 | 265.7
2008-10-20 | 17:41 | 0 | -0.0005 | -5.1585 | 2.3 | 265.7 | 268 | -0.6 | 0 | 10.3 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0254 | 0.0258 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 265.7
2008-10-20 | 17:41 | 0 | 0 | 0 | 2.3 | 265.5209 | 267.8 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025321 | 0.025721 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.7 | 27.3 | 265.5
2008-10-20 | 17:41 | 0 | -0.0002 | -1.6192 | 2.4 | 265.4 | 267.8 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2791 | 0.0253 | 0.0257 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6354 | 27.3 | 265.4
2008-10-20 | 17:41 | 0 | -0.001 | -9.7559 | 2.5 | 265.321 | 267.8 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0252 | 0.0256 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5582 | 27.3 | 265.3
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.6 | 265.3 | 267.9 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0251 | 0.025421 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.538 | 27.3 | 265.3
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.7 | 265.2 | 267.9 | -0.7 | 0 | 10.2 | 13 | 954 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.025 | 0.025321 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6923 | 27.3 | 265.2
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.7 | 265.2 | 267.9 | -0.7 | 0 | 10.2 | 13 | 954 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0249 | 0.0253 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6266 | 27.3 | 265.2
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.8 | 265.1 | 267.9 | -0.7 | 0 | 10.2 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0249 | 0.0252 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4989 | 27.3 | 265.1
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.8 | 265 | 267.8 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.024821 | 0.0252 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4554 | 27.3 | 265
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.9 | 264.921 | 267.8 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0248 | 0.0251 | 0.0154 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.9
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 2.9 | 264.9 | 267.8 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0247 | 0.025021 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 264.9
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 3 | 264.821 | 267.8 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.024621 | 0.024921 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6592 | 27.3 | 264.8
2008-10-20 | 17:41 | 0 | -0.001 | -10 | 3 | 264.7 | 267.7 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0246 | 0.0249 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6637 | 27.3 | 264.7
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3 | 264.7 | 267.7 | -0.7 | 0 | 10.2 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.0246 | 0.0249 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 264.7
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3 | 264.6 | 267.6 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.024521 | 0.0248 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.521 | 267.6 | -0.7 | 0 | 10.2 | 13 | 954 | 954 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.0248 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5321 | 27.3 | 264.5
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.421 | 267.5 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.3 | 0.0245 | 0.024721 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5901 | 27.3 | 264.4
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.3207 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0244 | 0.0247 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6333 | 27.3 | 264.3
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.3 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0244 | 0.0247 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4016 | 27.3 | 264.3
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.479 | 267.7 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0244 | 0.0246 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4765 | 27.3 | 264.5
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.5 | 267.6 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0243 | 0.0246 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5535 | 27.3 | 264.5
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.4 | 267.6 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0243 | 0.0246 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.4
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.3 | 267.5 | -0.7 | 0 | 10.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0243 | 0.024521 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.3
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.2211 | 267.4 | -0.7 | 0 | 10.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0243 | 0.0245 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.2
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.2 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0243 | 0.0245 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4758 | 27.3 | 264.2
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.121 | 267.2 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.024221 | 0.0245 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4787 | 27.3 | 264.1
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.1 | 264.1788 | 267.3 | -0.7 | 0 | 10.2 | 13 | 955 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0245 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.2
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.3 | 267.5 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.024421 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1004.2068 | 10.6 | 27.3 | 264.3
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.2 | 264.2208 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0242 | 0.0244 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.2
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.3 | 264.2 | 267.5 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.024121 | 0.0244 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.2
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.3 | 264.1 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0244 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 264.1
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.4 | 264 | 267.4 | -0.7 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0243 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5974 | 27.3 | 264
2008-10-20 | 17:42 | -0.008 | -0.0008 | -8.4082 | 3.4 | 264 | 267.4 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0241 | 0.0243 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5202 | 27.3 | 264
2008-10-20 | 17:42 | -0.004 | 0 | -0.2449 | 3.4 | 263.9 | 267.3 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.3 | 0.024021 | 0.0243 | 0.0153 | 0.0001 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.557 | 27.3 | 263.9
2008-10-20 | 17:42 | 0 | 0 | 0 | 3.4 | 263.8419 | 267.2 | -0.6 | 0 | 10.2 | 13.1 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.8
2008-10-20 | 17:42 | -0.003 | 0 | 0 | 3.5 | 263.9 | 267.4 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.9
2008-10-20 | 17:42 | -0.01 | 0 | 0 | 3.5 | 263.8 | 267.3 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.8
2008-10-20 | 17:42 | -0.01 | 0 | 0 | 3.5 | 263.7414 | 267.2 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.7
2008-10-20 | 17:42 | -0.01 | 0 | 0 | 3.5 | 263.7 | 267.2 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5144 | 27.3 | 263.7
2008-10-20 | 17:42 | -0.01 | 0 | 0 | 3.5 | 263.7 | 267.2 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.024 | 0.0242 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 263.7
2008-10-20 | 17:42 | -0.01 | 0 | 0 | 3.6 | 263.7 | 267.3 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.023921 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 263.7
2008-10-20 | 17:42 | -0.007 | 0 | 0 | 3.6 | 263.679 | 267.3 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5488 | 27.3 | 263.7
2008-10-20 | 17:42 | 0 | 0 | 0 | 3.6 | 263.7 | 267.3 | -0.5 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.7
2008-10-20 | 17:42 | 0 | -0.0003 | -3.1965 | 3.5 | 263.7 | 267.2 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5426 | 27.3 | 263.7
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.7 | 267.2 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4117 | 27.3 | 263.7
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.6 | 267.1 | -0.6 | 0 | 10.2 | 13.1 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.566 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.6 | 267.1 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4796 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.6 | 267.1 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4747 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.6 | 267.1 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.0241 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5854 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.6 | 263.6792 | 267.3 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0239 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5083 | 27.3 | 263.7
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.5 | 263.621 | 267.1 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0239 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5687 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.4 | 263.6 | 267 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.023821 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.4 | 263.6 | 267 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.4 | 263.6 | 267 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.6
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 3.4 | 263.5 | 266.9 | -0.6 | 0 | 10.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.024 | 0.0153 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 263.5
2008-10-20 | 17:42 | 0 | -0.0007 | -6.8167 | 3.3 | 255.8329 | 259.2 | -0.6 | 0 | 10.11 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0238 | 0.024 | 0.014826 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 255.8
2008-10-20 | 17:42 | 0 | 0 | 0 | 2.4 | 222.994 | 225.4 | -0.7 | 0 | 9.27 | 13.1 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.023005 | 0.023284 | 0.012926 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 223
2008-10-20 | 17:42 | 0 | 0 | 0 | 1.8 | 214.8892 | 216.7 | -0.7 | 0 | 8.61 | 13 | 953 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.021868 | 0.022168 | 0.012442 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5909 | 27.3 | 214.9
2008-10-20 | 17:42 | 0 | 0 | 0 | 1.7 | 212.4212 | 214.1 | -0.7 | 0 | 8.36 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.021363 | 0.021685 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5137 | 27.3 | 212.4
2008-10-20 | 17:42 | 0 | 0 | 0 | 1.9 | 212.3207 | 214.2 | -0.7 | 0 | 8.3 | 13 | 953 | 953 | 942 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.021121 | 0.0215 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 212.3
2008-10-20 | 17:42 | 0 | 0 | 0 | 1.9 | 212.3 | 214.2 | -0.7 | 0 | 8.3 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0211 | 0.0214 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4594 | 27.3 | 212.3
2008-10-20 | 17:42 | 0 | 0 | 0 | 2.1 | 212.221 | 214.3 | -0.7 | 0 | 8.2 | 13 | 953 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.021 | 0.021321 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 212.2
2008-10-20 | 17:42 | 0 | -0.0002 | -1.5643 | 2.3 | 212.2 | 214.5 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.020821 | 0.021121 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 212.2
2008-10-20 | 17:42 | -0.006 | -0.001 | -9.7543 | 2.3 | 212.1209 | 214.4 | -0.6 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.020721 | 0.021 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4721 | 27.3 | 212.1
2008-10-20 | 17:42 | -0.005 | -0.001 | -10 | 2.4 | 212.1209 | 214.5 | -0.6 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0207 | 0.0209 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 212.1
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.5 | 211.921 | 214.4 | -0.6 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0206 | 0.0208 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 211.9
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.5 | 211.8 | 214.3 | -0.6 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0205 | 0.0207 | 0.0123 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4963 | 27.3 | 211.8
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.5 | 211.021 | 213.5 | -0.6 | 0 | 8.2 | 13.1 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0205 | 0.0207 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4191 | 27.3 | 211
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.5 | 211 | 213.5 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.020421 | 0.020621 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.458 | 27.3 | 211
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.6 | 210.9207 | 213.5 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2793 | 0.0204 | 0.020521 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4648 | 27.3 | 210.9
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.7 | 210.9792 | 213.7 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0203 | 0.0205 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.437 | 27.3 | 211
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.7 | 210.9 | 213.6 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0203 | 0.020421 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6685 | 27.3 | 210.9
2008-10-20 | 17:42 | 0 | -0.001 | -10 | 2.7 | 210.9 | 213.6 | -0.7 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0202 | 0.0204 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6333 | 27.3 | 210.9
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.9 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0202 | 0.0204 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5123 | 27.3 | 210.9
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8212 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2788 | 0.0202 | 0.020321 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.442 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.020121 | 0.0203 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5964 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0201 | 0.0203 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0201 | 0.020221 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5474 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13.1 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.02 | 0.0202 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5298 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.6 | 210.8 | 213.4 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.02 | 0.0202 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5931 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.02 | 0.020121 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5159 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.7 | 210.8 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0199 | 0.0201 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5611 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.8 | 210.8 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0199 | 0.020021 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.8 | 210.9 | 213.7 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.019821 | 0.02 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.57 | 27.3 | 210.9
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.8 | 210.8 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.3 | 0.0198 | 0.02 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.416 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.8 | 213.7 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0198 | 0.019921 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.721 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0198 | 0.0199 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4469 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.7 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0198 | 0.0199 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5241 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.8 | 213.7 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0197 | 0.0199 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5985 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.8 | 213.7 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0197 | 0.0198 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5214 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 3 | 210.8 | 213.8 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0197 | 0.0198 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5558 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 3 | 210.7208 | 213.7 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -35 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0197 | 0.0198 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.567 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.9 | 210.721 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.279 | 0.0196 | 0.019721 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5102 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.8 | 210.7 | 213.5 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0196 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5873 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.8 | 210.779 | 213.6 | -0.8 | 0 | 8.2 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0196 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.5 | 210.7791 | 213.3 | -0.7 | 0 | 8.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.019521 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.8 | 213.2 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0195 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.8 | 213.2 | -0.7 | 0 | 8.2 | 13 | 956 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0195 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 210.8
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.8788 | 213.3 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0195 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2116 | 10.6 | 27.3 | 210.9
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.721 | 213.1 | -0.6 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4992 | 27.3 | 210.7
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.6 | 213 | -0.6 | 0 | 8.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4551 | 27.3 | 210.6
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 2.4 | 210.6 | 213 | -0.6 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5906 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.001 | -0.001 | -10 | 2.3 | 210.6 | 212.9 | -0.6 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4363 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.3 | 210.6 | 212.9 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.459 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.3 | 210.6 | 212.9 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4638 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.7 | 212.9 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0197 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4269 | 27.3 | 210.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.6211 | 212.8 | -0.7 | 0 | 8.2 | 13 | 956 | 954 | 945 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5812 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.6 | 212.8 | -0.7 | 0 | 8.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4645 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.6 | 212.8 | -0.8 | 0 | 8.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0194 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4898 | 27.3449 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.6788 | 212.9 | -0.8 | 0 | 8.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.019321 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5779 | 27.3779 | 210.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.7 | 212.9 | -0.8 | 0 | 8.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0193 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5007 | 27.3007 | 210.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.7 | 212.9 | -0.8 | 0 | 8.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2793 | 0.0193 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4235 | 27.3 | 210.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.7 | 212.9 | -0.8 | 0 | 8.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0193 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4534 | 27.3 | 210.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.621 | 212.8 | -0.8 | 0 | 8.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0193 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4696 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 2.2 | 210.6 | 212.8 | -0.8 | 0 | 8.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.3 | 0.0193 | 0.0196 | 0.0122 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4222 | 27.3 | 210.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 1.7 | 194.621 | 196.3 | -0.8 | 0 | 8.02 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.019221 | 0.019521 | 0.011252 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6533 | 27.3 | 194.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | -0.5 | 169.952 | 170 | -0.9 | 0 | 7.19 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.018505 | 0.019105 | 0.009847 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.638 | 27.3 | 170
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | -0.4 | 161.8765 | 161.9 | -0.9 | 0 | 6.61 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.017363 | 0.017984 | 0.009421 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 161.9
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | -0.1 | 161.421 | 161.4 | -0.9 | 0 | 6.36 | 13 | 953 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.016921 | 0.017421 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5833 | 27.3 | 161.4
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0 | 160.6 | 160.6 | -0.9 | 0 | 6.3 | 13 | 954 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0168 | 0.017321 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5062 | 27.3 | 160.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0.1 | 160.6 | 160.7 | -0.9 | 0 | 6.26 | 13 | 954 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0167 | 0.0172 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.571 | 27.3 | 160.6
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0.2 | 160.421 | 160.6 | -1 | 0 | 6.2 | 13 | 954 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.016521 | 0.017021 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 160.4
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0.4 | 160.4 | 160.8 | -1 | 0 | 6.2 | 13 | 954 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.016421 | 0.016921 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2084 | 10.5746 | 27.3 | 160.4
2008-10-20 | 17:43 | -0.004 | -0.001 | -10 | 0.5 | 160.3 | 160.8 | -1 | 0 | 6.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0163 | 0.0168 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5025 | 27.3 | 160.3
2008-10-20 | 17:43 | 0 | -0.001 | -10 | 0.5 | 160.3 | 160.8 | -1 | 0 | 6.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.016221 | 0.0167 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5797 | 27.3 | 160.3
2008-10-20 | 17:43 | -0.003 | -0.0004 | -3.8801 | 0.3 | 160.1 | 160.4 | -1 | 0 | 6.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2788 | 0.0161 | 0.0166 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 160.1
2008-10-20 | 17:43 | -0.01 | 0 | 0 | 0.3 | 160 | 160.3 | -1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0161 | 0.0165 | 0.0093 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.5319 | 27.3 | 160
2008-10-20 | 17:43 | -0.01 | 0 | 0 | 0.3 | 160 | 160.3 | -1 | 0 | 6.2 | 13 | 954 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.016 | 0.0165 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4113 | 27.3 | 160
2008-10-20 | 17:43 | -0.01 | -0.0001 | -1.1664 | 0.4 | 159.821 | 160.2 | -1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.015921 | 0.0164 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4884 | 27.3 | 159.8
2008-10-20 | 17:43 | -0.01 | -0.001 | -9.7427 | 0.5 | 159.721 | 160.2 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015821 | 0.016321 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.6312 | 27.3 | 159.7
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0.5 | 159.521 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0158 | 0.016221 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6573 | 27.3 | 159.5
2008-10-20 | 17:43 | -0.01 | -0.001 | -10 | 0.6 | 159.4205 | 160 | -1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0158 | 0.0162 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5602 | 27.3 | 159.4
2008-10-20 | 17:43 | -0.009 | -0.001 | -10 | 0.6 | 159.4 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0157 | 0.0162 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4059 | 27.3 | 159.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.6 | 159.3 | 159.9 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0157 | 0.0161 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 159.3
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.6 | 159.2 | 159.8 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0157 | 0.0161 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4515 | 27.3 | 159.2
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.7 | 159.1 | 159.8 | -1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0156 | 0.016 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4714 | 27.3 | 159.1
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.8 | 159.1 | 159.9 | -1 | 0 | 6.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015521 | 0.0159 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4117 | 27.3 | 159.1
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.9 | 159.1209 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0155 | 0.015821 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.566 | 27.3 | 159.1
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 0.9 | 159.1 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015421 | 0.0158 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 159.1
2008-10-20 | 17:44 | 0 | -0.001 | -10 | 1 | 159.021 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0154 | 0.015721 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 159
2008-10-20 | 17:44 | -0.008 | -0.001 | -10 | 1 | 159 | 160 | -1.1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0154 | 0.0157 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5855 | 27.3 | 159
2008-10-20 | 17:44 | -0.01 | -0.001 | -10 | 1.1 | 159.1 | 160.2 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0153 | 0.0157 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5083 | 27.3 | 159.1
2008-10-20 | 17:44 | -0.01 | -0.001 | -10 | 1.1 | 159 | 160.1 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0153 | 0.0156 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5 | 27.3 | 159
2008-10-20 | 17:44 | -0.01 | -0.001 | -10 | 1.2 | 159 | 160.2 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0153 | 0.0156 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4085 | 27.3 | 159
2008-10-20 | 17:44 | -0.01 | -0.0006 | -5.575 | 1.2 | 159 | 160.2 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0153 | 0.0156 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3227 | 27.3 | 159
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.2 | 158.921 | 160.1 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0153 | 0.0156 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3998 | 27.3 | 158.9
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.2 | 158.9 | 160.1 | -1.1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015221 | 0.0155 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4 | 27.3 | 158.9
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.3 | 158.8 | 160.1 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0152 | 0.0155 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3453 | 27.3 | 158.8
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.3 | 158.7 | 160 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0152 | 0.015421 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3318 | 27.3 | 158.7
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.3 | 158.7 | 160 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0152 | 0.0154 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4182 | 27.3 | 158.7
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.4 | 158.7 | 160.1 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0151 | 0.0154 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5725 | 27.3 | 158.7
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.4 | 158.621 | 160 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0151 | 0.0154 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.6 | 27.3 | 158.6
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.4 | 158.5 | 159.9 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0151 | 0.0153 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 158.5
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.4 | 158.5 | 159.9 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0151 | 0.0153 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 158.5
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.5 | 158.4 | 159.9 | -1 | 0 | 6.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0151 | 0.0153 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 158.4
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.5 | 158.3 | 159.8 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.015 | 0.0153 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5279 | 27.3 | 158.3
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.5 | 158.3 | 159.8 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.015 | 0.0153 | 0.009121 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5493 | 27.3 | 158.3
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.5 | 158.3 | 159.8 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015 | 0.015221 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 158.3
2008-10-20 | 17:44 | -0.002 | 0 | 0 | 1.4 | 158.3 | 159.7 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2791 | 0.015 | 0.0152 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6036 | 27.3 | 158.3
2008-10-20 | 17:44 | -0.006 | 0 | 0 | 1.4 | 158.2 | 159.6 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015 | 0.0152 | 0.009179 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6808 | 27.3 | 158.2
2008-10-20 | 17:44 | -0.01 | 0 | 0 | 1.4 | 158.2 | 159.6 | -1 | 0 | 6.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.015 | 0.0152 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.642 | 27.3 | 158.2
2008-10-20 | 17:44 | -0.007 | 0 | 0 | 1.4 | 158.2 | 159.6 | -1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.015 | 0.0152 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6352 | 27.3 | 158.2
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.121 | 159.5 | -1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0152 | 0.0092 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.663 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.1 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0152 | 0.009121 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4315 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.0792 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0152 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4667 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.1 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0149 | 0.0152 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4562 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.1 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0149 | 0.015121 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.379 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.121 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.3019 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.2 | 159.6 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5261 | 27.3 | 158.2
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.4 | 158.1209 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6525 | 27.3 | 158.1
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.5 | 158.021 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0149 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7904 | 10.6108 | 27.3 | 158
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.5 | 158 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.014821 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4139 | 27.3 | 158
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.5 | 158 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0148 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5682 | 27.3 | 158
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.5 | 158 | 159.5 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0148 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4778 | 27.3 | 158
2008-10-20 | 17:44 | 0 | 0 | 0 | 1.5 | 157.9 | 159.4 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0148 | 0.0151 | 0.0091 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 157.9
2008-10-20 | 17:44 | 0 | 0 | 0 | 0.9 | 149.6719 | 150.6 | -1.1 | 0 | 6.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0148 | 0.0151 | 0.008942 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 149.7
2008-10-20 | 17:44 | 0 | 0 | 0 | -1.5 | 130.584 | 130.6 | -1.2 | 0 | 5.77 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.014242 | 0.015 | 0.007526 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 130.6
2008-10-20 | 17:44 | 0 | 0 | 0 | -1 | 112.88 | 112.9 | -1.2 | 0 | 4.82 | 13 | 953 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.01311 | 0.013652 | 0.006505 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5385 | 27.3 | 112.9
2008-10-20 | 17:44 | 0 | 0 | 0 | -0.4 | 107.721 | 107.7 | -1.2 | 0 | 4.47 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.012521 | 0.013 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 107.7
2008-10-20 | 17:44 | 0 | 0 | 0 | -0.4 | 107.621 | 107.6 | -1.2 | 0 | 4.31 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.0124 | 0.012821 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 107.6
2008-10-20 | 17:44 | 0 | 0 | 0 | -0.3 | 106.9673 | 107 | -1.2 | 0 | 4.3 | 13 | 954 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.3 | 28 | 5 | 195 | 59.3 | 0.0123 | 0.012721 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5972 | 27.3 | 107
2008-10-20 | 17:44 | 0 | 0 | 0 | -0.3 | 106.7 | 106.7 | -1.2 | 0 | 4.3 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0122 | 0.012621 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4428 | 27.3 | 106.7
2008-10-20 | 17:44 | 0 | -0.0001 | -1.0679 | -0.3 | 106.621 | 106.6 | -1.2 | 0 | 4.2 | 13 | 954 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0121 | 0.0125 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5116 | 27.3 | 106.6
2008-10-20 | 17:44 | 0 | -0.0009 | -9.346 | -0.1 | 106.6 | 106.6 | -1.1 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2205 | 0.012 | 0.012321 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.6
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -0.7 | 106.5207 | 106.5 | -1 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.011741 | 0.012221 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5899 | 27.3 | 106.5
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1.1 | 106.5 | 106.5 | -0.7 | 0 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2792 | 0.0116 | 0.0122 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5127 | 27.3 | 106.5
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1 | 106.5 | 106.5 | -0.3 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0115 | 0.0121 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4355 | 27.3 | 106.5
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1 | 106.421 | 106.4 | 0 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.011421 | 0.012021 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4833 | 27.3 | 106.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1 | 106.4 | 106.4 | 0.2 | 0.2421 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0114 | 0.011921 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1.2 | 106.4 | 106.4 | 0.5 | 0.4545 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2212 | 0.0113 | 0.0119 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1.2 | 106.4 | 106.4 | 0.5 | 0.5353 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.011221 | 0.0118 | 0.0062 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4537 | 27.3 | 106.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1.2 | 106.4 | 106.4 | 0.6 | 0.6 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0112 | 0.0118 | 0.006121 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5006 | 27.3 | 106.4
2008-10-20 | 17:44 | 0 | -0.001 | -10 | -1.2 | 106.4 | 106.4 | 0.5 | 0.54 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2212 | 0.0112 | 0.0118 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -1.1 | 106.4 | 106.4 | 0.5 | 0.4604 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.011121 | 0.0117 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5953 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.0009 | -9.01 | -1.1 | 106.4 | 106.4 | 0.4 | 0.4 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.0111 | 0.0117 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5181 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.0001 | -0.69052 | -0.9 | 106.4 | 106.4 | 0.3 | 0.3 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2207 | 0.011021 | 0.0116 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.0008 | -7.6544 | -0.9 | 106.4 | 106.4 | 0.2 | 0.2 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.011 | 0.0115 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5724 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.8 | 106.4 | 106.4 | 0.1 | 0.1 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.010921 | 0.0115 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6867 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.8 | 106.4 | 106.4 | 0.1 | 0.1 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0109 | 0.0114 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.6095 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.8 | 106.4 | 106.4 | 0.1 | 0.1237 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0109 | 0.0114 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.7 | 106.4 | 106.4 | 0.3 | 0.3096 | 4.2 | 13 | 956 | 953 | 944 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2791 | 0.0109 | 0.0114 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7908 | 10.6 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.7 | 106.3 | 106.3 | 0.6 | 0.5594 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0108 | 0.0113 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5779 | 27.3 | 106.3
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.6 | 106.3 | 106.3 | 0.9 | 0.879 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0108 | 0.0113 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5008 | 27.3 | 106.3
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.6 | 106.4 | 106.4 | 1.1 | 1.1153 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.3 | 28 | 5 | 195 | 59.3 | 0.0108 | 0.0112 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4236 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.5 | 106.4 | 106.4 | 1.3 | 1.3386 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0107 | 0.0112 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 106.4
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.5 | 106.3207 | 106.3 | 1.5 | 1.5 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34.3 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0107 | 0.0112 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.461 | 27.3 | 106.3
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 1.5 | 1.5 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.26 | 0.01 | 989 | -34.7 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0107 | 0.011121 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.3
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 1.4 | 1.4304 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0107 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.3
2008-10-20 | 17:45 | 0 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 1.3 | 1.3354 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0107 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4767 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.003 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 1.2 | 1.1592 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0107 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4773 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.01 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 1.1 | 1.0529 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0107 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6159 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.01 | -0.001 | -10 | -0.5 | 106.2791 | 106.3 | 0.8 | 0.8443 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2791 | 0.010621 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6937 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.01 | -0.001 | -10 | -0.5 | 106.3 | 106.3 | 0.8 | 0.7604 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0106 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6291 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.01 | -0.001 | -10 | -0.4 | 106.3 | 106.3 | 0.6 | 0.5629 | 4.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0106 | 0.0111 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.3
2008-10-20 | 17:45 | -0.01 | -0.0002 | -2.1126 | -0.4 | 106.2 | 106.2 | 0.4 | 0.4452 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0106 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | 0.3 | 0.2643 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0106 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5951 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | 0.1 | 0.1 | 4.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0106 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4406 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.1 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0106 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4569 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.1 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191.3 | 28 | 5 | 195 | 59.3 | 0.0105 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.534 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.2 | 0 | 4.2 | 13 | 954 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0105 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5776 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.3 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0105 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4233 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.01 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.3 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0105 | 0.011 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.531 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.009 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.4 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0105 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.2
2008-10-20 | 17:45 | -0.001 | 0 | 0 | -0.4 | 106.2 | 106.2 | -0.4 | 0 | 4.2 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0105 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.2
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.1788 | 106.2 | -0.5 | 0 | 4.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2212 | 0.0105 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106.2
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.121 | 106.1 | -0.6 | 0 | 4.2 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0105 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.6 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.1 | 106.1 | -0.6 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2793 | 0.0105 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6513 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.1 | 106.1 | -0.6 | 0 | 4.2 | 13 | 955 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.0104 | 0.0109 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6145 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.1 | 106.1 | -0.7 | 0 | 4.2 | 13 | 954 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4114 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.3 | 106.1 | 106.1 | -0.7 | 0 | 4.2 | 13 | 955 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5657 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.2 | 106.1 | 106.1 | -0.8 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6601 | 27.3 | 106.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.2 | 106.0209 | 106 | -0.8 | 0 | 4.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5882 | 27.3 | 106
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.2 | 106 | 106 | -0.8 | 0 | 4.2 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2209 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4289 | 27.3 | 106
2008-10-20 | 17:45 | 0.003 | 0 | 0 | -0.2 | 106 | 106 | -0.8 | 0 | 4.2 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5832 | 27.3 | 106
2008-10-20 | 17:45 | 0.01 | 0 | 0 | -0.2 | 106 | 106 | -0.8 | 0 | 4.2 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0104 | 0.0108 | 0.0061 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 106
2008-10-20 | 17:45 | 0.01 | 0 | 0 | -0.4 | 94.5812 | 94.6 | -0.8 | 0 | 3.87 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2788 | 0.0104 | 0.010721 | 0.005448 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 94.6
2008-10-20 | 17:45 | 0.01 | 0 | 0 | 0 | 61.293 | 61.3 | -0.8 | 0 | 3.01 | 13.1 | 953 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.008815 | 0.009057 | 0.003573 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 61.3
2008-10-20 | 17:45 | 0.01 | 0 | 0 | -0.7 | 55.7655 | 55.8 | -0.8 | 0 | 2.52 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.008321 | 0.008621 | 0.003221 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 55.8
2008-10-20 | 17:45 | 0.006 | 0 | 0 | -1.1 | 54.8667 | 54.9 | -0.8 | 0 | 2.3 | 13 | 954 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2792 | 0.008021 | 0.008421 | 0.0032 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 54.9
2008-10-20 | 17:45 | 0 | 0 | 0 | -1 | 54.521 | 54.5 | -0.9 | 0 | 2.3 | 13 | 954 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.007921 | 0.008321 | 0.0032 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4918 | 27.3 | 54.5
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.9 | 54.5 | 54.5 | -0.9 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0078 | 0.0082 | 0.0032 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4627 | 27.3 | 54.5
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.9 | 54.5 | 54.5 | -0.9 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.007721 | 0.008042 | 0.003121 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.591 | 27.3 | 54.5
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.8 | 54.321 | 54.3 | -1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0076 | 0.007921 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5139 | 27.3 | 54.3
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.8 | 54.2211 | 54.2 | -1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2211 | 0.0075 | 0.007821 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5633 | 27.3 | 54.2
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.1 | 54.121 | 54.1 | -1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.007321 | 0.007721 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 54.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.2 | 54.1 | 54.1 | -1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2212 | 0.0072 | 0.007621 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5824 | 27.3 | 54.1
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.2 | 54 | 54 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0071 | 0.007521 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5052 | 27.3 | 54
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.1 | 54 | 54 | -1.1 | 0 | 2.2 | 13.1 | 955 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0071 | 0.0075 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5719 | 27.3 | 54
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.1 | 53.2667 | 53.3 | -1.1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.007 | 0.0074 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 53.3
2008-10-20 | 17:45 | 0 | 0 | 0 | -1.1 | 53 | 53 | -1.1 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.007 | 0.007321 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5475 | 27.3 | 53
2008-10-20 | 17:45 | 0 | 0 | 0 | -1 | 52.9 | 52.9 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0069 | 0.0073 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4068 | 27.3 | 52.9
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.9 | 52.9 | 52.9 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.006821 | 0.0072 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5612 | 27.3 | 52.9
2008-10-20 | 17:45 | 0 | 0 | 0 | -0.9 | 52.9 | 52.9 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0068 | 0.0072 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6578 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | 0 | 0 | -0.9 | 52.9 | 52.9 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2789 | 0.0068 | 0.0072 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.665 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | 0 | 0 | -0.9 | 52.821 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.006721 | 0.0071 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5758 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | 0 | 0 | -0.8 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0067 | 0.007 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4215 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | 0 | 0 | -0.8 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.006621 | 0.007 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5993 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.0006 | -5.9578 | -0.8 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2 | 0.0066 | 0.0069 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6564 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.7 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0066 | 0.0069 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.7 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0066 | 0.0069 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.7 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0065 | 0.006821 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.7 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 952 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.221 | 0.0065 | 0.0068 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4433 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.6 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2791 | 0.0065 | 0.0068 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3294 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.6 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0064 | 0.0067 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.6 | 52.879 | 52.9 | -1.2 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.0064 | 0.0067 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.5 | 52.9 | 52.9 | -1.2 | 0 | 2.2 | 13 | 955 | 952 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2212 | 0.0064 | 0.0067 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.522 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.5 | 52.9 | 52.9 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.279 | 0.006321 | 0.006621 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.5238 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.4 | 52.9 | 52.9 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0066 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2052 | 10.4 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.4 | 52.9 | 52.9 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0066 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 52.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.4 | 52.8208 | 52.8 | -1.1 | 0 | 2.2 | 13.1 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.2208 | 0.0063 | 0.0065 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5388 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.4 | 52.8 | 52.8 | -1.1 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0065 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5071 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.3 | 52.8 | 52.8 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0063 | 0.0065 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.447 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.3 | 52.8 | 52.8 | -1.1 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 192 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0065 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5989 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.3 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191.7 | 28 | 5 | 195 | 59.2 | 0.0062 | 0.0065 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4438 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.3 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.006421 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.3 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4326 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5098 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.587 | 27.3 | 52.8
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2207 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5359 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.006 | -0.001 | -10 | -0.2 | 52.8792 | 52.9 | -1.3 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2208 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2084 | 10.5 | 27.3 | 52.9
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.9 | 52.9 | -1.3 | 0 | 2.2 | 13.1 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4816 | 27.3 | 52.9
2008-10-20 | 17:46 | -0.01 | -0.0004 | -4.0354 | -0.2 | 52.9 | 52.9 | -1.3 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.3 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4044 | 27.3 | 52.9
2008-10-20 | 17:46 | -0.01 | -0.0004 | -4.1397 | -0.2 | 52.821 | 52.8 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 942 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5455 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.2 | 0.0062 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5002 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27.293 | 191 | 28 | 5 | 195 | 59.2 | 0.0061 | 0.0064 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4271 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 28 | 191 | 28 | 5 | 195 | 59.2211 | 0.0061 | 0.006321 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7888 | 10.4957 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27.707 | 191 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4185 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 191 | 28 | 5 | 195 | 59.221 | 0.0061 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3413 | 27.3587 | 52.8
2008-10-20 | 17:46 | -0.01 | -0.0006 | -5.875 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190.7 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4076 | 27.3641 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2207 | 0.0061 | 0.0063 | 0.0031 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.0061 | 0.0063 | 0.003021 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 957 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2 | 0.0061 | 0.0063 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5326 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.8 | 52.8 | -1.2 | 0 | 2.2 | 13 | 957 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.006 | 0.0063 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5445 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.8 | 52.8 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2212 | 0.006 | 0.0062 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5566 | 27.3 | 52.8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.2 | 52.7209 | 52.7 | -1.3 | 0 | 2.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2209 | 0.006 | 0.0062 | 0.003 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4023 | 27.3 | 52.7
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -0.8 | 38.6305 | 38.6 | -1.3 | 0 | 1.85 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.005921 | 0.0062 | 0.00221 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 38.6
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -3.5 | 8.032 | 8 | -1.4 | 0 | 1 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.221 | 0.00481 | 0.005468 | 0.000431 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5063 | 27.3 | 8
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -2.5 | 3.289 | 3.3 | -1.4 | 0 | 0.45 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.3 | 0.003684 | 0.004284 | 0.0002 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6303 | 27.3 | 3.3
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -2.4 | 2.2 | 2.2 | -1.5 | 0 | 0.3 | 13 | 955 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.279 | 0.003221 | 0.003721 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6927 | 27.3 | 2.2
2008-10-20 | 17:46 | -0.01 | 0 | 0 | -2.3 | 1.4655 | 1.5 | -1.5 | 0 | 0.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 190 | 28 | 5 | 195 | 59.2207 | 0.0032 | 0.003621 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6153 | 27.3 | 1.5
2008-10-20 | 17:46 | -0.01 | -0.0004 | -4.1118 | -2.2 | 1.3 | 1.3 | -1.5 | 0 | 0.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189.7 | 28 | 5 | 195 | 59.2208 | 0.0031 | 0.003521 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 1.3
2008-10-20 | 17:46 | -0.01 | -0.001 | -10 | -2.1 | 1.3 | 1.3 | -1.5 | 0 | 0.2 | 13 | 955 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.279 | 0.003021 | 0.003421 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 1.3
2008-10-20 | 17:46 | -0.002 | -0.001 | -10 | -2 | 1.221 | 1.2 | -1.5 | 0 | 0.2 | 13 | 955 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0029 | 0.003321 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.21 | 10.6 | 27.3 | 1.2
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.9 | 1.2 | 1.2 | -1.5 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.002821 | 0.003221 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.6 | 27.3 | 1.2
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.9 | 1.2 | 1.2 | -1.5 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.002721 | 0.003121 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4591 | 27.3 | 1.2
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.9 | 1.1 | 1.1 | -1.5 | 0 | 0.2 | 13 | 955 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2209 | 0.0027 | 0.0031 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4477 | 27.3 | 1.1
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.8 | 1.1 | 1.1 | -1.5 | 0 | 0.2 | 13 | 956 | 954 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0026 | 0.003 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5248 | 27.3 | 1.1
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.8 | 1.0212 | 1 | -1.6 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0026 | 0.002921 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 1
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.7 | 0.921 | 0.9 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0025 | 0.002821 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.9
2008-10-20 | 17:46 | 0 | -0.001 | -10 | -1.6 | 0.9 | 0.9 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.002421 | 0.0028 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4873 | 27.3 | 0.9
2008-10-20 | 17:46 | 0 | -0.0002 | -2.2985 | -1.6 | 0.9 | 0.9 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0024 | 0.0027 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5008 | 27.3 | 0.9
2008-10-20 | 17:46 | 0 | 0 | 0 | -1.6 | 0.9 | 0.9 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2208 | 0.002321 | 0.0027 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2084 | 10.6892 | 27.3 | 0.9
2008-10-20 | 17:46 | 0 | 0 | 0 | -1.6 | 0.9 | 0.9 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0023 | 0.0026 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.612 | 27.3 | 0.9
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.6 | 0.821 | 0.8 | -1.6 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.3 | 0.0023 | 0.0026 | 0.0001 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.5349 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.6 | 0.8 | 0.8 | -1.6 | 0 | 0.2 | 13 | 957 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.002221 | 0.0026 | 2.1e-05 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5423 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.6 | 0.8 | 0.8 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.7 | 28 | 5 | 195 | 59.221 | 0.0022 | 0.0025 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2209 | 0.002121 | 0.002421 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | -1.6 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0021 | 0.0024 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.5 | 0.8 | 0.8 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.0024 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.8
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.5 | 0.721 | 0.7 | -1.6 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0021 | 0.0024 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5438 | 27.3 | 0.7
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.4 | 0.7 | 0.7 | -1.6 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2793 | 0.0021 | 0.002321 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4105 | 27.3 | 0.7
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.4 | 0.7 | 0.7 | -1.7 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.002 | 0.0023 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5648 | 27.3 | 0.7
2008-10-20 | 17:47 | 0 | 0 | 0 | -1.4 | 0.7 | 0.7 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.279 | 0.002 | 0.0023 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.7
2008-10-20 | 17:47 | 0 | -0.0006 | -6.0531 | -1.4 | 0.62116 | 0.6 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2212 | 0.002 | 0.0023 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2116 | 10.6368 | 27.3 | 0.6
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.6 | 0.6 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.002 | 0.0022 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6581 | 27.3 | 0.6
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.6 | 0.6 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.002 | 0.0022 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4266 | 27.3 | 0.6
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.6 | 0.6 | -1.7 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0019 | 0.0022 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5366 | 27.3 | 0.6
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.0019 | 0.0022 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2096 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0019 | 0.002122 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.0019 | 0.0021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2795 | 0.0019 | 0.0021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2052 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2207 | 0.0019 | 0.0021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.5 | 0.5 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.001821 | 0.0021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.5
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.421 | 0.4 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.0018 | 0.0021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5834 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.2 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0018 | 0.002021 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4293 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.13 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0018 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5254 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.1 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0018 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2092 | 10.56 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.1 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.0018 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.5343 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0018 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.6886 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.001721 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.7 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2205 | 0.0017 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.2052 | 10.6514 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2207 | 0.0017 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.7928 | 10.5484 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0017 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0017 | 0.002 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0017 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5147 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.2 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0017 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.279 | 0.0017 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.3 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2209 | 0.0017 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.8 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.001542 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004 | 10.4679 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -1.9 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 954 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0015 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1004.79 | 10.3568 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.4 | 0.4 | 0.4 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.221 | 0.001421 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.3611 | 27.3 | 0.4
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.5 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2793 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5926 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.5 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.4 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.2 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.6519 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.4 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.671 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.4 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.5877 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.4 | 0.3 | 0.3 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 944 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 188.3 | 28 | 5 | 195 | 59.3 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4333 | 27.3 | 0.3
2008-10-20 | 17:47 | 0 | -0.001 | -10 | -2.3 | 0.2 | 0.2 | -1.7 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.2 | 0.0014 | 0.0019 | 0 | 0 |  | n/a |  | 33 | 55 | 1040 | 1005 | 10.4 | 27.3 | 0.2
2008-10-20 | 17:47 | 0.024 | -0.001 | -10 | -2.3 | 0.2 | 0.2 | -1.6 | 0 | 0.1 | 13 | 956 | 953 | 943 | 19 | 0.28 | 0.01 | 989 | -34 | 0 | 27 | 189 | 28 | 5 | 195 | 59.221 | 0.0014 | 0.001821 | 0 | 0 |  | n/a |  | 33 | 55 | 1035.7781 | 1005 | 10.4755 | 27.3 | 0.2
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-10-20
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 37249
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 1989
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 302
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.00
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 17:26
Test End Time | 17:47
Test Duration (s) | 1266
Average Ambient Temperature (deg C) | -34.537
Average Ambient Pressure (mbar) | 989
Average Relative Humidity (%) | 19
Average Absolute Humidity (grains/lb dry air) | 0.27
Average Kh Factor | 0.833
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 10/20/2008 17:26:10.695 - None Found
Warnings:
0X0000 - 10/20/2008 17:26:10.696 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1266
Brake Specific Dropout Limit Count | 1266
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-10-20 | 17:26 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
Cerion
light w/generator

## NO2 Linearity
Test Date:
Test Time:
10/20/2008
17:26:11
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## NDUV
Test Date:
Test Time:
10/20/2008
17:26:11
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## O2 Linearity
Test Date:
Test Time:
10/20/2008
17:26:11
SEMTECH Serial Number:
D06-SDS01

## O2
Test Date:
Test Time:
10/20/2008
17:26:11
SEMTECH Serial Number:
D06-SDS01