## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 26374
NDUV Serial Number: | 210
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 30.64 | Scott gas | ALM018306 | ±1% | 2008-06-27 | 91.92
200 ppm CO | 201.2
6% CO2 | 5.953
300 ppm NO | 302.6
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | CO2 - AMBII Serial Number: |  |  | 26374 |  |  | CO - AMBII Serial Number: |  |  | 26374 |  | HC - AMBII Serial Number: |  |  | 26374 |  | NO - NDUV Serial Number: |  |  | 210 |  | THC - FID Serial Number: |  |  | 66
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-05-06 |  | 12 | X | 0.1 | 3 |  |  | 1201 |  | 50 | 3 |  | 319 |  | 4 | 3 |  | 1495 |  | 15 | 3 |  | 957 | 1000 | 5 | 2
Test Time: | 16:21
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -2 | -2 |  |  | 0 | 996.1 | 996.1 |  |  | 0 | -0.3 | -0.3 |  |  | 0 | -0.9 | -0.9
90 | 105 | 1 | 1.2 | 1.184 | 1.226 | 0.042 | 3.55 |  | 120.1 | 129 | 9 | 7.49 |  | 31.9 | 996 | 964.1 | 3022.26 | Fail | 149.5 | 150 | 0.5 | 0.33 |  | 95.7 | 94.1 | -1.6 | -1.67
150 | 165 | 2 | 2.4 | 2.372 | 2.43 | 0.058 | 2.45 |  | 240.2 | 240 | 0 | 0 |  | 63.8 | 996.5 | 932.7 | 1461.91 | Fail | 299 | 298.6 | -0.4 | -0.13 |  | 191.4 | 189.2 | -2.2 | -1.15
210 | 225 | 3 | 3.6 | 3.563 | 3.613 | 0.05 | 1.4 |  | 360.3 | 365 | 5 | 1.39 |  | 95.7 | 996.6 | 900.9 | 941.38 | Fail | 448.5 | 447 | -1.5 | -0.33 |  | 287.1 | 284.7 | -2.4 | -0.84
270 | 285 | 4 | 4.8 | 4.758 | 4.78 | 0.022 | 0.46 |  | 480.4 | 480 | 0 | 0 |  | 127.6 | 996.5 | 868.9 | 680.96 | Fail | 598 | 596.9 | -1.1 | -0.18 |  | 382.8 | 378.9 | -3.9 | -1.02
330 | 345 | 5 | 6 | 5.956 | 5.95 | -0.006 | -0.1 |  | 600.5 | 603 | 3 | 0.5 |  | 159.5 | 995.9 | 836.4 | 524.39 | Fail | 747.5 | 746.6 | -0.9 | -0.12 |  | 478.5 | 473.5 | -5 | -1.04
390 | 405 | 6 | 7.2 | 7.158 | 7.197 | 0.039 | 0.54 |  | 720.6 | 726 | 5 | 0.69 |  | 191.4 | 996.3 | 804.9 | 420.53 | Fail | 897 | 897.2 | 0.2 | 0.02 |  | 574.2 | 569.1 | -5.1 | -0.89
450 | 465 | 7 | 8.4 | 8.363 | 8.375 | 0.012 | 0.14 |  | 840.7 | 839 | -2 | -0.24 |  | 223.3 | 996.4 | 773.1 | 346.22 | Fail | 1046.5 | 1046 | -0.5 | -0.05 |  | 669.9 | 664.3 | -5.6 | -0.84
510 | 525 | 8 | 9.6 | 9.572 | 9.58 | 0.008 | 0.08 |  | 960.8 | 960 | -1 | -0.1 |  | 255.2 | 996.2 | 741 | 290.36 | Fail | 1196 | 1197.5 | 1.5 | 0.13 |  | 765.6 | 761.5 | -4.1 | -0.54
570 | 585 | 9 | 10.8 | 10.784 | 10.787 | 0.003 | 0.03 |  | 1080.9 | 1080 | -1 | -0.09 |  | 287.1 | 996.3 | 709.2 | 247.02 | Fail | 1345.5 | 1346.9 | 1.4 | 0.1 |  | 861.3 | 857 | -4.3 | -0.5
630 | 645 | 10 | 12 | 12 | 11.972 | -0.028 | -0.23 |  | 1201 | 1199 | -2 | -0.17 |  | 319 | 996.6 | 677.6 | 212.41 | Fail | 1495 | 1499.9 | 4.9 | 0.33 |  | 957 | 954.5 | -2.5 | -0.26
690 | 705 | 9 | 10.8 | 10.784 | 10.805 | 0.021 | 0.19 |  | 1080.9 | 1081 | 0 | 0 |  | 287.1 | 996.2 | 709.1 | 246.99 | Fail | 1345.5 | 1352.7 | 7.2 | 0.54 |  | 861.3 | 857.1 | -4.2 | -0.49
750 | 765 | 8 | 9.6 | 9.572 | 9.595 | 0.023 | 0.24 |  | 960.8 | 960 | -1 | -0.1 |  | 255.2 | 996 | 740.8 | 290.28 | Fail | 1196 | 1202.6 | 6.6 | 0.55 |  | 765.6 | 760.8 | -4.8 | -0.63
810 | 825 | 7 | 8.4 | 8.363 | 8.396 | 0.033 | 0.39 |  | 840.7 | 841 | 0 | 0 |  | 223.3 | 996.1 | 772.8 | 346.08 | Fail | 1046.5 | 1048.4 | 1.9 | 0.18 |  | 669.9 | 663.7 | -6.2 | -0.93
870 | 885 | 6 | 7.2 | 7.158 | 7.167 | 0.009 | 0.13 |  | 720.6 | 720 | -1 | -0.14 |  | 191.4 | 996 | 804.6 | 420.38 | Fail | 897 | 900.3 | 3.3 | 0.37 |  | 574.2 | 568.3 | -5.9 | -1.03
930 | 945 | 5 | 6 | 5.956 | 5.955 | -0.001 | -0.02 |  | 600.5 | 600 | -1 | -0.17 |  | 159.5 | 996 | 836.5 | 524.45 | Fail | 747.5 | 750 | 2.5 | 0.33 |  | 478.5 | 472 | -6.5 | -1.36
990 | 1005 | 4 | 4.8 | 4.758 | 4.777 | 0.019 | 0.4 |  | 480.4 | 486 | 6 | 1.25 |  | 127.6 | 996 | 868.4 | 680.56 | Fail | 598 | 598.1 | 0.1 | 0.02 |  | 382.8 | 376.5 | -6.3 | -1.65
1050 | 1065 | 3 | 3.6 | 3.563 | 3.616 | 0.053 | 1.49 |  | 360.3 | 360 | 0 | 0 |  | 95.7 | 995.7 | 900 | 940.44 | Fail | 448.5 | 448.5 | 0 | 0 |  | 287.1 | 281.9 | -5.2 | -1.81
1110 | 1125 | 2 | 2.4 | 2.372 | 2.433 | 0.061 | 2.57 |  | 240.2 | 240 | 0 | 0 |  | 63.8 | 995.6 | 931.8 | 1460.5 | Fail | 299 | 299.2 | 0.2 | 0.07 |  | 191.4 | 186.8 | -4.6 | -2.4
1170 | 1185 | 1 | 1.2 | 1.184 | 1.24 | 0.056 | 4.73 |  | 120.1 | 124 | 4 | 3.33 |  | 31.9 | 995.7 | 963.8 | 3021.32 | Fail | 149.5 | 152 | 2.5 | 1.67 |  | 95.7 | 91.2 | -4.5 | -4.7
1230 | 1245 | 0 | 0 | 0 | -0.002 | -0.002 |  |  | 0 | -6 | -6 |  |  | 0 | 995.8 | 995.8 |  |  | 0 | -0.1 | -0.1 |  |  | 0 | -3.7 | -3.7
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | 0.039 | 0.12 |  |  |  | Intercept | 2 | 12 |  |  | Intercept | 996 | 3 | Fail |  | Intercept | -0.9 | 15 |  |  | Intercept | -3.4 | 9.6
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 0 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0227 | 0.12 |  |  |  | SEE | 3.3697 | 12.01 |  |  | SEE | 0.2752 | 3.19 |  |  | SEE | 1.9987 | 14.95 |  |  | SEE | 1.5047 | 9.57
|  |  | r2 | 1 | 0.998 |  |  |  | r2 | 0.9999 | 0.998 |  |  | r2 | 0.1789 | 0.998 | Fail |  | r2 | 1 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.5 | 0.7 | 1.2 | -0.8 | 0 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0004 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.2 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | -1 | 0 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.000366 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.947 | 27.2 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | -0.9 | 0 | 0 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.3666 | 0.0003 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.2 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3663 | 0.0003 | 0.000266 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8785 | 27.1696 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.4 | 0.7 | 1.1 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.333 | 0.0003 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.6354 | 27.1089 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.3 | 0.7 | 1 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3665 | 0.0003 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9489 | 27.1872 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.3 | 0.7 | 1 | -0.9 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.000267 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8687 | 27.1344 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.2 | 0.7 | 0.9 | -0.9 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.2 | 0.7 | 0.9 | -0.9 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.7
2008-05-06 | 16:21 | 0.01 | -0.001 | -10 | 0.1 | 0.7 | 0.8 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0002 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8009 | 27.1009 | 0.7
2008-05-06 | 16:22 | 0.009 | -0.001 | -10 | 0.1 | 0.7 | 0.8 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0002 | 0.000168 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8793 | 27.1793 | 0.7
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0.1 | 0.6 | 0.7 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3338 | 0.000166 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8423 | 27.1423 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0.1 | 0.6 | 0.7 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3666 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8146 | 27.1 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8931 | 27.1 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3354 | 25.0429 | 27.1715 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | 0 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0501 | 27.1501 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.1 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0283 | 27.1283 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.1 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3329 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.6707 | 25.0933 | 27.1933 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -1 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3273 | 25.0148 | 27.1148 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.3 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0636 | 27.1636 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3338 | 0 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0582 | 27.1582 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9799 | 27.1201 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3333 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9012 | 27.1988 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.8 | 925 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1319 | 27.2 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.6713 | 25.1443 | 27.2 | 0.6
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3324 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.2 | 0.6
2008-05-06 | 16:22 | 0 | -0.0004 | -4.0098 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3333 | 25.0881 | 27.1881 | 0.6
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.6707 | 25.0091 | 27.1091 | 0.6
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.6 | 0.6 | -0.9 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3673 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3273 | 25 | 27.1 | 0.6
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.56747 | 0.6 | -0.9 | 0 | 0 | 12.8 | 925 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.6
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3338 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9477 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.664 | 24.809 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9661 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3669 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9385 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9183 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9963 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.0008 | -7.774 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.3669 | -3.3e-05 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.3 | 0.5 | 0.5 | -0.9 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6694 | 25 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.3 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.3673 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9359 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.0007 | -7.4333 | -0.2 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3663 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.8 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3336 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.664 | 24.8671 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.3 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.3666 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9915 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.0006 | -5.9805 | -0.3 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0758 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.367 | -0.0001 | 0 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0027 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.5 | 0.5 | -0.8 | 0 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0809 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.46667 | 0.5 | -0.8 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9218 | 27.1 | 0.5
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3331 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8758 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 925 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0326 | 27.1163 | 0.4
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3673 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1895 | 27.1947 | 0.4
2008-05-06 | 16:22 | 0 | -0.0006 | -5.807 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3325 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0543 | 27.1272 | 0.4
2008-05-06 | 16:22 | 0 | -0.0003 | -2.5188 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6621 | 24.8973 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3664 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.86 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.9 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | -0.001 | -10 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3671 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | -0.0004 | -4.0407 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9352 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3 | -0.0002 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8567 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | 0 | 0 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.8 | 926 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3669 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.6687 | 24.8 | 27.1 | 0.4
2008-05-06 | 16:22 | 0 | 0.001 | 10.0328 | -0.2 | 0.4 | 0.4 | -0.8 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3 | -0.0002 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8 | 27.1 | 0.4
2008-05-06 | 16:22 | 0.669 | 0.0108 | 108.3934 | 6.3 | 0.4 | 6.7 | -0.8 | 0 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.3673 | 0.000913 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.167 | 0.0149 | 149.4 | 55.2 | 0.4 | 55.6 | -0.6 | 0 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3675 | 0.009224 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3253 | 24.9718 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.222 | 0.016 | 160 | 131 | 0.4 | 131.4 | 23.6 | 23.6348 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.021976 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.0148 | 147.5223 | 144.6 | 0.4 | 145 | 73.8 | 73.8406 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.024201 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 995.3353 | 25.0857 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.013 | 130 | 148.2 | 0.4 | 148.6 | 88.9 | 88.8976 | 0 | 12.9 | 927 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3666 | 0.024733 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0073 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.013 | 130 | 149.1 | 0.4 | 149.5 | 92.4 | 92.384 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3668 | 0.024933 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.013 | 130 | 149.7 | 0.4 | 150.1 | 93.1 | 93.0547 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.025033 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.013 | 130 | 150 | 0.4 | 150.4 | 93.3 | 93.2726 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.4
2008-05-06 | 16:22 | 1.23 | 0.013 | 130 | 150 | 0.4 | 150.4 | 93.6 | 93.5745 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9871 | 27.1 | 0.4
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 0.4 | 150.4 | 93.7 | 93.6714 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.8303 | 27.1 | 0.4
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 0.5 | 150.6 | 93.8 | 93.8 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3671 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.5
2008-05-06 | 16:23 | 1.23 | 0.0129 | 129.2027 | 150.1 | 0.6 | 150.7 | 93.8 | 93.8 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8826 | 27.1 | 0.6
2008-05-06 | 16:23 | 1.23 | 0.0121 | 120.897 | 150.2 | 0.63379 | 150.8 | 93.9 | 93.9135 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0194 | 27.1 | 0.6
2008-05-06 | 16:23 | 1.23 | 0.012 | 120 | 150.2 | 0.7 | 150.9 | 94 | 94 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.3335 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0985 | 27.1 | 0.7
2008-05-06 | 16:23 | 1.23 | 0.012 | 120 | 150.2 | 0.8 | 151 | 94.1 | 94.0747 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.946 | 27.1 | 0.8
2008-05-06 | 16:23 | 1.23 | 0.012 | 120 | 150.2 | 0.83307 | 151 | 94 | 94 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3331 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9555 | 27.1 | 0.8
2008-05-06 | 16:23 | 1.23 | 0.0122 | 122.3186 | 150.2 | 0.9 | 151.1 | 94.1 | 94.0779 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 0.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 0.9 | 151.1 | 94.1 | 94.1 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9761 | 27.1 | 0.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 0.9 | 151.1 | 94.2 | 94.1848 | 0 | 12.9 | 927 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8186 | 27.1 | 0.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 0.93313 | 151.1 | 94.2 | 94.2 | 0 | 12.8 | 927 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 1 | 151.2 | 94.2 | 94.2 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8951 | 27.1 | 1
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 1.1 | 151.3 | 94.2 | 94.1774 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.1
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 1.1325 | 151.3 | 94.2 | 94.1791 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3325 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9924 | 27.1 | 1.1
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.2 | 1.2338 | 151.4 | 94.2 | 94.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8346 | 27.1 | 1.2
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 6.1407 | 156.2 | 94.1 | 94.1207 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3336 | 0.0251 | -0.0002 | 0.000234 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8613 | 27.1 | 6.1
2008-05-06 | 16:23 | 1.23 | 0.0123 | 122.7099 | 150 | 8.9352 | 158.9 | 94.2 | 94.1559 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0002 | 0.000333 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9398 | 27.1 | 8.9
2008-05-06 | 16:23 | 1.23 | 0.0126 | 125.6926 | 149.9 | 10.9 | 160.8 | 94.1 | 94.1 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3332 | 0.0251 | -0.0001 | 0.0004 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 10.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 149.9 | 10.9 | 160.8 | 94.1 | 94.1 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0004 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 10.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 149.9 | 10.9667 | 160.9 | 94.1 | 94.1 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.3 | 0.0251 | -0.0001 | 0.0004 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.925 | 27.1 | 11
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 149.9 | 10.4702 | 160.4 | 94.1 | 94.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.000367 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0068 | 27.1 | 10.5
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 149.9 | 9.0737 | 159 | 94.1 | 94.1 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3328 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 9.1
2008-05-06 | 16:23 | 1.226 | 0.013 | 130 | 150 | 8.1 | 158.1 | 94.1 | 94.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3674 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0795 | 27.1 | 8.1
2008-05-06 | 16:23 | 1.22 | 0.013 | 130 | 150 | 7.9675 | 158 | 94.2 | 94.1774 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9234 | 27.1 | 8
2008-05-06 | 16:23 | 1.22 | 0.013 | 130 | 150 | 7.8663 | 157.9 | 94.2 | 94.1779 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.833 | 27.1 | 7.9
2008-05-06 | 16:23 | 1.22 | 0.013 | 130 | 150 | 7.8 | 157.8 | 94.1 | 94.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9368 | 27.1 | 7.8
2008-05-06 | 16:23 | 1.22 | 0.013 | 130 | 150 | 7.8 | 157.8 | 94.2 | 94.1768 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0759 | 27.1 | 7.8
2008-05-06 | 16:23 | 1.22 | 0.013 | 130 | 150 | 7.7 | 157.7 | 94.1 | 94.1 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 7.7
2008-05-06 | 16:23 | 1.224 | 0.0123 | 122.6435 | 150 | 7.6 | 157.6 | 94.1 | 94.1171 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 7.6
2008-05-06 | 16:23 | 1.23 | 0.012 | 120 | 150 | 7.6 | 157.6 | 94.1 | 94.0762 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9406 | 27.1 | 7.6
2008-05-06 | 16:23 | 1.23 | 0.0124 | 123.9333 | 150 | 7.6 | 157.6 | 94.2 | 94.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3667 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9378 | 27.1 | 7.6
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7.5669 | 157.6 | 94.1 | 94.1159 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9838 | 27.1 | 7.6
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7.4671 | 157.5 | 94.1 | 94.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9053 | 27.1 | 7.5
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7.2673 | 157.3 | 94.1 | 94.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0003 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9729 | 27.1 | 7.3
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7.0675 | 157.1 | 94.2 | 94.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8973 | 27.1 | 7.1
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7 | 157 | 94.1 | 94.1 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3662 | 0.0251 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8598 | 27.1 | 7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7 | 157 | 94.2 | 94.2 | 0 | 12.9 | 926 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3 | 0.0251 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9832 | 27.1 | 7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 7 | 157 | 94.2 | 94.2 | 0 | 12.8 | 927 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0251 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.8264 | 27.1 | 7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 6.9 | 156.9 | 94.1 | 94.1222 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.8648 | 27.1 | 6.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 6.867 | 156.9 | 94.2 | 94.1794 | 0 | 12.8 | 927 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9433 | 27.1 | 6.9
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 6.8 | 156.8 | 94.2 | 94.2 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.3335 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.1 | 6.8
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150 | 6.7667 | 156.8 | 94.1 | 94.1212 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3333 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.0995 | 6.8
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 6.7 | 156.8 | 94.1 | 94.0823 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.021 | 6.7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 6.7 | 156.8 | 94.2 | 94.2 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.0574 | 6.7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 6.7 | 156.8 | 94.2 | 94.1905 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.7
2008-05-06 | 16:23 | 1.23 | 0.013 | 130 | 150.1 | 6.6675 | 156.8 | 94.1 | 94.1 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.7
2008-05-06 | 16:23 | 1.234 | 0.013 | 130 | 150.1 | 6.6 | 156.7 | 94.2 | 94.2 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3336 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.6
2008-05-06 | 16:23 | 1.24 | 0.013 | 130 | 150.1 | 6.5659 | 156.7 | 94.2 | 94.2 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3659 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9288 | 27.1 | 6.6
2008-05-06 | 16:23 | 1.293 | 0.0138 | 138.0065 | 150.1 | 6.5 | 156.6 | 94.2 | 94.2 | 0 | 12.8 | 927 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.3 | 0.0252 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9494 | 27.1 | 6.5
2008-05-06 | 16:23 | 2.068 | 0.017 | 169.9334 | 154.7 | 6.467 | 161.2 | 94.2 | 94.2 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.367 | 0.025926 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.5
2008-05-06 | 16:23 | 2.31 | 0.0239 | 239.4496 | 205.8 | 6.4335 | 212.2 | 94.2 | 94.1827 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.033918 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.4
2008-05-06 | 16:23 | 2.387 | 0.0231 | 230.9822 | 273 | 6.5 | 279.5 | 120.3 | 120.297 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3333 | 0.044033 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.5
2008-05-06 | 16:23 | 2.404 | 0.023 | 230 | 289 | 6.6 | 295.6 | 170.8 | 170.8076 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.046299 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 6.6
2008-05-06 | 16:23 | 2.41 | 0.023 | 230 | 292.5 | 6.6 | 299.1 | 184.5 | 184.5161 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0469 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0418 | 27.1 | 6.6
2008-05-06 | 16:23 | 2.41 | 0.023 | 230 | 293.6 | 6.6 | 300.2 | 187.7 | 187.72 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0471 | -0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6747 | 25.1 | 27.1 | 6.6
2008-05-06 | 16:23 | 2.41 | 0.023 | 230 | 294 | 6.6 | 300.6 | 188.5 | 188.4623 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3675 | 0.0471 | -6.8e-05 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27.1 | 6.6
2008-05-06 | 16:23 | 2.41 | 0.0231 | 230.5495 | 294.4 | 6.6 | 301 | 188.6 | 188.6 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3338 | 0.0472 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0228 | 27.0228 | 6.6
2008-05-06 | 16:23 | 2.41 | 0.0239 | 238.8759 | 295 | 6.4664 | 301.5 | 188.8 | 188.78 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.047334 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0556 | 27 | 6.5
2008-05-06 | 16:23 | 2.414 | 0.024 | 240 | 295.8 | 6.4 | 302.2 | 189 | 188.986 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.047433 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0661 | 27.0339 | 6.4
2008-05-06 | 16:23 | 2.42 | 0.024 | 240 | 296.2 | 6.4 | 302.6 | 189 | 188.9766 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0475 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.1 | 6.4
2008-05-06 | 16:24 | 2.42 | 0.0244 | 244.0916 | 296.3 | 6.3 | 302.6 | 189.1 | 189.0787 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0475 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.1 | 6.3
2008-05-06 | 16:24 | 2.42 | 0.025 | 250 | 296.4 | 6.3 | 302.7 | 189 | 189.0192 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3335 | 0.0475 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.1 | 6.3
2008-05-06 | 16:24 | 2.42 | 0.025 | 250 | 296.6 | 6.3 | 302.9 | 189.2 | 189.2 | 0 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.047533 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6687 | 25 | 27.1 | 6.3
2008-05-06 | 16:24 | 2.42 | 0.025 | 250 | 296.6 | 6.2671 | 302.9 | 189.1 | 189.1223 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0476 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.026 | 27.1 | 6.3
2008-05-06 | 16:24 | 2.42 | 0.025 | 250 | 296.7 | 6.2 | 302.9 | 189.1 | 189.1 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3673 | 0.0476 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0961 | 27.1 | 6.2
2008-05-06 | 16:24 | 2.422 | 0.0244 | 244.3159 | 296.7 | 6.2 | 302.9 | 189.2 | 189.2 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3325 | 0.0476 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3253 | 25.0171 | 27.1 | 6.2
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 296.7 | 6.2 | 302.9 | 189.2 | 189.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0476 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0615 | 27.1 | 6.2
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 296.7 | 6.1664 | 302.9 | 189.2 | 189.2 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0476 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.06 | 27.1 | 6.2
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 296.9 | 6.1 | 303 | 189.2 | 189.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3334 | 0.047633 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.0815 | 6.1
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 297.7 | 6 | 303.7 | 189.3 | 189.2877 | 0 | 12.8 | 927 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.047733 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.0031 | 6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 6 | 304.4 | 189.2 | 189.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3671 | 0.0478 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3287 | 25.0754 | 27 | 6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.7665 | 304.2 | 189.2 | 189.1783 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3665 | 0.0478 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3347 | 25.0462 | 27.0538 | 5.8
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.7 | 304.1 | 189.3 | 189.3 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3 | 0.0478 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9678 | 27.1 | 5.7
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.5669 | 304 | 189.3 | 189.2806 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3331 | 0.0479 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9107 | 27.1 | 5.6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.4 | 303.8 | 189.2 | 189.2195 | 0 | 12.8 | 927 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3 | 0.0479 | 0 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6707 | 24.9891 | 27.1 | 5.4
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.4 | 303.8 | 189.3 | 189.2714 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0479 | 3.3e-05 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6727 | 25.0675 | 27.1 | 5.4
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 5.4 | 303.8 | 189.2 | 189.2226 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 5.4
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 5.3 | 303.8 | 189.3 | 189.2846 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0755 | 5.3
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 5.3 | 303.8 | 189.3 | 189.2565 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0971 | 27.0029 | 5.3
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 5.2 | 303.7 | 189.3 | 189.3 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3336 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0187 | 27.0813 | 5.2
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 5.1 | 303.6 | 189.1 | 189.1092 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.333 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0598 | 27.1 | 5.1
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 4.9663 | 303.5 | 189.3 | 189.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0618 | 5
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 4.8667 | 303.4 | 189.2 | 189.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3 | 0.0479 | 0.0001 | 0.000167 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0834 | 27 | 4.9
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 4.7669 | 303.3 | 189.2 | 189.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.3669 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0049 | 27 | 4.8
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 4.7 | 303.2 | 189.3 | 189.3 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.853 | 27.0735 | 4.7
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 4.6 | 303 | 189.2 | 189.2238 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.3327 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3273 | 24.9038 | 27.0481 | 4.6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.4 | 4.6 | 303 | 189.4 | 189.3562 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6761 | 25.0304 | 27.0304 | 4.6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.5 | 4.6 | 303.1 | 189.2 | 189.2401 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0823 | 27.0911 | 4.6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.6 | 4.6 | 303.2 | 189.3 | 189.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3336 | 0.0479 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9254 | 27.0127 | 4.6
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.6 | 4.5 | 303.1 | 189.3 | 189.2786 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9658 | 27 | 4.5
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.6 | 4.5 | 303.1 | 189.3 | 189.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0442 | 27 | 4.5
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.6 | 4.4 | 303 | 189.2 | 189.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0547 | 27 | 4.4
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.7 | 4.4 | 303.1 | 189.2 | 189.2 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3354 | 24.9011 | 27.0011 | 4.4
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.7 | 4.3 | 303 | 189.1 | 189.1 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9795 | 27.0795 | 4.3
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.7 | 4.1671 | 302.9 | 189.1 | 189.122 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25 | 27.042 | 4.2
2008-05-06 | 16:24 | 2.43 | 0.024 | 240 | 298.7 | 3.9673 | 302.7 | 189.2 | 189.1786 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3673 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9636 | 27 | 4
2008-05-06 | 16:24 | 2.43 | 0.0244 | 243.7615 | 298.7 | 3.9 | 302.6 | 189.1 | 189.0791 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.3 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9297 | 27 | 3.9
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.9 | 302.6 | 189.1 | 189.1156 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0867 | 27 | 3.9
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.9 | 302.6 | 189.1 | 189.0508 | 0 | 12.8 | 927 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27.0718 | 3.9
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.6 | 3.8666 | 302.5 | 189 | 188.9821 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3334 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0498 | 27.0498 | 3.9
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.8 | 302.5 | 189 | 188.979 | 0 | 12.8 | 927 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0287 | 27 | 3.8
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.7671 | 302.5 | 189 | 189 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0929 | 27 | 3.8
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.6 | 3.7 | 302.3 | 189.2 | 189.1881 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0144 | 27 | 3.7
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.7 | 302.4 | 189.1 | 189.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3667 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8721 | 27 | 3.7
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.7 | 302.4 | 188.9 | 188.9424 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9272 | 27 | 3.7
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.6672 | 302.4 | 189.1 | 189.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0001 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.7
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.6 | 3.6 | 302.2 | 189.1 | 189.0582 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.6
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.7 | 3.5 | 302.2 | 189.2 | 189.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.5
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.6 | 3.5 | 302.1 | 189.2 | 189.2 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9875 | 27 | 3.5
2008-05-06 | 16:24 | 2.43 | 0.025 | 250 | 298.5 | 3.5 | 302 | 189.2 | 189.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.334 | 0.048 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3399 | 24.9 | 27 | 3.5
2008-05-06 | 16:24 | 2.667 | 0.0254 | 253.7615 | 298.5 | 3.4 | 301.9 | 189.1 | 189.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.048 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9264 | 27 | 3.4
2008-05-06 | 16:24 | 3.22 | 0.0346 | 345.5046 | 319 | 3.4 | 322.4 | 189.1 | 189.1218 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.050896 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0832 | 27 | 3.4
2008-05-06 | 16:24 | 3.57 | 0.0339 | 338.5021 | 379.7 | 3.4 | 383.1 | 189.6 | 189.5745 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3671 | 0.059032 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6713 | 25.03 | 27 | 3.4
2008-05-06 | 16:24 | 3.579 | 0.0336 | 335.6621 | 434.5 | 3.4 | 437.9 | 228.3 | 228.327 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.066135 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 3.4
2008-05-06 | 16:24 | 3.61 | 0.0348 | 347.8617 | 440.3 | 3.4 | 443.7 | 273 | 273.049 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0669 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0268 | 27 | 3.4
2008-05-06 | 16:24 | 3.61 | 0.036 | 360 | 441.9 | 3.4 | 445.3 | 281 | 280.9583 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.067099 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.4
2008-05-06 | 16:24 | 3.61 | 0.036 | 360 | 445.3 | 3.4 | 448.7 | 283 | 283.0003 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.067533 | 0.0002 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.4
2008-05-06 | 16:24 | 3.61 | 0.036 | 360 | 446.3 | 3.4 | 449.7 | 284.1 | 284.0942 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3327 | 0.0677 | 0.000233 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9754 | 27 | 3.4
2008-05-06 | 16:24 | 3.61 | 0.036 | 360 | 446.4 | 3.4 | 449.8 | 284.1 | 284.1442 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0677 | 0.0003 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9408 | 27 | 3.4
2008-05-06 | 16:24 | 3.612 | 0.036 | 360 | 446.4 | 3.4 | 449.8 | 284.2 | 284.1838 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0677 | 0.0003 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0192 | 27 | 3.4
2008-05-06 | 16:25 | 3.62 | 0.0364 | 363.92 | 446.4 | 3.4 | 449.8 | 284.2 | 284.2 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0678 | 0.0003 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0977 | 27 | 3.4
2008-05-06 | 16:25 | 3.62 | 0.0368 | 368 | 446.4 | 3.4 | 449.8 | 284.3 | 284.3 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0678 | 0.0003 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9477 | 27 | 3.4
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.3 | 3.533 | 449.9 | 284.2 | 284.2246 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0678 | 0.0004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0092 | 27 | 3.5
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.3 | 3.7 | 450 | 284.3 | 284.3 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3666 | 0.0678 | 0.0004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.7
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.2 | 3.7 | 449.9 | 284.7 | 284.7139 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0678 | 0.0004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6667 | 25.0667 | 27 | 3.7
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.2 | 3.7 | 449.9 | 284.7 | 284.7257 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0678 | 0.000433 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8303 | 27 | 3.7
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.2 | 3.7 | 449.9 | 284.6 | 284.5798 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9368 | 27 | 3.7
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.2 | 3.7 | 449.9 | 284.6 | 284.5766 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 3.7
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 446.3 | 3.7 | 450 | 284.6 | 284.6226 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 3.7
2008-05-06 | 16:25 | 3.614 | 0.036 | 360 | 446.2 | 3.7 | 449.9 | 284.5 | 284.5446 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 3.7
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.3 | 3.6664 | 450 | 284.6 | 284.6 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 3.7
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.4 | 3.6 | 450 | 284.6 | 284.6 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.334 | 25.0607 | 27 | 3.6
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.4 | 3.6 | 450 | 284.8 | 284.7921 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.4 | 3.6 | 450 | 284.5 | 284.5226 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.3 | 3.5659 | 449.9 | 284.7 | 284.7209 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0679 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.3 | 3.5 | 449.8 | 284.6 | 284.6173 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.3333 | 0.067933 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.5
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.4 | 3.4 | 449.8 | 284.8 | 284.7776 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3669 | 0.068 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.4
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.9 | 3.4 | 450.3 | 284.9 | 284.8505 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.068 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0687 | 27 | 3.4
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 447 | 3.4 | 450.4 | 284.7 | 284.7 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.068 | 0.0005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0099 | 27 | 3.4
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 447 | 3.4 | 450.4 | 284.8 | 284.8226 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.068 | 0.000534 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0882 | 27 | 3.4
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 447 | 3.3 | 450.3 | 284.6 | 284.6409 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.068034 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.3
2008-05-06 | 16:25 | 3.61 | 0.036 | 360 | 446.9 | 3.3 | 450.2 | 284.8 | 284.7822 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 26.9548 | 3.3
2008-05-06 | 16:25 | 3.614 | 0.036 | 360 | 447 | 3.3 | 450.3 | 284.6 | 284.6253 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 26.9 | 3.3
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 447 | 3.2 | 450.2 | 284.7 | 284.7 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 26.902 | 3.2
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 447 | 3.2 | 450.2 | 284.7 | 284.7 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 26.9801 | 3.2
2008-05-06 | 16:25 | 3.62 | 0.036 | 360 | 447 | 3.2 | 450.2 | 284.9 | 284.8771 | 0 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.2
2008-05-06 | 16:25 | 3.62 | 0.036 | 360.3356 | 447 | 3.2 | 450.2 | 284.8 | 284.7641 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3331 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6687 | 25.1 | 27 | 3.2
2008-05-06 | 16:25 | 3.614 | 0.0369 | 368.7248 | 447.1 | 3.2 | 450.3 | 284.6 | 284.5757 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0006 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.2
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447 | 3.1684 | 450.2 | 284.7 | 284.6803 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.000632 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 3.2
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447 | 3.0675 | 450.1 | 284.5 | 284.5224 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6747 | 25.1 | 27 | 3.1
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447.1 | 2.9 | 450 | 284.6 | 284.562 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0681 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.9
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447.1 | 2.8664 | 450 | 284.6 | 284.5791 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.068134 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.9
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447 | 2.8 | 449.8 | 284.5 | 284.5 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1082 | 26.9918 | 2.8
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447 | 2.8 | 449.8 | 284.6 | 284.6 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3666 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1866 | 26.9134 | 2.8
2008-05-06 | 16:25 | 3.61 | 0.037 | 370 | 447 | 2.7 | 449.7 | 284.6 | 284.5869 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3335 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0706 | 26.9647 | 2.7
2008-05-06 | 16:25 | 3.614 | 0.037 | 370 | 447.1 | 2.6 | 449.7 | 284.5 | 284.5 | 0 | 12.9 | 927 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0432 | 27 | 2.6
2008-05-06 | 16:25 | 3.63 | 0.037 | 370 | 447 | 2.5 | 449.5 | 284.5 | 284.5 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.078 | 27 | 2.5
2008-05-06 | 16:25 | 3.63 | 0.037 | 370 | 447 | 2.4 | 449.4 | 284.3 | 284.3241 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6707 | 25.0004 | 27 | 2.4
2008-05-06 | 16:25 | 3.63 | 0.037 | 370 | 447 | 2.3 | 449.3 | 284.7 | 284.7 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0789 | 27 | 2.3
2008-05-06 | 16:25 | 3.63 | 0.037 | 370 | 447 | 2.3 | 449.3 | 284.7 | 284.6772 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3325 | 0.0682 | 0.0007 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0428 | 27 | 2.3
2008-05-06 | 16:25 | 3.626 | 0.037 | 370 | 446.9 | 2.2662 | 449.2 | 284.9 | 284.9301 | 0 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.000734 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0357 | 27 | 2.3
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 2.1664 | 449.1 | 285.2 | 285.1773 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 2.2
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 2.1 | 449 | 285.1 | 285.1217 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 2.1
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 2.1 | 449 | 285.3 | 285.2663 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0682 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0289 | 27 | 2.1
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.8 | 2.1 | 448.9 | 285.5 | 285.5 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.068233 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0493 | 27 | 2.1
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.8 | 2.0665 | 448.9 | 285.7 | 285.7 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0683 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6646 | 25.0724 | 27 | 2.1
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 2 | 448.9 | 285.9 | 285.8612 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0683 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0065 | 27 | 2
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 1.9 | 448.8 | 286.3 | 286.3 | 0 | 12.8 | 926 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0683 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0849 | 27 | 1.9
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 446.9 | 1.9 | 448.8 | 286.5 | 286.4646 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3 | 0.0683 | 0.0008 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0366 | 27 | 1.9
2008-05-06 | 16:25 | 3.62 | 0.037 | 370 | 447 | 1.9 | 448.9 | 286.9 | 286.9 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0683 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0418 | 27 | 1.9
2008-05-06 | 16:25 | 3.627 | 0.0373 | 373.1167 | 446.9 | 1.9 | 448.8 | 287 | 286.9769 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0683 | 0.0008 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.9
2008-05-06 | 16:25 | 4.234 | 0.0465 | 464.7833 | 456.9 | 1.9 | 458.8 | 287.1 | 287.078 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.06955 | 0.000834 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.9
2008-05-06 | 16:25 | 4.716 | 0.0466 | 465.6083 | 522.7 | 1.9 | 524.6 | 286.9 | 286.8758 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.077429 | 0.0009 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.662 | 25.1 | 27 | 1.9
2008-05-06 | 16:25 | 4.742 | 0.046 | 460 | 578 | 1.9 | 579.9 | 322.4 | 322.4164 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.083767 | 0.0009 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0443 | 27 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.046 | 460 | 589.1 | 1.9 | 591 | 366.8 | 366.7697 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0851 | 0.0009 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 26.9661 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.046 | 460 | 591.1 | 1.9 | 593 | 377.3 | 377.3145 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0853 | 0.001 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.334 | 25.0121 | 26.9121 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.046 | 460 | 594.3 | 1.9 | 596.2 | 379.5 | 379.5317 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.3665 | 0.085734 | 0.001 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.091 | 26.991 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.046 | 460 | 595.5 | 1.9 | 597.4 | 379.9 | 379.8785 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0859 | 0.001 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8915 | 27 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.0482 | 481.5803 | 595.4 | 1.9 | 597.3 | 379.9 | 379.9 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0859 | 0.0011 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8958 | 27 | 1.9
2008-05-06 | 16:25 | 4.75 | 0.049 | 490 | 595.5 | 1.9329 | 597.4 | 379.9 | 379.9138 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0859 | 0.0011 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0791 | 26.9736 | 1.9
2008-05-06 | 16:26 | 4.75 | 0.049 | 490 | 595.5 | 2.0327 | 597.5 | 380 | 380 | 0 | 12.9 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3673 | 0.0859 | 0.0011 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.2903 | 26.9049 | 2
2008-05-06 | 16:26 | 4.75 | 0.049 | 490 | 595.5 | 2.1 | 597.6 | 379.9 | 379.9 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0859 | 0.0011 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1333 | 26.9834 | 2.1
2008-05-06 | 16:26 | 4.75 | 0.049 | 490 | 595.6 | 2.2 | 597.8 | 379.9 | 379.9 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.086 | 0.0011 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.336 | 25.0382 | 27 | 2.2
2008-05-06 | 16:26 | 4.766 | 0.049 | 490 | 595.6 | 2.2 | 597.8 | 379.9 | 379.92 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.086 | 0.0011 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0402 | 27 | 2.2
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 595.8 | 2.1668 | 598 | 379.8 | 379.8112 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.086 | 0.001133 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0813 | 27 | 2.2
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 595.9 | 2.1 | 598 | 379.8 | 379.8163 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0861 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0029 | 27 | 2.1
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596 | 2.1 | 598.1 | 380 | 379.9539 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0861 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0756 | 27 | 2.1
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.3 | 2.1 | 598.4 | 379.8 | 379.804 | 0 | 12.8 | 927 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.086133 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.946 | 2.1
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.5 | 2.1 | 598.6 | 379.6 | 379.6229 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0675 | 26.9325 | 2.1
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.6 | 2.0672 | 598.7 | 379.7 | 379.7 | 0 | 12.8 | 927 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 26.9891 | 2.1
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.7 | 2 | 598.7 | 379.6 | 379.6 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 26.9105 | 2
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.8 | 2 | 598.8 | 379.6 | 379.6 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3253 | 25.0679 | 26.9679 | 2
2008-05-06 | 16:26 | 4.78 | 0.049 | 490 | 596.8 | 2 | 598.8 | 379.4 | 379.4 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9536 | 2
2008-05-06 | 16:26 | 4.78 | 0.0488 | 488.0182 | 596.8 | 2 | 598.8 | 379.3 | 379.3 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9248 | 2
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 2 | 599 | 379.2 | 379.2222 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0012 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0969 | 27 | 2
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597.1 | 2 | 599.1 | 379.1 | 379.1455 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.001233 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0183 | 27 | 2
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597.1 | 2 | 599.1 | 379.1 | 379.1 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0862 | 0.0013 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9398 | 27 | 2
2008-05-06 | 16:26 | 4.782 | 0.048 | 480 | 597.1 | 1.9 | 599 | 379 | 379 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 6.7e-05 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9386 | 27 | 1.9
2008-05-06 | 16:26 | 4.79 | 0.048 | 480 | 597 | 1.8 | 598.8 | 379.1 | 379.1 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.017 | 27 | 1.8
2008-05-06 | 16:26 | 4.781 | 0.048 | 480 | 597 | 1.8 | 598.8 | 379.2 | 379.1545 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3329 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0954 | 27 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.8 | 598.8 | 379 | 379.0227 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.026 | 27 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597.1 | 1.8 | 598.9 | 378.9 | 378.9123 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 26.9475 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.8 | 598.8 | 378.9 | 378.9241 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3373 | 25.0309 | 26.9 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.8 | 598.7 | 378.9 | 378.8547 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.664 | 25.1 | 26.9093 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.8 | 598.6 | 378.9 | 378.9 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9875 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.7669 | 598.6 | 378.7 | 378.7255 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.8
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.7 | 1.7 | 598.4 | 378.7 | 378.7472 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.7
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.7 | 1.7 | 598.4 | 378.8 | 378.8159 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0013 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0769 | 27 | 1.7
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.7 | 1.6 | 598.3 | 378.9 | 378.9215 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0016 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.6 | 598.4 | 378.9 | 378.8933 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0799 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.6 | 598.4 | 378.9 | 378.8737 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.6 | 598.5 | 378.9 | 378.9 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.063 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.6 | 598.6 | 378.9 | 378.8846 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0155 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.6 | 598.6 | 378.8 | 378.8 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.336 | 25.094 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.6 | 598.6 | 378.7 | 378.7242 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.6 | 598.6 | 378.6 | 378.6211 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.332 | 25.0494 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.6 | 598.4 | 378.8 | 378.7736 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.8 | 1.6 | 598.4 | 378.8 | 378.8 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0863 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0077 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.6 | 598.5 | 378.7 | 378.7268 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0862 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.6 | 598.5 | 378.7 | 378.6755 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.0014 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9356 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.6 | 598.5 | 378.7 | 378.7 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.0014 | 3.3e-05 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0571 | 26.9429 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.6 | 598.5 | 378.6 | 378.6222 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.0014 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6734 | 25.0216 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 597 | 1.5675 | 598.5 | 378.7 | 378.6774 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.0014 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3246 | 25.1 | 27 | 1.6
2008-05-06 | 16:26 | 4.78 | 0.048 | 480 | 596.9 | 1.5 | 598.4 | 378.7 | 378.7 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0864 | 0.001431 | 6.9e-05 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.5
2008-05-06 | 16:26 | 4.794 | 0.048 | 480 | 596.9 | 1.5 | 598.4 | 378.7 | 378.7 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.086434 | 0.0015 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1565 | 27 | 1.5
2008-05-06 | 16:26 | 4.795 | 0.048 | 480 | 596.9 | 1.5 | 598.4 | 378.6 | 378.5756 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.001533 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1651 | 27 | 1.5
2008-05-06 | 16:26 | 4.79 | 0.048 | 480 | 596.9 | 1.5 | 598.4 | 378.7 | 378.6803 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.5
2008-05-06 | 16:26 | 4.79 | 0.048 | 480 | 596.9 | 1.467 | 598.4 | 378.6 | 378.6223 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.67 | 25.1 | 27 | 1.5
2008-05-06 | 16:26 | 4.79 | 0.048 | 480 | 597 | 1.4 | 598.4 | 378.7 | 378.7 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3354 | 25.0293 | 27 | 1.4
2008-05-06 | 16:26 | 4.79 | 0.048 | 480 | 597 | 1.3667 | 598.4 | 378.6 | 378.6161 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.4
2008-05-06 | 16:26 | 5.287 | 0.0485 | 485.1369 | 597 | 1.3 | 598.3 | 378.6 | 378.577 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0865 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6687 | 25.0274 | 27 | 1.3
2008-05-06 | 16:26 | 5.718 | 0.0574 | 573.7037 | 626.2 | 1.3 | 627.5 | 378.5 | 378.5 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.089704 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.3
2008-05-06 | 16:26 | 5.9 | 0.0583 | 583.2113 | 705.1 | 1.3 | 706.4 | 383 | 382.9666 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.098009 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.3
2008-05-06 | 16:26 | 5.9 | 0.0585 | 585.23 | 736.7 | 1.3 | 738 | 434.3 | 434.2518 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.10113 | 0.0016 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9739 | 27 | 1.3
2008-05-06 | 16:26 | 5.9 | 0.059 | 590 | 740.9 | 1.3 | 742.2 | 464.3 | 464.3421 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.10163 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8588 | 27 | 1.3
2008-05-06 | 16:26 | 5.909 | 0.059 | 590 | 741.6 | 1.3 | 742.9 | 471.2 | 471.2392 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.10173 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.664 | 24.8579 | 27 | 1.3
2008-05-06 | 16:26 | 5.917 | 0.059 | 590 | 742.7 | 1.3 | 744 | 472.7 | 472.6729 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.10183 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.0953 | 27 | 1.3
2008-05-06 | 16:26 | 5.92 | 0.059 | 590 | 743.6 | 1.3 | 744.9 | 473 | 472.9798 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.102 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.1 | 27 | 1.3
2008-05-06 | 16:26 | 5.92 | 0.059 | 590 | 743.4 | 1.4 | 744.8 | 473.2 | 473.2221 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.102 | 0.0017 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9447 | 1.4
2008-05-06 | 16:26 | 5.922 | 0.059 | 590 | 743.5 | 1.5 | 745 | 473.6 | 473.5988 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.102 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0328 | 26.9336 | 1.5
2008-05-06 | 16:27 | 5.93 | 0.0594 | 593.7 | 743.8 | 1.5 | 745.3 | 473.6 | 473.6146 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.10207 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 1.5
2008-05-06 | 16:27 | 5.938 | 0.0598 | 598.0163 | 744.9 | 1.5 | 746.4 | 473.6 | 473.6 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.10223 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 1.5
2008-05-06 | 16:27 | 5.94 | 0.059 | 590 | 745.7 | 1.5 | 747.2 | 473.9 | 473.8932 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1023 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9692 | 27 | 1.5
2008-05-06 | 16:27 | 5.94 | 0.059 | 590 | 745.6 | 1.5 | 747.1 | 473.8 | 473.8 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1023 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 1.5
2008-05-06 | 16:27 | 5.94 | 0.0597 | 596.772 | 745.8 | 1.5 | 747.3 | 473.7 | 473.7231 | 0 | 12.9 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1023 | 0.0018 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 1.5
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 745.9 | 1.4664 | 747.4 | 473.8 | 473.811 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.10233 | 0.001834 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0045 | 27.0045 | 1.5
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746 | 1.4 | 747.4 | 473.9 | 473.9 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.666 | 25.083 | 27.083 | 1.4
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.1 | 1.4 | 747.5 | 474 | 473.977 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.332 | 25.1 | 27.0386 | 1.4
2008-05-06 | 16:27 | 5.94 | 0.06 | 600.1015 | 746.1 | 1.4 | 747.5 | 473.8 | 473.8 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0602 | 27 | 1.4
2008-05-06 | 16:27 | 5.94 | 0.0609 | 608.5618 | 746.1 | 1.4 | 747.5 | 473.9 | 473.9 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0182 | 27 | 1.4
2008-05-06 | 16:27 | 5.944 | 0.061 | 610 | 746.1 | 1.4 | 747.5 | 473.9 | 473.9 | 0 | 12.8 | 926 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0963 | 27 | 1.4
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.2 | 1.4 | 747.6 | 473.9 | 473.9151 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0248 | 27 | 1.4
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.3 | 1.4 | 747.7 | 474 | 473.9765 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0537 | 27 | 1.4
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.3 | 1.4 | 747.7 | 473.8 | 473.8 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1024 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9714 | 27 | 1.4
2008-05-06 | 16:27 | 5.957 | 0.061 | 610 | 746.4 | 1.4 | 747.8 | 473.7 | 473.7226 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.10243 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7315 | 27 | 1.4
2008-05-06 | 16:27 | 5.96 | 0.061 | 610 | 746.5 | 1.3677 | 747.8 | 473.8 | 473.7752 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.3232 | 24.9648 | 27 | 1.4
2008-05-06 | 16:27 | 5.96 | 0.061 | 610 | 746.4 | 1.3 | 747.7 | 473.6 | 473.6097 | 0 | 12.9 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.0019 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:27 | 5.96 | 0.061 | 610 | 746.4 | 1.3 | 747.7 | 473.7 | 473.7 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.001933 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.046 | 27 | 1.3
2008-05-06 | 16:27 | 5.96 | 0.061 | 610 | 746.3 | 1.3 | 747.6 | 473.7 | 473.7452 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.3
2008-05-06 | 16:27 | 5.952 | 0.061 | 610 | 746.5 | 1.3 | 747.8 | 473.8 | 473.7768 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.6646 | 25.0972 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.6 | 1.3 | 747.9 | 473.7 | 473.7228 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.3333 | 25.0187 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.5 | 1.3 | 747.8 | 473.6 | 473.6 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0597 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.061 | 610 | 746.5 | 1.3 | 747.8 | 473.6 | 473.6 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0618 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.0605 | 604.834 | 746.6 | 1.3 | 747.9 | 473.6 | 473.6 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9833 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.8 | 1.3 | 748.1 | 473.4 | 473.4163 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.3253 | 24.9049 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.7 | 1.3 | 748 | 473.5 | 473.476 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0471 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.8 | 1.3 | 748.1 | 473.4 | 473.4453 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1025 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.048 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.8 | 1.3 | 748.1 | 473.4 | 473.4255 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.334 | 25 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.7 | 1.3 | 748 | 473.5 | 473.5 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.002 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0089 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.6 | 1.3 | 747.9 | 473.5 | 473.458 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0874 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.4 | 473.4376 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9684 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.4 | 473.3811 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9886 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.6 | 473.5793 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6694 | 25.0772 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.3 | 473.3272 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3259 | 25 | 27 | 1.3
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.3 | 473.3 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:27 | 5.946 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.3 | 473.3166 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3662 | 0.1026 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:27 | 5.93 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.4 | 473.4 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.10263 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0365 | 27 | 1.3
2008-05-06 | 16:27 | 5.93 | 0.06 | 600 | 746.5 | 1.3 | 747.8 | 473.4 | 473.4 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 1.3
2008-05-06 | 16:27 | 5.937 | 0.06 | 600 | 746.4 | 1.3 | 747.7 | 473.5 | 473.4987 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.668 | 25.1 | 27 | 1.3
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.4 | 1.267 | 747.7 | 473.5 | 473.479 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.3
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.4 | 1.2 | 747.6 | 473.3 | 473.3226 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0497 | 27 | 1.2
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.5 | 1.2 | 747.7 | 473.3 | 473.3204 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3333 | 25.0288 | 27 | 1.2
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.4 | 1.2 | 747.6 | 473.3 | 473.3234 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0927 | 27 | 1.2
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.4 | 1.2 | 747.6 | 473.1 | 473.1212 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0143 | 27 | 1.2
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.5 | 1.2 | 747.7 | 473.4 | 473.3541 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0021 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3273 | 25 | 27.0642 | 1.2
2008-05-06 | 16:27 | 5.945 | 0.06 | 600 | 746.6 | 1.2 | 747.8 | 473.3 | 473.3 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3246 | 25 | 27.0574 | 1.2
2008-05-06 | 16:27 | 5.95 | 0.06 | 600 | 746.5 | 1.1 | 747.6 | 473.3 | 473.3 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.979 | 27 | 1.1
2008-05-06 | 16:27 | 5.948 | 0.06 | 600 | 746.6 | 1.1 | 747.7 | 473.4 | 473.3659 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6653 | 24.9005 | 27 | 1.1
2008-05-06 | 16:27 | 5.94 | 0.06 | 600 | 746.7 | 1.1 | 747.8 | 473.2 | 473.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 1.1
2008-05-06 | 16:27 | 5.94 | 0.0602 | 601.9324 | 746.8 | 1.1 | 747.9 | 473.2 | 473.2 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1027 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9562 | 27 | 1.1
2008-05-06 | 16:27 | 6.601 | 0.061 | 609.9839 | 747.1 | 1.1 | 748.2 | 473 | 473.0449 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.10273 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0346 | 27.0346 | 1.1
2008-05-06 | 16:27 | 7.026 | 0.0693 | 693.1667 | 786.1 | 1.1 | 787.2 | 473.1 | 473.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.10651 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27.0867 | 1.1
2008-05-06 | 16:27 | 7.134 | 0.0703 | 703.3555 | 862.3 | 1.1 | 863.4 | 484 | 483.9518 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1137 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27.0082 | 1.1
2008-05-06 | 16:27 | 7.14 | 0.07 | 700 | 888.4 | 1.1 | 889.5 | 542.1 | 542.0731 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11613 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8894 | 27 | 1.1
2008-05-06 | 16:27 | 7.14 | 0.07 | 700 | 892.5 | 1.1 | 893.6 | 562.9 | 562.902 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11653 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 27.0484 | 1.1
2008-05-06 | 16:27 | 7.14 | 0.0702 | 701.9582 | 893.5 | 1.1 | 894.6 | 566.9 | 566.9108 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11663 | 0.0022 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3273 | 24.8804 | 27.0732 | 1.1
2008-05-06 | 16:27 | 7.14 | 0.071 | 710 | 893.5 | 1.1328 | 894.6 | 568.1 | 568.0548 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1167 | 0.0023 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.1
2008-05-06 | 16:27 | 7.14 | 0.071 | 710 | 893.5 | 1.2 | 894.7 | 568.6 | 568.5897 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1167 | 0.0023 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 1.2
2008-05-06 | 16:27 | 7.14 | 0.071 | 710 | 893.4 | 1.3 | 894.7 | 568.9 | 568.8765 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11673 | 0.002334 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0376 | 27 | 1.3
2008-05-06 | 16:27 | 7.14 | 0.071 | 710 | 893.6 | 1.3 | 894.9 | 569.1 | 569.0793 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1168 | 0.0024 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.14 | 0.071 | 710 | 894.8 | 1.3 | 896.1 | 569 | 569 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1169 | 0.0024 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.981 | 27 | 1.3
2008-05-06 | 16:28 | 7.14 | 0.071 | 710 | 894.9 | 1.3 | 896.2 | 569.2 | 569.158 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.11693 | 0.0024 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6646 | 24.9024 | 27 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.071 | 710 | 895.2 | 1.3 | 896.5 | 569.1 | 569.1212 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9762 | 27 | 1.3
2008-05-06 | 16:28 | 7.17 | 0.071 | 710 | 895.2 | 1.3 | 896.5 | 569.2 | 569.1647 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27.0547 | 1.3
2008-05-06 | 16:28 | 7.174 | 0.071 | 710 | 895.2 | 1.3 | 896.5 | 569.1 | 569.1 | 0 | 12.9 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9668 | 27.0668 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.071 | 710 | 895.2 | 1.3 | 896.5 | 569.3 | 569.3209 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9232 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.071 | 710 | 895.3 | 1.3 | 896.6 | 569.2 | 569.1698 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3253 | 25.0801 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.071 | 710 | 895.2 | 1.3 | 896.5 | 569.1 | 569.1254 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.117 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0315 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.071 | 710 | 895.3 | 1.3336 | 896.6 | 569.1 | 569.1 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.11703 | 0.0025 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.664 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.17 | 0.071 | 710 | 895.4 | 1.4 | 896.8 | 569.3 | 569.2795 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1171 | 0.002533 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.666 | 25 | 27.0253 | 1.4
2008-05-06 | 16:28 | 7.16 | 0.071 | 710 | 895.4 | 1.3 | 896.7 | 569.3 | 569.3 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1171 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27.0965 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.071 | 710 | 895.5 | 1.3 | 896.8 | 569.4 | 569.385 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1171 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27.0178 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.071 | 710 | 895.6 | 1.3 | 896.9 | 569.3 | 569.3 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.11713 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9392 | 27 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.0712 | 711.9067 | 896 | 1.3 | 897.3 | 569.2 | 569.2228 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3 | 0.1172 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9784 | 27 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.072 | 719.9517 | 896.5 | 1.3 | 897.8 | 569.1 | 569.1 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1172 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0825 | 27 | 1.3
2008-05-06 | 16:28 | 7.16 | 0.0728 | 728.4596 | 896.6 | 1.3 | 897.9 | 569.3 | 569.2773 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.11723 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0047 | 27 | 1.3
2008-05-06 | 16:28 | 7.171 | 0.073 | 730 | 896.8 | 1.3 | 898.1 | 569.2 | 569.19 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1173 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0744 | 27 | 1.3
2008-05-06 | 16:28 | 7.188 | 0.073 | 730 | 896.8 | 1.3 | 898.1 | 569.1 | 569.1452 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1173 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.047 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.073 | 730 | 897 | 1.3 | 898.3 | 569.1 | 569.1 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3 | 0.1173 | 0.0026 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.073 | 730 | 897.1 | 1.3 | 898.4 | 569.2 | 569.1868 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1173 | 0.002633 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0098 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.073 | 730 | 896.9 | 1.3 | 898.2 | 569.1 | 569.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3331 | 0.11733 | 0.0027 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.088 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.073 | 730 | 897.1 | 1.3 | 898.4 | 569.3 | 569.2624 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1174 | 0.0027 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0334 | 27 | 1.3
2008-05-06 | 16:28 | 7.191 | 0.073 | 730 | 896.9 | 1.3 | 898.2 | 569.1 | 569.0847 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1174 | 0.0027 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.21 | 0.073 | 730 | 896.9 | 1.3 | 898.2 | 569 | 568.9811 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1174 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9525 | 27 | 1.3
2008-05-06 | 16:28 | 7.21 | 0.073 | 730 | 897 | 1.3 | 898.3 | 569.1 | 569.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11743 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8022 | 27 | 1.3
2008-05-06 | 16:28 | 7.202 | 0.073 | 730 | 897 | 1.3 | 898.3 | 569 | 569.0184 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8806 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.073 | 730 | 897.3 | 1.3 | 898.6 | 569 | 568.9774 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9591 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.073 | 730 | 897.2 | 1.2 | 898.4 | 569 | 568.9747 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.2
2008-05-06 | 16:28 | 7.2 | 0.0723 | 723.4167 | 897.4 | 1.2 | 898.6 | 569.1 | 569.0639 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9841 | 27 | 1.2
2008-05-06 | 16:28 | 7.2 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 569 | 568.963 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9056 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 569 | 569 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 569.1 | 569.0772 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9511 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.072 | 720 | 897.4 | 1.3 | 898.7 | 569.1 | 569.1391 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9703 | 27 | 1.3
2008-05-06 | 16:28 | 7.2 | 0.072 | 720 | 897.4 | 1.3 | 898.7 | 569.2 | 569.1856 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0028 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9165 | 27.0082 | 1.3
2008-05-06 | 16:28 | 7.193 | 0.072 | 720 | 897.4 | 1.3 | 898.7 | 569.1 | 569.1 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.002833 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0733 | 27.0867 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.3 | 1.3 | 898.6 | 569.1 | 569.0539 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0354 | 27.0354 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.3 | 1.3 | 898.6 | 569.1 | 569.1184 | 0 | 12.9 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 569.1 | 569.121 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1175 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.3 | 1.3 | 898.6 | 569.3 | 569.3 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3323 | 0.1175 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9995 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.4 | 1.3 | 898.7 | 569 | 569.0354 | 0 | 12.8 | 926 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9211 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 568.9 | 568.9227 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3334 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8429 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.5 | 1.3 | 898.8 | 569 | 568.9882 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.871 | 27 | 1.3
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.6 | 1.4 | 899 | 569 | 569 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9716 | 27 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.6 | 1.4 | 899 | 569 | 569 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8145 | 27 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.6 | 1.4 | 899 | 569.2 | 569.1555 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9424 | 27 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.6 | 1.4 | 899 | 568.8 | 568.819 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0496 | 27.0496 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.5 | 1.4 | 898.9 | 569.1 | 569.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0719 | 27.0719 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.3 | 1.4 | 898.7 | 568.9 | 568.9384 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.4
2008-05-06 | 16:28 | 7.18 | 0.072 | 720 | 897.3 | 1.4 | 898.7 | 568.8 | 568.7618 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6754 | 25 | 27 | 1.4
2008-05-06 | 16:28 | 7.188 | 0.072 | 720 | 897.2 | 1.4 | 898.6 | 568.9 | 568.8507 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1176 | 0.0029 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8733 | 27 | 1.4
2008-05-06 | 16:28 | 7.804 | 0.072 | 720 | 898.2 | 1.4 | 899.6 | 568.8 | 568.8144 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1177 | 0.002934 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8833 | 27 | 1.4
2008-05-06 | 16:28 | 8.236 | 0.0827 | 826.5397 | 932.8 | 1.4 | 934.2 | 568.6 | 568.62 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.12077 | 0.003 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.666 | 25.02 | 27 | 1.4
2008-05-06 | 16:28 | 8.347 | 0.085 | 850 | 1016.5 | 1.4 | 1017.9 | 573.9 | 573.9466 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.12796 | 0.003 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0987 | 27 | 1.4
2008-05-06 | 16:28 | 8.36 | 0.085 | 850 | 1040.6 | 1.4 | 1042 | 637.2 | 637.1788 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.12993 | 0.003 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.33 | 24.8683 | 27 | 1.4
2008-05-06 | 16:28 | 8.36 | 0.085 | 850 | 1043.6 | 1.4 | 1045 | 658.3 | 658.3168 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.13023 | 0.003 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8 | 27 | 1.4
2008-05-06 | 16:28 | 8.36 | 0.0848 | 848.4661 | 1044 | 1.4 | 1045.4 | 662.7 | 662.6656 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1303 | 0.003 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8 | 27 | 1.4
2008-05-06 | 16:28 | 8.36 | 0.084 | 840.0838 | 1044.3 | 1.5331 | 1045.9 | 663.6 | 663.5784 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1303 | 0.0031 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3313 | 24.8237 | 27 | 1.5
2008-05-06 | 16:28 | 8.365 | 0.084 | 840 | 1044.3 | 1.7329 | 1046 | 664 | 663.9788 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.13033 | 0.0031 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9818 | 27 | 1.7
2008-05-06 | 16:28 | 8.37 | 0.084 | 840 | 1044.4 | 1.8 | 1046.2 | 664.2 | 664.182 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1304 | 0.003133 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9306 | 27 | 1.8
2008-05-06 | 16:28 | 8.37 | 0.084 | 840 | 1044.5 | 1.8 | 1046.3 | 664.3 | 664.2769 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3325 | 0.1304 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9957 | 27 | 1.8
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1044.7 | 1.8336 | 1046.5 | 664.3 | 664.2782 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1304 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0211 | 27 | 1.8
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1044.8 | 1.9 | 1046.7 | 664.5 | 664.484 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3666 | 0.1304 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8096 | 27 | 1.9
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045 | 1.9 | 1046.9 | 664.4 | 664.4224 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1305 | 0.0032 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9664 | 27 | 1.9
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1044.9 | 1.9 | 1046.8 | 664.5 | 664.5 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1305 | 0.003233 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 1.9
2008-05-06 | 16:29 | 8.406 | 0.084 | 840 | 1044.9 | 1.9335 | 1046.8 | 664.5 | 664.5382 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1305 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9598 | 27 | 1.9
2008-05-06 | 16:29 | 8.4 | 0.084 | 840 | 1045.1 | 2 | 1047.1 | 664.7 | 664.6802 | 0 | 12.9 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3667 | 0.1305 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27.0186 | 2
2008-05-06 | 16:29 | 8.4 | 0.084 | 840 | 1045.1 | 2 | 1047.1 | 664.5 | 664.5274 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195.1 | 59.4 | 0.1306 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27.097 | 2
2008-05-06 | 16:29 | 8.392 | 0.084 | 840 | 1045.1 | 2 | 1047.1 | 664.5 | 664.5 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195.4 | 59.4 | 0.1306 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8245 | 27.0245 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.1 | 2 | 1047.1 | 664.6 | 664.5795 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195.8 | 59.4 | 0.1306 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9616 | 27.0539 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.2 | 2 | 1047.2 | 664.4 | 664.4217 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195.8 | 59.4 | 0.1306 | 0.0033 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0677 | 27.0677 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.2 | 2 | 1047.2 | 664.5 | 664.5 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195.5 | 59.4 | 0.1306 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.4 | 2 | 1047.4 | 664.4 | 664.4 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195.1 | 59.4 | 0.1306 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.5 | 2.0334 | 1047.6 | 664.6 | 664.584 | 0 | 12.8 | 926 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.13063 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8646 | 27 | 2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.4 | 2.2 | 1047.6 | 664.5 | 664.5 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.1 | 2.2 | 1047.3 | 664.5 | 664.5 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.33 | 24.8 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.2 | 2.2 | 1047.4 | 664.5 | 664.5 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8061 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.2 | 2.2 | 1047.4 | 664.5 | 664.4791 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9631 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.3 | 2.2 | 1047.5 | 664.5 | 664.4801 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.94 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.4 | 2.2 | 1047.6 | 664.4 | 664.3762 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1307 | 0.0034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8616 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.6 | 2.2 | 1047.8 | 664.2 | 664.2231 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.13073 | 0.003433 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6747 | 24.8169 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.8 | 2.2 | 1048 | 664.2 | 664.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8953 | 27 | 2.2
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.3 | 664.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7525 | 27 | 2.1
2008-05-06 | 16:29 | 8.384 | 0.084 | 840 | 1046 | 2.1 | 1048.1 | 664.3 | 664.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8566 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.7 | 2.1 | 1047.8 | 664.2 | 664.2161 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9693 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.6 | 2.1 | 1047.7 | 664.3 | 664.2769 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8908 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.7 | 2.1 | 1047.8 | 664.3 | 664.3112 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8125 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.8 | 2.1 | 1047.9 | 664.3 | 664.3421 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3668 | 0.1308 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.866 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.3 | 664.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.13083 | 0.0035 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.3 | 664.3056 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.003533 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8312 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.0835 | 835.221 | 1046.1 | 2.0673 | 1048.2 | 664.4 | 664.4 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.6056 | 27 | 2.1
2008-05-06 | 16:29 | 8.37 | 0.083 | 830 | 1046.3 | 2 | 1048.3 | 664.4 | 664.4 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9186 | 27 | 2
2008-05-06 | 16:29 | 8.37 | 0.0831 | 831.4912 | 1046.4 | 2 | 1048.4 | 664.4 | 664.3823 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9418 | 27.0582 | 2
2008-05-06 | 16:29 | 8.37 | 0.084 | 839.9158 | 1046.4 | 2 | 1048.4 | 664.4 | 664.4226 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9367 | 27.0633 | 2
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1046.4 | 2.0332 | 1048.4 | 664.5 | 664.4716 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9697 | 27 | 2
2008-05-06 | 16:29 | 8.37 | 0.084 | 840 | 1046.2 | 2.1 | 1048.3 | 664.6 | 664.5814 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8127 | 27 | 2.1
2008-05-06 | 16:29 | 8.372 | 0.084 | 840 | 1046.1 | 2.1 | 1048.2 | 664.5 | 664.4806 | 0 | 12.9 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3354 | 24.9437 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.5 | 664.5 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1045.8 | 2.1 | 1047.9 | 664.6 | 664.5788 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.971 | 27 | 2.1
2008-05-06 | 16:29 | 8.386 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.5 | 664.5124 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1309 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9075 | 27 | 2.1
2008-05-06 | 16:29 | 8.39 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.5 | 664.5 | 0 | 12.9 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.13093 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9859 | 27 | 2.1
2008-05-06 | 16:29 | 8.387 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.7 | 664.6536 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9358 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1045.9 | 2.1 | 1048 | 664.7 | 664.7 | 0 | 12.8 | 926 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8574 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.1 | 2.1 | 1048.2 | 664.5 | 664.5 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8425 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.1 | 2.1 | 1048.2 | 664.6 | 664.6 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9994 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.1 | 2.1 | 1048.2 | 664.6 | 664.6 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.3 | 2.1 | 1048.4 | 664.8 | 664.7663 | 0 | 12.8 | 926 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0036 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9437 | 27.0563 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.4 | 2.1 | 1048.5 | 664.5 | 664.5452 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0037 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8306 | 27.0653 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.8 | 2.1 | 1048.9 | 664.7 | 664.687 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7135 | 27 | 2.1
2008-05-06 | 16:29 | 8.38 | 0.084 | 840 | 1046.9 | 2.1 | 1049 | 664.6 | 664.6229 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.792 | 27 | 2.1
2008-05-06 | 16:29 | 8.404 | 0.0857 | 857.2131 | 1046.9 | 2.1 | 1049 | 664.4 | 664.4427 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.131 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9409 | 27 | 2.1
2008-05-06 | 16:29 | 9.126 | 0.0939 | 939.1803 | 1057 | 2.1 | 1059.1 | 664.5 | 664.5 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.13182 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 2.1
2008-05-06 | 16:29 | 9.539 | 0.0924 | 923.8235 | 1140 | 2.1 | 1142.1 | 664.6 | 664.5774 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.13842 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6747 | 24.9729 | 27 | 2.1
2008-05-06 | 16:29 | 9.553 | 0.0933 | 932.7667 | 1181.6 | 2.1 | 1183.7 | 708.1 | 708.1485 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14157 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9057 | 27 | 2.1
2008-05-06 | 16:29 | 9.56 | 0.094 | 940 | 1192.2 | 2.1 | 1194.3 | 746.9 | 746.8669 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14237 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9842 | 27 | 2.1
2008-05-06 | 16:29 | 9.56 | 0.0943 | 943.2461 | 1194.1 | 2.1 | 1196.2 | 758 | 757.958 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14253 | 0.0037 | 0 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9373 | 27 | 2.1
2008-05-06 | 16:29 | 9.56 | 0.0951 | 951.4696 | 1194.2 | 2.2 | 1196.4 | 760.4 | 760.3747 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1426 | 0.0037 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9409 | 27.0409 | 2.2
2008-05-06 | 16:29 | 9.573 | 0.096 | 959.9155 | 1194.3 | 3.433 | 1197.7 | 761 | 760.9573 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1426 | 0.0038 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9619 | 27.081 | 3.4
2008-05-06 | 16:29 | 9.575 | 0.096 | 960 | 1194.2 | 3.5 | 1197.7 | 761.1 | 761.0811 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1426 | 0.0038 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8039 | 27.002 | 3.5
2008-05-06 | 16:29 | 9.57 | 0.096 | 960 | 1196 | 3.5 | 1199.5 | 761.2 | 761.2 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1428 | 0.0038 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9529 | 27.0765 | 3.5
2008-05-06 | 16:29 | 9.57 | 0.096 | 960 | 1196.3 | 3.5 | 1199.8 | 761.4 | 761.3776 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1428 | 0.003833 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9451 | 27.1 | 3.5
2008-05-06 | 16:30 | 9.57 | 0.096 | 960 | 1196.4 | 3.5327 | 1200 | 761.4 | 761.4 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1428 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27.0667 | 3.5
2008-05-06 | 16:30 | 9.57 | 0.096 | 960 | 1196.6 | 3.6 | 1200.2 | 761.5 | 761.5 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8884 | 27 | 3.6
2008-05-06 | 16:30 | 9.57 | 0.096 | 960 | 1196.4 | 3.6 | 1200 | 761.4 | 761.4226 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8097 | 27 | 3.6
2008-05-06 | 16:30 | 9.57 | 0.096 | 960 | 1196.5 | 3.6 | 1200.1 | 761.3 | 761.3222 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8687 | 27 | 3.6
2008-05-06 | 16:30 | 9.57 | 0.096 | 960 | 1196.6 | 3.6 | 1200.2 | 761.4 | 761.3784 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9 | 27 | 3.6
2008-05-06 | 16:30 | 9.571 | 0.096 | 960 | 1196.7 | 3.6 | 1200.3 | 761.5 | 761.5 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9255 | 27 | 3.6
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1196.8 | 3.5 | 1200.3 | 761.5 | 761.5 | 0 | 12.9 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1196.8 | 3.5 | 1200.3 | 761.4 | 761.4164 | 0 | 12.8 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1429 | 0.0039 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1196.9 | 3.5 | 1200.4 | 761.3 | 761.3233 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6667 | 24.878 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197 | 3.5 | 1200.5 | 761.3 | 761.3135 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8395 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.2 | 3.5 | 1200.7 | 761.4 | 761.3769 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.4 | 3.5 | 1200.9 | 761.3 | 761.3355 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.4 | 3.5 | 1200.9 | 761.4 | 761.4 | 0 | 12.9 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.3 | 3.5 | 1200.8 | 761.4 | 761.4 | 0 | 12.8 | 926 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.143 | 0.004 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9533 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.1 | 3.4 | 1200.5 | 761.3 | 761.3222 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.004034 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9683 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197 | 3.4 | 1200.4 | 761.3 | 761.3212 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8899 | 27.0101 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1196.8 | 3.4 | 1200.2 | 761.4 | 761.3767 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.812 | 27.088 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197 | 3.4 | 1200.4 | 761.2 | 761.2231 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8 | 27.0331 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1196.9 | 3.4 | 1200.3 | 761.2 | 761.2167 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8911 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197 | 3.4 | 1200.4 | 761.3 | 761.3239 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9279 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.2 | 3.5 | 1200.7 | 761.5 | 761.4968 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6707 | 24.7025 | 27 | 3.5
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.4 | 1200.9 | 761.4 | 761.3777 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7806 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.6 | 3.4 | 1201 | 761.6 | 761.5776 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3325 | 0.1431 | 0.0041 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9184 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.7 | 3.4 | 1201.1 | 761.4 | 761.4 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14313 | 0.004134 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.7 | 3.4 | 1201.1 | 761.4 | 761.4 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9673 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.4 | 3.4 | 1200.8 | 761.3 | 761.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.3 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8104 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.4 | 1200.9 | 761.5 | 761.4566 | 0 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7267 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.4 | 1200.9 | 761.4 | 761.4 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8035 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.4 | 1200.9 | 761.5 | 761.4791 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9302 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.4 | 3.4 | 1200.8 | 761.6 | 761.5857 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3327 | 24.9827 | 27 | 3.4
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.3 | 3.3 | 1200.6 | 761.6 | 761.6 | 0 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8257 | 27 | 3.3
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.3 | 1200.8 | 761.7 | 761.6917 | 0 | 12.8 | 926 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9311 | 27 | 3.3
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.5 | 3.2 | 1200.7 | 761.4 | 761.4241 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1432 | 0.0042 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.6 | 3.2 | 1200.8 | 761.5 | 761.4744 | 0 | 12.8 | 926 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.14323 | 0.004234 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.7 | 3.2 | 1200.9 | 761.7 | 761.7 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9981 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.7 | 3.2 | 1200.9 | 761.5 | 761.541 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8411 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.7 | 3.2 | 1200.9 | 761.4 | 761.4201 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9159 | 26.9421 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.8 | 3.2 | 1201 | 761.4 | 761.4 | 0 | 12.8 | 926 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9272 | 26.9364 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.9 | 3.2 | 1201.1 | 761.4 | 761.4 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8148 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.9 | 3.2 | 1201.1 | 761.3 | 761.3 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8933 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.096 | 960 | 1197.8 | 3.2 | 1201 | 761.3 | 761.3 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8282 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.0955 | 955.2083 | 1197.8 | 3.2 | 1201 | 761.3 | 761.3 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9005 | 27 | 3.2
2008-05-06 | 16:30 | 9.58 | 0.095 | 950 | 1197.7 | 3.2 | 1200.9 | 761.2 | 761.2 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9426 | 27 | 3.2
2008-05-06 | 16:30 | 9.588 | 0.095 | 950 | 1197.7 | 3.2 | 1200.9 | 761.3 | 761.3 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8143 | 26.9929 | 3.2
2008-05-06 | 16:30 | 9.584 | 0.095 | 950 | 1197.7 | 3.2 | 1200.9 | 761.1 | 761.124 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9713 | 26.9144 | 3.2
2008-05-06 | 16:30 | 9.585 | 0.095 | 950 | 1197.8 | 3.2 | 1201 | 761.1 | 761.1 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.936 | 26.964 | 3.2
2008-05-06 | 16:30 | 9.59 | 0.095 | 950 | 1197.8 | 3.3 | 1201.1 | 761.3 | 761.2712 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1433 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9425 | 27 | 3.3
2008-05-06 | 16:30 | 9.59 | 0.095 | 950 | 1198 | 3.3 | 1201.3 | 761.1 | 761.1196 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1434 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6686 | 25 | 27 | 3.3
2008-05-06 | 16:30 | 9.59 | 0.095 | 950 | 1198.3 | 3.3 | 1201.6 | 761.3 | 761.3 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1434 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.3
2008-05-06 | 16:30 | 9.59 | 0.0964 | 964.4793 | 1198.5 | 3.3 | 1201.8 | 761.1 | 761.1 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1434 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.922 | 27 | 3.3
2008-05-06 | 16:30 | 10.199 | 0.1049 | 1049.1533 | 1199.7 | 3.3 | 1203 | 761.1 | 761.1 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1435 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9563 | 26.9438 | 3.3
2008-05-06 | 16:30 | 10.668 | 0.1074 | 1073.9516 | 1262.4 | 3.3 | 1265.7 | 761.2 | 761.1847 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14676 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9653 | 26.9347 | 3.3
2008-05-06 | 16:30 | 10.78 | 0.1074 | 1073.7965 | 1318.3 | 3.3 | 1321.6 | 778 | 777.9724 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.15189 | 0.0043 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9134 | 27 | 3.3
2008-05-06 | 16:30 | 10.78 | 0.1075 | 1074.6638 | 1340.6 | 3.3 | 1343.9 | 831.9 | 831.9211 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.15343 | 0.004333 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9918 | 27 | 3.3
2008-05-06 | 16:30 | 10.78 | 0.108 | 1080 | 1341.6 | 3.3 | 1344.9 | 851.8 | 851.825 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.15353 | 0.0044 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9297 | 26.9297 | 3.3
2008-05-06 | 16:30 | 10.772 | 0.108 | 1080 | 1341.6 | 4.5 | 1346.1 | 855.6 | 855.5898 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1536 | 0.0044 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.7541 | 26.9486 | 4.5
2008-05-06 | 16:30 | 10.77 | 0.108 | 1080 | 1342.5 | 4.6 | 1347.1 | 856.5 | 856.5337 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.15363 | 0.0044 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7081 | 27 | 4.6
2008-05-06 | 16:30 | 10.77 | 0.108 | 1080 | 1344.5 | 4.5667 | 1349 | 856.9 | 856.8765 | 0 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1538 | 0.0044 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 4.6
2008-05-06 | 16:30 | 10.773 | 0.108 | 1080 | 1344.8 | 4.4668 | 1349.2 | 857 | 856.9787 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1538 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 4.5
2008-05-06 | 16:30 | 10.78 | 0.108 | 1080 | 1344.9 | 4.4 | 1349.3 | 857.2 | 857.1794 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.15383 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9373 | 26.9373 | 4.4
2008-05-06 | 16:31 | 10.78 | 0.108 | 1080 | 1345.2 | 4.4 | 1349.6 | 857.1 | 857.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1539 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8588 | 26.9412 | 4.4
2008-05-06 | 16:31 | 10.78 | 0.108 | 1080 | 1345.1 | 4.4 | 1349.5 | 857.2 | 857.1895 | 0.1 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1539 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3313 | 24.8392 | 27 | 4.4
2008-05-06 | 16:31 | 10.799 | 0.108 | 1080 | 1344.9 | 4.4 | 1349.3 | 857.2 | 857.2 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1539 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9962 | 27 | 4.4
2008-05-06 | 16:31 | 10.81 | 0.108 | 1080 | 1345 | 4.3 | 1349.3 | 857.2 | 857.2 | 0.1 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3673 | 0.1539 | 0.0045 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9235 | 27 | 4.3
2008-05-06 | 16:31 | 10.81 | 0.108 | 1080 | 1344.9 | 4.3 | 1349.2 | 857.1 | 857.1 | 0.1 | 12.9 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.15393 | 0.004533 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9 | 26.945 | 4.3
2008-05-06 | 16:31 | 10.809 | 0.108 | 1080 | 1344.9 | 4.2 | 1349.1 | 857.2 | 857.2 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9333 | 26.9333 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1345 | 4.2 | 1349.2 | 857.1 | 857.1 | 0.1 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9885 | 27 | 4.2
2008-05-06 | 16:31 | 10.808 | 0.108 | 1080 | 1345.2 | 4.2 | 1349.4 | 857.2 | 857.2 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9096 | 27 | 4.2
2008-05-06 | 16:31 | 10.804 | 0.108 | 1080 | 1345.5 | 4.2 | 1349.7 | 857.2 | 857.2 | 0.1 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9 | 27 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1345.7 | 4.2 | 1349.9 | 857.2 | 857.2 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6713 | 24.8526 | 26.9526 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346 | 4.2 | 1350.2 | 857.3 | 857.3 | 0.1 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3335 | 0.154 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8516 | 26.9 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.1 | 4.2 | 1350.3 | 857.4 | 857.4 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.15403 | 0.0046 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9918 | 26.9 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.2 | 4.1669 | 1350.4 | 857.3 | 857.3 | 0.1 | 12.8 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8346 | 26.9 | 4.2
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.2 | 4.1 | 1350.3 | 857.3 | 857.3 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8612 | 26.9 | 4.1
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.3 | 4.1 | 1350.4 | 857.3 | 857.326 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9 | 26.9397 | 4.1
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.2 | 4.1 | 1350.3 | 857.3 | 857.3 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.882 | 27 | 4.1
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.1 | 4.1 | 1350.2 | 857.3 | 857.3 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8041 | 27 | 4.1
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.3 | 4.1 | 1350.4 | 857.3 | 857.3 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7251 | 27 | 4.1
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1346.5 | 4.1 | 1350.6 | 857.2 | 857.2 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1541 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7 | 26.9464 | 4.1
2008-05-06 | 16:31 | 10.784 | 0.108 | 1080 | 1346.6 | 4.1 | 1350.7 | 857.2 | 857.2 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.7321 | 26.9321 | 4.1
2008-05-06 | 16:31 | 10.78 | 0.108 | 1080 | 1346.5 | 4.1 | 1350.6 | 857.2 | 857.2 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 26.9895 | 4.1
2008-05-06 | 16:31 | 10.78 | 0.108 | 1080 | 1346.5 | 4.1 | 1350.6 | 857.2 | 857.2204 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 26.9114 | 4.1
2008-05-06 | 16:31 | 10.783 | 0.108 | 1080 | 1346.3 | 4.1 | 1350.4 | 857 | 857 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0047 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8 | 26.9 | 4.1
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1346.4 | 4.1 | 1350.5 | 857.2 | 857.1769 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.8918 | 26.9 | 4.1
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1346.5 | 4.1 | 1350.6 | 857.2 | 857.2 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1542 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9512 | 26.9244 | 4.1
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1346.7 | 4.1 | 1350.8 | 857.1 | 857.1 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.15423 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6747 | 24.8 | 27 | 4.1
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347 | 4.1 | 1351.1 | 857.1 | 857.0881 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3379 | 24.8 | 27 | 4.1
2008-05-06 | 16:31 | 10.785 | 0.108 | 1080 | 1347.3 | 4.1 | 1351.4 | 855 | 854.9982 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 26.9401 | 4.1
2008-05-06 | 16:31 | 10.783 | 0.108 | 1080 | 1347.4 | 4.0666 | 1351.5 | 855.5 | 855.4668 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8767 | 26.9 | 4.1
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.4 | 4 | 1351.4 | 856.3 | 856.3011 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 26.9168 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.4 | 4 | 1351.4 | 857.1 | 857.0528 | 0.1 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 26.9953 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.3 | 4 | 1351.3 | 857.3 | 857.2772 | 0.1 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8525 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.3 | 4 | 1351.3 | 857.4 | 857.4 | 0.1 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.0048 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.8 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.2 | 4 | 1351.2 | 857.7 | 857.6992 | 0.1 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1543 | 0.004833 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3293 | 24.8613 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.1 | 4 | 1351.1 | 858.4 | 858.4131 | 0.1 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.9727 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.1 | 4 | 1351.1 | 858.5 | 858.5 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.7378 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.2 | 4 | 1351.2 | 858.5 | 858.5 | 0.1 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 24.964 | 26.934 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.4 | 4 | 1351.4 | 858.5 | 858.5 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0554 | 26.9446 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.4 | 4 | 1351.4 | 858.5 | 858.4763 | 0.1 | 12.8 | 926 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.3 | 4 | 1351.3 | 858.4 | 858.4 | 0.1 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.332 | 24.9984 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.5 | 4 | 1351.5 | 857.9 | 857.8609 | 0.1 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6713 | 24.9203 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.7 | 4 | 1351.7 | 857.2 | 857.2411 | 0.1 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3281 | 24.9583 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.6 | 4 | 1351.6 | 857.2 | 857.2 | 0.1 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 4
2008-05-06 | 16:31 | 10.79 | 0.108 | 1080 | 1347.8 | 4 | 1351.8 | 857.2 | 857.2 | 0.1 | 12.8 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0155 | 26.9845 | 4
2008-05-06 | 16:31 | 10.796 | 0.108 | 1080 | 1347.9 | 4 | 1351.9 | 857.3 | 857.2737 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6707 | 25.0939 | 26.9061 | 4
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1348 | 3.9673 | 1352 | 857.1 | 857.1231 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1348 | 3.9 | 1351.9 | 857.1 | 857.1 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0494 | 26.9506 | 3.9
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1348 | 3.9 | 1351.9 | 857 | 857 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1544 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9413 | 27 | 3.9
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1348 | 3.9 | 1351.9 | 857.1 | 857.1 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1545 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.336 | 24.8235 | 27 | 3.9
2008-05-06 | 16:31 | 10.8 | 0.108 | 1080 | 1347.9 | 3.9 | 1351.8 | 857.1 | 857.1 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1545 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.059 | 27 | 3.9
2008-05-06 | 16:31 | 11.057 | 0.108 | 1080 | 1348.1 | 3.9 | 1352 | 857.1 | 857.0775 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3664 | 0.1545 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.336 | 25.0353 | 27 | 3.9
2008-05-06 | 16:31 | 11.549 | 0.1181 | 1181.2552 | 1355.1 | 3.9 | 1359 | 857.1 | 857.0908 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.15496 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 3.9
2008-05-06 | 16:31 | 11.96 | 0.1198 | 1197.5593 | 1433.1 | 3.9 | 1437 | 857.2 | 857.2308 | 0.1 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.16004 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25 | 27 | 3.9
2008-05-06 | 16:31 | 11.959 | 0.119 | 1190 | 1480.7 | 3.9 | 1484.6 | 899.5 | 899.5024 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.16303 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6687 | 25.0002 | 26.9998 | 3.9
2008-05-06 | 16:31 | 11.95 | 0.119 | 1190 | 1490.2 | 3.9 | 1494.1 | 940 | 940.0396 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3331 | 0.16363 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3306 | 25.0787 | 26.9213 | 3.9
2008-05-06 | 16:31 | 11.958 | 0.119 | 1190 | 1490.4 | 3.9 | 1494.3 | 951.4 | 951.3543 | 0.1 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1637 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3293 | 25.0433 | 26.9 | 3.9
2008-05-06 | 16:31 | 11.96 | 0.119 | 1190 | 1490.7 | 5.1 | 1495.8 | 953.3 | 953.2987 | 0.1 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1637 | 0.005 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0352 | 26.9352 | 5.1
2008-05-06 | 16:31 | 11.96 | 0.119 | 1190 | 1492.2 | 5.1 | 1497.3 | 953.9 | 953.9486 | 0.1 | 12.9 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.16383 | 0.0051 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0858 | 27 | 5.1
2008-05-06 | 16:31 | 11.96 | 0.1196 | 1196.1143 | 1494.7 | 5 | 1499.7 | 954 | 954 | 0.1 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.16403 | 0.0051 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0074 | 27 | 5
2008-05-06 | 16:31 | 11.96 | 0.12 | 1200 | 1495.5 | 5 | 1500.5 | 954.1 | 954.0898 | 0.1 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1641 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 5
2008-05-06 | 16:32 | 11.96 | 0.12 | 1200 | 1495.9 | 5 | 1500.9 | 954.3 | 954.3 | 0.1 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3332 | 0.16413 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0494 | 27 | 5
2008-05-06 | 16:32 | 11.96 | 0.12 | 1200 | 1496.4 | 4.9 | 1501.3 | 954.3 | 954.2812 | 0.1 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1642 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1277 | 27 | 4.9
2008-05-06 | 16:32 | 11.954 | 0.12 | 1200 | 1496.8 | 4.8665 | 1501.7 | 954.1 | 954.1119 | 0.1 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.16423 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3354 | 25.1935 | 26.9935 | 4.9
2008-05-06 | 16:32 | 11.954 | 0.12 | 1200 | 1497 | 4.8 | 1501.8 | 954.2 | 954.2 | 0.1 | 12.9 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1643 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.115 | 26.915 | 4.8
2008-05-06 | 16:32 | 11.963 | 0.12 | 1200 | 1496.9 | 4.8 | 1501.7 | 954.2 | 954.2 | 0.1 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1643 | 0.0052 | 0.0002 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9 | 4.8
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1497.3 | 4.7671 | 1502.1 | 954.2 | 954.2 | 0.1 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1643 | 0.0052 | 0.000167 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1419 | 26.9418 | 4.8
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1497.9 | 4.5 | 1502.4 | 954.3 | 954.3373 | 0.1 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.16433 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1594 | 27 | 4.5
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1498.6 | 4.4675 | 1503.1 | 954.4 | 954.4 | 0.1 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3325 | 0.1644 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.002 | 27 | 4.5
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1499.1 | 4.4 | 1503.5 | 954.5 | 954.4818 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3663 | 0.1644 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0774 | 27 | 4.4
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1499.4 | 4.3664 | 1503.8 | 954.4 | 954.42 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1644 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0441 | 27 | 4.4
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1499.8 | 4.3 | 1504.1 | 954.3 | 954.3 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0344 | 26.9656 | 4.3
2008-05-06 | 16:32 | 11.97 | 0.1195 | 1195.4167 | 1499.9 | 4.3 | 1504.2 | 954.4 | 954.4 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9129 | 4.3
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.9 | 4.2 | 1504.1 | 954.5 | 954.4774 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3671 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9913 | 4.2
2008-05-06 | 16:32 | 11.962 | 0.119 | 1190 | 1500 | 4.1676 | 1504.1 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 4.2
2008-05-06 | 16:32 | 11.96 | 0.119 | 1190 | 1499.9 | 4.1 | 1504 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0518 | 26.9518 | 4.1
2008-05-06 | 16:32 | 11.965 | 0.119 | 1190 | 1499.9 | 4.0671 | 1504 | 954.5 | 954.4771 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3329 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3293 | 25 | 26.9267 | 4.1
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.5 | 4 | 1503.5 | 954.6 | 954.6 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.3327 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0052 | 27 | 4
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.6 | 3.9 | 1503.5 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0837 | 27 | 3.9
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.6 | 3.9 | 1503.5 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.9
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.6 | 3.9 | 1503.5 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9594 | 3.9
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.6 | 3.8 | 1503.4 | 954.4 | 954.4224 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1645 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9191 | 3.8
2008-05-06 | 16:32 | 11.97 | 0.119 | 1190 | 1499.8 | 3.8 | 1503.6 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9976 | 3.8
2008-05-06 | 16:32 | 11.97 | 0.1196 | 1196.0661 | 1500 | 3.8 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 3.8
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500 | 3.766 | 1503.8 | 954.6 | 954.6 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.34 | 25.1 | 26.9454 | 3.8
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500.2 | 3.7 | 1503.9 | 954.5 | 954.5377 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0671 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500.2 | 3.7 | 1503.9 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0111 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500.1 | 3.7 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0899 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500.1 | 3.7 | 1503.8 | 954.4 | 954.4197 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.237 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1500.1 | 3.7 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.8 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3324 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.206 | 26.947 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.12 | 1200 | 1499.9 | 3.7 | 1503.6 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.1203 | 1202.7879 | 1499.7 | 3.7 | 1503.4 | 954.4 | 954.4 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.7
2008-05-06 | 16:32 | 11.97 | 0.1209 | 1208.639 | 1499.9 | 3.7 | 1503.6 | 954.5 | 954.5 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.7
2008-05-06 | 16:32 | 11.976 | 0.12 | 1200.489 | 1499.9 | 3.7 | 1503.6 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1646 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1609 | 26.9391 | 3.7
2008-05-06 | 16:32 | 11.98 | 0.12 | 1200 | 1499.8 | 3.6 | 1503.4 | 954.4 | 954.4206 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1607 | 26.9 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.12 | 1200 | 1500 | 3.6 | 1503.6 | 954.5 | 954.4758 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.6667 | 25.0822 | 26.9178 | 3.6
2008-05-06 | 16:32 | 11.981 | 0.12 | 1200 | 1500.1 | 3.6 | 1503.7 | 954.5 | 954.5 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0038 | 26.9962 | 3.6
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.2 | 3.6 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0747 | 26.9253 | 3.6
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.4 | 3.6 | 1504 | 954.6 | 954.6 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3333 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9532 | 3.6
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.4 | 3.6325 | 1504 | 954.5 | 954.5112 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9683 | 3.6
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.3 | 3.7 | 1504 | 954.6 | 954.6 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.1 | 3.7 | 1503.8 | 954.5 | 954.5102 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1500.1 | 3.7 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.7
2008-05-06 | 16:32 | 11.99 | 0.12 | 1200 | 1499.9 | 3.6 | 1503.5 | 954.6 | 954.6 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1912 | 26.9 | 3.6
2008-05-06 | 16:32 | 11.982 | 0.12 | 1200 | 1500.1 | 3.6 | 1503.7 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2278 | 26.9241 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.12 | 1200 | 1500 | 3.6 | 1503.6 | 954.5 | 954.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3367 | 25 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.12 | 1200 | 1500.2 | 3.6 | 1503.8 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.1206 | 1206.0556 | 1500.1 | 3.6 | 1503.7 | 954.5 | 954.4895 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1647 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3313 | 25.0594 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.121 | 1210 | 1499.8 | 3.6 | 1503.4 | 954.5 | 954.4791 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.16473 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.121 | 1210 | 1500 | 3.6 | 1503.6 | 954.7 | 954.6774 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1648 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.121 | 1210 | 1500.1 | 3.6 | 1503.7 | 954.6 | 954.6192 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1648 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.6
2008-05-06 | 16:32 | 11.98 | 0.121 | 1210 | 1500.1 | 3.6 | 1503.7 | 954.6 | 954.6 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1648 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9266 | 3.6
2008-05-06 | 16:32 | 11.5 | 0.1147 | 1147.4834 | 1499.1 | 3.6 | 1502.7 | 954.5 | 954.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.16473 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.867 | 0.1112 | 1111.8868 | 1446.8 | 3.6 | 1450.4 | 954.4 | 954.4155 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1621 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.86 | 0.11 | 1100 | 1378.9 | 3.6 | 1382.5 | 935.1 | 935.1319 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.15707 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.86 | 0.11 | 1100 | 1362.1 | 3.6 | 1365.7 | 881.8 | 881.7891 | 0.1 | 12.9 | 921 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.15597 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.852 | 0.11 | 1100 | 1359 | 3.6 | 1362.6 | 861.9 | 861.9494 | 0.1 | 12.9 | 921 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.15577 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.85 | 0.11 | 1100 | 1358.3 | 3.6 | 1361.9 | 858.9 | 858.8593 | 0.1 | 12.9 | 922 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1557 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.6
2008-05-06 | 16:32 | 10.85 | 0.11 | 1100 | 1356.9 | 3.6 | 1360.5 | 858.1 | 858.0605 | 0.1 | 12.8 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1556 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0773 | 26.9227 | 3.6
2008-05-06 | 16:32 | 10.842 | 0.11 | 1100 | 1356.8 | 3.6 | 1360.4 | 857.7 | 857.7234 | 0.1 | 12.9 | 922 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.15557 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0013 | 27 | 3.6
2008-05-06 | 16:32 | 10.82 | 0.11 | 1100 | 1354.6 | 3.6 | 1358.2 | 857.6 | 857.6243 | 0.1 | 12.8 | 922 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0797 | 27 | 3.6
2008-05-06 | 16:33 | 10.82 | 0.11 | 1100 | 1354.4 | 3.6 | 1358 | 857.4 | 857.4376 | 0.1 | 12.9 | 922 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 3.6
2008-05-06 | 16:33 | 10.804 | 0.11 | 1100 | 1354.2 | 3.6 | 1357.8 | 857.3 | 857.3 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9633 | 3.6
2008-05-06 | 16:33 | 10.8 | 0.11 | 1100 | 1353.9 | 3.6 | 1357.5 | 857.2 | 857.2 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9151 | 3.6
2008-05-06 | 16:33 | 10.8 | 0.11 | 1100 | 1353.9 | 3.6 | 1357.5 | 857.1 | 857.1153 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9936 | 3.6
2008-05-06 | 16:33 | 10.8 | 0.1088 | 1087.9475 | 1353.8 | 3.6 | 1357.4 | 857.1 | 857.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0282 | 27 | 3.6
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1353.6 | 3.7 | 1357.3 | 857.2 | 857.1888 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.7
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1353.5 | 3.7 | 1357.2 | 857.1 | 857.1 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.058 | 26.971 | 3.7
2008-05-06 | 16:33 | 10.792 | 0.108 | 1080 | 1353.5 | 3.7 | 1357.2 | 857.1 | 857.1 | 0.1 | 12.8 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1851 | 26.9 | 3.7
2008-05-06 | 16:33 | 10.796 | 0.108 | 1080 | 1353.4 | 3.7329 | 1357.1 | 857.2 | 857.2 | 0.1 | 12.8 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.028 | 26.9 | 3.7
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1353.3 | 3.8 | 1357.1 | 857.2 | 857.2 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0643 | 26.9 | 3.8
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1353.2 | 3.8 | 1357 | 857.2 | 857.2 | 0.1 | 12.8 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3325 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.8
2008-05-06 | 16:33 | 10.8 | 0.1089 | 1088.5762 | 1353.2 | 3.8337 | 1357 | 857.1 | 857.1204 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1554 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1214 | 26.9 | 3.8
2008-05-06 | 16:33 | 10.8 | 0.11 | 1100 | 1353.2 | 3.9 | 1357.1 | 857.1 | 857.1226 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1999 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.1099 | 1099.0281 | 1353.3 | 3.9 | 1357.2 | 857.1 | 857.1 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0432 | 26.9784 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.1091 | 1090.503 | 1353.3 | 3.9 | 1357.2 | 857.2 | 857.179 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195.3 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 27 | 3.9
2008-05-06 | 16:33 | 10.804 | 0.109 | 1090 | 1353.2 | 3.9 | 1357.1 | 857.1 | 857.1131 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195.6 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0353 | 26.9647 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.109 | 1090 | 1353.1 | 3.9 | 1357 | 857.2 | 857.2 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 196 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9138 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.109 | 1090 | 1353 | 3.9 | 1356.9 | 857.1 | 857.1217 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195.7 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9923 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.109 | 1090 | 1352.9 | 3.9 | 1356.8 | 857.2 | 857.1787 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195.3 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9292 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.1089 | 1088.6831 | 1352.8 | 3.9 | 1356.7 | 857.2 | 857.1883 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9489 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080.4527 | 1352.8 | 3.9 | 1356.7 | 857.2 | 857.1764 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 27 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1352.9 | 3.8 | 1356.7 | 857.1 | 857.1 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3366 | 25.0875 | 27 | 3.8
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1352.9 | 3.8 | 1356.7 | 857.1 | 857.1 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.664 | 24.9306 | 27 | 3.8
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1353.1 | 3.9 | 1357 | 857 | 857 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 24.9632 | 26.9368 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1353.1 | 3.9 | 1357 | 857.2 | 857.1745 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25 | 26.9415 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1353.1 | 3.9 | 1357 | 857.1 | 857.118 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0199 | 27 | 3.9
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1353 | 3.9 | 1356.9 | 857 | 857.0226 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0987 | 27 | 3.9
2008-05-06 | 16:33 | 10.806 | 0.108 | 1080 | 1352.8 | 3.9 | 1356.7 | 857 | 856.996 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0229 | 26.9229 | 3.9
2008-05-06 | 16:33 | 10.797 | 0.108 | 1080 | 1352.9 | 3.9 | 1356.8 | 857 | 857 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0556 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.791 | 0.108 | 1080 | 1352.6 | 3.9 | 1356.5 | 857 | 857.0391 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1341 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352.4 | 3.9 | 1356.3 | 857.1 | 857.1 | 0.1 | 12.8 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3327 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.175 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352.2 | 3.9 | 1356.1 | 857.1 | 857.1 | 0.1 | 12.8 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0179 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352.2 | 3.9 | 1356.1 | 856.9 | 856.9444 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0695 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352.3 | 3.9333 | 1356.3 | 856.9 | 856.9 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352 | 4 | 1356 | 857 | 856.9771 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.334 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 857 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0951 | 26.9049 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 857 | 0.1 | 12.8 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.0166 | 26.9834 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 856.9811 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1858 | 27 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 856.9 | 856.9 | 0.1 | 12.8 | 923 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.2596 | 27 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352.1 | 4 | 1356.1 | 857 | 857.0209 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1811 | 26.9811 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857.1 | 857.0762 | 0.1 | 12.8 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1026 | 26.9026 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 857 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0241 | 26.9759 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352 | 4 | 1356 | 857.1 | 857.0767 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0544 | 27 | 4
2008-05-06 | 16:33 | 10.8 | 0.108 | 1080 | 1352 | 4 | 1356 | 857.1 | 857.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 4
2008-05-06 | 16:33 | 10.806 | 0.108 | 1080 | 1352 | 4 | 1356 | 857.1 | 857.1 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9887 | 4
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 857.0445 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9102 | 4
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1352 | 4 | 1356 | 857 | 857 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1351.9 | 4 | 1355.9 | 857 | 857 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1352 | 4 | 1356 | 857.2 | 857.2365 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 10.81 | 0.108 | 1080 | 1352 | 4 | 1356 | 857.2 | 857.2122 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 10.319 | 0.108 | 1080 | 1351.8 | 4 | 1355.8 | 857 | 857 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1553 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 9.818 | 0.1002 | 1001.7681 | 1326.5 | 4 | 1330.5 | 856.9 | 856.9476 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.15421 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3259 | 25.1 | 26.9608 | 4
2008-05-06 | 16:33 | 9.58 | 0.0986 | 985.9741 | 1250.4 | 4.0325 | 1254.4 | 852.5 | 852.517 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.14821 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1392 | 27 | 4
2008-05-06 | 16:33 | 9.581 | 0.0986 | 985.8132 | 1215.8 | 4.1 | 1219.9 | 798.1 | 798.0532 | 0.1 | 12.9 | 922 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.14577 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3379 | 25.1823 | 27 | 4.1
2008-05-06 | 16:33 | 9.59 | 0.0972 | 972.4124 | 1206.8 | 4.0664 | 1210.8 | 768.8 | 768.8296 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3664 | 0.1451 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1041 | 27 | 4.1
2008-05-06 | 16:33 | 9.59 | 0.095 | 950 | 1205.5 | 4 | 1209.5 | 763.1 | 763.1413 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.145 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9254 | 4
2008-05-06 | 16:33 | 9.59 | 0.095 | 950 | 1205.6 | 4 | 1209.6 | 761.8 | 761.7678 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.145 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1 | 26.9 | 4
2008-05-06 | 16:33 | 9.594 | 0.095 | 950 | 1205.6 | 4 | 1209.6 | 761.4 | 761.4417 | 0.1 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.145 | 0.0056 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0684 | 26.9 | 4
2008-05-06 | 16:33 | 9.6 | 0.095 | 950 | 1205.4 | 4 | 1209.4 | 761.3 | 761.3139 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1449 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.01 | 26.91 | 4
2008-05-06 | 16:33 | 9.6 | 0.095 | 950 | 1205.4 | 4 | 1209.4 | 761.1 | 761.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1449 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.0885 | 26.9885 | 4
2008-05-06 | 16:34 | 9.6 | 0.0953 | 952.7639 | 1205.3 | 4 | 1209.3 | 761.3 | 761.2932 | 0.1 | 12.8 | 924 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1449 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9329 | 4
2008-05-06 | 16:34 | 9.592 | 0.096 | 960 | 1205.1 | 4 | 1209.1 | 761.1 | 761.1222 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1449 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1204.9 | 4 | 1208.9 | 761.1 | 761.1 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.14487 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.924 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1204.7 | 4 | 1208.7 | 760.8 | 760.8385 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3338 | 0.1448 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1204.5 | 4 | 1208.5 | 760.9 | 760.9 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1448 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 27 | 4
2008-05-06 | 16:34 | 9.591 | 0.096 | 960 | 1204.3 | 4 | 1208.3 | 760.8 | 760.8209 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1448 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996.3333 | 25.2184 | 26.9408 | 4
2008-05-06 | 16:34 | 9.599 | 0.096 | 960 | 1203.8 | 4 | 1207.8 | 760.9 | 760.8804 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.14477 | 0.0055 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.3 | 26.9 | 4
2008-05-06 | 16:34 | 9.592 | 0.096 | 960 | 1203.3 | 3.9 | 1207.2 | 760.8 | 760.8 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1447 | 0.005467 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2672 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.596 | 0.096 | 960 | 1203.3 | 3.9 | 1207.2 | 760.7 | 760.7 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1447 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1102 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1203 | 3.9 | 1206.9 | 760.8 | 760.8 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1447 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9734 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.7 | 3.9 | 1206.6 | 760.7 | 760.7 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.152 | 26.948 | 3.9
2008-05-06 | 16:34 | 9.599 | 0.096 | 960 | 1202.7 | 3.9 | 1206.6 | 760.8 | 760.7879 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1696 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.591 | 0.096 | 960 | 1202.8 | 3.9 | 1206.7 | 760.8 | 760.8243 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.598 | 0.096 | 960 | 1202.8 | 3.9 | 1206.7 | 760.8 | 760.76 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.8 | 3.9 | 1206.7 | 760.7 | 760.7275 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1203 | 3.9 | 1206.9 | 761 | 760.9762 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1203 | 3.9 | 1206.9 | 760.9 | 760.9 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.1455 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.599 | 0.096 | 960 | 1203 | 3.9 | 1206.9 | 760.8 | 760.8224 | 0.1 | 12.9 | 922 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 995 | 25.2975 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.591 | 0.096 | 960 | 1202.7 | 3.9 | 1206.6 | 760.9 | 760.9 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.3381 | 25.1408 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.6 | 3.9 | 1206.5 | 760.8 | 760.7722 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9582 | 3.9
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.6 | 3.9 | 1206.5 | 760.9 | 760.8656 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1736 | 26.9632 | 3.9
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.5 | 3.9 | 1206.4 | 760.8 | 760.846 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2695 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.593 | 0.096 | 960 | 1202.7 | 3.9 | 1206.6 | 760.9 | 760.8966 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1126 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.6 | 3.9 | 1206.5 | 760.8 | 760.8231 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.7 | 3.9 | 1206.6 | 760.8 | 760.8226 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 3.9
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.9 | 4 | 1206.9 | 760.7 | 760.7132 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3664 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 995.336 | 25.1 | 26.9 | 4
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.7 | 4 | 1206.7 | 760.8 | 760.8 | 0.1 | 12.8 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3336 | 0.1446 | 0.0054 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:34 | 9.596 | 0.096 | 960 | 1202.6 | 4 | 1206.6 | 760.8 | 760.7773 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3668 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035.7228 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.5 | 4 | 1206.5 | 760.7 | 760.7 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1645 | 26.9 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.6 | 4.0337 | 1206.7 | 760.9 | 760.9 | 0.1 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2431 | 26.9 | 4
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.6 | 4.1 | 1206.7 | 760.8 | 760.8 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3667 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2567 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.5 | 4.1 | 1206.6 | 760.7 | 760.7153 | 0.1 | 12.8 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1002 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.596 | 0.096 | 960 | 1202.7 | 4.1 | 1206.8 | 760.7 | 760.7 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2571 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1202.4 | 4.1 | 1206.5 | 760.6 | 760.6 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3675 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1867 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.598 | 0.096 | 960 | 1202.3 | 4.1 | 1206.4 | 760.7 | 760.7223 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1351 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.589 | 0.096 | 960 | 1202.3 | 4.1 | 1206.4 | 760.6 | 760.6 | 0.1 | 12.8 | 923 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.214 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.581 | 0.096 | 960 | 1202.2 | 4.1 | 1206.3 | 760.7 | 760.7 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2925 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.58 | 0.096 | 960 | 1202.2 | 4.1 | 1206.3 | 760.8 | 760.7762 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.229 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.586 | 0.096 | 960 | 1202.2 | 4.1 | 1206.3 | 760.7 | 760.7115 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2492 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.1 | 4.1 | 1206.2 | 760.8 | 760.7808 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202.2 | 4.1333 | 1206.3 | 760.5 | 760.5228 | 0.1 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3667 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2871 | 26.9064 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1202 | 4.1 | 1206.1 | 760.6 | 760.5771 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1301 | 26.9849 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1201.8 | 4.2 | 1206 | 760.5 | 760.5165 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3279 | 25.1 | 27 | 4.2
2008-05-06 | 16:34 | 9.598 | 0.096 | 960 | 1201.7 | 4.2 | 1205.9 | 760.4 | 760.4236 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3673 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 4.2
2008-05-06 | 16:34 | 9.6 | 0.096 | 960 | 1201.7 | 4.1 | 1205.8 | 760.5 | 760.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1445 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1404 | 26.9798 | 4.1
2008-05-06 | 16:34 | 9.596 | 0.096 | 960 | 1201.8 | 4.1 | 1205.9 | 760.5 | 760.5 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6865 | 25.2976 | 26.9012 | 4.1
2008-05-06 | 16:34 | 9.59 | 0.096 | 960 | 1201.5 | 4.1 | 1205.6 | 760.6 | 760.6 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1454 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.592 | 0.096 | 960 | 1201.4 | 4.1 | 1205.5 | 760.6 | 760.5891 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3332 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 997 | 25.1558 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.609 | 0.096 | 960 | 1201.4 | 4.1 | 1205.5 | 760.6 | 760.5757 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2 | 26.9 | 4.1
2008-05-06 | 16:34 | 9.61 | 0.096 | 960 | 1200.9 | 4.1 | 1205 | 760.6 | 760.6226 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2 | 26.9128 | 4.1
2008-05-06 | 16:34 | 9.392 | 0.096 | 960 | 1200.8 | 4.1 | 1204.9 | 760.6 | 760.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1444 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2 | 26.9913 | 4.1
2008-05-06 | 16:34 | 8.921 | 0.0875 | 875.2627 | 1185.7 | 4.1 | 1189.8 | 760.7 | 760.6806 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.14361 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.3396 | 26.9302 | 4.1
2008-05-06 | 16:34 | 8.441 | 0.0862 | 861.7108 | 1121.7 | 4.1 | 1125.8 | 760.5 | 760.4737 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.13832 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035.1795 | 996 | 25.3035 | 26.9 | 4.1
2008-05-06 | 16:34 | 8.35 | 0.0866 | 865.8023 | 1061.6 | 4.1 | 1065.7 | 717.6 | 717.5919 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.13361 | 0.0053 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2267 | 26.9 | 4.1
2008-05-06 | 16:34 | 8.35 | 0.086 | 860 | 1054.1 | 4.1 | 1058.2 | 677.6 | 677.6173 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3675 | 0.133 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2895 | 26.9 | 4.1
2008-05-06 | 16:34 | 8.35 | 0.086 | 860 | 1052.5 | 4 | 1056.5 | 666.8 | 666.8132 | 0.1 | 12.8 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.13283 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1326 | 26.9 | 4
2008-05-06 | 16:34 | 8.356 | 0.086 | 860 | 1050.9 | 4 | 1054.9 | 665 | 665.0132 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.13267 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9 | 4
2008-05-06 | 16:34 | 8.36 | 0.086 | 860 | 1050.3 | 4 | 1054.3 | 664.4 | 664.4419 | 0.1 | 12.8 | 923 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1326 | 0.0052 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1406 | 26.9 | 4
2008-05-06 | 16:34 | 8.36 | 0.086 | 860 | 1050.1 | 4 | 1054.1 | 664.3 | 664.3 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1326 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.238 | 26.9 | 4
2008-05-06 | 16:34 | 8.361 | 0.086 | 860 | 1050.1 | 4 | 1054.1 | 664.3 | 664.3 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.13257 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3951 | 26.9 | 4
2008-05-06 | 16:35 | 8.369 | 0.086 | 860 | 1049.9 | 4 | 1053.9 | 664 | 664.0238 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1325 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1716 | 26.9 | 4
2008-05-06 | 16:35 | 8.37 | 0.086 | 860 | 1049.8 | 4 | 1053.8 | 664.1 | 664.0779 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1325 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9546 | 4
2008-05-06 | 16:35 | 8.37 | 0.086 | 860 | 1049.7 | 4 | 1053.7 | 663.9 | 663.9 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1325 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9671 | 4
2008-05-06 | 16:35 | 8.37 | 0.086 | 860 | 1049.5 | 4 | 1053.5 | 663.9 | 663.9 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.3671 | 0.1325 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.911 | 4
2008-05-06 | 16:35 | 8.37 | 0.0854 | 854.2678 | 1049.3 | 4 | 1053.3 | 663.9 | 663.9 | 0.1 | 12.9 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1325 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.99 | 4
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1049.3 | 4 | 1053.3 | 663.9 | 663.9495 | 0.1 | 12.8 | 923 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.0051 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9315 | 4
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1049.1 | 4 | 1053.1 | 664.1 | 664.1 | 0.1 | 12.8 | 924 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.005066 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.053 | 26.947 | 4
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1048.9 | 4 | 1052.9 | 664 | 664 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.336 | 25.0508 | 26.9746 | 4
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1048.7 | 4 | 1052.7 | 663.9 | 663.9251 | 0.1 | 12.8 | 924 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2 | 26.9035 | 4
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1048.7 | 3.9 | 1052.6 | 664 | 663.9876 | 0.1 | 12.9 | 923 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2 | 26.9823 | 3.9
2008-05-06 | 16:35 | 8.37 | 0.085 | 850 | 1048.7 | 3.9 | 1052.6 | 663.8 | 663.8227 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1324 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.139 | 26.939 | 3.9
2008-05-06 | 16:35 | 8.371 | 0.085 | 850 | 1048.7 | 3.9 | 1052.6 | 663.8 | 663.8 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.13237 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1395 | 26.9 | 3.9
2008-05-06 | 16:35 | 8.379 | 0.085 | 850 | 1048.9 | 3.8 | 1052.7 | 663.9 | 663.8621 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2357 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.085 | 850 | 1048.8 | 3.8 | 1052.6 | 663.8 | 663.8211 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.3313 | 25.3925 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.085 | 850 | 1048.7 | 3.8 | 1052.5 | 663.8 | 663.7805 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.005 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.3252 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.085 | 850 | 1048.8 | 3.8 | 1052.6 | 663.8 | 663.8 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.6754 | 25.3 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.0844 | 844.2821 | 1048.9 | 3.8 | 1052.7 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2682 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.084 | 840 | 1048.9 | 3.8 | 1052.7 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 996.336 | 25.2102 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.084 | 840 | 1048.6 | 3.8 | 1052.4 | 663.8 | 663.7509 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | 0 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2881 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.084 | 840 | 1048.6 | 3.8 | 1052.4 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | -3.3e-05 |  | n/a |  | 32 | 59 | 1035 | 996.668 | 25.233 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.084 | 840 | 1048.8 | 3.8 | 1052.6 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2915 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.38 | 0.084 | 840 | 1048.7 | 3.8 | 1052.5 | 663.8 | 663.7762 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3515 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.387 | 0.0846 | 845.8503 | 1048.8 | 3.8 | 1052.6 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1973 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.409 | 0.0846 | 845.9571 | 1048.6 | 3.8 | 1052.4 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1323 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3313 | 25.1188 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.401 | 0.084 | 840 | 1048.3 | 3.8 | 1052.1 | 663.8 | 663.8 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.13227 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6714 | 25.1597 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.4 | 3.8 | 1052.2 | 663.8 | 663.8232 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1619 | 26.9381 | 3.8
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.3 | 3.7676 | 1052 | 663.8 | 663.8 | 0.1 | 12.9 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.8
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.8 | 663.8251 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.7 | 663.7 | 0.1 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2471 | 26.9264 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.0844 | 844.0396 | 1048.4 | 3.7 | 1052.1 | 663.7 | 663.7232 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1959 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.0848 | 847.6917 | 1048.4 | 3.7 | 1052.1 | 663.7 | 663.7 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.333 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.8 | 663.7884 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1091 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.8 | 663.7763 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0049 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1875 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.4 | 3.7 | 1052.1 | 663.7 | 663.7212 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.4 | 3.7 | 1052.1 | 663.7 | 663.7 | 0.1 | 12.8 | 924 | 943 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1555 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.5 | 3.7 | 1052.2 | 663.7 | 663.7 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1919 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.5 | 3.7 | 1052.2 | 663.8 | 663.8 | 0.1 | 12.8 | 924 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4958 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.5 | 3.7 | 1052.2 | 663.6 | 663.6298 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.261 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.5 | 3.7 | 1052.2 | 663.8 | 663.7774 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3748 | 26.9583 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.4 | 3.7 | 1052.1 | 663.8 | 663.8 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4263 | 26.9631 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.3 | 3.7 | 1052 | 663.6 | 663.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3153 | 26.9153 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.1 | 3.7 | 1051.8 | 663.6 | 663.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1322 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.33 | 25.3939 | 26.9939 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048 | 3.7 | 1051.7 | 663.5 | 663.5122 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2555 | 26.9277 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.7 | 663.6786 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2507 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.2 | 3.7 | 1051.9 | 663.8 | 663.8 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2414 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.4 | 3.7 | 1052.1 | 663.7 | 663.7196 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.1155 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.084 | 840 | 1048.5 | 3.7 | 1052.2 | 663.6 | 663.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.3675 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2724 | 26.9 | 3.7
2008-05-06 | 16:35 | 8.4 | 0.0836 | 836.0616 | 1048.5 | 3.8 | 1052.3 | 663.7 | 663.6747 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.3643 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.4 | 0.083 | 830 | 1048.5 | 3.8 | 1052.3 | 663.6 | 663.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.3144 | 26.9 | 3.8
2008-05-06 | 16:35 | 8.4 | 0.0829 | 829.0328 | 1048.2 | 3.8 | 1052 | 663.7 | 663.7 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3666 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.2433 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.928 | 0.0821 | 820.8361 | 1048.2 | 3.8 | 1052 | 663.5 | 663.5132 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1321 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1036 | 996 | 25.3998 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.429 | 0.0747 | 746.8134 | 1022.5 | 3.8 | 1026.3 | 663.7 | 663.7221 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.13095 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1641 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.22 | 0.0726 | 725.6816 | 963.1 | 3.8 | 966.9 | 661.6 | 661.5896 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.12503 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9569 | 3.8
2008-05-06 | 16:35 | 7.22 | 0.073 | 730 | 913 | 3.8 | 916.8 | 611.4 | 611.4214 | 0.1 | 12.9 | 923 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1206 | 0.0048 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1706 | 26.9647 | 3.8
2008-05-06 | 16:35 | 7.22 | 0.073 | 730 | 905.3 | 3.8 | 909.1 | 577.3 | 577.3064 | 0.1 | 12.9 | 923 | 942 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3669 | 0.11987 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.213 | 0.073 | 730 | 902.7 | 3.8 | 906.5 | 570.6 | 570.5567 | 0.1 | 12.8 | 923 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.11967 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.204 | 0.073 | 730 | 901.9 | 3.8 | 905.7 | 569.2 | 569.2151 | 0.1 | 12.8 | 923 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1196 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.158 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.2 | 0.073 | 730 | 901.7 | 3.8 | 905.5 | 568.9 | 568.9206 | 0.1 | 12.8 | 923 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.11957 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1988 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.2 | 0.073 | 730 | 901.6 | 3.8 | 905.4 | 568.7 | 568.7238 | 0.1 | 12.8 | 924 | 944 | 930 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3662 | 0.1195 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3557 | 26.9 | 3.8
2008-05-06 | 16:35 | 7.199 | 0.0726 | 725.8824 | 901.5 | 3.8 | 905.3 | 568.5 | 568.5 | 0.1 | 12.9 | 923 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1195 | 0.0047 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4746 | 26.9 | 3.8
2008-05-06 | 16:36 | 7.182 | 0.072 | 720 | 901.3 | 3.7661 | 905 | 568.4 | 568.4227 | 0.1 | 12.8 | 924 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1195 | 0.0046 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1606 | 26.9 | 3.8
2008-05-06 | 16:36 | 7.187 | 0.072 | 720 | 901.2 | 3.7 | 904.9 | 568.4 | 568.4119 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1194 | 0.0046 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1634 | 26.9634 | 3.7
2008-05-06 | 16:36 | 7.201 | 0.072 | 720 | 901.1 | 3.7 | 904.8 | 568.3 | 568.2723 | 0.1 | 12.8 | 923 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1194 | 0.0046 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2416 | 26.9584 | 3.7
2008-05-06 | 16:36 | 7.21 | 0.072 | 720 | 901.1 | 3.7 | 904.8 | 568.1 | 568.1234 | 0.1 | 12.9 | 923 | 944 | 930 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1194 | 0.0046 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.201 | 0.072 | 720 | 901.3 | 3.7 | 905 | 568.3 | 568.3 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1194 | 0.0046 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 901.2 | 3.7 | 904.9 | 568.2 | 568.2189 | 0.1 | 12.9 | 923 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1194 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 901.1 | 3.7 | 904.8 | 568.2 | 568.2 | 0.1 | 12.8 | 923 | 942 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2443 | 26.9557 | 3.7
2008-05-06 | 16:36 | 7.148 | 0.0721 | 720.5579 | 901 | 3.7 | 904.7 | 568.2 | 568.1843 | 0.1 | 12.8 | 924 | 944 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2342 | 27 | 3.7
2008-05-06 | 16:36 | 7.146 | 0.0729 | 728.8843 | 901.1 | 3.7 | 904.8 | 568.3 | 568.2731 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2874 | 26.9874 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.0723 | 722.7531 | 901.1 | 3.7 | 904.8 | 568.2 | 568.1782 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2089 | 26.9089 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 900.9 | 3.7 | 904.6 | 568.1 | 568.0769 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 900.9 | 3.7 | 904.6 | 568.1 | 568.1222 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1193 | 0.0045 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2481 | 26.9481 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 901 | 3.7 | 904.7 | 568.2 | 568.1774 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9734 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 901 | 3.7 | 904.7 | 568 | 568 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 900.9 | 3.7 | 904.6 | 568.1 | 568.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.15 | 0.072 | 720 | 900.8 | 3.7 | 904.5 | 568.3 | 568.2814 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.155 | 0.072 | 720 | 900.7 | 3.7 | 904.4 | 568.2 | 568.2231 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1787 | 26.9404 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.7 | 3.7 | 904.4 | 568.2 | 568.2 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0754 | 26.9811 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.7 | 3.7 | 904.4 | 568.3 | 568.2796 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1192 | 0.0044 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3896 | 26.9026 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.6 | 3.7 | 904.3 | 568.4 | 568.38 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3242 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.5 | 3.7 | 904.2 | 568.3 | 568.2813 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9543 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.5 | 3.7 | 904.2 | 568.4 | 568.3546 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3656 | 26.9672 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.3 | 3.7 | 904 | 568.2 | 568.2238 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4775 | 26.9112 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.3 | 3.7 | 904 | 568.2 | 568.2156 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3204 | 26.9898 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.3 | 3.7 | 904 | 568.4 | 568.4 | 0.1 | 12.8 | 924 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1635 | 26.9318 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.3 | 3.7 | 904 | 568.2 | 568.2103 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1191 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2401 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.17 | 0.072 | 720 | 900.3 | 3.7 | 904 | 568.3 | 568.3 | 0.1 | 12.8 | 924 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3749 | 26.9 | 3.7
2008-05-06 | 16:36 | 7.166 | 0.072 | 720 | 900.3 | 3.6 | 903.9 | 568.4 | 568.4 | 0.1 | 12.9 | 924 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0043 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2968 | 26.9 | 3.6
2008-05-06 | 16:36 | 7.16 | 0.072 | 720 | 900.2 | 3.5 | 903.7 | 568.2 | 568.154 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.004267 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.218 | 26.9 | 3.5
2008-05-06 | 16:36 | 7.16 | 0.072 | 720 | 900.1 | 3.5 | 903.6 | 568.2 | 568.2 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3211 | 26.9 | 3.5
2008-05-06 | 16:36 | 7.16 | 0.072 | 720 | 900.2 | 3.5 | 903.7 | 568.3 | 568.3 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4 | 26.9 | 3.5
2008-05-06 | 16:36 | 7.16 | 0.072 | 720 | 900.3 | 3.4 | 903.7 | 568.4 | 568.3845 | 0.1 | 12.9 | 924 | 942 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3476 | 26.9175 | 3.4
2008-05-06 | 16:36 | 7.16 | 0.072 | 720 | 900.3 | 3.4 | 903.7 | 568.4 | 568.4 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1127 | 26.9958 | 3.4
2008-05-06 | 16:36 | 7.168 | 0.072 | 720 | 900.4 | 3.4 | 903.8 | 568.3 | 568.2892 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2489 | 27 | 3.4
2008-05-06 | 16:36 | 7.178 | 0.072 | 720 | 900.4 | 3.3 | 903.7 | 568.2 | 568.2 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2471 | 27 | 3.3
2008-05-06 | 16:36 | 7.171 | 0.072 | 720 | 900.5 | 3.3 | 903.8 | 568.1 | 568.0891 | 0.1 | 12.8 | 924 | 943 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.119 | 0.0042 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 27 | 3.3
2008-05-06 | 16:36 | 7.179 | 0.072 | 720 | 900.4 | 3.3 | 903.7 | 568.2 | 568.1877 | 0.1 | 12.9 | 924 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3331 | 0.119 | 0.004167 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 27 | 3.3
2008-05-06 | 16:36 | 7.173 | 0.072 | 720 | 900.3 | 3.3 | 903.6 | 568.3 | 568.2529 | 0.1 | 12.8 | 924 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3279 | 25.2 | 27 | 3.3
2008-05-06 | 16:36 | 7.176 | 0.072 | 720 | 900.2 | 3.3 | 903.5 | 568.3 | 568.3 | 0.1 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2665 | 26.9335 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.3 | 3.3 | 903.6 | 568.1 | 568.1196 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2547 | 26.9453 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.3 | 3.3 | 903.6 | 568.2 | 568.18 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3135 | 25.2474 | 26.9763 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.3 | 3.3 | 903.6 | 568.1 | 568.1232 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3979 | 26.9 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.4 | 3.3 | 903.7 | 567.9 | 567.9298 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 6.7e-05 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3197 | 26.9 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.7 | 3.3 | 904 | 568 | 567.975 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.3 | 26.9589 | 3.3
2008-05-06 | 16:36 | 7.18 | 0.072 | 720 | 900.6 | 3.3 | 903.9 | 567.9 | 567.9061 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3287 | 25.3376 | 27 | 3.3
2008-05-06 | 16:36 | 7.184 | 0.072 | 720 | 900.7 | 3.3 | 904 | 568 | 568 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.4 | 26.9839 | 3.3
2008-05-06 | 16:36 | 7.19 | 0.072 | 720 | 900.8 | 3.3 | 904.1 | 567.9 | 567.9228 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.4 | 26.9055 | 3.3
2008-05-06 | 16:36 | 7.19 | 0.072 | 720 | 900.5 | 3.3 | 903.8 | 567.9 | 567.9213 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4 | 26.9 | 3.3
2008-05-06 | 16:36 | 7.19 | 0.072 | 720 | 900.5 | 3.3 | 903.8 | 568 | 567.9646 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1189 | 0.0041 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3486 | 26.9 | 3.3
2008-05-06 | 16:36 | 7.19 | 0.072 | 720 | 900.4 | 3.3 | 903.7 | 568 | 567.987 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.11887 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9299 | 3.3
2008-05-06 | 16:36 | 7.19 | 0.072 | 720 | 900.4 | 3.3 | 903.7 | 568 | 568 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1188 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9916 | 3.3
2008-05-06 | 16:36 | 6.743 | 0.0638 | 638.363 | 900.3 | 3.3 | 903.6 | 568.1 | 568.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1188 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9132 | 3.3
2008-05-06 | 16:36 | 6.033 | 0.0627 | 627.3134 | 851.1 | 3.2667 | 854.4 | 567.9 | 567.9 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.11437 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9 | 3.3
2008-05-06 | 16:36 | 5.931 | 0.0632 | 632.1095 | 774 | 3.2 | 777.2 | 549.9 | 549.9007 | 0.1 | 12.9 | 922 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.107 | 0.004 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2566 | 26.9434 | 3.2
2008-05-06 | 16:36 | 5.948 | 0.062 | 620 | 759.5 | 3.2 | 762.7 | 496.7 | 496.7266 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.1056 | 0.003967 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2223 | 27 | 3.2
2008-05-06 | 16:36 | 5.95 | 0.062 | 620 | 754.9 | 3.2 | 758.1 | 478 | 477.9978 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.10507 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6646 | 25.3 | 26.9993 | 3.2
2008-05-06 | 16:36 | 5.95 | 0.062 | 620 | 754.6 | 3.2 | 757.8 | 473.8 | 473.7747 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.105 | 0.0039 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9208 | 3.2
2008-05-06 | 16:36 | 5.95 | 0.062 | 620 | 754.4 | 3.2 | 757.6 | 472.9 | 472.9036 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.10497 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3313 | 25.3 | 26.9 | 3.2
2008-05-06 | 16:36 | 5.95 | 0.0614 | 614.3038 | 753.9 | 3.1 | 757 | 472.6 | 472.5898 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.10487 | 0.0038 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6721 | 25.336 | 26.9 | 3.1
2008-05-06 | 16:36 | 5.95 | 0.061 | 610 | 753 | 3.1 | 756.1 | 472.6 | 472.5774 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1048 | 0.003767 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.3708 | 26.9146 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.0608 | 607.9091 | 752.5 | 3.1 | 755.6 | 472.4 | 472.3716 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1047 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.214 | 26.993 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 751.6 | 3.1 | 754.7 | 472.3 | 472.3 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1046 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.343 | 26.9285 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 751.4 | 3.1 | 754.5 | 472.2 | 472.2233 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1045 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.35 | 26.9 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 751.3 | 3.1 | 754.4 | 472.4 | 472.3779 | 0.1 | 12.8 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1045 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2431 | 26.9285 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 751.1 | 3.1 | 754.2 | 472.2 | 472.2304 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.1045 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1276 | 26.9931 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 751.1 | 3.1 | 754.2 | 472.3 | 472.2888 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1045 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.4416 | 26.9146 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.9 | 3.1 | 754 | 472.2 | 472.2113 | 0.1 | 12.9 | 924 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1045 | 0.0037 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3354 | 25.3085 | 26.9638 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.8 | 3.0667 | 753.9 | 472.1 | 472.1237 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.10447 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.2423 | 27 | 3.1
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.6 | 3 | 753.6 | 472.1 | 472.15 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.1044 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.3 | 27 | 3
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.7 | 3 | 753.7 | 472 | 472.0145 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3328 | 0.1044 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.3 | 27 | 3
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.5 | 3 | 753.5 | 472.1 | 472.1204 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1044 | 0.0036 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.326 | 25.3777 | 26.9223 | 3
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.4 | 3 | 753.4 | 472.2 | 472.2 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.10437 | 0.003567 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6747 | 25.2878 | 26.9561 | 3
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.3 | 2.9 | 753.2 | 471.9 | 471.8546 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1043 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6865 | 25.2 | 27 | 2.9
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.4 | 2.9 | 753.3 | 472.2 | 472.1763 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1043 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.213 | 26.987 | 2.9
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.3 | 2.9 | 753.2 | 472.2 | 472.2464 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1043 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2916 | 26.9084 | 2.9
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.1 | 2.9 | 753 | 472.2 | 472.2 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1043 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1599 | 26.97 | 2.9
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.2 | 2.9 | 753.1 | 472.1 | 472.0766 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.9
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.1 | 2.8 | 752.9 | 471.9 | 471.9179 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9732 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.2 | 2.8 | 753 | 472.1 | 472.1 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1108 | 26.9 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.1 | 2.8 | 752.9 | 472 | 472.0496 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2678 | 26.9 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.2 | 2.8 | 753 | 472 | 471.9768 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1753 | 26.9 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.2 | 2.8 | 753 | 471.9 | 471.8779 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0035 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9408 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.2 | 2.7662 | 752.9 | 472.1 | 472.1079 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1577 | 26.9808 | 2.8
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 750.1 | 2.7 | 752.8 | 472.2 | 472.1659 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.1042 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3932 | 26.9023 | 2.7
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.9 | 2.7 | 752.6 | 471.9 | 471.9309 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.10417 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.334 | 25.3238 | 26.9762 | 2.7
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.9 | 2.7 | 752.6 | 471.7 | 471.741 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1359 | 27 | 2.7
2008-05-06 | 16:37 | 5.956 | 0.06 | 600 | 750 | 2.7 | 752.7 | 471.9 | 471.9491 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.066 | 27 | 2.7
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 749.9 | 2.6666 | 752.6 | 472 | 471.9774 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0034 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 27 | 2.7
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 750 | 2.6 | 752.6 | 471.9 | 471.8915 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.003367 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 27 | 2.6
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 750 | 2.6 | 752.6 | 472 | 472 | 0.1 | 12.8 | 924 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2685 | 26.9315 | 2.6
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 750 | 2.6 | 752.6 | 471.9 | 471.9214 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.206 | 26.947 | 2.6
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 750 | 2.6 | 752.6 | 471.9 | 471.9191 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.1041 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1508 | 26.9746 | 2.6
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 749.8 | 2.6 | 752.4 | 472.1 | 472.1 | 0.1 | 12.9 | 924 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9038 | 2.6
2008-05-06 | 16:37 | 5.96 | 0.06 | 600 | 749.7 | 2.6 | 752.3 | 471.9 | 471.9 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3 | 26.9823 | 2.6
2008-05-06 | 16:37 | 5.958 | 0.06 | 600 | 749.5 | 2.6 | 752.1 | 472 | 472 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1782 | 27 | 2.6
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.5 | 2.5665 | 752.1 | 472.2 | 472.1555 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1787 | 27 | 2.6
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.4 | 2.5 | 751.9 | 472 | 472 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2646 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.4 | 2.5 | 751.9 | 472.1 | 472.1 | 0.1 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1089 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.3 | 2.5 | 751.8 | 472 | 472 | 0.1 | 12.8 | 924 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3982 | 26.9255 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.3 | 2.5 | 751.8 | 472.1 | 472.0911 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3331 | 0.104 | 0.0033 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.3936 | 26.9 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.4 | 2.5 | 751.9 | 472.1 | 472.0742 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.10397 | 0.003267 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2368 | 26.9316 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.4 | 2.5 | 751.9 | 472 | 472.0197 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.12 | 26.99 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600 | 749.4 | 2.5 | 751.9 | 471.9 | 471.9 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2765 | 26.9117 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.06 | 600.3553 | 749.4 | 2.5 | 751.9 | 472 | 471.9791 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1662 | 26.9669 | 2.5
2008-05-06 | 16:37 | 5.956 | 0.0609 | 609.0208 | 749.6 | 2.5 | 752.1 | 472 | 471.9786 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3347 | 25.1455 | 27 | 2.5
2008-05-06 | 16:37 | 5.956 | 0.061 | 610 | 749.5 | 2.5 | 752 | 471.9 | 471.945 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.176 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.061 | 610 | 749.5 | 2.5 | 752 | 471.7 | 471.741 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.061 | 610 | 749.4 | 2.5 | 751.9 | 471.8 | 471.7527 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.061 | 610 | 749.4 | 2.5333 | 751.9 | 471.9 | 471.8843 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.5
2008-05-06 | 16:37 | 5.95 | 0.0609 | 609.4309 | 749.3 | 2.5669 | 751.9 | 471.9 | 471.9212 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1756 | 27 | 2.6
2008-05-06 | 16:37 | 5.532 | 0.0601 | 601.3008 | 749.2 | 2.5 | 751.7 | 472 | 472 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.1039 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6707 | 25.2513 | 27 | 2.5
2008-05-06 | 16:37 | 5.055 | 0.0506 | 506.4267 | 739.7 | 2.5 | 742.2 | 471.8 | 471.8452 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.10295 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.016 | 27 | 2.5
2008-05-06 | 16:37 | 4.81 | 0.0481 | 481.2825 | 665.3 | 2.5 | 667.8 | 470.1 | 470.056 | 0.1 | 12.9 | 922 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.095258 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3253 | 25.1456 | 27 | 2.5
2008-05-06 | 16:37 | 4.81 | 0.049 | 490 | 608.6 | 2.5 | 611.1 | 419.3 | 419.2814 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.089399 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1486 | 27 | 2.5
2008-05-06 | 16:37 | 4.81 | 0.049 | 490 | 601.6 | 2.5 | 604.1 | 386.7 | 386.6972 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.088533 | 0.0032 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 26.9699 | 2.5
2008-05-06 | 16:37 | 4.81 | 0.049 | 490 | 599.1 | 2.5 | 601.6 | 379.2 | 379.2029 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0882 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9086 | 2.5
2008-05-06 | 16:37 | 4.81 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 377.9 | 377.9098 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0881 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.987 | 2.4
2008-05-06 | 16:37 | 4.806 | 0.049 | 490 | 598.6 | 2.467 | 601 | 377.4 | 377.3769 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0881 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.5
2008-05-06 | 16:37 | 4.8 | 0.0484 | 484.3106 | 598.6 | 2.4 | 601 | 377.2 | 377.2 | 0.1 | 12.8 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0881 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:37 | 4.8 | 0.0484 | 483.9621 | 598.6 | 2.4 | 601 | 377 | 377.017 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0881 | 0.0031 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1224 | 27 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 377 | 377 | 0.1 | 12.8 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.088 | 0.003067 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 26.9991 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 377 | 377.0116 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.088 | 0.003 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2 | 26.9207 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 377 | 377 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1423 | 26.9577 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 376.9 | 376.8763 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9638 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 377 | 377.0496 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9147 | 2.4
2008-05-06 | 16:38 | 4.8 | 0.0486 | 486.1241 | 598.5 | 2.4 | 600.9 | 376.8 | 376.7778 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 26.9932 | 2.4
2008-05-06 | 16:38 | 4.791 | 0.048 | 480 | 598.6 | 2.4 | 601 | 376.8 | 376.8288 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0029 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1717 | 26.9283 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.048 | 480 | 598.5 | 2.4 | 600.9 | 376.9 | 376.9 | 0.1 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0879 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2501 | 26.9 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.048 | 480 | 598.5 | 2.4 | 600.9 | 376.7 | 376.7 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.2429 | 26.9285 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.048 | 480 | 598.4 | 2.4 | 600.8 | 376.8 | 376.7577 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.048 | 480 | 598.4 | 2.5 | 600.9 | 376.8 | 376.8 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.5
2008-05-06 | 16:38 | 4.79 | 0.0484 | 483.8684 | 598.4 | 2.5 | 600.9 | 376.6 | 376.6229 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1639 | 26.9361 | 2.5
2008-05-06 | 16:38 | 4.79 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 376.8 | 376.7929 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1576 | 26.9424 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 376.7 | 376.6614 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.049 | 490 | 598.5 | 2.4 | 600.9 | 376.6 | 376.6 | 0.1 | 12.9 | 923 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.79 | 0.049 | 490 | 598.4 | 2.3 | 600.7 | 376.7 | 376.6969 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0222 | 26.9222 | 2.3
2008-05-06 | 16:38 | 4.79 | 0.0484 | 484.4213 | 598.5 | 2.3 | 600.8 | 376.5 | 376.52 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 26.9562 | 2.3
2008-05-06 | 16:38 | 4.79 | 0.048 | 480 | 598.4 | 2.333 | 600.8 | 376.7 | 376.7 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 2.3
2008-05-06 | 16:38 | 4.781 | 0.048 | 480 | 598.5 | 2.4 | 600.9 | 376.5 | 376.5214 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0878 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0132 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.048 | 480.2931 | 598.4 | 2.4 | 600.8 | 376.5 | 376.5 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3665 | 0.087766 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0916 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.0489 | 488.9138 | 598.2 | 2.4 | 600.6 | 376.6 | 376.6 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.3333 | 0.0877 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.049 | 490 | 598.3 | 2.4 | 600.7 | 376.4 | 376.435 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0516 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.0484 | 484.3467 | 598.2 | 2.4 | 600.6 | 376.5 | 376.4742 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0268 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.048 | 480 | 598.2 | 2.4 | 600.6 | 376.5 | 376.5 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1054 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.048 | 480 | 598 | 2.4 | 600.4 | 376.6 | 376.5796 | 0.1 | 12.8 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0028 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1838 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.048 | 480 | 598 | 2.4 | 600.4 | 376.4 | 376.4495 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.002766 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1377 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.048 | 480 | 598 | 2.4 | 600.4 | 376.6 | 376.5713 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.0487 | 487.2035 | 598 | 2.4 | 600.4 | 376.4 | 376.3772 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.5 | 376.4887 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3668 | 0.0877 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.78 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.5 | 376.5 | 0.1 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0877 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.771 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.4 | 376.3806 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3667 | 0.0876 | 0.0027 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0454 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598 | 2.4 | 600.4 | 376.6 | 376.5802 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0876 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0329 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598 | 2.4 | 600.4 | 376.5 | 376.5177 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0876 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5 | 0.1 | 12.9 | 924 | 943 | 934 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0876 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5141 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0876 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598 | 2.4 | 600.4 | 376.4 | 376.4 | 0.1 | 12.8 | 924 | 941 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.087569 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.0488 | 488.0149 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5224 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0747 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598.2 | 2.4 | 600.6 | 376.6 | 376.578 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3327 | 25 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598.1 | 2.4 | 600.5 | 376.5 | 376.5 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0391 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598 | 2.4 | 600.4 | 376.4 | 376.4 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0026 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480 | 598 | 2.4 | 600.4 | 376.4 | 376.4 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.002567 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.048 | 480.2333 | 597.9 | 2.4 | 600.3 | 376.7 | 376.6512 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0875 | 0.0025 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.0257 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.0489 | 488.5667 | 598 | 2.4 | 600.4 | 376.5 | 376.5 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9471 | 26.9471 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.5 | 376.5 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9627 | 26.9313 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 597.8 | 2.4 | 600.2 | 376.4 | 376.4204 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.5 | 376.4898 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 598 | 2.4 | 600.4 | 376.4 | 376.4 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 24.9667 | 27 | 2.4
2008-05-06 | 16:38 | 4.77 | 0.049 | 490 | 597.9 | 2.4 | 600.3 | 376.4 | 376.4 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6687 | 24.9452 | 27 | 2.4
2008-05-06 | 16:38 | 4.359 | 0.049 | 490 | 598 | 2.4 | 600.4 | 376.3 | 376.3226 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0874 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0236 | 27 | 2.4
2008-05-06 | 16:38 | 3.88 | 0.0411 | 410.5197 | 588.9 | 2.4 | 591.3 | 376.4 | 376.4 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.086371 | 0.0024 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.098 | 27 | 2.4
2008-05-06 | 16:38 | 3.64 | 0.0385 | 385.3735 | 517.9 | 2.4 | 520.3 | 375.7 | 375.6912 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.078188 | 0.002367 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.02 | 27 | 2.4
2008-05-06 | 16:38 | 3.639 | 0.039 | 390 | 458.3 | 2.4 | 460.7 | 323.2 | 323.1957 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.071168 | 0.0023 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1174 | 27 | 2.4
2008-05-06 | 16:38 | 3.597 | 0.039 | 390 | 452.4 | 2.4 | 454.8 | 292.3 | 292.3104 | 0.1 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.070267 | 0.0023 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1252 | 27 | 2.4
2008-05-06 | 16:38 | 3.611 | 0.039 | 390 | 451.1 | 2.4 | 453.5 | 284.7 | 284.6755 | 0.1 | 12.9 | 923 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0701 | 0.0022 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0317 | 27 | 2.4
2008-05-06 | 16:38 | 3.62 | 0.039 | 390 | 450.8 | 2.4 | 453.2 | 283.1 | 283.1237 | 0.1 | 12.8 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.07 | 0.0022 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1887 | 27 | 2.4
2008-05-06 | 16:38 | 3.62 | 0.0383 | 382.8596 | 450.5 | 2.4 | 452.9 | 282.5 | 282.5 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.07 | 0.0022 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1274 | 27 | 2.4
2008-05-06 | 16:38 | 3.622 | 0.038 | 380 | 450.5 | 2.4 | 452.9 | 282.6 | 282.5695 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.07 | 0.0022 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.0511 | 2.4
2008-05-06 | 16:38 | 3.63 | 0.038 | 380 | 450.3 | 2.4 | 452.7 | 282.4 | 282.4 | 0.1 | 12.8 | 924 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0699 | 0.002167 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0704 | 27.1 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.038 | 380 | 450.1 | 2.4 | 452.5 | 282.5 | 282.4779 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.069867 | 0.002067 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.33 | 25.0081 | 27.0919 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.038 | 380 | 450 | 2.4 | 452.4 | 282.2 | 282.2464 | 0.1 | 12.8 | 924 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.069766 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0866 | 27.0134 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.038 | 380 | 449.8 | 2.4 | 452.2 | 282.2 | 282.2201 | 0.1 | 12.9 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0697 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.038 | 380 | 449.8 | 2.4 | 452.2 | 282.4 | 282.3765 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0697 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0138 | 27 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.0375 | 374.5558 | 449.7 | 2.4 | 452.1 | 282.1 | 282.1231 | 0.1 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0697 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.037 | 370 | 449.4 | 2.4 | 451.8 | 282.1 | 282.1 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0697 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3266 | 24.9006 | 27 | 2.4
2008-05-06 | 16:39 | 3.621 | 0.037 | 370 | 449.3 | 2.4 | 451.7 | 282.1 | 282.1 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3672 | 0.0696 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0577 | 27 | 2.4
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 449.3 | 2.4 | 451.7 | 282.1 | 282.1161 | 0.1 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3663 | 0.0696 | 0.002 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:39 | 3.625 | 0.037 | 370 | 449.2 | 2.4 | 451.6 | 282.1 | 282.1419 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0696 | 0.001966 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27 | 2.4
2008-05-06 | 16:39 | 3.63 | 0.037 | 370 | 449.2 | 2.4 | 451.6 | 282.1 | 282.1201 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0696 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0142 | 2.4
2008-05-06 | 16:39 | 3.628 | 0.037 | 370 | 449.2 | 2.4 | 451.6 | 282.2 | 282.2 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0696 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0926 | 2.4
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 449.1 | 2.4 | 451.5 | 282.1 | 282.1226 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0696 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0289 | 2.4
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 449.1 | 2.4 | 451.5 | 281.9 | 281.9259 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0696 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1492 | 27 | 2.4
2008-05-06 | 16:39 | 3.627 | 0.037 | 370 | 449.1 | 2.4 | 451.5 | 282 | 282 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.069567 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1725 | 27 | 2.4
2008-05-06 | 16:39 | 3.625 | 0.037 | 370 | 448.8 | 2.4 | 451.2 | 282.2 | 282.1519 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0695 | 0.0019 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0871 | 27 | 2.4
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 448.6 | 2.4 | 451 | 281.9 | 281.9206 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.069467 | 0.001867 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9303 | 27 | 2.4
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 448.4 | 2.5 | 450.9 | 282 | 282 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6747 | 25.0267 | 27 | 2.5
2008-05-06 | 16:39 | 3.62 | 0.037 | 370 | 448.5 | 2.5 | 451 | 281.9 | 281.9209 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3431 | 25.0583 | 27 | 2.5
2008-05-06 | 16:39 | 3.62 | 0.0368 | 368.036 | 448.5 | 2.5 | 451 | 281.9 | 281.9 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 2.5
2008-05-06 | 16:39 | 3.613 | 0.036 | 360 | 448.4 | 2.5 | 450.9 | 281.8 | 281.8224 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27 | 2.5
2008-05-06 | 16:39 | 3.61 | 0.036 | 360 | 448.4 | 2.5 | 450.9 | 282 | 281.9762 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0771 | 27.0771 | 2.5
2008-05-06 | 16:39 | 3.61 | 0.036 | 360 | 448.4 | 2.5 | 450.9 | 282 | 281.9725 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0444 | 27.0444 | 2.5
2008-05-06 | 16:39 | 3.61 | 0.036 | 360 | 448.5 | 2.5 | 451 | 281.9 | 281.881 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0338 | 27.0338 | 2.5
2008-05-06 | 16:39 | 3.61 | 0.036 | 360 | 448.6 | 2.5 | 451.1 | 282.1 | 282.1 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.5
2008-05-06 | 16:39 | 3.61 | 0.036 | 360 | 448.5 | 2.5 | 451 | 281.9 | 281.9 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.5
2008-05-06 | 16:39 | 3.617 | 0.036 | 360 | 448.6 | 2.5 | 451.1 | 281.9 | 281.9213 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0307 | 27.0307 | 2.5
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.6 | 2.5327 | 451.1 | 281.8 | 281.8225 | 0.1 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0694 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0478 | 27.0478 | 2.5
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.6 | 2.6 | 451.2 | 281.8 | 281.8 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0018 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0738 | 27.1 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.6 | 2.6 | 451.2 | 282 | 282 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3379 | 25 | 27.0954 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.5 | 2.6 | 451.1 | 282 | 282 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.0169 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.5 | 2.6 | 451.1 | 281.9 | 281.8533 | 0.1 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.668 | 25 | 27.0616 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.6 | 2.6 | 451.2 | 281.8 | 281.8 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.33 | 25 | 27.1 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.6 | 2.6 | 451.2 | 281.8 | 281.8 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3354 | 25.0184 | 27.0816 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.5 | 2.6 | 451.1 | 281.9 | 281.8755 | 0.1 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0969 | 27.0031 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.5 | 2.6 | 451.1 | 281.8 | 281.8458 | 0.1 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0693 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0246 | 27 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.5 | 2.6 | 451.1 | 281.8 | 281.8209 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.069267 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.0538 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.3 | 2.6 | 450.9 | 281.8 | 281.8 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0323 | 27.0677 | 2.6
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.2 | 2.7 | 450.9 | 281.7 | 281.7213 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0893 | 27 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.2 | 2.7 | 450.9 | 281.8 | 281.7829 | 0.1 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3662 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0108 | 27 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.1 | 2.7 | 450.8 | 282 | 281.9774 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.334 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9324 | 27.0676 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.1 | 2.7 | 450.8 | 281.7 | 281.7138 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0381 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448 | 2.7 | 450.7 | 281.8 | 281.8 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.668 | 25.1511 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448.1 | 2.7 | 450.8 | 281.8 | 281.8 | 0.1 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.334 | 25 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 448 | 2.7 | 450.7 | 281.8 | 281.7911 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.3236 | 25 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 447.8 | 2.7 | 450.5 | 281.7 | 281.7 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.3665 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1195 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 447.8 | 2.7 | 450.5 | 281.9 | 281.9 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3329 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1618 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 447.7 | 2.7 | 450.4 | 281.8 | 281.8383 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 997 | 25.1 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 447.8 | 2.7 | 450.5 | 281.6 | 281.6206 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0692 | 0.0017 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.036 | 360 | 447.9 | 2.7 | 450.6 | 281.8 | 281.8 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3662 | 0.0692 | 0.001666 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.7
2008-05-06 | 16:39 | 3.62 | 0.0359 | 358.6261 | 447.8 | 2.7 | 450.5 | 281.8 | 281.8 | 0.1 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.069166 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0479 | 2.7
2008-05-06 | 16:39 | 3.038 | 0.0234 | 233.7302 | 443.6 | 2.7 | 446.3 | 281.7 | 281.7466 | 0.1 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.068593 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.0305 | 2.7
2008-05-06 | 16:39 | 2.492 | 0.023 | 230.2331 | 405.7 | 2.7 | 408.4 | 281 | 281.0144 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.063805 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0911 | 27.1 | 2.7
2008-05-06 | 16:39 | 2.458 | 0.024 | 240 | 338.4 | 2.7 | 341.1 | 251.3 | 251.3158 | 0.1 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.333 | 0.054886 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0126 | 27.1 | 2.7
2008-05-06 | 16:39 | 2.45 | 0.024 | 240 | 300.7 | 2.7 | 303.4 | 201.5 | 201.5042 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0496 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0658 | 27.1 | 2.7
2008-05-06 | 16:39 | 2.45 | 0.024 | 240 | 300.6 | 2.6667 | 303.3 | 190.4 | 190.3688 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0496 | 0.0016 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0558 | 27.1 | 2.7
2008-05-06 | 16:39 | 2.45 | 0.024 | 240 | 300.6 | 2.6 | 303.2 | 187.9 | 187.8719 | 0.1 | 12.8 | 924 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0496 | 0.0016 | 0 | -0.000133 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0227 | 27.1 | 2.6
2008-05-06 | 16:39 | 2.439 | 0.024 | 240 | 300.5 | 2.6 | 303.1 | 187.2 | 187.2495 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0496 | 0.0016 | 0 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:39 | 2.43 | 0.024 | 240 | 300.5 | 2.6 | 303.1 | 187 | 187.0196 | 0.1 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0496 | 0.0016 | 0 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:39 | 2.432 | 0.024 | 240 | 300.5 | 2.6 | 303.1 | 186.8 | 186.8 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0495 | 0.001568 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.042 | 27.1 | 2.6
2008-05-06 | 16:39 | 2.44 | 0.024 | 240 | 300.4 | 2.6 | 303 | 186.6 | 186.62 | 0.1 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0495 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0365 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 300.1 | 2.6 | 302.7 | 186.7 | 186.7164 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.049466 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0851 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.42 | 0.024 | 240 | 300 | 2.6 | 302.6 | 186.5 | 186.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0494 | 0.0015 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0067 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.426 | 0.024 | 240 | 299.9 | 2.5668 | 302.5 | 186.6 | 186.6 | 0.1 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0494 | 0.0015 | -3.3e-05 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0718 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.436 | 0.024 | 240 | 299.9 | 2.5 | 302.4 | 186.5 | 186.5234 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0494 | 0.0015 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.8 | 2.5 | 302.3 | 186.5 | 186.5246 | 0.1 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0494 | 0.001466 | -6.6e-05 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.6646 | 25.0714 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.8 | 2.5 | 302.3 | 186.5 | 186.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0493 | 0.0014 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3333 | 25.0071 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.7 | 2.5 | 302.2 | 186.5 | 186.5 | 0.1 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0492 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0855 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.7 | 2.5 | 302.2 | 186.6 | 186.5809 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0492 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0361 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.7 | 2.5 | 302.2 | 186.5 | 186.5 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0492 | 0.0013 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.7 | 2.5 | 302.2 | 186.6 | 186.5528 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.049167 | 0.001267 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0209 | 27.0791 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.5 | 302.1 | 186.4 | 186.4 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0993 | 27.0007 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.5 | 2.5 | 302 | 186.5 | 186.4847 | 0 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3336 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0223 | 27.0777 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.5 | 302.1 | 186.4 | 186.3773 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.5 | 186.4871 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.5 | 186.4811 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.4 | 186.4215 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.5 | 186.471 | 0 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.4 | 186.3787 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.6 | 2.6 | 302.2 | 186.3 | 186.3253 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0491 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0267 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.5 | 2.6 | 302.1 | 186.4 | 186.4 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.049067 | 0.0012 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.4 | 2.6 | 302 | 186.4 | 186.4 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.4 | 2.6336 | 302 | 186.4 | 186.4 | 0 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9758 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.7 | 301.9 | 186.4 | 186.4 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1 | 2.7
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.7 | 301.9 | 186.5 | 186.4767 | 0 | 12.8 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3332 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.938 | 27.1 | 2.7
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.7 | 301.9 | 186.5 | 186.506 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0948 | 27.1 | 2.7
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.6 | 301.8 | 186.6 | 186.5774 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.6 | 301.8 | 186.8 | 186.7772 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.6667 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.6 | 301.8 | 186.9 | 186.9 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.6 | 301.8 | 187 | 186.9787 | 0 | 12.8 | 924 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.049 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.2 | 2.6 | 301.8 | 187.1 | 187.1 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.438 | 0.024 | 240 | 299.2 | 2.5663 | 301.8 | 187.1 | 187.1 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.1 | 27.1 | 2.6
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299.2 | 2.5 | 301.7 | 187.1 | 187.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3386 | 25.1 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299.1 | 2.5 | 301.6 | 187.2 | 187.1762 | 0 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.666 | 25.1 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299.1 | 2.5 | 301.6 | 187.1 | 187.1398 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.332 | 25.0969 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299.1 | 2.5 | 301.6 | 187.2 | 187.2 | 0 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0182 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299 | 2.5 | 301.5 | 187.1 | 187.1 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.5
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299 | 2.4 | 301.4 | 187 | 187.0206 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.0011 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.433 | 0.024 | 240 | 299 | 2.4 | 301.4 | 186.9 | 186.9 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0489 | 0.001 | 0 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.67 | 25 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.435 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.9 | 186.887 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.001 | -0.0001 | -0.000133 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.44 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.9 | 186.8774 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9261 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.438 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.9 | 186.9227 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8476 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.8 | 186.8 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6621 | 24.8925 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.7 | 186.7132 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.336 | 25.0908 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.43 | 0.024 | 240 | 299.1 | 2.4 | 301.5 | 186.8 | 186.82 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0125 | 27.1 | 2.4
2008-05-06 | 16:40 | 2.445 | 0.024 | 240 | 299.1 | 2.3 | 301.4 | 186.6 | 186.616 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.3
2008-05-06 | 16:40 | 2.46 | 0.024 | 240 | 299 | 2.3 | 301.3 | 186.7 | 186.7 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.3
2008-05-06 | 16:40 | 2.458 | 0.024 | 240 | 299.1 | 2.3 | 301.4 | 186.6 | 186.6 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.977 | 27.1 | 2.3
2008-05-06 | 16:40 | 2.45 | 0.024 | 240 | 299.1 | 2.2 | 301.3 | 186.4 | 186.4467 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9014 | 27.1 | 2.2
2008-05-06 | 16:40 | 2.442 | 0.024 | 240 | 299 | 2.2 | 301.2 | 186.5 | 186.5213 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9798 | 27.1 | 2.2
2008-05-06 | 16:40 | 2.314 | 0.024 | 240 | 298.9 | 2.2 | 301.1 | 186.4 | 186.4119 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3327 | 0.0488 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0583 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.766 | 0.0138 | 138.3986 | 283 | 2.2 | 285.2 | 186.4 | 186.4215 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.047399 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.267 | 0.0147 | 146.7319 | 244.1 | 2.2 | 246.3 | 186.4 | 186.4 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0408 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0849 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.26 | 0.0145 | 144.5841 | 179.5 | 2.2 | 181.7 | 147.1 | 147.1345 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.030861 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0064 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 155 | 2.2 | 157.2 | 105 | 105.019 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.027233 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 153.4 | 2.2 | 155.6 | 94.9 | 94.8693 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.026634 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0503 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 152.6 | 2.167 | 154.7 | 92.8 | 92.7728 | 0 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0265 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.1 | 27.1 | 2.2
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 152.5 | 2.1 | 154.6 | 92.3 | 92.327 | 0 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.026466 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6646 | 25.0929 | 27.1 | 2.1
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 152.4 | 2.1 | 154.5 | 91.9 | 91.9112 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3667 | 0.0264 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0143 | 27.1 | 2.1
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 152.4 | 2.1 | 154.5 | 91.9 | 91.9194 | 0 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0264 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6694 | 24.936 | 27.1 | 2.1
2008-05-06 | 16:40 | 1.26 | 0.013 | 130 | 152.3 | 2.1 | 154.4 | 91.7 | 91.7 | 0 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0264 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6707 | 24.9425 | 27.1 | 2.1
2008-05-06 | 16:41 | 1.252 | 0.013 | 130 | 152.3 | 2.1 | 154.4 | 91.7 | 91.7 | 0 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3327 | 0.026367 | 0.000767 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6727 | 25.021 | 27.1 | 2.1
2008-05-06 | 16:41 | 1.243 | 0.013 | 130 | 152.3 | 2.0675 | 154.4 | 91.6 | 91.6284 | 0 | 12.8 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0995 | 27.1 | 2.1
2008-05-06 | 16:41 | 1.24 | 0.0122 | 121.7333 | 152.3 | 2 | 154.3 | 91.5 | 91.5257 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995.6634 | 25.1 | 27.1 | 2
2008-05-06 | 16:41 | 1.24 | 0.0127 | 126.6 | 152.3 | 2 | 154.3 | 91.4 | 91.4222 | 0 | 12.8 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25.0438 | 27.1 | 2
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.3 | 1.9667 | 154.3 | 91.5 | 91.5 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 2
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.3 | 1.9 | 154.2 | 91.4 | 91.4 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.2 | 1.9 | 154.1 | 91.3 | 91.3167 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.9 | 154 | 91.4 | 91.3788 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0263 | 0.0007 | -0.0001 | -0.0002 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026267 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9517 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3212 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0262 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9267 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3214 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0262 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9893 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.1 | 91.1232 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3327 | 0.0262 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8325 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3 | 0 | 12.9 | 924 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3325 | 0.0262 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8622 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.2742 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0262 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9405 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0262 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.2 | 91.2392 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.3 | 91.3 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.2 | 91.2164 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.1 | 91.1464 | 0 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.1 | 91.1 | 0 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.000567 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.989 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152.1 | 1.8 | 153.9 | 91.2 | 91.2 | 0 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3329 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3293 | 24.9104 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3327 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.6727 | 24.8319 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8929 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.8 | 925 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0248 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.9 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0969 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.013 | 130 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.8 | 925 | 944 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.666 | 25.0183 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.0113 | 112.9163 | 152 | 1.8 | 153.8 | 91.2 | 91.2 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1602 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.011 | 110 | 152 | 1.8 | 153.8 | 91.2 | 91.1811 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0261 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9614 | 27.1614 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120.1211 | 152 | 1.8 | 153.8 | 91.1 | 91.1152 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3347 | 24.917 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.0127 | 126.5597 | 152 | 1.8 | 153.8 | 91.2 | 91.18 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3313 | 24.9955 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.9 | 1.8 | 153.7 | 91.1 | 91.1 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3293 | 25.0739 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.9 | 1.8 | 153.7 | 91.2 | 91.2 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0477 | 27.1523 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.9 | 1.8 | 153.7 | 91.1 | 91.1253 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9693 | 27.1693 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.2 | 91.1848 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8908 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.9 | 1.8 | 153.7 | 91 | 91.0209 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8125 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.9 | 1.8 | 153.7 | 91.1 | 91.1 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3333 | 24.734 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.1 | 91.0812 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.7444 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.9 | 153.7 | 91 | 91.0238 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8456 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.9 | 153.7 | 91 | 91.0243 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9975 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.9 | 153.7 | 91.1 | 91.1 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8408 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.8669 | 153.7 | 91 | 91.0234 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8581 | 27.1 | 1.9
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.1 | 91.0823 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.6707 | 24.9 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.24 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.1 | 91.1 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.3273 | 24.885 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.233 | 0.012 | 120 | 151.7 | 1.8 | 153.5 | 91 | 91.0217 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.026 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8066 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91 | 91.0123 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.3337 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9436 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.7 | 1.8 | 153.5 | 91.2 | 91.1739 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.7 | 1.8 | 153.5 | 91.1 | 91.1 | 0 | 12.9 | 924 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9426 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.7 | 1.8 | 153.5 | 91.1 | 91.1 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8071 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.7 | 1.8 | 153.5 | 91.1 | 91.1263 | 0 | 12.8 | 924 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8855 | 27.1 | 1.8
2008-05-06 | 16:41 | 1.23 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.1 | 91.1 | 0 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8361 | 27.1639 | 1.8
2008-05-06 | 16:41 | 0.786 | 0.012 | 120 | 151.8 | 1.8 | 153.6 | 91.1 | 91.1 | 0 | 12.9 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0259 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1576 | 1.8
2008-05-06 | 16:41 | 0.242 | 0.0036 | 35.825 | 142.6 | 1.8 | 144.4 | 91.1 | 91.1 | 0 | 12.8 | 925 | 944 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.0244 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8208 | 27.1 | 1.8
2008-05-06 | 16:41 | 0.043 | -0.0005 | -4.9297 | 78.3 | 1.8 | 80.1 | 90.3 | 90.3258 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.01374 | 0.0004 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8992 | 27.1 | 1.8
2008-05-06 | 16:41 | 0 | 0 | 0 | 7.5 | 1.8 | 9.3 | 32.9 | 32.8848 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0.003731 | 0.000368 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1 | 1.8
2008-05-06 | 16:41 | 0 | -0.0002 | -1.6209 | 4.3 | 1.7659 | 6 | 3.6 | 3.5975 | 0 | 12.9 | 923 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3341 | 0.001032 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995.6595 | 24.8439 | 27.1 | 1.8
2008-05-06 | 16:41 | 0.007 | -0.001 | -9.9335 | 2.6 | 1.6664 | 4.2 | -1.3 | 0 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | 0.0007 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8345 | 27.1 | 1.7
2008-05-06 | 16:41 | 0.005 | -0.001 | -10 | 2.2 | 1.5666 | 3.8 | -2.5 | 0 | 0 | 12.9 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0006 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8871 | 27.1 | 1.6
2008-05-06 | 16:41 | 0 | -0.001 | -10 | 2.1 | 1.4 | 3.5 | -3 | 0 | 0 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0006 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8087 | 27.1 | 1.4
2008-05-06 | 16:41 | 0.002 | -0.001 | -10 | 2.1 | 1.334 | 3.4 | -3.1 | 0 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0.0006 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9395 | 27.1 | 1.3
2008-05-06 | 16:41 | 0.01 | -0.001 | -10 | 1.4 | 1.1665 | 2.6 | -3.2 | 0 | 0 | 12.8 | 924 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | 0.0005 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9036 | 27.1 | 1.2
2008-05-06 | 16:42 | 0.002 | -0.001 | -10 | 1.3 | 1 | 2.3 | -3.3 | 0 | 0 | 12.8 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0005 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8267 | 27.1267 | 1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 1.2 | 0.8 | 2 | -3.3 | 0 | 0 | 12.8 | 924 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0004 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9051 | 27.2 | 0.8
2008-05-06 | 16:42 | 0.005 | -0.001 | -10 | 0.5 | 0.6 | 1.1 | -3.4 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.3328 | 0.000367 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9835 | 27.2 | 0.6
2008-05-06 | 16:42 | 0.01 | -0.001 | -10 | 0 | 0.6 | 0.6 | -3.4 | 0 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0002 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8762 | 27.1381 | 0.6
2008-05-06 | 16:42 | 0.008 | -0.001 | -10 | 0 | 0.6 | 0.6 | -3.4 | 0 | 0 | 12.9 | 923 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3325 | 0.0002 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8 | 27.1 | 0.6
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.5 | 0.5 | -3.5 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.6 | 27 | 5 | 195 | 59.3338 | 0.0002 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8187 | 27.1 | 0.5
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.4 | 0.4 | -3.5 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8971 | 27.1 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.4 | 0.4 | -3.5 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1755 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.4 | 0.4 | -3.5 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 193 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.146 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.4 | 0.4 | -3.5 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192.4 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8676 | 27.1324 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.4 | 0.4 | -3.6 | 0 | 0 | 12.9 | 923 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8216 | 27.2 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.36667 | 0.4 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3667 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9785 | 27.2 | 0.4
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.3 | 0.3 | -3.6 | 0 | 0 | 12.9 | 923 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8646 | 27.1323 | 0.3
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.267 | 0.3 | -3.6 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.33 | 24.8461 | 27.1 | 0.3
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996.6747 | 24.9245 | 27.1245 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.1 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | 0 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9971 | 27.2 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3336 | 0 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9187 | 27.2 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3334 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8403 | 27.1403 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.367 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1034 | 996 | 24.8331 | 27.1166 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9901 | 27.1951 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.2 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.9 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 25 | 27.1482 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.7 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9697 | 27.1 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.7 | 0 | 0 | 12.8 | 925 | 944 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3327 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.2 | 0.2 | -3.6 | 0 | 0 | 12.8 | 925 | 943 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9 | 27.1 | 0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | 0 | 0.16627 | 0.2 | -3.7 | 0 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.6 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8349 | 27.1 | 0.2
2008-05-06 | 16:42 | -0.003 | -0.001 | -10 | -0.1 | 0.1 | 0.1 | -3.7 | 0 | 0 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.1435 | 0.1
2008-05-06 | 16:42 | -0.01 | -0.0005 | -4.9221 | -0.1 | 0.1 | 0.1 | -3.7 | 0 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | -0.0001 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8668 | 27.2 | 0.1
2008-05-06 | 16:42 | -0.01 | 0 | 0 | -0.1 | 0.1 | 0.1 | -3.7 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3668 | -0.000133 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0993 | 27.2 | 0.1
2008-05-06 | 16:42 | -0.002 | 0 | 0 | -0.1 | 0.1 | 0.1 | -3.7 | 0 | 0 | 12.8 | 925 | 943 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.6 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25.0209 | 27.2 | 0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.1 | 0.066667 | 0.1 | -3.6 | 0 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3333 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.2 | 0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.1 | 0 | 0 | -3.7 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.928 | 27.164 | 0
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.2 | 0 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.6 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8144 | 27.1144 | 0
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.2 | 0 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3673 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.8928 | 27.1928 | 0
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.2 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0001 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 995 | 24.9712 | 27.1288 | -0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.2 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3662 | -0.0002 | -0.0001 | -0.000134 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.901 | 27.1495 | -0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.3 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.3 | -0.1 | 0 | -3.5 | 0 | 0 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.3 | -0.1332 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3668 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.3 | -0.2 | 0 | -3.5 | 0 | 0 | 12.9 | 923 | 941 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9899 | 27.2 | -0.2
2008-05-06 | 16:42 | 0 | 0 | 0 | -0.3 | -0.1 | 0 | -3.5 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.3666 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.934 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.0002 | -1.5339 | -0.4 | -0.1 | 0 | -3.5 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 192 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.7591 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -9.9162 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191.4 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9957 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.3673 | -0.0002 | -0.0002 | -0.000167 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 924 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 191 | 27 | 5 | 195 | 59.4 | -0.0002 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.8893 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190.4 | 27 | 5 | 195 | 59.3 | -0.000234 | -0.0002 | -0.000134 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.7665 | 27.2 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.9 | 925 | 942 | 931 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.4 | -0.0003 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.7231 | 27.1885 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.1 | 0 | -3.6 | 0 | 0 | 12.8 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 190 | 27 | 5 | 195 | 59.3334 | -0.0003 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.881 | 27.1095 | -0.1
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -3.6 | 0 | 0 | 12.9 | 925 | 941 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189.4 | 27 | 5 | 195 | 59.4 | -0.0003 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.969 | 27.169 | -0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.3 | -0.2 | 0 | -3.6 | 0 | 0 | 12.8 | 925 | 942 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | -0.0003 | -0.0002 | -0.0002 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 25 | 27.2 | -0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -3.6 | 0 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.4 | -0.0003 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9742 | 27.2 | -0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.2 | 0 | -3.7 | 0 | 0 | 12.8 | 925 | 943 | 933 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3667 | -0.0003 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1035 | 996 | 24.9082 | 27.2 | -0.2
2008-05-06 | 16:42 | 0 | -0.001 | -10 | -0.4 | -0.23313 | 0 | -3.7 | 0 | 0 | 12.9 | 925 | 942 | 932 | 0 | 0 | 0 | 985 | -40 | 0 | 28 | 189 | 27 | 5 | 195 | 59.3331 | -0.0003 | -0.0002 | -0.0001 | -0.0001 |  | n/a |  | 32 | 59 | 1034 | 995 | 25.0654 | 27.2 | -0.2
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-05-06
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 26374
Version | 6.907
CO Span(%) | 0.12
CO2 Span(%) | 12
C6H14 Span(ppm) | 2004
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 210
Version | 1.611
NO Span(ppm) | 1495
NO2 Span(ppm) | 504
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 957.0
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | CAT TH83
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 4
Torque (ecm or calc) | ecm
Use Frictional Torque | NO
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 16:21
Test End Time | 16:42
Test Duration (s) | 1264
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 985
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
iVEH_SPEED | iSCB_EAI1_XF
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/06/2008 16:15:08.909 - None Found
0X0000 - 05/06/2008 16:18:12.801 - None Found
0X0000 - 05/06/2008 16:21:20.671 - None Found
Warnings:
0X0000 - 05/06/2008 16:15:08.910 - None Found
0X0000 - 05/06/2008 16:18:12.801 - None Found
0X0000 - 05/06/2008 16:21:20.672 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 124
Brake Specific Dropout Limit Count | 1264
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 | iVEH_SPEED | Vehicle Speed (Dyno) | MPH | 1 | x^0: 0; x^1: 6
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-05-06
Time | 16:15
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
CO(ppm) | 40 | -0.66667 | 40.6667
CO2(%) | 0.01 | 0 | 0.01
NO(ppm) | 3.1812 | 0.44667 | 2.7346
NO2(ppm) | 0.725 | -0.36 | 1.085
THC(ppmC) | -1.1625 | -0.4 | -0.7625
Span
InfoVer | 1
Date | 2008-05-06
Time | 16:18
Purge Delay | 30
Gas Path | Span
| CO | CO2 | O2 | HC | x^0: NO; x^1: NO2; x^2: CH4; x^3: THC
Bottle Values | 1200 | 12 | 20.9 | 2004 | x^0: 1495; x^1: 504; x^2: 0; x^3: 957
Gas | Previous | Current | Difference
CO(ppm) | 1199.375 | 1200.6667 | -1.2917
CO2(%) | 11.9625 | 12.0013 | -0.038833
NO(ppm) | 1490.4062 | 1498.2467 | -7.8404
THC(ppmC) | 956.975 | 957.82 | -0.845
Test Information:
SEMTECH_DATA_FILE | 2008-05-06 | 16:15 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
No Information Available

## THC Lin
Test Date:
Test Time:
5/6/2008
16:21:21
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## CO2 Lin
Test Date:
Test Time:
5/6/2008
16:21:21
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## CO Lin
Test Date:
Test Time:
5/6/2008
16:21:21
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## NO Lin
Test Date:
Test Time:
5/6/2008
16:21:21
NDUV Serial No:
210
SEMTECH Serial Number:
G05-SDS02

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
5/6/2008
16:21:21
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
5/6/2008
16:21:21
NDUV Serial No:
210
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
12 | Rm = | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554 | 0.98554
| C.F.= | 0.098697 | 0.19768 | 0.29695 | 0.39651 | 0.49636 | 0.5965 | 0.69693 | 0.79766 | 0.89868
| gas = | 1.1844 | 2.3722 | 3.5634 | 4.7581 | 5.9563 | 7.158 | 8.3632 | 9.5719 | 10.7842
Ro for CO2= | 0.8795