## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
0 | 0 | 0 | 0
Horiba SGD-A10 | GDU-003 | 2008-08-08 | 2009-08-08
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | GDU-000 | 39577 | 39942
SEMTECH test software: | 2.018 66
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 28261
NDUV Serial Number: | 131
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM058694 | ±1% | 2011-09-10 | 957
1200 ppm CO | 1196
12% CO2 | 11.99
1500 ppm NO | 1504
Audit bottle | [5 empty] | THC
33 ppm Prop | 31 | Scott gas | AAL8755 | ±1% | 2010-07-07 | 93
200 ppm CO | 201
6% CO2 | 6
300 ppm NO | 302
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3214 |  |  |  |  | 9642
80000 ppm CO | 80460
12% CO2 | 11.99
3000 ppm NO | 3004
BAR97 Low | [5 empty] | THC
200 ppm Prop | 198.1 |  |  |  |  | 594.3
5000 ppm CO | 5008
6% CO2 | 6
300 ppm NO | 302
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13030 |  |  |  |  | 39090
NO2 bottles
80 - 150 ppm | 114
500 ppm | 518.6 | Scott gas | ALM034853 | ±1% | 2010-04-15
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | CO2 - AMBII Serial Number: |  |  | 28261 |  |  | CO - AMBII Serial Number: |  |  | 28261 |  | HC - AMBII Serial Number: |  |  | 28261 |  | NO - NDUV Serial Number: |  |  | 131 |  | THC - FID Serial Number: |  |  | 66
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-12-02 |  | 6 | X | 0.1 | 3 |  |  | 201 |  | 50 | 3 |  | 31 |  | 4 | 3 |  | 302 |  | 15 | 3 |  | 93 | 100 | 5 | 2
Test Time: | 13:18
Test Software: | 2.018 66
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0.001 | 0.001 |  |  | 0 | -2 | -2 |  |  | 0 | 1013 | 1013 |  |  | 0 | 0.4 | 0.4 |  |  | 0 | -0.11 | -0.11
90 | 105 | 1 | 0.6 | 0.596 | 0.57 | -0.026 | -4.36 |  | 20.1 | 30 | 10 | 49.75 |  | 3.1 | 1013.8 | 1010.7 | 32603.23 | Fail | 30.2 | 31 | 0.8 | 2.65 |  | 9.3 | 9.3 | 0 | 0
150 | 165 | 2 | 1.2 | 1.193 | 1.146 | -0.047 | -3.94 |  | 40.2 | 50 | 10 | 24.88 |  | 6.2 | 1014 | 1007.8 | 16254.84 | Fail | 60.4 | 61.7 | 1.3 | 2.15 |  | 18.6 | 18.7 | 0.1 | 0.54
210 | 225 | 3 | 1.8 | 1.791 | 1.705 | -0.086 | -4.8 |  | 60.3 | 75 | 15 | 24.88 |  | 9.3 | 1014.3 | 1005 | 10806.45 | Fail | 90.6 | 92.3 | 1.7 | 1.88 |  | 27.9 | 27.87 | -0.03 | -0.11
270 | 285 | 4 | 2.4 | 2.39 | 2.302 | -0.088 | -3.68 |  | 80.4 | 90 | 10 | 12.44 |  | 12.4 | 1014.5 | 1002.1 | 8081.45 | Fail | 120.8 | 121.2 | 0.4 | 0.33 |  | 37.2 | 36.9 | -0.3 | -0.81
330 | 345 | 5 | 3 | 2.989 | 2.89 | -0.099 | -3.31 |  | 100.5 | 108 | 8 | 7.96 |  | 15.5 | 1014.4 | 998.9 | 6444.52 | Fail | 151 | 150.8 | -0.2 | -0.13 |  | 46.5 | 46.1 | -0.4 | -0.86
390 | 405 | 6 | 3.6 | 3.59 | 3.5 | -0.09 | -2.51 |  | 120.6 | 122 | 1 | 0.83 |  | 18.6 | 1014 | 995.4 | 5351.61 | Fail | 181.2 | 180.7 | -0.5 | -0.28 |  | 55.8 | 55.39 | -0.41 | -0.73
450 | 465 | 7 | 4.2 | 4.191 | 4.126 | -0.065 | -1.55 |  | 140.7 | 147 | 6 | 4.26 |  | 21.7 | 1014.9 | 993.2 | 4576.96 | Fail | 211.4 | 210.7 | -0.7 | -0.33 |  | 65.1 | 64.76 | -0.34 | -0.52
510 | 525 | 8 | 4.8 | 4.793 | 4.745 | -0.048 | -1 |  | 160.8 | 170 | 9 | 5.6 |  | 24.8 | 1014 | 989.2 | 3988.71 | Fail | 241.6 | 240.9 | -0.7 | -0.29 |  | 74.4 | 74 | -0.4 | -0.54
570 | 585 | 9 | 5.4 | 5.396 | 5.373 | -0.023 | -0.43 |  | 180.9 | 176 | -5 | -2.76 |  | 27.9 | 1014.4 | 986.5 | 3535.84 | Fail | 271.8 | 270.9 | -0.9 | -0.33 |  | 83.7 | 83.39 | -0.31 | -0.37
630 | 645 | 10 | 6 | 6 | 6.009 | 0.009 | 0.15 |  | 201 | 199 | -2 | -1 |  | 31 | 1014.1 | 983.1 | 3171.29 | Fail | 302 | 301.4 | -0.6 | -0.2 |  | 93 | 92.89 | -0.11 | -0.12
690 | 705 | 9 | 5.4 | 5.396 | 5.367 | -0.029 | -0.54 |  | 180.9 | 176 | -5 | -2.76 |  | 27.9 | 1013.9 | 986 | 3534.05 | Fail | 271.8 | 272.7 | 0.9 | 0.33 |  | 83.7 | 83.29 | -0.41 | -0.49
750 | 765 | 8 | 4.8 | 4.793 | 4.759 | -0.034 | -0.71 |  | 160.8 | 160 | -1 | -0.62 |  | 24.8 | 1014.3 | 989.5 | 3989.92 | Fail | 241.6 | 242.9 | 1.3 | 0.54 |  | 74.4 | 73.83 | -0.57 | -0.77
810 | 825 | 7 | 4.2 | 4.191 | 4.126 | -0.065 | -1.55 |  | 140.7 | 140 | -1 | -0.71 |  | 21.7 | 1014 | 992.3 | 4572.81 | Fail | 211.4 | 212.4 | 1 | 0.47 |  | 65.1 | 64.5 | -0.6 | -0.92
870 | 885 | 6 | 3.6 | 3.59 | 3.508 | -0.082 | -2.28 |  | 120.6 | 124 | 3 | 2.49 |  | 18.6 | 1013.9 | 995.3 | 5351.08 | Fail | 181.2 | 182.4 | 1.2 | 0.66 |  | 55.8 | 55.17 | -0.63 | -1.13
930 | 945 | 5 | 3 | 2.989 | 2.892 | -0.097 | -3.25 |  | 100.5 | 108 | 8 | 7.96 |  | 15.5 | 1013.5 | 998 | 6438.71 | Fail | 151 | 152.6 | 1.6 | 1.06 |  | 46.5 | 45.98 | -0.52 | -1.12
990 | 1005 | 4 | 2.4 | 2.39 | 2.294 | -0.096 | -4.02 |  | 80.4 | 80 | 0 | 0 |  | 12.4 | 1013.9 | 1001.5 | 8076.61 | Fail | 120.8 | 123.1 | 2.3 | 1.9 |  | 37.2 | 36.76 | -0.44 | -1.18
1050 | 1065 | 3 | 1.8 | 1.791 | 1.71 | -0.081 | -4.52 |  | 60.3 | 67 | 7 | 11.61 |  | 9.3 | 1013.7 | 1004.4 | 10800 | Fail | 90.6 | 93.2 | 2.6 | 2.87 |  | 27.9 | 27.39 | -0.51 | -1.83
1110 | 1125 | 2 | 1.2 | 1.193 | 1.139 | -0.054 | -4.53 |  | 40.2 | 44 | 4 | 9.95 |  | 6.2 | 1013.7 | 1007.5 | 16250 | Fail | 60.4 | 63.7 | 3.3 | 5.46 |  | 18.6 | 18.29 | -0.31 | -1.67
1170 | 1185 | 1 | 0.6 | 0.596 | 0.565 | -0.031 | -5.2 |  | 20.1 | 22 | 2 | 9.95 |  | 3.1 | 1013.2 | 1010.1 | 32583.87 | Fail | 30.2 | 32.5 | 2.3 | 7.62 |  | 9.3 | 8.9 | -0.4 | -4.3
1230 | 1245 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -2 | -2 |  |  | 0 | 1013.2 | 1013.2 |  |  | 0 | 0.7 | 0.7 |  |  | 0 | -0.48 | -0.48
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | -0.052 | 0.06 |  |  |  | Intercept | 6 | 2 | Fail |  | Intercept | 1014 | 0 | Fail |  | Intercept | 1.8 | 3 |  |  | Intercept | -0.24 | 0.93
|  |  | Slope | 1 | 0.99 - 1.01 |  |  |  | Slope | 0.97 | 0.99 - 1.01 | Fail |  | Slope | 0.03 | 0.99 - 1.01 | Fail |  | Slope | 0.99 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0356 | 0.06 |  |  |  | SEE | 5.3963 | 2.01 | Fail |  | SEE | 0.3763 | 0.31 | Fail |  | SEE | 1.0451 | 3.02 |  |  | SEE | 0.1986 | 0.93
|  |  | r2 | 0.9997 | 0.998 |  |  |  | r2 | 0.9926 | 0.998 | Fail |  | r2 | 0.3697 | 0.998 | Fail |  | r2 | 0.9999 | 0.998 |  |  | r2 | 1 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 939 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8422 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 939 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8421 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.4206 | 11.3521 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3266 | 28.2734 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.205 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28.2834 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3384 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.4177 | 11.3186 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.397 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.4229 | 11.2491 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.5782 | 11.2538 | 28.2462 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.2322 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8419 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2893 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2109 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.6 | 0 | 0.6 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.4222 | 11.2675 | 28.2325 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.5 | 0 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3459 | 28.2459 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.5 | 0 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8581 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.5 | 0 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.3 | 0
2008-12-02 | 13:18 | 0 | 0 | 0 | 0.5 | -0.041767 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.9 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.3 | 0
2008-12-02 | 13:18 | -0.001 | 0 | 0 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3404 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0.0003 | 2.959 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8577 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.5771 | 11.338 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0.001 | 10 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0002 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3837 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0.001 | 10 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0001 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3052 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0.0002 | 1.8419 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2484 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.858 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.26 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8418 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.5 | -0.1 | 0.5 | -0.1 | 0 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.4 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8418 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013.5823 | 11.4 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.001 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28.2563 | -0.1
2008-12-02 | 13:19 | 0.007 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8577 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28.2221 | -0.1
2008-12-02 | 13:19 | 0 | -0.0005 | -4.6373 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3995 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | -0.001 | -10 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3211 | 28.3 | 0
2008-12-02 | 13:19 | 0 | -0.0009 | -8.8593 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.8581 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3573 | 28.3 | 0
2008-12-02 | 13:19 | 0 | -0.0001 | -0.53289 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3643 | 28.3 | 0
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | 0
2008-12-02 | 13:19 | 0.003 | 0 | 0 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | 0
2008-12-02 | 13:19 | 0.009 | 0 | 0 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3709 | 28.3 | 0
2008-12-02 | 13:19 | 0.001 | 0 | 0 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3507 | 28.3 | 0
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | 0 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8426 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2722 | 28.3 | 0
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8579 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2412 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8424 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.858 | 0.0001 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.005 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3547 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3009 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.2 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8582 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1237 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2805 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.002 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8577 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3687 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8578 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3533 | 28.3 | -0.1
2008-12-02 | 13:19 | 0 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2502 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.003 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8574 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1078 | 28.2961 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2646 | 28.2177 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2393 | 28.2607 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.858 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.4 | -0.1 | 0.4 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.01 | 0 | 0 | 0.3 | -0.1 | 0.3 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | -0.007 | 0 | 0 | 0.3 | -0.1 | 0.3 | -0.1 | 0 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8418 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.063 | 0.0027 | 27.1779 | 0.3 | -0.1 | 0.3 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8584 | 0 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1051.3358 | 1014 | 11.1472 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.537 | 0.0052 | 52.1295 | 8.6 | -0.1 | 8.6 | -0.1 | 0 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.001392 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.56 | 0.003 | 30 | 26.2 | -0.1 | 26.2 | 0.3 | 0.3255 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.003806 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.578 | 0.003 | 30 | 27.9 | -0.1 | 27.9 | 6.3 | 6.3175 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.004642 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.003 | 30 | 29.5 | -0.1 | 29.5 | 8.7 | 8.7198 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.004942 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1664 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.003 | 30 | 30.5 | -0.1 | 30.5 | 9.1 | 9.1044 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8424 | 0.005042 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.003 | 30 | 30.5 | -0.1 | 30.5 | 9.3 | 9.3 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1769 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.0027 | 27.3204 | 30.6 | -0.1 | 30.6 | 9.3 | 9.3 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1011 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.0022 | 22.186 | 30.6 | -0.1 | 30.6 | 9.3 | 9.3163 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1799 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.0038 | 38.4992 | 30.6 | -0.057312 | 30.6 | 9.4 | 9.4 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.3 | -0.1
2008-12-02 | 13:19 | 0.59 | 0.0032 | 32.3569 | 30.6 | 0 | 30.6 | 9.4 | 9.4 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2369 | 28.3 | 0
2008-12-02 | 13:19 | 0.59 | 0.003 | 30 | 30.6 | 0 | 30.6 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2697 | 28.3 | 0
2008-12-02 | 13:19 | 0.589 | 0.003 | 30 | 30.6 | 0 | 30.6 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8422 | 0.0051 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1142 | 28.3 | 0
2008-12-02 | 13:19 | 0.58 | 0.003 | 30 | 30.7 | 0.042063 | 30.7 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8579 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3156 | 28.3 | 0
2008-12-02 | 13:19 | 0.58 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2486 | 28.3 | 0.1
2008-12-02 | 13:19 | 0.58 | 0.003 | 30 | 30.6 | 0.1 | 30.7 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8576 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:19 | 0.58 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1146 | 28.3 | 0.1
2008-12-02 | 13:19 | 0.58 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2712 | 28.3 | 0.1
2008-12-02 | 13:19 | 0.579 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.571 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4234 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.215 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.5 | 9.4962 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8427 | 0.0052 | 5.7e-05 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.7 | 0.1 | 30.8 | 9.4 | 9.4 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.4 | 9.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1775 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.4 | 9.3921 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3119 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8426 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3655 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.858 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2305 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.8 | 0.1 | 30.9 | 9.3 | 9.3 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2958 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.9 | 0.1 | 31 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.347 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.9 | 0.1 | 31 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.2951 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.9 | 0.1 | 31 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8575 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.2166 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.9 | 0.1 | 31 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8423 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3231 | 28.2616 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 30.9 | 0.1 | 31 | 9.3 | 9.3 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2801 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28.2814 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28.203 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1754 | 28.2754 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8422 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1462 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1675 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8577 | 0.0052 | 0 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1542 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8422 | 0.0052 | 4.2e-05 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0758 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.2 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0053 | 28.3 | 0.2
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.5 | 59.8 | 0.0052 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1622 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.005242 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1405 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.842 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31 | 0.1 | 31.1 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8419 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.562 | 0.003 | 30 | 31.1 | 0.14185 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8419 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1731 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.566 | 0.003 | 30 | 31.1 | 0.15731 | 31.3 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1485 | 28.3 | 0.2
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8425 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.13 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1832 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8578 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0265 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8421 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.4206 | 11 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3069 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8426 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.0436 | 28.3 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.1 | 28.2781 | 0.1
2008-12-02 | 13:20 | 0.57 | 0.003 | 30 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3025 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.1003 | 28.2003 | 0.1
2008-12-02 | 13:20 | 0.595 | 0.0024 | 24.1427 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.1787 | 28.2787 | 0.1
2008-12-02 | 13:20 | 0.72 | 0.0041 | 40.625 | 31.1 | 0.1 | 31.2 | 9.3 | 9.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0053 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1428 | 28.3 | 0.1
2008-12-02 | 13:20 | 1.091 | 0.0065 | 64.9915 | 40.8 | 0.1 | 40.9 | 9.3 | 9.3086 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8 | 0.00686 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0644 | 28.3 | 0.1
2008-12-02 | 13:20 | 1.13 | 0.005 | 50 | 56.2 | 0.1 | 56.3 | 10.4 | 10.4442 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8 | 0.009467 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.014 | 28.3 | 0.1
2008-12-02 | 13:20 | 1.13 | 0.005 | 50 | 59.4 | 0.2 | 59.6 | 16.2 | 16.2029 | 0.1 | 13.2 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8425 | 0.01 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0925 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.0065 | 65.217 | 59.7 | 0.2 | 59.9 | 18.2 | 18.1665 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8577 | 0.010042 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0294 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.0064 | 64.0628 | 59.9 | 0.2 | 60.1 | 18.6 | 18.6043 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8421 | 0.0101 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0491 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.006 | 60 | 60.3 | 0.2 | 60.5 | 18.7 | 18.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8574 | 0.0101 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.006 | 60 | 60.3 | 0.2 | 60.5 | 18.7 | 18.7 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0101 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.106 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.0059 | 58.9202 | 60.3 | 0.2 | 60.5 | 18.7 | 18.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.8422 | 0.010142 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1845 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.13 | 0.0051 | 50.8622 | 60.4 | 0.2 | 60.6 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4202 | 11.1375 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.134 | 0.0058 | 57.5614 | 60.5 | 0.2 | 60.7 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5806 | 11.1 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.006 | 60 | 60.6 | 0.2 | 60.8 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4185 | 11.1197 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.0056 | 55.5929 | 60.7 | 0.2 | 60.9 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1981 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.005 | 50 | 60.7 | 0.2 | 60.9 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8425 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1235 | 28.2235 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.005 | 50 | 60.8 | 0.2 | 61 | 18.7 | 18.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8423 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5771 | 11.155 | 28.255 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.005 | 50 | 60.8 | 0.24218 | 61 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8422 | 0.0102 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2668 | 28.3 | 0.2
2008-12-02 | 13:20 | 1.14 | 0.005 | 50 | 60.9 | 0.3 | 61.2 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3882 | 28.3 | 0.3
2008-12-02 | 13:20 | 1.14 | 0.005 | 50 | 60.9 | 0.3 | 61.2 | 18.7 | 18.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.010258 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3098 | 28.3 | 0.3
2008-12-02 | 13:20 | 1.139 | 0.005 | 50 | 61 | 0.3 | 61.3 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3686 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.131 | 0.005 | 50 | 61.1 | 0.3 | 61.4 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8422 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.1 | 0.3 | 61.4 | 18.7 | 18.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3747 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.2 | 0.3 | 61.5 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.3 | 0.3 | 61.6 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.8427 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.4 | 0.3 | 61.7 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2393 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.4 | 0.3 | 61.7 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8423 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5771 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.5 | 0.3 | 61.8 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0103 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2828 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.5 | 0.3 | 61.8 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2041 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.5 | 0.3 | 61.8 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8419 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2743 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.6 | 0.3 | 61.9 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.138 | 0.005 | 50 | 61.6 | 0.3 | 61.9 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.16 | 0.005 | 50 | 61.6 | 0.3 | 61.9 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.16 | 0.005 | 50 | 61.6 | 0.3 | 61.9 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.152 | 0.005 | 50 | 61.7 | 0.3 | 62 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1336 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.15 | 0.005 | 50 | 61.7 | 0.3 | 62 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1448 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.15 | 0.005 | 50 | 61.7 | 0.3 | 62 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2232 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.147 | 0.005 | 50 | 61.7 | 0.3 | 62 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2984 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.7 | 0.3 | 62 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.22 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2632 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2152 | 28.2152 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2934 | 28.2934 | 0.3
2008-12-02 | 13:21 | 1.142 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.228 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.15 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.15 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4177 | 11.2289 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.15 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8575 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4249 | 11.2927 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.144 | 0.005 | 50 | 61.8 | 0.3 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2147 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2639 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3425 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3372 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8424 | 0.0104 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1021 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 62 | 0.3 | 62.3 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.134 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0105 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1558 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0105 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1309 | 28.3 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.9 | 0.3 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0105 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.026 | 28.287 | 0.3
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 61.9 | 0.2 | 62.1 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1828 | 28.2086 | 0.2
2008-12-02 | 13:21 | 1.13 | 0.005 | 50 | 62 | 0.2 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8584 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2697 | 28.2 | 0.2
2008-12-02 | 13:21 | 1.138 | 0.005 | 50 | 62 | 0.2 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2036 | 28.2482 | 0.2
2008-12-02 | 13:21 | 1.14 | 0.005 | 50 | 62 | 0.2 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.136 | 0.005 | 50 | 62 | 0.2 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8579 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.105 | 28.295 | 0.2
2008-12-02 | 13:21 | 1.249 | 0.005 | 50 | 62 | 0.2 | 62.2 | 18.7 | 18.7 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8419 | 0.0105 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1835 | 28.2165 | 0.2
2008-12-02 | 13:21 | 1.636 | 0.0058 | 58.4228 | 63.4 | 0.2 | 63.6 | 18.7 | 18.7 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.010712 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2617 | 0.2
2008-12-02 | 13:21 | 1.711 | 0.007 | 70 | 77.3 | 0.2 | 77.5 | 19.1 | 19.0913 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.012944 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.72 | 0.007 | 70 | 87.5 | 0.2 | 87.7 | 23.6 | 23.5583 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.858 | 0.014652 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1813 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.72 | 0.007 | 70 | 89.7 | 0.2 | 89.9 | 26.9 | 26.9299 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8581 | 0.015 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1028 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.712 | 0.007 | 70 | 89.8 | 0.2 | 90 | 27.6 | 27.6068 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.015 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5815 | 11.2511 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.007 | 70 | 89.8 | 0.2 | 90 | 27.8 | 27.8 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.015 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2462 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.007 | 70 | 89.9 | 0.2 | 90.1 | 27.8 | 27.8 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.015 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2679 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.007 | 70 | 90.1 | 0.2 | 90.3 | 27.8 | 27.8 | 0.1 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.015042 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2107 | 28.2107 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.007 | 70 | 91 | 0.2 | 91.2 | 27.8 | 27.8 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.015184 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5794 | 11.2892 | 28.2892 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.007 | 70 | 91.8 | 0.2 | 92 | 27.8 | 27.8 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.015342 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2324 | 0.2
2008-12-02 | 13:21 | 1.7 | 0.0085 | 85.0989 | 92 | 0.2 | 92.2 | 27.8 | 27.8 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0154 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4242 | 11.3459 | 28.2459 | 0.2
2008-12-02 | 13:21 | 1.702 | 0.009 | 90 | 92 | 0.2 | 92.2 | 27.8 | 27.8 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0154 | 0.0001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3758 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.71 | 0.009 | 90 | 92 | 0.2 | 92.2 | 27.8 | 27.8067 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0154 | 0.000142 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.71 | 0.009 | 90 | 92.1 | 0.2 | 92.3 | 27.9 | 27.8935 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.71 | 0.009 | 90 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8419 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3597 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.704 | 0.009 | 90 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5823 | 11.362 | 28.3 | 0.2
2008-12-02 | 13:21 | 1.704 | 0.009 | 90 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8425 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2837 | 28.2837 | 0.2
2008-12-02 | 13:21 | 1.71 | 0.0078 | 78.4996 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.1 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2051 | 28.2051 | 0.2
2008-12-02 | 13:21 | 1.71 | 0.007 | 70 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3466 | 28.2733 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8421 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3483 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8574 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3301 | 28.2699 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.1 | 0.2 | 92.3 | 27.9 | 27.9 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8423 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5766 | 11.3831 | 28.2085 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.0078 | 77.528 | 92.1 | 0.2 | 92.3 | 27.9 | 27.8977 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5798 | 11.2262 | 28.2869 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.008 | 80 | 92.2 | 0.2 | 92.4 | 27.8 | 27.836 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2653 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.0068 | 68.0101 | 92.2 | 0.2 | 92.4 | 27.9 | 27.8932 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3438 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.0052 | 52.3684 | 92.2 | 0.2 | 92.4 | 27.8 | 27.802 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3778 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.006 | 60 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2995 | 28.2995 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.006 | 60 | 92.2 | 0.2 | 92.4 | 27.9 | 27.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8423 | 0.0154 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.221 | 28.221 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.006 | 60 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.015442 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3148 | 28.2574 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.006 | 60 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3642 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.0068 | 68.3279 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2716 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.008 | 80 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1147 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.008 | 80 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.242 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.008 | 80 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.842 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2505 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.008 | 80 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8419 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3279 | 28.2279 | 0.2
2008-12-02 | 13:22 | 1.708 | 0.008 | 80 | 92.3 | 0.2 | 92.5 | 27.8 | 27.8043 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3937 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.7 | 0.008 | 80 | 92.3 | 0.3 | 92.6 | 27.9 | 27.9 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.9 | 59.8427 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3153 | 28.3 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.008 | 80 | 92.3 | 0.3 | 92.6 | 27.9 | 27.8917 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5759 | 11.2369 | 28.3 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.008 | 80 | 92.3 | 0.3 | 92.6 | 27.8 | 27.806 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2585 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.008 | 80 | 92.3 | 0.3 | 92.6 | 27.9 | 27.8879 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8421 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2199 | 28.2199 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.008 | 80 | 92.3 | 0.3 | 92.6 | 27.8 | 27.8 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8581 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5805 | 11.2984 | 28.2984 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.0074 | 74.2579 | 92.3 | 0.3 | 92.6 | 27.8 | 27.8 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3767 | 28.3 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.007 | 70 | 92.3 | 0.3 | 92.6 | 27.8 | 27.8309 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 28.3 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.007 | 70 | 92.3 | 0.3 | 92.6 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3336 | 28.3 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.007 | 70 | 92.3 | 0.25806 | 92.6 | 27.8 | 27.8048 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.212 | 28.288 | 0.3
2008-12-02 | 13:22 | 1.7 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2904 | 28.2096 | 0.2
2008-12-02 | 13:22 | 1.704 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.8936 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.8 | 27.8042 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2469 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3251 | 28.2749 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3878 | 28.2041 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8421 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1527 | 28.2824 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0155 | 0.000243 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2217 | 28.2391 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.9 | 0.1 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8576 | 0.0155 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.879 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8422 | 0.015542 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2648 | 28.2176 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.8 | 27.8 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.015558 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1077 | 28.2962 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.8 | 27.8 | 0.1 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1746 | 28.2254 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.3 | 0.2 | 92.5 | 27.9 | 27.8659 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.253 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.4 | 0.2 | 92.6 | 27.8 | 27.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.71 | 0.007 | 70 | 92.4 | 0.2 | 92.6 | 27.9 | 27.8984 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1906 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.712 | 0.007 | 70 | 92.4 | 0.2 | 92.6 | 27.8 | 27.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1118 | 28.3 | 0.2
2008-12-02 | 13:22 | 1.76 | 0.0086 | 85.7823 | 92.4 | 0.2 | 92.6 | 27.8 | 27.8 | 0.1 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8579 | 0.0155 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.2334 | 0.2
2008-12-02 | 13:22 | 2.2 | 0.0105 | 105.188 | 98.4 | 0.2 | 98.6 | 27.8 | 27.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.016489 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.19 | 28.245 | 0.2
2008-12-02 | 13:22 | 2.26 | 0.009 | 90 | 112.4 | 0.2 | 112.6 | 28.4 | 28.4328 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8426 | 0.018712 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2767 | 28.2767 | 0.2
2008-12-02 | 13:22 | 2.271 | 0.009 | 90 | 118.1 | 0.2 | 118.3 | 34.5 | 34.5138 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.019642 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2013 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.28 | 0.009 | 90 | 119.3 | 0.2 | 119.5 | 36.3 | 36.3426 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.019842 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2802 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.28 | 0.009 | 90 | 119.5 | 0.2 | 119.7 | 36.8 | 36.8 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0199 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.281 | 0.009 | 90 | 119.5 | 0.2 | 119.7 | 36.8 | 36.8045 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0199 | 0.0002 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3371 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.289 | 0.009 | 90 | 119.6 | 0.2 | 119.8 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8573 | 0.0199 | 0.000243 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3534 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.312 | 0.009 | 90 | 120 | 0.2 | 120.2 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8575 | 0.02 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5751 | 11.1181 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.32 | 0.009 | 90 | 120.2 | 0.2 | 120.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8577 | 0.02 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2447 | 28.2723 | 0.2
2008-12-02 | 13:22 | 2.316 | 0.009 | 90 | 120.2 | 0.2 | 120.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.8422 | 0.02 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5782 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:22 | 2.31 | 0.009 | 90 | 120.2 | 0.2 | 120.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.02 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:22 | 2.309 | 0.009 | 90 | 120.3 | 0.2 | 120.5 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:22 | 2.301 | 0.009 | 90 | 120.4 | 0.2 | 120.6 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:22 | 2.3 | 0.009 | 90 | 120.5 | 0.2 | 120.7 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8435 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2356 | 28.3 | 0.2
2008-12-02 | 13:22 | 2.294 | 0.009 | 90 | 120.5 | 0.2 | 120.7 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2429 | 28.2571 | 0.2
2008-12-02 | 13:22 | 2.294 | 0.0097 | 97.1583 | 120.6 | 0.2 | 120.8 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4194 | 11.2787 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.3 | 0.01 | 100 | 120.6 | 0.2 | 120.8 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2003 | 28.2 | 0.2
2008-12-02 | 13:22 | 2.3 | 0.01 | 100 | 120.6 | 0.2 | 120.8 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8584 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2777 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.01 | 100 | 120.7 | 0.2 | 120.9 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8425 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3125 | 28.2437 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.01 | 100 | 120.7 | 0.2 | 120.9 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2953 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.01 | 100 | 120.8 | 0.2 | 121 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0201 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1133 | 28.2133 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.0093 | 92.6684 | 120.9 | 0.2 | 121.1 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.020142 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1917 | 28.2917 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 120.9 | 0.2 | 121.1 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2304 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.0094 | 93.9197 | 120.9 | 0.2 | 121.1 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.0098 | 97.5839 | 121 | 0.2 | 121.2 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8567 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.307 | 0.009 | 90 | 121 | 0.2 | 121.2 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2054 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121 | 0.2 | 121.2 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2838 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4177 | 11.1379 | 28.2379 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1811 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121.1 | 0.2 | 121.3 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.31 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2759 | 0.2
2008-12-02 | 13:23 | 2.305 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3541 | 28.2459 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3351 | 28.2 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2111 | 28.2111 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.2 | 0.2 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2896 | 28.2896 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.2 | 0.15815 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.3 | 0.2
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.3 | 0.1 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8582 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2078 | 28.2539 | 0.1
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.3 | 0.1 | 121.4 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8425 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1243 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.3 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2032 | 0.1
2008-12-02 | 13:23 | 2.298 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0202 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2816 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8579 | 0.0203 | 0.0003 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.14 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0203 | 0.0003 | -0.0001 | -0.000142 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1768 | 28.2616 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0003 | -0.000142 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2832 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0203 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2047 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.000342 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2522 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.4 | 0.1 | 121.5 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0203 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2391 | 28.2305 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0203 | 0.0003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1087 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8575 | 0.0203 | 0.0003 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1873 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0203 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2658 | 28.2342 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3442 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8579 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3325 | 28.2225 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.0094 | 93.8832 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.01 | 100 | 121.5 | 0.1 | 121.6 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.020342 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.0099 | 99.2388 | 121.6 | 0.1 | 121.7 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0204 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2158 | 28.3 | 0.1
2008-12-02 | 13:23 | 2.296 | 0.0091 | 91.226 | 121.6 | 0.1 | 121.7 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0204 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2637 | 28.2637 | 0.1
2008-12-02 | 13:23 | 2.296 | 0.009 | 90 | 121.6 | 0.1 | 121.7 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.29 | 0.009 | 90 | 121.6 | 0.1 | 121.7 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.32 | 0.009 | 90 | 121.6 | 0.1 | 121.7 | 36.9 | 36.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1284 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.65 | 0.0097 | 97.2241 | 121.7 | 0.1 | 121.8 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8575 | 0.0203 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.15 | 28.25 | 0.1
2008-12-02 | 13:23 | 2.813 | 0.012 | 120 | 128.7 | 0.1 | 128.8 | 36.9 | 36.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.02143 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1716 | 28.2716 | 0.1
2008-12-02 | 13:23 | 2.865 | 0.012 | 120 | 146.8 | 0.1 | 146.9 | 38.7 | 38.6846 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.024258 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1136 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.87 | 0.012 | 120 | 149.4 | 0.1 | 149.5 | 44.4 | 44.3574 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0247 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2701 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.87 | 0.012 | 120 | 149.6 | 0.1 | 149.7 | 45.6 | 45.6256 | 0.1 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0247 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2364 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.871 | 0.012 | 120 | 149.6 | 0.1 | 149.7 | 46 | 45.9524 | 0.1 | 13.1 | 921 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0247 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1579 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.888 | 0.012 | 120 | 149.7 | 0.1 | 149.8 | 46 | 46 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.842 | 0.0247 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0795 | 28.2205 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 149.8 | 0.1 | 149.9 | 46 | 46 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.8 | 0.024742 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0011 | 28.2989 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 149.9 | 0.1 | 150 | 46 | 46 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8583 | 0.0248 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4173 | 11.2313 | 28.2229 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 149.9 | 0.1 | 150 | 46 | 46.0187 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8582 | 0.0248 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5823 | 11.2444 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 150 | 0.1 | 150.1 | 46 | 46.0104 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0248 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 150.1 | 0.1 | 150.2 | 46.1 | 46.1 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8422 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2125 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 150.1 | 0.1 | 150.2 | 46.1 | 46.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8578 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.291 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 150.2 | 0.1 | 150.3 | 46.1 | 46.1 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8579 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5794 | 11.2307 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.012 | 120 | 150.3 | 0.1 | 150.4 | 46.1 | 46.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4195 | 11.1523 | 28.2477 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.0113 | 112.8457 | 150.3 | 0.1 | 150.4 | 46.1 | 46.1 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5758 | 11.1524 | 28.2738 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.011 | 110 | 150.4 | 0.1 | 150.5 | 46.1 | 46.1 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0249 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:23 | 2.89 | 0.011 | 110 | 150.4 | 0.1 | 150.5 | 46.1 | 46.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0249 | 0.0004 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.4 | 0.1 | 150.5 | 46.1 | 46.1 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2385 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.5 | 0.1 | 150.6 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4185 | 11.2399 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.5 | 0.1 | 150.6 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.5 | 0.1 | 150.6 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8423 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.5 | 0.1 | 150.6 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.5 | 0.1 | 150.6 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3536 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.6 | 0.1 | 150.7 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8419 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.368 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.6 | 0.1 | 150.7 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8582 | 0.0249 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2104 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.7 | 0.1 | 150.8 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.8422 | 0.025 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2889 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.7 | 0.1 | 150.8 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.025 | 0.000458 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2327 | 28.2327 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.7 | 0.1 | 150.8 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.025 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2915 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.7 | 0.1 | 150.8 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8581 | 0.025 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5815 | 11.3517 | 28.2242 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.7 | 0.1 | 150.8 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.025 | 0.0004 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2974 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.219 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4218 | 11.2 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2162 | 28.2162 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8419 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2946 | 28.2946 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.8 | 0.1 | 150.9 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8422 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2269 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.0109 | 109.2828 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.842 | 0.025 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5798 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.0101 | 101.1328 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8581 | 0.025042 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5806 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.01 | 100 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8419 | 0.0251 | 0.000542 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.0105 | 105.3289 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0251 | 0.0006 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0251 | 0.0006 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8577 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8422 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.256 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 150.9 | 0.1 | 151 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.8421 | 0.025142 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1004 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.1 | 151.1 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1788 | 28.2 | 0.1
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2573 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5778 | 11.3357 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3859 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3075 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8582 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2012 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8577 | 0.0252 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4229 | 11.1124 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.025158 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2693 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2369 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.89 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2415 | 28.2 | 0.2
2008-12-02 | 13:24 | 2.959 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4175 | 11.2801 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.163 | 0.011 | 110 | 151 | 0.2 | 151.2 | 46.1 | 46.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0251 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2016 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.49 | 0.011 | 110 | 156.8 | 0.2 | 157 | 46.1 | 46.1043 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.842 | 0.025982 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4202 | 11.2768 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.49 | 0.011 | 110 | 175 | 0.2 | 175.2 | 49.3 | 49.2861 | 0.1 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.6 | 59.9 | 0.028777 | 0.0006 | -0.0001 | -0.000142 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2452 | 28.2548 | 0.2
2008-12-02 | 13:24 | 3.49 | 0.011 | 110 | 178.7 | 0.2 | 178.9 | 53.9 | 53.9135 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8581 | 0.0294 | 0.0006 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2331 | 28.2669 | 0.2
2008-12-02 | 13:24 | 3.49 | 0.011 | 110 | 179.2 | 0.2 | 179.4 | 55 | 55.0162 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8573 | 0.0294 | 0.000557 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.288 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.494 | 0.011 | 110 | 179.2 | 0.2 | 179.4 | 55.2 | 55.2062 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5759 | 11.2096 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.011 | 110 | 179.2 | 0.2 | 179.4 | 55.3 | 55.3 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2689 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.011 | 110 | 179.3 | 0.2 | 179.5 | 55.3 | 55.3 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2056 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.0112 | 112.0173 | 179.3 | 0.2 | 179.5 | 55.3 | 55.3 | 0.1 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8581 | 0.0294 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1256 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 179.3 | 0.2 | 179.5 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1958 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 179.4 | 0.2 | 179.6 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1174 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 179.4 | 0.2 | 179.6 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0294 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.161 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 179.7 | 0.2 | 179.9 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.029442 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2785 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 179.9 | 0.2 | 180.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3826 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 180 | 0.2 | 180.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3038 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 180.1 | 0.2 | 180.3 | 55.4 | 55.406 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3747 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 180.1 | 0.2 | 180.3 | 55.5 | 55.5 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8577 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3469 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.012 | 120 | 180.1 | 0.2 | 180.3 | 55.5 | 55.5 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.0125 | 125.4374 | 180.2 | 0.2 | 180.4 | 55.5 | 55.5 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2815 | 28.2 | 0.2
2008-12-02 | 13:24 | 3.5 | 0.013 | 130 | 180.2 | 0.2 | 180.4 | 55.5 | 55.4912 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1234 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.013 | 130 | 180.1 | 0.2 | 180.3 | 55.4 | 55.4342 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3005 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.013 | 130 | 180.2 | 0.2 | 180.4 | 55.5 | 55.4975 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3096 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.013 | 130 | 180.2 | 0.2 | 180.4 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2236 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.013 | 130 | 180.2 | 0.2 | 180.4 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2981 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.0125 | 124.804 | 180.2 | 0.2 | 180.4 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195 | 59.9 | 0.0295 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2196 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.2 | 0.2 | 180.4 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.9 | 0.029542 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2589 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.2 | 0.2 | 180.4 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.8423 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2254 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.8422 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4218 | 11.1312 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.495 | 0.012 | 120 | 180.3 | 0.2 | 180.5 | 55.4 | 55.4 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2874 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.49 | 0.0127 | 126.9036 | 180.4 | 0.2 | 180.6 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 195 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5805 | 11.2275 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.492 | 0.0125 | 124.7862 | 180.5 | 0.2 | 180.7 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3021 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.0123 | 123.4862 | 180.5 | 0.2 | 180.7 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3116 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.5 | 0.013 | 130 | 180.6 | 0.2 | 180.8 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.858 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1233 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.486 | 0.0129 | 129.4928 | 180.6 | 0.2 | 180.8 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3568 | 28.2 | 0.2
2008-12-02 | 13:25 | 3.491 | 0.0121 | 121.4412 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8581 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2711 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8427 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8577 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8579 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2782 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3568 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0296 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3647 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.504 | 0.012 | 120 | 180.7 | 0.3 | 181 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0296 | 0.000558 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2863 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.3 | 55.3341 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2079 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1299 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1487 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2547 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8425 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3826 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1475 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2242 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.029658 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2188 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.8 | 0.3 | 181.1 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0296 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0805 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.002 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.505 | 0.012 | 120 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1527 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.5 | 0.012 | 120 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0005 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2546 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.502 | 0.012 | 120 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.000542 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.267 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.012 | 120 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.0122 | 121.8685 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.013 | 130 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0297 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.013 | 130 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8577 | 0.029742 | 0.0006 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2466 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.51 | 0.013 | 130 | 181 | 0.3 | 181.3 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.3 | 59.9 | 0.0298 | 0.000642 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.2 | 0.3
2008-12-02 | 13:25 | 3.625 | 0.013 | 130 | 180.9 | 0.3 | 181.2 | 55.4 | 55.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8419 | 0.0298 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2925 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.072 | 0.013 | 130 | 182.6 | 0.3 | 182.9 | 55.4 | 55.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.030052 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1356 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.104 | 0.013 | 130 | 200 | 0.3 | 200.3 | 55.7 | 55.6738 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.032633 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.131 | 0.0131 | 130.7299 | 208.4 | 0.3 | 208.7 | 61.1 | 61.1405 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0339 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.0147 | 146.9505 | 208.6 | 0.3 | 208.9 | 64 | 64.0269 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.842 | 0.0339 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5798 | 11.1168 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 208.9 | 0.3 | 209.2 | 64.5 | 64.5055 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.033942 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1958 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.2 | 0.3 | 209.5 | 64.7 | 64.7 | 0.1 | 13.1 | 922 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.034 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1257 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.2 | 0.3 | 209.5 | 64.7 | 64.7 | 0.1 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.034 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.2 | 0.3 | 209.5 | 64.7 | 64.7 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8 | 0.034 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.162 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.2 | 0.3 | 209.5 | 64.7 | 64.7 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.034042 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2907 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.7 | 0.3 | 210 | 64.7 | 64.7 | 0.1 | 13.1 | 921 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8579 | 0.0341 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4206 | 11.212 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.8 | 0.3 | 210.1 | 64.7 | 64.7 | 0.1 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0341 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2664 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.9 | 0.3 | 210.2 | 64.7 | 64.7203 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0341 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1654 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 209.9 | 0.3 | 210.2 | 64.8 | 64.8 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4222 | 11.0231 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.14 | 0.015 | 150 | 210 | 0.3 | 210.3 | 64.8 | 64.8 | 0.1 | 13.2 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.858 | 0.0342 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1011 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.133 | 0.015 | 150 | 210 | 0.3 | 210.3 | 64.8 | 64.8 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.18 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.13 | 0.015 | 150 | 210.1 | 0.3 | 210.4 | 64.8 | 64.8 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2585 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.13 | 0.015 | 150 | 210.1 | 0.3 | 210.4 | 64.8 | 64.8 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0007 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.13 | 0.015 | 150 | 210.2 | 0.3 | 210.5 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8425 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2693 | 28.2 | 0.3
2008-12-02 | 13:25 | 4.13 | 0.0147 | 146.5638 | 210.2 | 0.3 | 210.5 | 64.8 | 64.8 | 0.1 | 13.2 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5771 | 11.1124 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.3 | 0.3 | 210.6 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.144 | 0.014 | 140 | 210.3 | 0.3 | 210.6 | 64.8 | 64.8 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8419 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1507 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.139 | 0.014 | 140 | 210.4 | 0.3 | 210.7 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2291 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.4 | 0.3 | 210.7 | 64.8 | 64.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5778 | 11.2925 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.4 | 0.3 | 210.7 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2143 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.4 | 0.3 | 210.7 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8419 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.5 | 0.3 | 210.8 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8419 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.6 | 0.3 | 210.9 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.6 | 0.3 | 210.9 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8575 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.6 | 0.3 | 210.9 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8423 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.6 | 0.3 | 210.9 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8578 | 0.0343 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.014 | 140 | 210.6 | 0.3 | 210.9 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8421 | 0.034342 | 0.0008 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053.1073 | 1015 | 11.235 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.0142 | 141.7539 | 210.7 | 0.3 | 211 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8426 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5737 | 11.2867 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2082 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.12 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.8 | 64.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.11 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.8 | 64.7977 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.11 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.8 | 64.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8419 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2271 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.111 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5835 | 11.3 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.127 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8425 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.3 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.7 | 64.7041 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1758 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.8 | 0.3 | 211.1 | 64.8 | 64.8 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1406 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.8 | 0.2 | 211 | 64.7 | 64.7077 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2192 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.0145 | 144.9291 | 210.7 | 0.2 | 210.9 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2976 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.0144 | 143.5473 | 210.7 | 0.24242 | 210.9 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1478 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.7 | 0.3 | 211 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2082 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.7 | 0.25798 | 211 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2352 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.8 | 0.2 | 211 | 64.8 | 64.7911 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1226 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.8 | 0.2 | 211 | 64.8 | 64.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5815 | 11.2795 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8573 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2318 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8577 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1751 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.0343 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1069 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.034342 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5794 | 11.2638 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.0143 | 143.1596 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1794 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.0145 | 145.0458 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1386 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.1664 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.13 | 0.015 | 150 | 210.9 | 0.2 | 211.1 | 64.7 | 64.7132 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.0088 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.137 | 0.015 | 150 | 211 | 0.2 | 211.2 | 64.8 | 64.7958 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5806 | 11.074 | 28.126 | 0.2
2008-12-02 | 13:26 | 4.14 | 0.015 | 150 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.1524 | 0.2
2008-12-02 | 13:26 | 4.14 | 0.015 | 150 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1306 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.155 | 0.015 | 150 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.24 | 0.0164 | 163.524 | 211.1 | 0.2 | 211.3 | 64.7 | 64.7117 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0344 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.632 | 0.0188 | 188.1995 | 215.5 | 0.2 | 215.7 | 64.7 | 64.7 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.035006 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2662 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.717 | 0.018 | 180 | 229.6 | 0.2 | 229.8 | 65.5 | 65.5113 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.037031 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2554 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.725 | 0.018 | 180 | 237.6 | 0.3 | 237.9 | 71.5 | 71.5273 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.038142 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.223 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.733 | 0.018 | 180 | 238.1 | 0.3 | 238.4 | 73.5 | 73.5081 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.038242 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.4222 | 11.2987 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.018 | 180 | 238.7 | 0.3 | 239 | 73.9 | 73.9014 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0384 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1015 | 11.2202 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.0176 | 176.4604 | 238.8 | 0.3 | 239.1 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0384 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5806 | 11.2583 | 28.1417 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.017 | 170 | 238.9 | 0.3 | 239.2 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0384 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2633 | 28.1367 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.017 | 169.9917 | 238.9 | 0.3 | 239.2 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0384 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1699 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.0162 | 161.6514 | 239.5 | 0.3 | 239.8 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.038442 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0139 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.1 | 0.3 | 240.4 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.038543 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1437 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.2 | 0.3 | 240.5 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8578 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.2 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.3 | 0.3 | 240.6 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1711 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.3 | 0.3 | 240.6 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8581 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1073 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.4 | 0.25758 | 240.7 | 74 | 74 | 0.1 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8424 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1857 | 0.3
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.4 | 0.2 | 240.6 | 74 | 74 | 0.1 | 13.1 | 920 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.3283 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.4 | 0.2 | 240.6 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.3574 | 28.2 | 0.2
2008-12-02 | 13:26 | 4.74 | 0.016 | 160 | 240.4 | 0.2 | 240.6 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.6 | 28 | 5 | 195 | 59.8 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.279 | 28.179 | 0.2
2008-12-02 | 13:26 | 4.74 | 0.0163 | 163.3557 | 240.4 | 0.2 | 240.6 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.4 | 59.8419 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2006 | 28.1006 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.0168 | 168.2115 | 240.5 | 0.2 | 240.7 | 74.1 | 74.1 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 195.7 | 59.8582 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2776 | 28.1776 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.016 | 160 | 240.5 | 0.24229 | 240.7 | 74.1 | 74.0652 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.3 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.016 | 160 | 240.5 | 0.3 | 240.8 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2306 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.743 | 0.016 | 160 | 240.5 | 0.3 | 240.8 | 74 | 74.0035 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.752 | 0.016 | 160 | 240.6 | 0.3 | 240.9 | 74.1 | 74.1 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 189.4 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.016 | 160 | 240.7 | 0.3 | 241 | 74.1 | 74.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1695 | 28.1305 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.016 | 160 | 240.7 | 0.3 | 241 | 74.1 | 74.0961 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.2 | 28.1481 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.016 | 160 | 240.7 | 0.3 | 241 | 74.1 | 74.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0386 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1732 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.0168 | 168.3083 | 240.7 | 0.3 | 241 | 74 | 74.0473 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8582 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014.5818 | 11.1053 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.017 | 170 | 240.7 | 0.25815 | 241 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.1837 | 28.2 | 0.3
2008-12-02 | 13:27 | 4.76 | 0.017 | 170 | 240.8 | 0.2 | 241 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1054 | 1014 | 11.0759 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.76 | 0.017 | 170 | 240.8 | 0.2 | 241 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053.0125 | 1014 | 11.1215 | 28.1595 | 0.2
2008-12-02 | 13:27 | 4.777 | 0.017 | 170 | 240.8 | 0.2 | 241 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.281 | 28.119 | 0.2
2008-12-02 | 13:27 | 4.766 | 0.017 | 170 | 240.8 | 0.2 | 241 | 74 | 74 | 0.1 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2026 | 28.1974 | 0.2
2008-12-02 | 13:27 | 4.755 | 0.017 | 170 | 240.9 | 0.2 | 241.1 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0387 | 0.000842 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.017 | 170 | 240.9 | 0.2 | 241.1 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8419 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2541 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.72 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2677 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.72 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74.0119 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1891 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.72 | 0.017 | 170 | 240.9 | 0.2 | 241.1 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8581 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1105 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.734 | 0.017 | 170 | 240.9 | 0.2 | 241.1 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.168 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8582 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2461 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.74 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2514 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.742 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0968 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0183 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.741 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1805 | 28.1398 | 0.2
2008-12-02 | 13:27 | 4.747 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.223 | 28.1385 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8424 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1832 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241.1 | 0.2 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0008 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4222 | 11.1 | 28.1046 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241 | 0.2 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0387 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5806 | 11.1 | 28.1523 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241 | 0.24185 | 241.2 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.038742 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0695 | 28.2 | 0.2
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241 | 0.3 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0173 | 28.1913 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.017 | 170 | 241 | 0.3 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.8575 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1749 | 28.1125 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.0163 | 163.1634 | 241 | 0.3 | 241.3 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8422 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5782 | 11.068 | 28.1 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.1 | 0.3 | 241.4 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8579 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0445 | 28.1445 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.1 | 0.3 | 241.4 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.1771 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.1 | 0.3 | 241.4 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1013 | 28.1013 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.1 | 0.3 | 241.4 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1797 | 28.1797 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.2 | 0.3 | 241.5 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.858 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1419 | 28.1419 | 0.3
2008-12-02 | 13:27 | 4.75 | 0.016 | 160 | 241.2 | 0.3 | 241.5 | 74 | 74 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 4.833 | 0.016 | 160 | 241.2 | 0.25827 | 241.5 | 74 | 74 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5827 | 11.0851 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.215 | 0.0167 | 166.6155 | 241.2 | 0.2 | 241.4 | 74.1 | 74.1 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0388 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0069 | 28.1 | 0.2
2008-12-02 | 13:27 | 5.352 | 0.018 | 180 | 247.5 | 0.2 | 247.7 | 74.1 | 74.1114 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8575 | 0.039695 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0718 | 28.1 | 0.2
2008-12-02 | 13:27 | 5.366 | 0.018 | 180 | 267.7 | 0.3 | 268 | 76.9 | 76.9197 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8577 | 0.0425 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.36 | 0.018 | 180 | 267.9 | 0.3 | 268.2 | 81.9 | 81.8501 | 0.1 | 13.2 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.042542 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.36 | 0.018 | 180 | 269.9 | 0.3 | 270.2 | 83 | 82.9839 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.042784 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.36 | 0.018 | 180 | 270.1 | 0.3 | 270.4 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0429 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.36 | 0.018 | 180 | 270.1 | 0.3 | 270.4 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0429 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.368 | 0.018 | 180 | 270.1 | 0.3 | 270.4 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0429 | 0.0009 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0576 | 28.1424 | 0.3
2008-12-02 | 13:27 | 5.37 | 0.018 | 180 | 270.2 | 0.3 | 270.5 | 83.3 | 83.3 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.842 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0417 | 28.1791 | 0.3
2008-12-02 | 13:27 | 5.37 | 0.018 | 180 | 270.2 | 0.3 | 270.5 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.9 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1986 | 28.1007 | 0.3
2008-12-02 | 13:27 | 5.37 | 0.018 | 180 | 270.3 | 0.3 | 270.6 | 83.3 | 83.3 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8581 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5815 | 11.0447 | 28.1777 | 0.3
2008-12-02 | 13:27 | 5.374 | 0.018 | 180 | 270.3 | 0.3 | 270.6 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1682 | 28.1439 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.3 | 0.3 | 270.6 | 83.3 | 83.3 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8575 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2655 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.3 | 0.3 | 270.6 | 83.3 | 83.3 | 0.1 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8423 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.187 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.3 | 0.3 | 270.6 | 83.3 | 83.3096 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8578 | 0.0429 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1086 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.4 | 0.3 | 270.7 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.4 | 0.3 | 270.7 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1962 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.4 | 0.3 | 270.7 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8576 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2733 | 28.1 | 0.3
2008-12-02 | 13:27 | 5.39 | 0.018 | 180 | 270.4 | 0.3 | 270.7 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8422 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.4 | 0.3 | 270.7 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8581 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.5 | 0.25815 | 270.8 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3233 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3599 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.5 | 83.4978 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3188 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3973 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.39 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.4 | 83.4 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.4 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.382 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8421 | 0.043 | 0.001 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3459 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.018 | 180 | 270.6 | 0.2 | 270.8 | 83.4 | 83.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8581 | 0.043 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.3 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.018 | 180 | 270.7 | 0.2 | 270.9 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2779 | 28.1111 | 0.2
2008-12-02 | 13:28 | 5.377 | 0.018 | 180 | 270.8 | 0.2 | 271 | 83.4 | 83.4 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.043 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.121 | 28.1895 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.8 | 0.2 | 271 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.9 | 0.043 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1679 | 28.1321 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.8 | 0.2 | 271 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.043 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1461 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.043042 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1757 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8573 | 0.0431 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8575 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 271 | 0.2 | 271.2 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8577 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1399 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1385 | 28.1385 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.0177 | 176.5961 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4195 | 11.1661 | 28.1831 | 0.2
2008-12-02 | 13:28 | 5.378 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0093 | 28.1046 | 0.2
2008-12-02 | 13:28 | 5.373 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0431 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2214 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2478 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.4 | 83.3935 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2305 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.372 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.4 | 83.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2913 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.017 | 170 | 270.9 | 0.2 | 271.1 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8418 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2125 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.0173 | 173.4733 | 271 | 0.2 | 271.2 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.134 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.019 | 189.6284 | 271 | 0.2 | 271.2 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.38 | 0.019 | 190 | 271.4 | 0.2 | 271.6 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 196 | 59.8 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1227 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.377 | 0.019 | 190 | 271.4 | 0.2 | 271.6 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1994 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.019 | 190 | 271.4 | 0.2 | 271.6 | 83.3 | 83.3 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1204 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.0187 | 186.7474 | 271.4 | 0.2 | 271.6 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.0182 | 181.8182 | 271.4 | 0.2 | 271.6 | 83.3 | 83.3 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1365 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.019 | 190 | 271.4 | 0.3 | 271.7 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.37 | 0.019 | 190 | 271.5 | 0.3 | 271.8 | 83.3 | 83.3 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.37 | 0.019 | 190 | 271.5 | 0.2 | 271.7 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8582 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1282 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.0185 | 185.2083 | 271.5 | 0.2 | 271.7 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8583 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.018 | 180 | 271.6 | 0.2 | 271.8 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1288 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.378 | 0.0177 | 176.559 | 271.6 | 0.2 | 271.8 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1928 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.373 | 0.016 | 160.4039 | 271.6 | 0.2 | 271.8 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1149 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.016 | 160 | 271.6 | 0.2 | 271.8 | 83.4 | 83.4 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.37 | 0.0173 | 173.2215 | 271.5 | 0.24242 | 271.7 | 83.3 | 83.3 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.2
2008-12-02 | 13:28 | 5.426 | 0.018 | 180 | 271.5 | 0.3 | 271.8 | 83.4 | 83.4 | 0.1 | 13.1 | 917 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0791 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.696 | 0.0186 | 186.3918 | 271.6 | 0.3 | 271.9 | 83.4 | 83.4 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0432 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0007 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.922 | 0.0198 | 198.3239 | 278.5 | 0.3 | 278.8 | 83.4 | 83.4 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.044187 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.977 | 0.019 | 190.2659 | 293.1 | 0.3 | 293.4 | 85.3 | 85.3225 | 0.1 | 13.1 | 920 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.046186 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0562 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.98 | 0.019 | 190 | 298.8 | 0.3 | 299.1 | 91 | 90.9641 | 0.1 | 13.1 | 922 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.047 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0654 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.98 | 0.019 | 190 | 298.9 | 0.3 | 299.2 | 92.4 | 92.4141 | 0.1 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.047 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.98 | 0.019 | 190 | 299 | 0.3 | 299.3 | 92.7 | 92.71 | 0.1 | 13.1 | 923 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.047 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.98 | 0.019 | 190 | 299.1 | 0.3 | 299.4 | 92.8 | 92.8 | 0.1 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.047 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0697 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.988 | 0.019 | 190 | 300 | 0.3 | 300.3 | 92.8 | 92.8 | 0.1 | 13.1 | 922 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.047142 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.99 | 0.019 | 190 | 300.4 | 0.3 | 300.7 | 92.8 | 92.8 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.047242 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0732 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.99 | 0.0198 | 198.2049 | 300.5 | 0.3 | 300.8 | 92.8 | 92.8027 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0105 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.99 | 0.0207 | 206.6163 | 300.6 | 0.3 | 300.9 | 92.9 | 92.9 | 0.1 | 13.1 | 921 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1674 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.99 | 0.0205 | 205.1424 | 300.7 | 0.3 | 301 | 92.9 | 92.9 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1383 | 28.1 | 0.3
2008-12-02 | 13:28 | 5.99 | 0.02 | 200 | 300.7 | 0.3 | 301 | 92.9 | 92.9 | 0.1 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.007 | 0.02 | 200 | 300.7 | 0.3 | 301 | 92.9 | 92.9 | 0.1 | 13.1 | 920 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.081 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.01 | 0.02 | 200 | 300.7 | 0.3 | 301 | 92.9 | 92.9007 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8573 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.5731 | 11.0025 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.01 | 0.02 | 200 | 300.8 | 0.3 | 301.1 | 93 | 93 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1518 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.01 | 0.02 | 200 | 300.8 | 0.3 | 301.1 | 92.9 | 92.906 | 0.1 | 13.1 | 920 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.01 | 0.02 | 200 | 300.9 | 0.3 | 301.2 | 93 | 93 | 0.1 | 13.1 | 919 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.3
2008-12-02 | 13:28 | 6.01 | 0.02 | 200 | 300.9 | 0.3 | 301.2 | 92.9 | 92.9177 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2109 | 28.1 | 0.3
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 300.9 | 0.3 | 301.2 | 93 | 93 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2895 | 28.1 | 0.3
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301 | 0.3 | 301.3 | 93 | 92.9975 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0473 | 0.0012 | -0.0001 | -0.000142 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.1 | 0.3
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 300.9 | 0.2 | 301.1 | 93 | 93 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.047342 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.5798 | 11.2538 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301 | 0.2 | 301.2 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8418 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301 | 0.2 | 301.2 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1967 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.0194 | 193.9667 | 301.1 | 0.2 | 301.3 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1182 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.002 | 0.0181 | 181.35 | 301.1 | 0.2 | 301.3 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1602 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.019 | 189.6833 | 301.2 | 0.2 | 301.4 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1614 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.019 | 190 | 301.2 | 0.2 | 301.4 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1341 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.003 | 0.0197 | 196.5178 | 301.2 | 0.2 | 301.4 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8421 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2912 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8426 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.226 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0474 | 0.0012 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2525 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 919 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0474 | 0.001242 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2692 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.997 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1911 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.99 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8581 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1123 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.996 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1662 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.018 | 0.0199 | 198.6689 | 301.3 | 0.2 | 301.5 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.02 | 0.019 | 190.3494 | 301.3 | 0.2 | 301.5 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.7 | 59.8421 | 0.0474 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.02 | 0.0198 | 198.188 | 301.4 | 0.2 | 301.6 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8419 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2584 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.017 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2631 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1848 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.8 | 92.807 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1066 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.9 | 92.9 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.172 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.01 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0475 | 0.0013 | -0.0001 | -0.000158 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0987 | 28.1 | 0.2
2008-12-02 | 13:29 | 6.005 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.9 | 92.8931 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8427 | 0.0475 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.0582 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014.4249 | 11.1925 | 28.0925 | 0.2
2008-12-02 | 13:29 | 5.999 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.001258 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1147 | 28.0147 | 0.2
2008-12-02 | 13:29 | 5.99 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1641 | 28.0641 | 0.2
2008-12-02 | 13:29 | 5.998 | 0.02 | 200 | 301.5 | 0.2 | 301.7 | 92.8 | 92.8 | 0.1 | 13.1 | 918 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8579 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.8 | 92.8 | 0.1 | 13.2 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2213 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.6 | 0.2 | 301.8 | 92.8 | 92.8 | 0.1 | 13.1 | 917 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2997 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.7 | 0.2 | 301.9 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2221 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.7 | 0.2 | 301.9 | 92.8 | 92.8 | 0.1 | 13.1 | 917 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2564 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.7 | 0.2 | 301.9 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0475 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.23 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.8 | 0.2 | 302 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8582 | 0.0475 | 0.0013 | -0.000142 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.127 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.8 | 0.2 | 302 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.0475 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2838 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.8 | 0.2 | 302 | 92.8 | 92.8035 | 0.1 | 13.2 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8577 | 0.0475 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.3 | 28.1 | 0.2
2008-12-02 | 13:29 | 6 | 0.02 | 200 | 301.7 | 0.2 | 301.9 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.4 | 59.9 | 0.0475 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2031 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.933 | 0.02 | 200 | 301.7 | 0.2 | 301.9 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8421 | 0.0475 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1273 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.498 | 0.0191 | 190.7334 | 301.8 | 0.2 | 302 | 92.8 | 92.8 | 0.1 | 13.2 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0475 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052.105 | 1014 | 11.2057 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.464 | 0.017 | 170 | 293.4 | 0.2 | 293.6 | 92.8 | 92.7536 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.046367 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2841 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.401 | 0.017 | 170 | 276 | 0.2 | 276.2 | 89.8 | 89.7515 | 0.1 | 13.2 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.044 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2376 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 275.1 | 0.2 | 275.3 | 84.8 | 84.7863 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0439 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 274.8 | 0.2 | 275 | 83.7 | 83.693 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.043858 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.4194 | 11.2195 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.369 | 0.017 | 170 | 274.1 | 0.2 | 274.3 | 83.6 | 83.5659 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0438 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.298 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.36 | 0.017 | 170 | 274.1 | 0.2 | 274.3 | 83.5 | 83.5 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0438 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2236 | 28.1 | 0.2
2008-12-02 | 13:29 | 5.36 | 0.017 | 170 | 273.9 | 0.3 | 274.2 | 83.4 | 83.4 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8575 | 0.0437 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1455 | 28.1 | 0.3
2008-12-02 | 13:29 | 5.36 | 0.017 | 170 | 273.8 | 0.3 | 274.1 | 83.4 | 83.4036 | 0.1 | 13.2 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0437 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1328 | 28.1 | 0.3
2008-12-02 | 13:29 | 5.36 | 0.017 | 170 | 273.7 | 0.3 | 274 | 83.5 | 83.5 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0437 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2117 | 28.1 | 0.3
2008-12-02 | 13:29 | 5.363 | 0.017 | 170 | 273.6 | 0.3 | 273.9 | 83.4 | 83.4 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0437 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2901 | 28.1 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.5 | 0.3 | 273.8 | 83.5 | 83.4962 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0437 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3685 | 28.1 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.4 | 0.3 | 273.7 | 83.4 | 83.4 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3061 | 28.0531 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.4 | 0.3 | 273.7 | 83.4 | 83.4142 | 0.1 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8578 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2507 | 28 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.3 | 0.3 | 273.6 | 83.4 | 83.4 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3923 | 28 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.3 | 0.3 | 273.6 | 83.5 | 83.5 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8419 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2353 | 28 | 0.3
2008-12-02 | 13:29 | 5.37 | 0.017 | 170 | 273.3 | 0.3 | 273.6 | 83.5 | 83.4654 | 0.1 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0.3
2008-12-02 | 13:29 | 5.371 | 0.017 | 170 | 273.4 | 0.3 | 273.7 | 83.4 | 83.4 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8418 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.039 | 0.3
2008-12-02 | 13:29 | 5.38 | 0.017 | 170 | 273.3 | 0.2 | 273.5 | 83.4 | 83.4 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.38 | 0.017 | 170 | 273.3 | 0.2 | 273.5 | 83.4 | 83.4 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8578 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.38 | 0.017 | 170 | 273.2 | 0.2 | 273.4 | 83.4 | 83.4 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8421 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2745 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.371 | 0.017 | 170 | 273.2 | 0.2 | 273.4 | 83.4 | 83.4 | 0.1 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3529 | 28.0471 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.0177 | 176.5941 | 273.1 | 0.2 | 273.3 | 83.4 | 83.4 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3686 | 28.0314 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 273 | 0.2 | 273.2 | 83.4 | 83.4 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8422 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3098 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 273 | 0.2 | 273.2 | 83.4 | 83.4 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3883 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.9 | 0.2 | 273.1 | 83.4 | 83.4 | 0.1 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8419 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3332 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.2 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.5815 | 11.3 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3236 | 28.0764 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.000142 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3983 | 28.0017 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8423 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013.5771 | 11.3195 | 28.0805 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.241 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.36 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1626 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.367 | 0.018 | 180 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1472 | 28.0843 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.0172 | 171.826 | 272.8 | 0.2 | 273 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.043558 | 0.001258 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3806 | 28.0065 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.017 | 170 | 272.7 | 0.2 | 272.9 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8578 | 0.0435 | 0.0012 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.255 | 28.0725 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.017 | 170 | 272.7 | 0.2 | 272.9 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0435 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2512 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.017 | 170 | 272.7 | 0.2 | 272.9 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0435 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2704 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.017 | 170 | 272.7 | 0.2 | 272.9 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0435 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.017 | 170 | 272.7 | 0.2 | 272.9 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.0178 | 178.1724 | 272.6 | 0.15771 | 272.8 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2649 | 28.1 | 0.2
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.3 | 83.2934 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2565 | 28.0565 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.3 | 83.3 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0219 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.3 | 83.2931 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1997 | 28.1 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 913 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1213 | 28.1 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.3 | 83.2933 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0436 | 0.0013 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.1 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1356 | 28.1 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8581 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0859 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8419 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0075 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.342 | 28 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8575 | 0.0436 | 0.0013 | -0.0001 | -0.000158 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3506 | 28 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.7 | 0.1 | 272.8 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2722 | 28.0278 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1937 | 28.0937 | 0.1
2008-12-02 | 13:30 | 5.376 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1153 | 28.0153 | 0.1
2008-12-02 | 13:30 | 5.39 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0369 | 28.0631 | 0.1
2008-12-02 | 13:30 | 5.39 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 913 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0832 | 28.0584 | 0.1
2008-12-02 | 13:30 | 5.39 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.2 | 83.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0436 | 0.0013 | -0.000142 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2201 | 28.0201 | 0.1
2008-12-02 | 13:30 | 5.377 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.1 | 83.1 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2986 | 28.0986 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.1 | 83.1 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.146 | 28.023 | 0.1
2008-12-02 | 13:30 | 5.37 | 0.018 | 180 | 272.6 | 0.1 | 272.7 | 83.1 | 83.1 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0435 | 0.0013 | -0.0001 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1555 | 28 | 0.1
2008-12-02 | 13:30 | 5.3 | 0.0161 | 161.0746 | 272.6 | 0.1 | 272.7 | 83.1 | 83.1 | 0.1 | 13.1 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0436 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1661 | 28.0339 | 0.1
2008-12-02 | 13:30 | 4.808 | 0.0146 | 145.7024 | 271.4 | 0.1 | 271.5 | 83.1 | 83.1 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8577 | 0.043431 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1122 | 28.0878 | 0.1
2008-12-02 | 13:30 | 4.78 | 0.016 | 160 | 257.4 | 0.1 | 257.5 | 81.1 | 81.0793 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.04142 | 0.0013 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1907 | 28.0093 | 0.1
2008-12-02 | 13:30 | 4.78 | 0.016 | 160 | 247.2 | 0.1 | 247.3 | 76.4 | 76.4138 | 0.1 | 13.1 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.040174 | 0.0013 | -0.000158 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014.4198 | 11.1308 | 28.0692 | 0.1
2008-12-02 | 13:30 | 4.78 | 0.016 | 160 | 246.1 | 0.1 | 246.2 | 74.5 | 74.4743 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.039858 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.0524 | 28.0524 | 0.1
2008-12-02 | 13:30 | 4.774 | 0.016 | 160 | 245 | 0.1 | 245.1 | 74 | 74 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0397 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.5758 | 11.0521 | 28 | 0.1
2008-12-02 | 13:30 | 4.76 | 0.016 | 160 | 244.3 | 0.1 | 244.4 | 73.9 | 73.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0396 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1909 | 28.0046 | 0.1
2008-12-02 | 13:30 | 4.76 | 0.016 | 160 | 244 | 0.1 | 244.1 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0396 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0339 | 28.083 | 0.1
2008-12-02 | 13:30 | 4.76 | 0.0162 | 161.5076 | 244 | 0.1 | 244.1 | 73.9 | 73.9 | 0.1 | 13.1 | 917 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.039557 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0615 | 28.0385 | 0.1
2008-12-02 | 13:30 | 4.753 | 0.017 | 169.5269 | 243.9 | 0.1 | 244 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0395 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1399 | 28 | 0.1
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.9 | 0.1 | 244 | 73.8 | 73.8038 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0395 | 0.0013 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.5823 | 11.2 | 28.0184 | 0.1
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.8 | 0.1 | 243.9 | 73.8 | 73.8 | 0.1 | 13.1 | 916 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0395 | 0.001258 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0968 | 0.1
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.8 | 0 | 243.8 | 73.8 | 73.8 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.039458 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.7 | 0 | 243.7 | 73.8 | 73.8099 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.7 | 0 | 243.7 | 73.8 | 73.8 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8579 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1679 | 28.0679 | 0
2008-12-02 | 13:30 | 4.75 | 0.017 | 170 | 243.6 | 0 | 243.6 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1106 | 28 | 0
2008-12-02 | 13:30 | 4.745 | 0.0162 | 161.8631 | 243.5 | 0.1 | 243.6 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.189 | 28 | 0.1
2008-12-02 | 13:30 | 4.743 | 0.016 | 160 | 243.5 | 0.1 | 243.6 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1325 | 28.0675 | 0.1
2008-12-02 | 13:30 | 4.75 | 0.016 | 160 | 243.4 | 0.1 | 243.5 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28.0541 | 0.1
2008-12-02 | 13:30 | 4.75 | 0.016 | 160 | 243.4 | 0.1 | 243.5 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | 0.1
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 243.4 | 0.1 | 243.5 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8418 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1028 | 28 | 0.1
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 243.2 | 0.1 | 243.3 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1813 | 28 | 0.1
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 243.2 | 0.057708 | 243.3 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0394 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0.1
2008-12-02 | 13:31 | 4.753 | 0.0173 | 172.7006 | 243.1 | 0.042178 | 243.1 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.039358 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.238 | 28 | 0
2008-12-02 | 13:31 | 4.76 | 0.0171 | 170.701 | 243.1 | 0.1 | 243.2 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8579 | 0.0393 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2837 | 28 | 0.1
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243.1 | 0.1 | 243.2 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0393 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2049 | 28 | 0.1
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243.1 | 0 | 243.1 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0393 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243.1 | 0 | 243.1 | 73.9 | 73.8975 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 196 | 59.9 | 0.0393 | 0.0012 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243.1 | 0 | 243.1 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.842 | 0.0393 | 0.001158 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0304 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243.1 | 0 | 243.1 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243 | 0 | 243 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8419 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014.5815 | 11.2 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243 | 0 | 243 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.2 | 28.0342 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243 | 0 | 243 | 73.9 | 73.9 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8425 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.2 | 28.0442 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 243 | 0 | 243 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8423 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.2 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.9 | 0 | 242.9 | 73.8 | 73.8 | 0.1 | 13.1 | 916 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0393 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1015 | 11.2 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.9 | 0 | 242.9 | 73.9 | 73.8973 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.8 | 0 | 242.8 | 73.8 | 73.8 | 0.1 | 13.1 | 916 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8424 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.042 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.8 | 0 | 242.8 | 73.9 | 73.9 | 0.1 | 13.1 | 915 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28.0365 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.8 | 0 | 242.8 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.185 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.8 | 0 | 242.8 | 73.8 | 73.8 | 0.1 | 13.1 | 915 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1067 | 28.1 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.8 | 0 | 242.8 | 73.8 | 73.8 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1718 | 28.0282 | 0
2008-12-02 | 13:31 | 4.76 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.759 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.2 | 28.0287 | 0
2008-12-02 | 13:31 | 4.751 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 915 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1929 | 28.0929 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.1148 | 28.0148 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.1 | 914 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.0361 | 28.0639 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4195 | 11.085 | 28.0575 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 196 | 59.8424 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1015 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014.4222 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8017 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1222 | 28 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.7 | 0 | 242.7 | 73.8 | 73.8028 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1563 | 28.0563 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8581 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1652 | 28.0652 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1133 | 28 | 0
2008-12-02 | 13:31 | 4.75 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1917 | 28 | 0
2008-12-02 | 13:31 | 4.752 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8579 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1302 | 28 | 0
2008-12-02 | 13:31 | 4.769 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28 | 0
2008-12-02 | 13:31 | 4.77 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8581 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.127 | 28.027 | 0
2008-12-02 | 13:31 | 4.77 | 0.016 | 160 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1946 | 28.0946 | 0
2008-12-02 | 13:31 | 4.756 | 0.016 | 160 | 242.5 | 0 | 242.5 | 73.9 | 73.9 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1161 | 28.0161 | 0
2008-12-02 | 13:31 | 4.74 | 0.016 | 160 | 242.5 | 0 | 242.5 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.9 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1053 | 1014 | 11.1 | 28.0623 | 0
2008-12-02 | 13:31 | 4.678 | 0.0155 | 155.4459 | 242.6 | 0 | 242.6 | 73.9 | 73.9 | 0.1 | 13.1 | 914 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8419 | 0.0392 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052.6427 | 1014 | 11.1408 | 28.0592 | 0
2008-12-02 | 13:31 | 4.217 | 0.015 | 150 | 239.3 | 0 | 239.3 | 73.9 | 73.9 | 0.1 | 13.2 | 914 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8581 | 0.03874 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.157 | 0.0148 | 147.7074 | 228.7 | 0 | 228.7 | 73.7 | 73.6702 | 0.1 | 13.2 | 912 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8418 | 0.037223 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.15 | 0.0131 | 130.5988 | 217.3 | 0 | 217.3 | 68.3 | 68.3257 | 0.1 | 13.2 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.035573 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.5751 | 11.2 | 28.0762 | 0
2008-12-02 | 13:31 | 4.15 | 0.0138 | 137.7448 | 214.6 | 0 | 214.6 | 65.5 | 65.4614 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.035158 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4229 | 11.2543 | 28.0457 | 0
2008-12-02 | 13:31 | 4.15 | 0.014 | 140 | 213.8 | 0 | 213.8 | 64.8 | 64.7942 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.4 | 59.9 | 0.035058 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3 | 28.0327 | 0
2008-12-02 | 13:31 | 4.149 | 0.014 | 140 | 213.3 | 0 | 213.3 | 64.6 | 64.6 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2887 | 28.0887 | 0
2008-12-02 | 13:31 | 4.141 | 0.014 | 140 | 213.2 | 0 | 213.2 | 64.6 | 64.6 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8581 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2101 | 28.0101 | 0
2008-12-02 | 13:31 | 4.14 | 0.014 | 140 | 213.2 | 0 | 213.2 | 64.6 | 64.6 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:31 | 4.14 | 0.014 | 140 | 213.1 | 0 | 213.1 | 64.6 | 64.6 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8578 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1531 | 28 | 0
2008-12-02 | 13:31 | 4.14 | 0.014 | 140 | 213.1 | 0 | 213.1 | 64.6 | 64.6 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.034958 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1253 | 28 | 0
2008-12-02 | 13:31 | 4.137 | 0.014 | 140 | 213.1 | 0 | 213.1 | 64.6 | 64.6 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1962 | 28 | 0
2008-12-02 | 13:31 | 4.13 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1177 | 28 | 0
2008-12-02 | 13:31 | 4.13 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1607 | 28 | 0
2008-12-02 | 13:31 | 4.138 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.2 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0349 | 0.0011 | -0.000142 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2779 | 28 | 0
2008-12-02 | 13:31 | 4.134 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0349 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3656 | 28 | 0
2008-12-02 | 13:31 | 4.13 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8578 | 0.0349 | 0.0011 | -0.000158 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.208 | 28 | 0
2008-12-02 | 13:31 | 4.127 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2744 | 28 | 0
2008-12-02 | 13:31 | 4.12 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8574 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1943 | 28 | 0
2008-12-02 | 13:31 | 4.12 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8577 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1313 | 28 | 0
2008-12-02 | 13:32 | 4.128 | 0.014 | 140 | 213 | 0 | 213 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.842 | 0.034942 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.2098 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.2 | 59.9 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4194 | 11.2883 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4185 | 11.2333 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.9 | 0 | 212.9 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8582 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.8 | 0 | 212.8 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8575 | 0.035 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.247 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.8 | 0 | 212.8 | 64.5 | 64.5466 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.034958 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.3953 | 28 | 0
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.8 | 0 | 212.8 | 64.5 | 64.5116 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1587 | 28 | 0
2008-12-02 | 13:32 | 4.124 | 0.014 | 140 | 212.7 | -0.042063 | 212.7 | 64.6 | 64.6 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8579 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.6 | -0.057374 | 212.6 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8574 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1374 | 28 | -0.1
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.5 | 0 | 212.5 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2158 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.5 | 0 | 212.5 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2943 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.5 | 0 | 212.5 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.842 | 0.0349 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2272 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.5 | 0 | 212.5 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0349 | 0.0011 | -0.000142 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1487 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | 0 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0349 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1297 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | 0 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0349 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | 0 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0349 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | 0 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0349 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | 0 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.034858 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | -0.041948 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | 0
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | -0.1 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8424 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2004 | 28 | -0.1
2008-12-02 | 13:32 | 4.12 | 0.014 | 140 | 212.4 | -0.1 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8578 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2789 | 28 | -0.1
2008-12-02 | 13:32 | 4.128 | 0.014 | 140 | 212.4 | -0.1 | 212.4 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.858 | 0.0348 | 0.0011 | -0.000158 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2427 | 28 | -0.1
2008-12-02 | 13:32 | 4.15 | 0.014 | 140 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.15 | 0.014 | 140 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2142 | 28 | -0.1
2008-12-02 | 13:32 | 4.142 | 0.014 | 140 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2927 | 28 | -0.1
2008-12-02 | 13:32 | 4.134 | 0.014 | 140 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2289 | 28 | -0.1
2008-12-02 | 13:32 | 4.135 | 0.0148 | 147.6815 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.14 | 0.015 | 150 | 212.3 | -0.1 | 212.3 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 4.6 | 195.7 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.139 | 0.015 | 150 | 212.2 | -0.1 | 212.2 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 4.4 | 196 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.131 | 0.015 | 150 | 212.2 | -0.1 | 212.2 | 64.5 | 64.5 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8419 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.145 | 0.015 | 150 | 212.2 | -0.1 | 212.2 | 64.5 | 64.5 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.138 | 0.015 | 150 | 212.1 | -0.1 | 212.1 | 64.5 | 64.4658 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:32 | 4.13 | 0.015 | 150 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0348 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1797 | 28 | -0.1
2008-12-02 | 13:32 | 4.13 | 0.0144 | 143.7706 | 212.2 | -0.1 | 212.2 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8419 | 0.0348 | 0.001 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1013 | 28 | -0.1
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.2 | -0.1 | 212.2 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0348 | 0.001 | -0.000142 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 4.13 | 0.014 | 140 | 212.2 | -0.1 | 212.2 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 4.138 | 0.014 | 140 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8578 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 4.14 | 0.014 | 140 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8422 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1125 | 28 | -0.1
2008-12-02 | 13:32 | 4.14 | 0.014 | 140 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.191 | 28 | -0.1
2008-12-02 | 13:32 | 4.137 | 0.014 | 140 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0348 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1306 | 28 | -0.1
2008-12-02 | 13:32 | 4.087 | 0.014 | 140 | 212.1 | -0.1 | 212.1 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1958 | 28 | -0.1
2008-12-02 | 13:32 | 3.769 | 0.0131 | 131.2776 | 211.6 | -0.1 | 211.6 | 64.4 | 64.4 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.034716 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2473 | 28 | -0.1
2008-12-02 | 13:32 | 3.558 | 0.011 | 110 | 209.1 | -0.1 | 209.1 | 64.4 | 64.4 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.034348 | 0.0011 | -0.0001 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.011 | 110 | 191.9 | -0.1 | 191.9 | 62.7 | 62.6787 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.032097 | 0.0011 | -0.000142 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.5806 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.0118 | 117.656 | 185.4 | -0.1 | 185.4 | 57 | 56.961 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8581 | 0.030858 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 185 | -0.1 | 185 | 55.5 | 55.4944 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0308 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013 | 11.1 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 184.7 | -0.1 | 184.7 | 55.2 | 55.1993 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8425 | 0.030758 | 0.0011 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4249 | 11.0814 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 183.7 | -0.1 | 183.7 | 55.1 | 55.1 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0306 | 0.0011 | -0.0002 | -0.000142 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0029 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 183.5 | -0.1 | 183.5 | 55.2 | 55.1996 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.030558 | 0.0011 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 183.2 | -0.1 | 183.2 | 55.1 | 55.1 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8579 | 0.0305 | 0.001 | -0.0002 | -0.000158 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.054 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 183.2 | -0.1 | 183.2 | 55.1 | 55.1 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0305 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0676 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.012 | 120 | 183.2 | -0.1 | 183.2 | 55.1 | 55.1 | 0.1 | 13.1 | 912 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.0305 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0108 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.0124 | 124.373 | 183.1 | -0.1 | 183.1 | 55.1 | 55.1 | 0.1 | 13.1 | 913 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0894 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.013 | 130 | 183 | -0.1 | 183 | 55.1 | 55.1 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9643 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.013 | 130 | 183 | -0.1 | 183 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8581 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 10.9921 | 28 | -0.1
2008-12-02 | 13:32 | 3.5 | 0.013 | 130 | 183 | -0.1 | 183 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0756 | 28 | -0.1
2008-12-02 | 13:32 | 3.504 | 0.013 | 130 | 182.9 | -0.1 | 182.9 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.8 | 59.8584 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0032 | 28 | -0.1
2008-12-02 | 13:32 | 3.51 | 0.013 | 130 | 182.9 | -0.1 | 182.9 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0816 | 28 | -0.1
2008-12-02 | 13:32 | 3.51 | 0.013 | 130 | 182.9 | -0.1 | 182.9 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1601 | 28 | -0.1
2008-12-02 | 13:32 | 3.51 | 0.013 | 130 | 182.9 | -0.1 | 182.9 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8579 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1615 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.9 | -0.1 | 182.9 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.8 | -0.1 | 182.8 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.506 | 0.013 | 130 | 182.8 | -0.1 | 182.8 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0304 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.174 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.7 | -0.1 | 182.7 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0951 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.7 | -0.1 | 182.7 | 55.2 | 55.2 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0614 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.7 | -0.1 | 182.7 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1912 | 28 | -0.1
2008-12-02 | 13:33 | 3.508 | 0.013 | 130 | 182.7 | -0.1 | 182.7 | 55.2 | 55.2 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1123 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.6 | -0.1 | 182.6 | 55.2 | 55.2 | 0.1 | 13.1 | 912 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.000958 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0338 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.6 | -0.1 | 182.6 | 55.2 | 55.2 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0303 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0447 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.6 | -0.14218 | 182.6 | 55.2 | 55.2 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4218 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.0126 | 125.6466 | 182.6 | -0.2 | 182.6 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.8 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0986 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.5 | -0.1 | 182.5 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8581 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.02 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.5 | -0.1 | 182.5 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8424 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.5758 | 11.0586 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.5 | -0.2 | 182.5 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.5 | -0.2 | 182.5 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0846 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.4 | -0.2 | 182.4 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8581 | 0.030258 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0067 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.4 | -0.2 | 182.4 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.5 | 59.9 | 0.0302 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.4 | -0.2 | 182.4 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0302 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0508 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.4 | -0.2 | 182.4 | 55.1 | 55.1079 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0302 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.3 | -0.2 | 182.3 | 55.2 | 55.1901 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8422 | 0.030242 | 0.0009 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 3.506 | 0.0128 | 127.6118 | 182.3 | -0.2 | 182.3 | 55.1 | 55.1 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.4 | -0.2 | 182.4 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0354 | 28 | -0.2
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.4 | -0.2 | 182.4 | 55.1 | 55.1 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.4 | -0.1 | 182.4 | 55.2 | 55.1874 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.1
2008-12-02 | 13:33 | 3.508 | 0.0129 | 128.9132 | 182.3 | -0.1 | 182.3 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0002 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.0121 | 120.8029 | 182.3 | -0.1 | 182.3 | 55.1 | 55.1 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1558 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.012 | 120 | 182.3 | -0.1 | 182.3 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.2 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.0126 | 125.876 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8418 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1292 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8425 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0139 | 28 | -0.1
2008-12-02 | 13:33 | 3.51 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1059 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8423 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.5771 | 11.0923 | 28 | -0.1
2008-12-02 | 13:33 | 3.502 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0001 |  | n/a |  | 35 | 61 | 1052 | 1013.4218 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.8419 | 0.0302 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.1 | 55.1 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0302 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1 | 28 | -0.1
2008-12-02 | 13:33 | 3.5 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8578 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0938 | 28 | -0.1
2008-12-02 | 13:33 | 3.499 | 0.013 | 130 | 182.2 | -0.1 | 182.2 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0154 | 28 | -0.1
2008-12-02 | 13:33 | 3.491 | 0.013 | 130 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8581 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.1
2008-12-02 | 13:33 | 3.498 | 0.013 | 130 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.8419 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.1
2008-12-02 | 13:33 | 3.494 | 0.013 | 130 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.04 | 28 | -0.1
2008-12-02 | 13:33 | 3.49 | 0.013 | 130 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.197 | 28 | -0.1
2008-12-02 | 13:33 | 3.49 | 0.0124 | 123.995 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0461 | 28 | -0.1
2008-12-02 | 13:33 | 3.438 | 0.012 | 120 | 182.1 | -0.1 | 182.1 | 55.2 | 55.2 | 0.1 | 13.2 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8422 | 0.0303 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11 | 28 | -0.1
2008-12-02 | 13:33 | 3.05 | 0.0115 | 114.8925 | 181.3 | -0.2 | 181.3 | 55.2 | 55.2 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.030174 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0677 | 28 | -0.2
2008-12-02 | 13:33 | 2.911 | 0.0101 | 101.0815 | 177.6 | -0.2 | 177.6 | 55.2 | 55.2 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.029581 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1752 | 28 | -0.2
2008-12-02 | 13:33 | 2.88 | 0.0109 | 109.1525 | 159.9 | -0.1 | 159.9 | 51.6 | 51.6191 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8424 | 0.026888 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.0184 | 28 | -0.1
2008-12-02 | 13:33 | 2.884 | 0.011 | 110 | 154.9 | -0.1 | 154.9 | 47.6 | 47.6003 | 0.1 | 13.1 | 907 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.026158 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1052 | 1014 | 11.1385 | 28 | -0.1
2008-12-02 | 13:33 | 2.893 | 0.0104 | 104.1475 | 154.7 | -0.1 | 154.7 | 46.3 | 46.2896 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0261 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1046 | 28 | -0.1
2008-12-02 | 13:33 | 2.898 | 0.01 | 100 | 154.5 | -0.1 | 154.5 | 46.1 | 46.1 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.026058 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5806 | 11 | 28 | -0.1
2008-12-02 | 13:33 | 2.882 | 0.01 | 100 | 154.4 | -0.1 | 154.4 | 46.1 | 46.1 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8417 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0046 | 28 | -0.1
2008-12-02 | 13:33 | 2.88 | 0.01 | 100 | 154.4 | -0.1 | 154.4 | 46.1 | 46.1 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0831 | 28 | -0.1
2008-12-02 | 13:33 | 2.892 | 0.01 | 100 | 154.3 | -0.2 | 154.3 | 46 | 46 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.026 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 2.9 | 0.01 | 100 | 153.8 | -0.2 | 153.8 | 46.1 | 46.1 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0259 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 2.9 | 0.01 | 100 | 153.7 | -0.2 | 153.7 | 46.1 | 46.094 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 196 | 59.9 | 0.0259 | 0.001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0815 | 28 | -0.2
2008-12-02 | 13:33 | 2.899 | 0.01 | 100 | 153.6 | -0.2 | 153.6 | 46 | 46 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8421 | 0.0259 | 0.000958 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.003 | 28 | -0.2
2008-12-02 | 13:33 | 2.891 | 0.01 | 100 | 153.5 | -0.2 | 153.5 | 46 | 46 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0258 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0755 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.5 | -0.2 | 153.5 | 46 | 46 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0258 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.4 | -0.2 | 153.4 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0258 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.3 | -0.2 | 153.3 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0258 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1108 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.3 | -0.2 | 153.3 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8 | 0.025758 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1893 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.2 | -0.2 | 153.2 | 46.1 | 46.095 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0646 | 28 | -0.2
2008-12-02 | 13:33 | 2.89 | 0.01 | 100 | 153.1 | -0.2 | 153.1 | 46 | 46.0066 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0459 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.01 | 100 | 153.1 | -0.2 | 153.1 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8577 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.01 | 100 | 153 | -0.2 | 153 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.886 | 0.01 | 100 | 153 | -0.2 | 153 | 46 | 46 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8579 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.01 | 100 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2202 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.0104 | 104.2926 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8576 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1847 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.011 | 110 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.011 | 110 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.011 | 110 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0739 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.011 | 110 | 152.9 | -0.2 | 152.9 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1524 | 28 | -0.2
2008-12-02 | 13:34 | 2.88 | 0.011 | 110 | 152.8 | -0.2 | 152.8 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8418 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1388 | 28 | -0.2
2008-12-02 | 13:34 | 2.881 | 0.011 | 110 | 152.8 | -0.2 | 152.8 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8425 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.889 | 0.011 | 110 | 152.8 | -0.2 | 152.8 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.7 | -0.2 | 152.7 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0662 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.7 | -0.2 | 152.7 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8421 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0553 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.7 | -0.2 | 152.7 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.7 | -0.2 | 152.7 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0013 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.6 | -0.2 | 152.6 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0257 | 0.0009 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.08 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.6 | -0.2 | 152.6 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0256 | 0.000858 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0415 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.5 | -0.2 | 152.5 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.896 | 0.011 | 110 | 152.5 | -0.2 | 152.5 | 45.9 | 45.9064 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8417 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0155 | 28 | -0.2
2008-12-02 | 13:34 | 2.9 | 0.0102 | 102.3353 | 152.5 | -0.2 | 152.5 | 46 | 46 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0939 | 28 | -0.2
2008-12-02 | 13:34 | 2.9 | 0.01 | 100 | 152.5 | -0.2 | 152.5 | 46 | 46 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0276 | 28 | -0.2
2008-12-02 | 13:34 | 2.899 | 0.01 | 100 | 152.5 | -0.2 | 152.5 | 46 | 46 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.891 | 0.0102 | 102.4957 | 152.5 | -0.2 | 152.5 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0292 | 28 | -0.2
2008-12-02 | 13:34 | 2.898 | 0.011 | 110 | 152.6 | -0.2 | 152.6 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0853 | 28 | -0.2
2008-12-02 | 13:34 | 2.894 | 0.011 | 110 | 152.6 | -0.2 | 152.6 | 46 | 46 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 10.9285 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.5 | -0.2 | 152.5 | 45.9 | 45.911 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 10.9645 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.5 | -0.2 | 152.5 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0256 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.011 | 110 | 152.4 | -0.2 | 152.4 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.025558 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.0105 | 105.0811 | 152.4 | -0.2 | 152.4 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.009 | 90 | 152.4 | -0.2 | 152.4 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.009 | 90 | 152.4 | -0.2 | 152.4 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8423 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1139 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.0098 | 97.5815 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8422 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.129 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.01 | 100 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.9 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4283 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.891 | 0.01 | 100 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.899 | 0.01 | 100 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.071 | 28 | -0.2
2008-12-02 | 13:34 | 2.9 | 0.01 | 100 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0506 | 28 | -0.2
2008-12-02 | 13:34 | 2.894 | 0.01 | 100 | 152.3 | -0.2 | 152.3 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.858 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0279 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.01 | 100 | 152.2 | -0.2 | 152.2 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0936 | 28 | -0.2
2008-12-02 | 13:34 | 2.89 | 0.01 | 100 | 152.2 | -0.2 | 152.2 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0255 | 0.0008 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0152 | 28 | -0.2
2008-12-02 | 13:34 | 2.841 | 0.0092 | 91.521 | 152.2 | -0.2 | 152.2 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8582 | 0.0255 | 0.000758 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5823 | 11.0633 | 28 | -0.2
2008-12-02 | 13:34 | 2.418 | 0.0085 | 84.8649 | 152.2 | -0.2 | 152.2 | 45.9 | 45.9 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8425 | 0.0255 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0583 | 28 | -0.2
2008-12-02 | 13:34 | 2.319 | 0.01 | 100 | 142.2 | -0.2 | 142.2 | 45.9 | 45.8875 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.023946 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4229 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 128.6 | -0.2 | 128.6 | 42.1 | 42.1342 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.021847 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124.5 | -0.2 | 124.5 | 37.9 | 37.8676 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.021158 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0771 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124.3 | -0.2 | 124.3 | 37 | 36.9983 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0211 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0444 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124.1 | -0.2 | 124.1 | 36.8 | 36.8 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195.2 | 59.9 | 0.0211 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124.1 | -0.2 | 124.1 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0211 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0126 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124 | -0.2 | 124 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0211 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0911 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 124 | -0.2 | 124 | 36.7 | 36.7065 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0307 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 123.9 | -0.2 | 123.9 | 36.7 | 36.7 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 10.9522 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 123.9 | -0.2 | 123.9 | 36.7 | 36.7 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8427 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 10.9529 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 123.8 | -0.2 | 123.8 | 36.8 | 36.7898 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.9 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.01 | 100 | 123.7 | -0.2 | 123.7 | 36.7 | 36.7 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.0099 | 99.0732 | 123.7 | -0.2 | 123.7 | 36.7 | 36.7344 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.0091 | 90.9431 | 123.6 | -0.2 | 123.6 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8579 | 0.021 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.020958 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8579 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5786 | 11.0243 | 28 | -0.2
2008-12-02 | 13:34 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0322 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.009 | 90 | 123.5 | -0.2 | 123.5 | 36.8 | 36.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4177 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.009 | 90 | 123.4 | -0.2 | 123.4 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4249 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.295 | 0.0078 | 77.99 | 123.4 | -0.2 | 123.4 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.28 | 0.0074 | 74.2255 | 123.4 | -0.2 | 123.4 | 36.8 | 36.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0536 | 28 | -0.2
2008-12-02 | 13:35 | 2.28 | 0.008 | 80 | 123.4 | -0.2 | 123.4 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195.9 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4206 | 10.9752 | 28 | -0.2
2008-12-02 | 13:35 | 2.288 | 0.008 | 80 | 123.3 | -0.2 | 123.3 | 36.8 | 36.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 10.9035 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.3 | -0.2 | 123.3 | 36.8 | 36.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 10.982 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.3 | -0.2 | 123.3 | 36.8 | 36.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0209 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.2 | -0.2 | 123.2 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0209 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.2 | -0.2 | 123.2 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0209 | 0.000642 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.2 | -0.2 | 123.2 | 36.8 | 36.8 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8 | 0.0209 | 0.0007 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.2 | -0.2 | 123.2 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0209 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 10.9257 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8425 | 0.020858 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0057 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.2 | -0.2 | 123.2 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8423 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0689 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.8 | 36.8 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8421 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.8 | 36.8 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8581 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.7 | 36.7064 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8424 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0667 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.8 | 36.7923 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0548 | 28 | -0.2
2008-12-02 | 13:35 | 2.296 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.7 | 36.7 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.2 | 59.8 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5798 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.6 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0038 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.7 | 36.7 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.161 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 123.1 | -0.2 | 123.1 | 36.7 | 36.7 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.141 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 123 | -0.2 | 123 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0625 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.016 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8422 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0945 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.7 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0271 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5758 | 11.0299 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.8 | 59.8 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4185 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.3 | 0.008 | 80 | 122.9 | -0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8418 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.298 | 0.0103 | 103.2216 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.078 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.0116 | 115.9383 | 122.9 | -0.2 | 122.9 | 36.7 | 36.694 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.0208 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0001 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.0108 | 107.61 | 122.9 | -0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.9 | 0.020758 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.079 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.01 | 100 | 122.9 | -0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8 | 0.0207 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.01 | 100 | 122.9 | -0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8419 | 0.0207 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 2.29 | 0.0092 | 92.4697 | 122.8 | -0.2 | 122.8 | 36.6 | 36.6 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0207 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5758 | 11.1146 | 28 | -0.2
2008-12-02 | 13:35 | 2.252 | 0.009 | 90 | 122.9 | -0.2 | 122.9 | 36.6 | 36.6 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8422 | 0.0207 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1929 | 28 | -0.2
2008-12-02 | 13:35 | 2.099 | 0.0082 | 82.078 | 122.9 | -0.2 | 122.9 | 36.7 | 36.7 | 0.1 | 13.1 | 911 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.858 | 0.0207 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1285 | 28 | -0.2
2008-12-02 | 13:35 | 1.756 | 0.0075 | 74.8363 | 112.3 | -0.2 | 112.3 | 36.6 | 36.6 | 0.1 | 13.1 | 911 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195 | 59.8581 | 0.019019 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.15 | 28 | -0.2
2008-12-02 | 13:35 | 1.71 | 0.009 | 90 | 100.6 | -0.2 | 100.6 | 34.1 | 34.0793 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0172 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1715 | 28 | -0.2
2008-12-02 | 13:35 | 1.704 | 0.009 | 90 | 96.7 | -0.2 | 96.7 | 29.5 | 29.5055 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.016616 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.107 | 28 | -0.2
2008-12-02 | 13:35 | 1.704 | 0.0076 | 75.5713 | 95.2 | -0.2 | 95.2 | 28 | 27.9893 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8425 | 0.016315 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1853 | 28 | -0.2
2008-12-02 | 13:35 | 1.705 | 0.007 | 70 | 94.5 | -0.2 | 94.5 | 27.6 | 27.6 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0162 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.2
2008-12-02 | 13:35 | 1.69 | 0.007 | 70 | 94.2 | -0.2 | 94.2 | 27.6 | 27.5796 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.016158 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4218 | 11.1151 | 28 | -0.2
2008-12-02 | 13:35 | 1.69 | 0.007 | 70 | 94.1 | -0.2 | 94.1 | 27.5 | 27.5 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0209 | 28 | -0.2
2008-12-02 | 13:35 | 1.69 | 0.007 | 70 | 94.1 | -0.2 | 94.1 | 27.5 | 27.5 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0994 | 28 | -0.2
2008-12-02 | 13:35 | 1.696 | 0.007 | 70 | 94 | -0.2 | 94 | 27.5 | 27.5 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8576 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2545 | 28 | -0.2
2008-12-02 | 13:35 | 1.7 | 0.007 | 70 | 93.9 | -0.2 | 93.9 | 27.5 | 27.5 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8422 | 0.0161 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1881 | 28 | -0.2
2008-12-02 | 13:35 | 1.7 | 0.007 | 70 | 93.9 | -0.2 | 93.9 | 27.5 | 27.5 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0161 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.2
2008-12-02 | 13:35 | 1.7 | 0.007 | 70 | 93.9 | -0.3 | 93.9 | 27.5 | 27.5 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8581 | 0.0161 | 0.000542 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1265 | 28 | -0.3
2008-12-02 | 13:35 | 1.7 | 0.007 | 70 | 93.8 | -0.25815 | 93.8 | 27.5 | 27.5 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8419 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2834 | 28 | -0.3
2008-12-02 | 13:35 | 1.7 | 0.007 | 70 | 93.8 | -0.2 | 93.8 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2301 | 28 | -0.2
2008-12-02 | 13:35 | 1.706 | 0.007 | 70 | 93.8 | -0.3 | 93.8 | 27.5 | 27.4983 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1515 | 28 | -0.3
2008-12-02 | 13:35 | 1.71 | 0.007 | 70 | 93.8 | -0.3 | 93.8 | 27.4 | 27.4 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1272 | 28 | -0.3
2008-12-02 | 13:35 | 1.71 | 0.007 | 70 | 93.8 | -0.3 | 93.8 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8578 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.3
2008-12-02 | 13:35 | 1.71 | 0.007 | 70 | 93.8 | -0.3 | 93.8 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0161 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.3
2008-12-02 | 13:35 | 1.71 | 0.007 | 70 | 93.7 | -0.3 | 93.7 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8582 | 0.016058 | 0.000558 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.0069 | 69.2445 | 93.7 | -0.3 | 93.7 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.241 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.0061 | 61.121 | 93.7 | -0.3 | 93.7 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2805 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.7 | -0.34193 | 93.7 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.202 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.6 | -0.4 | 93.6 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2765 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.6 | -0.4 | 93.6 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8573 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.6 | -0.4 | 93.6 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4249 | 11.3 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.5 | -0.3 | 93.5 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2761 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.4 | -0.4 | 93.4 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.015958 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1192 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.4 | -0.4 | 93.4 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8579 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5794 | 11.1689 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.3 | -0.4 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.3 | -0.4 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1742 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.3 | -0.35778 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8578 | 0.015942 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5778 | 11.1085 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2655 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1776 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.0067 | 67.1547 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8417 | 0.015942 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8425 | 0.016 | 0.000542 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4249 | 11.1182 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.016 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1967 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.3 | -0.3 | 93.3 | 27.4 | 27.4 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8422 | 0.016 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.2 | -0.3 | 93.2 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.016 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.2 | -0.34263 | 93.2 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8574 | 0.016 | 0.0006 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1679 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.2 | -0.4 | 93.2 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.016 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.2 | -0.34222 | 93.2 | 27.4 | 27.4 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8578 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5778 | 11.1 | 28 | -0.3
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.2 | -0.4 | 93.2 | 27.4 | 27.3683 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.842 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4202 | 11.1675 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.0072 | 72.406 | 93.1 | -0.4 | 93.1 | 27.4 | 27.393 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2459 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.0079 | 79.2632 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2514 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.0071 | 71.166 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5823 | 11.103 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1815 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8423 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1401 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1763 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8579 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2672 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93.1 | -0.4 | 93.1 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1093 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.007 | 70 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0262 | 28 | -0.4
2008-12-02 | 13:36 | 1.702 | 0.007 | 70 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11 | 28 | -0.4
2008-12-02 | 13:36 | 1.692 | 0.007 | 70 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0612 | 28 | -0.4
2008-12-02 | 13:36 | 1.7 | 0.0066 | 66.1116 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5815 | 11.1912 | 28 | -0.4
2008-12-02 | 13:36 | 1.7 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.015857 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1123 | 28 | -0.4
2008-12-02 | 13:36 | 1.714 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0158 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1662 | 28 | -0.4
2008-12-02 | 13:36 | 1.715 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.8423 | 0.015842 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1553 | 28 | -0.4
2008-12-02 | 13:36 | 1.71 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.4
2008-12-02 | 13:36 | 1.712 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8579 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.4
2008-12-02 | 13:36 | 1.719 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.4
2008-12-02 | 13:36 | 1.711 | 0.006 | 60 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1586 | 28 | -0.4
2008-12-02 | 13:36 | 1.398 | 0.0061 | 60.7779 | 93 | -0.4 | 93 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0159 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.163 | 28 | -0.4
2008-12-02 | 13:36 | 1.2 | 0.0069 | 68.7971 | 90.9 | -0.4 | 90.9 | 27.3 | 27.3 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.842 | 0.015564 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0846 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.0049 | 48.75 | 73.5 | -0.4 | 73.5 | 25.8 | 25.8312 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.012719 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0067 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 66 | -0.4 | 66 | 20.5 | 20.5067 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.011458 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1446 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 65.4 | -0.4 | 65.4 | 18.7 | 18.7385 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8582 | 0.0114 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.149 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 65.1 | -0.4 | 65.1 | 18.3 | 18.3 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0113 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.0705 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.7 | -0.4 | 64.7 | 18.3 | 18.2938 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0113 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.0159 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.7 | -0.4 | 64.7 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.011258 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1721 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.6 | -0.4 | 64.6 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0112 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.6 | -0.4 | 64.6 | 18.2 | 18.2 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0112 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1566 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.3 | -0.4 | 64.3 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.011158 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64.1 | -0.4 | 64.1 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.8 | 0.0111 | 0.000458 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1003 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0111 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1788 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.004 | 40 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8581 | 0.0111 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1427 | 28 | -0.4
2008-12-02 | 13:36 | 1.144 | 0.004 | 40 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0111 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1717 | 28 | -0.4
2008-12-02 | 13:36 | 1.148 | 0.004 | 40 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8418 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2714 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.0044 | 43.8515 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8423 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1144 | 28 | -0.4
2008-12-02 | 13:36 | 1.14 | 0.005 | 50 | 64 | -0.4 | 64 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1713 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2282 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8576 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2867 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1301 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2085 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.842 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1635 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2479 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8582 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2421 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2587 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.9 | -0.4 | 63.9 | 18.2 | 18.2 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1017 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.8 | -0.4 | 63.8 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.8 | -0.4 | 63.8 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8423 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.8 | -0.4 | 63.8 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8422 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1345 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.8 | -0.4 | 63.8 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 196 | 59.8421 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5794 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.0049 | 49.3076 | 63.8 | -0.4 | 63.8 | 18.2 | 18.2 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.0041 | 41.256 | 63.8 | -0.4 | 63.8 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.3 | 59.8 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.061 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195 | 59.858 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4202 | 11.0481 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1269 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1946 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8573 | 0.0111 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1161 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.9 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 1.135 | 0.004 | 40 | 63.7 | -0.4 | 63.7 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.9 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1406 | 28 | -0.4
2008-12-02 | 13:37 | 1.134 | 0.004 | 40 | 63.6 | -0.4 | 63.6 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8578 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.6 | -0.4 | 63.6 | 18.3 | 18.3 | 0.1 | 13.2 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8579 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.0044 | 43.8248 | 63.6 | -0.44187 | 63.6 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 196 | 59.8419 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1237 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.005 | 50 | 63.5 | -0.5 | 63.5 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8576 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1547 | 28 | -0.5
2008-12-02 | 13:37 | 1.14 | 0.005 | 49.6589 | 63.5 | -0.4 | 63.5 | 18.3 | 18.3 | 0.1 | 13.2 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.3 | 59.9 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.0041 | 41.3394 | 63.5 | -0.4 | 63.5 | 18.3 | 18.3 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.14 | 0.004 | 40 | 63.5 | -0.4 | 63.5 | 18.3 | 18.3 | 0.1 | 13.1 | 910 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.138 | 0.004 | 40 | 63.4 | -0.4 | 63.4 | 18.3 | 18.3 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1313 | 28 | -0.4
2008-12-02 | 13:37 | 1.131 | 0.0051 | 50.975 | 63.4 | -0.4 | 63.4 | 18.3 | 18.3 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8418 | 0.011 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 1.139 | 0.007 | 70 | 63.4 | -0.4 | 63.4 | 18.3 | 18.294 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.5 | 59.8 | 0.011 | 0.000442 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1503 | 28 | -0.4
2008-12-02 | 13:37 | 1.133 | 0.0069 | 69.3559 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.0061 | 61.3043 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8579 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.006 | 60 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8581 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1778 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.006 | 60 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1396 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.006 | 60 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8422 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.006 | 60 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.006 | 59.6833 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.125 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.0051 | 51.35 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1536 | 28 | -0.4
2008-12-02 | 13:37 | 1.13 | 0.005 | 50 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8582 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1679 | 28 | -0.4
2008-12-02 | 13:37 | 1.053 | 0.005 | 50 | 63.3 | -0.4 | 63.3 | 18.2 | 18.2 | 0.1 | 13.1 | 911 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.011 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 0.741 | 0.0035 | 35.4667 | 62.3 | -0.4 | 62.3 | 18.2 | 18.2 | 0.1 | 13.1 | 910 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.010827 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4323 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 0.598 | 0.0012 | 12.038 | 48.1 | -0.44218 | 48.1 | 18.2 | 18.1981 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.008499 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1 | 28 | -0.4
2008-12-02 | 13:37 | 0.573 | 0.002 | 20 | 37 | -0.5 | 37 | 15.3 | 15.341 | 0.1 | 13.1 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.006674 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5794 | 11.1916 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 34.9 | -0.45805 | 34.9 | 10.2 | 10.1515 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.8419 | 0.006374 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2758 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.9 | -0.4 | 33.9 | 9.2 | 9.1853 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8424 | 0.0061 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5758 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.8 | -0.4 | 33.8 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0061 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.6 | -0.44308 | 33.6 | 8.9 | 8.9 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.006057 | 0.0005 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.4
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.3 | -0.5 | 33.3 | 8.9 | 8.9 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.006 | 0.000458 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.3 | -0.5 | 33.3 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.006 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.5
2008-12-02 | 13:37 | 0.575 | 0.002 | 20 | 33.3 | -0.5 | 33.3 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8575 | 0.006 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.5
2008-12-02 | 13:37 | 0.576 | 0.002 | 20 | 33.2 | -0.5 | 33.2 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.006 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1262 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.2 | -0.5 | 33.2 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.857 | 0.0059 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33.1 | -0.5 | 33.1 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0059 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1612 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.002 | 20 | 33 | -0.5 | 33 | 8.9 | 8.9 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0059 | 0.0004 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.0021 | 20.625 | 33 | -0.5 | 33 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0059 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.5
2008-12-02 | 13:37 | 0.581 | 0.0029 | 28.6378 | 32.9 | -0.5 | 32.9 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0058 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2339 | 28 | -0.5
2008-12-02 | 13:37 | 0.59 | 0.003 | 30 | 32.9 | -0.5 | 32.9 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0058 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2446 | 28 | -0.5
2008-12-02 | 13:37 | 0.585 | 0.003 | 30 | 32.9 | -0.5 | 32.9 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0058 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.003 | 30 | 32.9 | -0.5 | 32.9 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0058 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3009 | 28 | -0.5
2008-12-02 | 13:37 | 0.57 | 0.003 | 30 | 32.8 | -0.5 | 32.8 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.005758 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3799 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 29.7417 | 32.8 | -0.5 | 32.8 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8417 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3414 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.0021 | 21.4083 | 32.8 | -0.5 | 32.8 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.8577 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5771 | 11.263 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.8 | -0.5 | 32.8 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.5 | 59.8 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5782 | 11.2155 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.9 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2937 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3509 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.9 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3411 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0057 | 0.0003 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2079 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2864 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.7 | -0.5 | 32.7 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3649 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.6 | -0.5 | 32.6 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3566 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.6 | -0.5 | 32.6 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8425 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5751 | 11.3 | 28 | -0.5
2008-12-02 | 13:38 | 0.563 | 0.002 | 20 | 32.6 | -0.5 | 32.6 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8423 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3003 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.6 | -0.5 | 32.6 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8571 | 0.0057 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3787 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.6 | -0.5 | 32.6 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.9 | 0.005658 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3428 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.5 | -0.5 | 32.5 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.8 | 59.8 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.5 | -0.5 | 32.5 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2858 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.5 | -0.5 | 32.5 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.1 | 59.858 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2073 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.4 | -0.5 | 32.4 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.5
2008-12-02 | 13:38 | 0.56 | 0.002 | 20 | 32.4 | -0.5 | 32.4 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8581 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.5
2008-12-02 | 13:38 | 0.564 | 0.002 | 20 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8418 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2563 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.0025 | 25.2353 | 32.4 | -0.5 | 32.4 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8575 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3934 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.4 | 28 | 5 | 195 | 59.8423 | 0.0056 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.315 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8422 | 0.0056 | 0.000158 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3635 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.005557 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.358 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3205 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.3 | -0.5 | 32.3 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8424 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.399 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.003 | 30 | 32.2 | -0.54222 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2457 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.0027 | 26.5221 | 32.2 | -0.6 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2558 | 28 | -0.6
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.6 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8581 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2311 | 28 | -0.6
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.55823 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8582 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.6
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8425 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3092 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8578 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4218 | 11.3035 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2268 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.2 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3053 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.9 | 8.9 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8576 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5758 | 11.3838 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.5 | 32.2 | 8.8 | 8.8 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4222 | 11.4 | 28 | -0.5
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.6 | 32.2 | 8.9 | 8.9 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.858 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2792 | 28 | -0.6
2008-12-02 | 13:38 | 0.57 | 0.002 | 20 | 32.2 | -0.6 | 32.2 | 8.9 | 8.8878 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8419 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.6
2008-12-02 | 13:38 | 0.552 | 0.002 | 20 | 32.2 | -0.6 | 32.2 | 8.9 | 8.9 | 0.1 | 13.2 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0055 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1 | 28 | -0.6
2008-12-02 | 13:38 | 0.13 | 0.0011 | 11.1073 | 27.2 | -0.6 | 27.2 | 8.9 | 8.9 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.004646 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1762 | 28 | -0.6
2008-12-02 | 13:38 | 0.016 | -0.0019 | -19.304 | 11 | -0.6 | 11 | 8.7 | 8.6591 | 0.1 | 13.1 | 907 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.4 | 28 | 5 | 195 | 59.8425 | 0.001933 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.6
2008-12-02 | 13:38 | 0 | -0.0011 | -11.304 | 5.5 | -0.6 | 5.5 | 3.8 | 3.8395 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8577 | 0.001004 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2327 | 28 | -0.6
2008-12-02 | 13:38 | 0 | -0.001 | -10 | 2.2 | -0.6 | 2.2 | 0.6 | 0.5921 | 0.1 | 13.1 | 905 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.000458 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2888 | 28 | -0.6
2008-12-02 | 13:38 | 0 | -0.001 | -10 | 2.1 | -0.6 | 2.1 | -0.3 | 0 | 0.1 | 13.2 | 904 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0004 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.21 | 28 | -0.6
2008-12-02 | 13:38 | 0 | -0.0006 | -6.2 | 1.9 | -0.6 | 1.9 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.5 | 59.8424 | 0.0004 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.4 | -0.6 | 1.4 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4222 | 11.2939 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.3 | -0.6 | 1.3 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.8 | 59.9 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4202 | 11.3746 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.2 | -0.6 | 1.2 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.9 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.296 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.2 | -0.6 | 1.2 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.1 | 59.9 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2175 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.2 | -0.6 | 1.2 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.2 | -0.6 | 1.2 | -0.5 | 0 | 0.1 | 13.2 | 905 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8575 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4249 | 11.2 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.2 | -0.64229 | 1.2 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195 | 59.8423 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5771 | 11.2172 | 28 | -0.6
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.1 | -0.7 | 1.1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.4 | 59.8578 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2964 | 28 | -0.7
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.1 | -0.7 | 1.1 | -0.5 | 0 | 0.1 | 13.1 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.7 | 59.8421 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3748 | 28 | -0.7
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.1 | -0.7 | 1.1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 192 | 28 | 5 | 195.9 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -0.7
2008-12-02 | 13:38 | 0 | 0 | 0 | 1.1 | -0.7 | 1.1 | -0.5 | 0 | 0.1 | 13.1 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191.6 | 28 | 5 | 195.5 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -0.7
2008-12-02 | 13:38 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.8578 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -0.7
2008-12-02 | 13:38 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.9 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.4 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8582 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2655 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2458 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8573 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 1 | -0.7 | 1 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8417 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.9 | -0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195 | 59.8423 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.9 | -0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.4 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2404 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.9 | -0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8421 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.238 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.9 | -0.7 | 0.9 | -0.5 | 0 | 0.1 | 13.2 | 906 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.0003 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2672 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.9 | -0.74242 | 0.9 | -0.5 | 0 | 0.1 | 13.1 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.5 | 59.8 | 0.000258 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1098 | 28 | -0.7
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.8 | -0.8 | 0.8 | -0.5 | 0 | 0.1 | 13.2 | 907 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.1 | 59.8 | 0.0002 | 0.0001 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2471 | 28 | -0.8
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.8 | -0.8 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.2 | 59.9 | 0.0003 | 0.000142 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1959 | 28 | -0.8
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.8 | -0.8 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.6 | 59.9 | 0.000342 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.5806 | 11.1609 | 28 | -0.8
2008-12-02 | 13:39 | 0 | -0.0001 | -0.61142 | 0.8 | -0.84185 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.9 | 59.8 | 0.0004 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.8
2008-12-02 | 13:39 | 0 | -0.0009 | -8.6565 | 0.8 | -0.9 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 191 | 28 | 5 | 195.7 | 59.8426 | 0.000357 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.8 | -0.9 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190.6 | 28 | 5 | 195.3 | 59.8423 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2341 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.8 | -0.9 | 0.8 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2444 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.0006 | -6.3947 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8421 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2987 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8576 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2201 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8422 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2584 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.5 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013.4177 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8575 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1986 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.1292 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.0005 | -5.4498 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2645 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8423 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.858 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2784 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 908 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.8419 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.7 | -0.9 | 0.7 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.2 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1013 | 11.2779 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.001 | -10 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.6 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2434 | 28 | -0.9
2008-12-02 | 13:39 | 0 | -0.0006 | -6.4578 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.9 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1645 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 908 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1278 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.3 | 59.8422 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2848 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.3 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 952 | 941 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.4 | 59.9 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2509 | 28 | -0.9
2008-12-02 | 13:39 | -0.002 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 952 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 195.7 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2 | 28 | -0.9
2008-12-02 | 13:39 | -0.01 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8422 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2063 | 28 | -0.9
2008-12-02 | 13:39 | -0.001 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 909 | 953 | 941 | 0 | 0 | 0 | 990 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.858 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.2848 | 28 | -0.9
2008-12-02 | 13:39 | 0 | 0 | 0 | 0.6 | -0.9 | 0.6 | -0.4 | 0 | 0.1 | 13.1 | 908 | 953 | 940 | 0 | 0 | 0 | 991 | -40 | 0 | 28 | 190 | 28 | 5 | 196 | 59.8 | 0.0003 | 0.0002 | -0.0002 | -0.0002 |  | n/a |  | 35 | 61 | 1051 | 1014 | 11.1736 | 28 | -0.9
Summary Information:
Post Processor DLL Version | 6.01
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-12-02
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 66
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 28261
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 2005
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 131
Version | 1.611
NO Span(ppm) | 302
NO2 Span(ppm) | 519
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.00
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | Backhoe J Deere MN 700212 310G
License Plate
Engine Displacement | 4.5
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 1
Torque (ecm or calc) | ecm
Use Frictional Torque | YES
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 13:18
Test End Time | 13:39
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 990.966
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iGPS_GROUND_SPEED
iENG_SPEED_USED | iAMBII_RPM
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
Faults:
0X5009 - 12/02/2008 11:06:44.976 - FID battery backed RAM corrupt
Warnings:
0X0000 - 12/02/2008 13:18:10.391 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 124
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-12-02 | 13:18 | [RELEASE_VER=2.018 BUILD=66 BDATE=10/21/2008 IP=10.10.1.55]
No Information Available

## THC Lin
Test Date:
Test Time:
12/2/2008
13:18:10
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## CO2 Lin
Test Date:
Test Time:
12/2/2008
13:18:10
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## CO Lin
Test Date:
Test Time:
12/2/2008
13:18:10
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## NO Lin
Test Date:
Test Time:
12/2/2008
13:18:10
NDUV Serial No:
131
SEMTECH Serial Number:
G05-SDS02

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
12/2/2008
13:18:10
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
12/2/2008
13:18:10
NDUV Serial No:
131
AMBII Serial No:
28261
SEMTECH Serial Number:
G05-SDS02

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6 | Rm = | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277
| C.F.= | 0.099349 | 0.19884 | 0.29848 | 0.39826 | 0.49819 | 0.59826 | 0.69847 | 0.79884 | 0.89935
| gas = | 0.59609 | 1.193 | 1.7909 | 2.3896 | 2.9891 | 3.5895 | 4.1908 | 4.793 | 5.3961
Ro for CO2= | 0.8795