## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-710C | GDU-002 | 2008-01-01 | 2009-01-01
0 | 0 | 0 | 0
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | D06-SDS01
AMBII Serial Number: | 50568
NDUV Serial Number: | 208
FID Serial Number: | 123
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.04 | Scott gas | ALM0315016 | ±1% | 2010-01-29 | 93.12
200 ppm CO | 200.6
6% CO2 | 6
300 ppm NO | 298.1
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300.5
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | D06-SDS01 |  | NO2 NDUV Serial Number: |  |  | 208 |  | O2 Sensor
|  |  | Bottle Value ppm |  | Absolute Limit ppm | Relative Limit % |  | Bottle Value % |  | Absolute Limit % | Relative Limit %
Test Date: | 2008-05-08 |  | 505 |  | 10 | 3 |  | 20.9 |  | 1 | 1
Test Time: | 5:32
Test software: | 10.09 SP1
Gas Divider Step |  |  | NO2 Accuracy Test Results |  |  |  |  | O2 Accuracy Test Results
Start | End | Step | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 1 | 1 |  |  | 0 | 0 | 0
90 | 105 | 1 | 50.5 | 48 | -2.5 | -4.95 |  | 2.09 | 2 | -0.09 | -4.31
150 | 165 | 2 | 101 | 100.2 | -0.8 | -0.79 |  | 4.18 | 4.1 | -0.08 | -1.91
210 | 225 | 3 | 151.5 | 150.8 | -0.7 | -0.46 |  | 6.27 | 6.2 | -0.07 | -1.12
270 | 285 | 4 | 202 | 203.1 | 1.1 | 0.54 |  | 8.36 | 8.3 | -0.06 | -0.72
330 | 345 | 5 | 252.5 | 255.5 | 3 | 1.19 |  | 10.45 | 10.39 | -0.06 | -0.57
390 | 405 | 6 | 303 | 308.3 | 5.3 | 1.75 |  | 12.54 | 12.5 | -0.04 | -0.32
450 | 465 | 7 | 353.5 | 356.5 | 3 | 0.85 |  | 14.63 | 14.6 | -0.03 | -0.21
510 | 525 | 8 | 404 | 407.7 | 3.7 | 0.92 |  | 16.72 | 16.7 | -0.02 | -0.12
570 | 585 | 9 | 454.5 | 460.9 | 6.4 | 1.41 |  | 18.81 | 18.9 | 0.09 | 0.48
630 | 645 | 10 | 505 | 509.5 | 4.5 | 0.89 |  | 20.9 | 21 | 0.1 | 0.48
690 | 705 | 9 | 454.5 | 462.5 | 8 | 1.76 |  | 18.81 | 18.9 | 0.09 | 0.48
750 | 765 | 8 | 404 | 409.7 | 5.7 | 1.41 |  | 16.72 | 16.8 | 0.08 | 0.48
810 | 825 | 7 | 353.5 | 360.1 | 6.6 | 1.87 |  | 14.63 | 14.6 | -0.03 | -0.21
870 | 885 | 6 | 303 | 309.7 | 6.7 | 2.21 |  | 12.54 | 12.5 | -0.04 | -0.32
930 | 945 | 5 | 252.5 | 257.1 | 4.6 | 1.82 |  | 10.45 | 10.4 | -0.05 | -0.48
990 | 1005 | 4 | 202 | 207.5 | 5.5 | 2.72 |  | 8.36 | 8.15 | -0.21 | -2.51
1050 | 1065 | 3 | 151.5 | 155.4 | 3.9 | 2.57 |  | 6.27 | 6.2 | -0.07 | -1.12
1110 | 1125 | 2 | 101 | 104 | 3 | 2.97 |  | 4.18 | 4.1 | -0.08 | -1.91
1170 | 1185 | 1 | 50.5 | 50.5 | 0 | 0 |  | 2.09 | 2.1 | 0.01 | 0.48
1230 | 1245 | 0 | 0 | 0.6 | 0.6 |  |  | 0 | 0 | 0
Comments |  |  |  | NO2 Linearity Test Results |  |  |  | O2 Linearity Test Results
[5 empty] | Actual | Criteria | Status (Pass if blank) |  | Actual | Criteria | Status (Pass if blank)
|  |  |  | Intercept | -0.1 | 5.1 |  | Intercept | -0.09 | 0.21
|  |  |  | Slope | 1.01 | 0.99 - 1.01 |  | Slope | 1.01 | 0.99 - 1.01
|  |  |  | SEE | 1.9031 | 5.05 |  | SEE | 0.0639 | 0.209
|  |  |  | r2 | 0.9999 | 0.998 |  | r2 | 0.9999 | 0.998

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.5 | 1.5 | 1.1 | 1.0506 | 0 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.002014 | 0.0022 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1 | 26.3 | 1.5
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.5 | 1.5 | 1.1 | 1.1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.002 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1 | 26.3 | 1.5
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.5 | 1.5 | 1 | 1 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 27 | 6 | 195 | 59.2 | 0.002 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1 | 26.3 | 1.5
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.4 | 1.4 | 1.1 | 1.0582 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.002 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0283 | 26.3 | 1.4
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.4 | 1.4 | 1 | 1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.002 | 0.0021 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0508 | 26.3508 | 1.4
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.4 | 1.4 | 1.1 | 1.051 | 0 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.002 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.37 | 1.4
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.4 | 1.4 | 1 | 1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2856 | 0.002 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0816 | 26.3092 | 1.4
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.3146 | 1.3 | 1 | 1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2146 | 0.001915 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9234 | 26.3883 | 1.3
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.3 | 1.3 | 1 | 1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.144 | 26.0346 | 26.3327 | 1.3
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.2144 | 1.2 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1 | 26.3465 | 1.2
2008-05-08 | 5:32 | 0 | 0.0005 | 4.8683 | -0.6 | 1.2 | 1.2 | 1 | 0.9555 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1258 | 26.3742 | 1.2
2008-05-08 | 5:32 | 0 | 0 | 0 | -0.6 | 1.2 | 1.2 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1747 | 26.3051 | 1.2
2008-05-08 | 5:32 | 0 | 0.0002 | 1.7081 | -0.6 | 1.2 | 1.2 | 1 | 1 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.7788 | 26.3842 | 1.2
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.6 | 1.2 | 1.2 | 0.9 | 0.8584 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0019 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.0802 | 26.3366 | 1.2
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.9 | 1.2 | 1.2 | 0.9 | 0.8574 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.001814 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.215 | 26.3425 | 1.2
2008-05-08 | 5:32 | 0 | 0.001 | 10 | -0.9 | 1.2 | 1.2 | 0.8 | 0.8415 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.3783 | 1.2
2008-05-08 | 5:32 | 0.002 | 0.001 | 10 | -0.9 | 1.2 | 1.2 | 0.8 | 0.8492 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.3009 | 1.2
2008-05-08 | 5:32 | 0.01 | 0.0006 | 6.4943 | -0.9 | 1.2 | 1.2 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.38 | 1.2
2008-05-08 | 5:32 | 0.01 | 0 | 0 | -0.9 | 1.2 | 1.2 | 0.8 | 0.8438 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.041 | 26.341 | 1.2
2008-05-08 | 5:32 | 0.01 | 0 | 0 | -0.9 | 1.2 | 1.2 | 0.9 | 0.9303 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2856 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0382 | 26.3 | 1.2
2008-05-08 | 5:32 | 0.01 | 0 | 0 | -0.9 | 1.2 | 1.2 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.3 | 1.2
2008-05-08 | 5:32 | 0.01 | 0 | 0 | -0.9 | 1.2 | 1.2 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.3 | 1.2
2008-05-08 | 5:32 | 0.008 | 0 | 0 | -0.9 | 1.1 | 1.1 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1 | 26.3757 | 1.1
2008-05-08 | 5:32 | 0 | 0 | 0 | -0.9 | 1.1 | 1.1 | 0.9 | 0.9 | 0 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2146 | 0.0018 | 0.002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.155 | 26.4 | 1.1
2008-05-08 | 5:32 | 0 | 0 | 0 | -0.8 | 1.1 | 1.1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2144 | 0.0018 | 0.001914 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1657 | 26.3657 | 1.1
2008-05-08 | 5:33 | 0 | 0 | 0 | -0.8 | 1.1 | 1.1 | 0.9 | 0.9111 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2 | 0.001714 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0865 | 26.3 | 1.1
2008-05-08 | 5:33 | 0 | 0 | 0 | -0.8 | 1.1 | 1.1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0074 | 26.3 | 1.1
2008-05-08 | 5:33 | 0 | 0 | 0 | -0.8 | 1 | 1 | 0.8 | 0.8448 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2856 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7846 | 26.3718 | 1
2008-05-08 | 5:33 | 0 | 0 | 0 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9035 | 26.4 | 1
2008-05-08 | 5:33 | 0 | 0 | 0 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.1 | 26.4 | 1
2008-05-08 | 5:33 | 0 | 0.0002 | 1.8234 | -0.8 | 1 | 1 | 0.9 | 0.9229 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 26.0813 | 26.4 | 1
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -0.8 | 1 | 1 | 0.9 | 0.8584 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.923 | 26.4 | 1
2008-05-08 | 5:33 | 0 | 0.0002 | 1.5966 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0017 | 0.0019 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1017 | 26.3328 | 1
2008-05-08 | 5:33 | 0 | 0.0013 | 13.1387 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0017 | 0.001814 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.1071 | 26.3465 | 1
2008-05-08 | 5:33 | 0 | 0.002 | 20 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.026 | 26.4 | 1
2008-05-08 | 5:33 | 0 | 0.002 | 20 | -0.8 | 1 | 1 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0896 | 26.3948 | 1
2008-05-08 | 5:33 | 0 | 0.002 | 20 | -0.7 | 1 | 1 | 0.9 | 0.9175 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9313 | 26.3157 | 1
2008-05-08 | 5:33 | 0 | 0.002 | 20 | -0.7 | 1 | 1 | 0.9 | 0.8585 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9 | 26.3635 | 1
2008-05-08 | 5:33 | 0 | 0.002 | 20 | -0.7 | 1 | 1 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9854 | 26.3573 | 1
2008-05-08 | 5:33 | 0.003 | 0.0014 | 13.5387 | -0.8 | 1 | 1 | 0.9 | 0.8787 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0781 | 26.3 | 1
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 1 | 1 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9967 | 26.3011 | 1
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 1 | 1 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7591 | 26.3803 | 1
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 1 | 1 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0017 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9373 | 26.4 | 1
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 0.9144 | 0.9 | 0.9 | 0.9 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.9846 | 26.4 | 0.9
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 0.9 | 0.9 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.8709 | 26.3823 | 0.9
2008-05-08 | 5:33 | 0.01 | 0.001 | 10 | -0.8 | 0.9 | 0.9 | 0.8 | 0.841 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2146 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1001 | 26.1876 | 26.3031 | 0.9
2008-05-08 | 5:33 | 0.009 | 0.001 | 10 | -0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.9733 | 26.3756 | 0.9
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -0.9 | 0.9 | 0.9 | 0.9 | 0.8608 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.9 | 26.3451 | 0.9
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -0.9 | 0.8144 | 0.8 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -0.9 | 0.8 | 0.8 | 0.7 | 0.7494 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1038 | 1002 | 25.8728 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.8 | 0.8 | 0.8 | 0.7572 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7144 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.8 | 0.8 | 0.7 | 0.7 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.8439 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.8 | 0.8 | 0.8 | 0.8 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.798 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.8 | 0.8 | 0.7 | 0.7 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7301 | 26.3 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.8 | 0.8 | 0.7 | 0.7413 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.144 | 25.8094 | 26.3094 | 0.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.7 | 0.7 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.8886 | 26.3886 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.7 | 0.7425 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.8323 | 26.3323 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.7 | 0.6561 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8 | 26.3469 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.7 | 0.66 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8523 | 26.3739 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.7 | 0.658 | 0 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0016 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9946 | 26.3054 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 0.7 | 0.7 | 0.6 | 0.6 | 0 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 26 | 5 | 195 | 59.2 | 0.001514 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.9154 | 26.3846 | 0.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1.1 | 0.61457 | 0.6 | 0.7 | 0.6561 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.0015 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.7726 | 26.4 | 0.6
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1.2 | 0.6 | 0.6 | 0.6 | 0.6 | 0 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0015 | 0.0018 | 0.0001 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7857 | 26.4 | 0.6
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -1 | 8.7616 | 8.8 | 0.7 | 0.7 | 0 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2856 | 0.0015 | 0.0018 | 0.000528 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.944 | 26.4 | 8.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | -0.5 | 27.5 | 27.5 | 0.7 | 0.7424 | 0.66 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.0016 | 0.0018 | 0.0017 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 26.0976 | 26.3988 | 27.5
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0 | 40.3648 | 40.4 | 1.1 | 1.0862 | 1.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2 | 0.001686 | 0.001886 | 0.002442 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9393 | 26.3196 | 40.4
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.1 | 45.7848 | 45.9 | 1.1 | 1.1478 | 1.45 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2144 | 0.002442 | 0.002642 | 0.002786 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9 | 26.3592 | 45.8
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.5 | 46.6848 | 47.2 | 1.1 | 1.1 | 1.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.002786 | 0.002886 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9767 | 26.4 | 46.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 46.9 | 47.2 | 1 | 1 | 1.75 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.002886 | 0.002986 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 26.0642 | 26.3821 | 46.9
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 47 | 47.3 | 1.1 | 1.0597 | 1.8 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2144 | 0.003086 | 0.003186 | 0.0028 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.9059 | 26.3029 | 47
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 47.1 | 47.4 | 1.1 | 1.1 | 1.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.003186 | 0.003286 | 0.0028 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9762 | 26.3762 | 47.1
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 47.1 | 47.4 | 1 | 1 | 1.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.0032 | 0.0033 | 0.0028 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8337 | 26.3446 | 47.1
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 47.2856 | 47.6 | 1.1 | 1.0592 | 1.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0033 | 0.003486 | 0.0028 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.7 | 26.3346 | 47.3
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.3 | 47.3856 | 47.7 | 0.9 | 0.9304 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0034 | 0.003586 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.7138 | 26.4 | 47.4
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.5 | 47.4 | 47.9 | 1 | 1 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0035 | 0.0036 | 0.0028 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1001 | 25.7929 | 26.4 | 47.4
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.8 | 47.5 | 48.3 | 0.9 | 0.9 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.003686 | 0.0037 | 0.0028 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.9441 | 26.4 | 47.5
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.7 | 47.5 | 48.2 | 1 | 1 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0037 | 0.0037 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.8977 | 26.3489 | 47.5
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.7 | 47.5 | 48.2 | 0.9 | 0.9489 | 2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2146 | 0.003785 | 0.0038 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.7697 | 26.3303 | 47.5
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.7 | 47.5 | 48.2 | 0.9 | 0.878 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0038 | 0.0038 | 0.0028 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.7097 | 26.3903 | 47.5
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.6 | 48.2 | 1 | 0.9581 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0039 | 0.0039 | 0.0028 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1001 | 25.7888 | 26.3112 | 47.6
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.6 | 48.2 | 0.9 | 0.9 | 2 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0039 | 0.003986 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.7325 | 26.3 | 47.6
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.6856 | 48.3 | 0.9 | 0.8581 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.003986 | 0.004 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7468 | 26.3 | 47.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.7 | 48.3 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.2 | 0.004 | 0.004 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7211 | 26.3263 | 47.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.7 | 48.3 | 1 | 0.9565 | 2 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.004 | 0.004086 | 0.0029 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5109 | 26.4 | 47.7
2008-05-08 | 5:33 | 0 | 0.001 | 10 | 0.6 | 47.7856 | 48.4 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.004086 | 0.0041 | 0.0029 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6693 | 26.4 | 47.8
2008-05-08 | 5:34 | 0 | 0.001 | 10 | 0.6 | 47.8 | 48.4 | 0.9 | 0.9 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0041 | 0.0041 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001 | 25.7638 | 26.4 | 47.8
2008-05-08 | 5:34 | 0 | 0.001 | 10 | 0.7 | 47.8856 | 48.6 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.004186 | 0.0042 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1001.856 | 25.843 | 26.4 | 47.9
2008-05-08 | 5:34 | 0 | 0.001 | 10 | 0.7 | 47.9 | 48.6 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0042 | 0.0042 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.9 | 26.4 | 47.9
2008-05-08 | 5:34 | 0 | 0 | 0.47697 | 0.7 | 47.9 | 48.6 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0042 | 0.0042 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.144 | 25.9 | 26.3987 | 47.9
2008-05-08 | 5:34 | 0 | -0.001 | -10 | 0.7 | 47.9856 | 48.7 | 0.9 | 0.8709 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.004286 | 0.004286 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.9 | 26.3195 | 48
2008-05-08 | 5:34 | 0 | -0.001 | -10 | 0.8 | 48 | 48.8 | 0.8 | 0.8 | 2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2146 | 0.0043 | 0.0043 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1001.8543 | 25.8404 | 26.3596 | 48
2008-05-08 | 5:34 | 0 | -0.001 | -10 | 0.8 | 48 | 48.8 | 0.8 | 0.8 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2854 | 0.0043 | 0.0043 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.8 | 26.3612 | 48
2008-05-08 | 5:34 | 0 | -0.001 | -10 | 0.8 | 48 | 48.8 | 0.9 | 0.8569 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0043 | 0.0043 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8361 | 26.3181 | 48
2008-05-08 | 5:34 | 0 | -0.0004 | -3.5299 | 0.8 | 48 | 48.8 | 0.8 | 0.8 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0043 | 0.0043 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9945 | 26.3972 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.9 | 0.8613 | 2 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0043 | 0.0043 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0763 | 26.3237 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.878 | 26.3555 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.9 | 0.876 | 2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.7694 | 26.4 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8861 | 26.3861 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8416 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8069 | 26.3069 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8479 | 2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 26.0169 | 26.3 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8379 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9976 | 26.3512 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.8 | 0.8185 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.8395 | 26.3698 | 48
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48 | 48.8 | 0.7 | 0.7 | 2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0044 | 0.0044 | 0.0029 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.856 | 25.7097 | 26.3097 | 48
2008-05-08 | 5:34 | 0.006 | 0 | 0 | 0.8 | 48.0863 | 48.9 | 0.8 | 0.7622 | 2 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0045 | 0.0045 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7889 | 26.3889 | 48.1
2008-05-08 | 5:34 | 0.005 | 0 | 0 | 0.8 | 48.1 | 48.9 | 0.7 | 0.7 | 2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.732 | 26.332 | 48.1
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48.1 | 48.9 | 0.8 | 0.8 | 2 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2856 | 0.0045 | 0.0045 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.8417 | 26.3 | 48.1
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 48.1 | 48.9 | 0.7 | 0.7204 | 1.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2854 | 0.0045 | 0.0045 | 0.0029 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.9204 | 26.3265 | 48.1
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.9 | 48.1 | 49 | 0.9 | 0.8653 | 1.8 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2144 | 0.0045 | 0.0045 | 0.0029 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7228 | 26.3943 | 48.1
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 47.4168 | 48.2 | 0.8 | 0.8 | 1.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.2146 | 0.0045 | 0.0045 | 0.002815 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.854 | 26.0395 | 26.3151 | 47.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 43.6626 | 44.4 | 0.8 | 0.7815 | 1.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.002629 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001.1446 | 25.844 | 26.364 | 43.7
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 41.5152 | 42.2 | 0.8 | 0.8 | 1.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0025 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7862 | 26.3569 | 41.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 40.5 | 41.2 | 0.8 | 0.8 | 1.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0024 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.9 | 26.3223 | 40.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 40.4 | 41.1 | 0.9 | 0.8583 | 1.8 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.8985 | 26.3985 | 40.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 40.3144 | 41 | 0.9 | 0.8603 | 1.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0024 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8193 | 26.3193 | 40.3
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.6 | 40.3 | 40.9 | 0.9 | 0.9227 | 1.8 | 13 | 956 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.6809 | 26.3 | 40.3
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.6 | 40.3856 | 41 | 0.8 | 0.8 | 1.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0024 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1038 | 1002 | 25.6775 | 26.3387 | 40.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 42.6256 | 43.3 | 0.9 | 0.9 | 1.9 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.002571 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8366 | 26.4 | 42.6
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 43.7848 | 44.5 | 0.8 | 0.8 | 1.9 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0026 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9948 | 26.4 | 43.8
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 43.9856 | 44.7 | 0.9 | 0.8665 | 1.9 | 12.9 | 956 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0026 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9234 | 26.4 | 44
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 45.4552 | 46.2 | 0.8 | 0.8 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.002686 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0115 | 26.3443 | 45.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.7 | 45.7856 | 46.5 | 0.8 | 0.8426 | 2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.0027 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9602 | 26.3349 | 45.8
2008-05-08 | 5:34 | 0 | 0 | 0 | 0.8 | 46.5704 | 47.4 | 0.9 | 0.8523 | 2 | 12.9 | 956 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0045 | 0.0045 | 0.002786 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7142 | 26.3858 | 46.6
2008-05-08 | 5:34 | 0 | 0 | 0 | 2.2 | 83.58 | 85.8 | 1 | 0.9664 | 2.42 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.2 | 0.005542 | 0.00537 | 0.004956 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7933 | 26.3067 | 83.6
2008-05-08 | 5:34 | 0 | 0 | 0 | 2.3 | 120.2248 | 122.5 | 1.2 | 1.2 | 3.8 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.008642 | 0.008642 | 0.007071 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8724 | 26.3 | 120.2
2008-05-08 | 5:34 | 0 | 0 | 0 | 2.1 | 113.8 | 115.9 | 1.2 | 1.2 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.3 | 0.0088 | 0.0088 | 0.0067 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7454 | 26.3515 | 113.8
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.8 | 104.4781 | 106.3 | 1.3 | 1.2563 | 4 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.008885 | 0.008885 | 0.006229 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8526 | 25.6614 | 26.4 | 104.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.7 | 100.5 | 102.2 | 1 | 1 | 3.9 | 12.9 | 957 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0089 | 0.0089 | 0.006 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.79 | 26.4 | 100.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.7 | 100.4856 | 102.2 | 1.1 | 1.1 | 3.9 | 13 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0089 | 0.0089 | 0.006 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002.1443 | 25.7108 | 26.4 | 100.5
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.7 | 100.4146 | 102.1 | 1 | 1 | 3.9 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0089 | 0.0089 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8543 | 25.7683 | 26.3317 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.4 | 102 | 1 | 1 | 4 | 12.9 | 957 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.008986 | 0.008986 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8 | 26.3 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.3856 | 102 | 1 | 1 | 4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.009 | 0.009 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8534 | 26.3 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.4 | 102 | 0.9 | 0.939 | 4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.009 | 0.009 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.9941 | 26.3 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.4 | 102 | 1 | 1 | 4 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.009 | 0.009 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.915 | 26.3 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.4 | 102 | 0.9 | 0.9419 | 4 | 12.9 | 957 | 962 | 945 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2144 | 0.009 | 0.009 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.9 | 26.3642 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.6 | 100.4 | 102 | 1.1 | 1.0572 | 4.04 | 13 | 957 | 961 | 946 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2856 | 0.009 | 0.009 | 0.006 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.9432 | 26.3568 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1 | 1 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.009 | 0.009086 | 0.006 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.9334 | 26.3 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1.1 | 1.0588 | 4.1 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.009 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.7055 | 26.3018 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1 | 1 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2144 | 0.009 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9428 | 26.3809 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1.1 | 1.0793 | 4.1 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.3 | 0.0091 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8801 | 26.34 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.3144 | 101.8 | 1 | 1 | 4.1 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0091 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.8 | 26.3 | 100.3
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.3 | 101.8 | 1 | 1.0053 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0091 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7816 | 26.3 | 100.3
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.3 | 101.8 | 1 | 1 | 4.1 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0091 | 0.0091 | 0.006 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7024 | 26.3 | 100.3
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1 | 1.0442 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0091 | 0.0092 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7768 | 26.3768 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.5 | 100.4 | 101.9 | 1 | 1.0476 | 4.1 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0091 | 0.0092 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6319 | 26.344 | 100.4
2008-05-08 | 5:34 | 0 | 0 | 0 | 1.4 | 100.4 | 101.8 | 1 | 1 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 26 | 5 | 195 | 59.2 | 0.009185 | 0.0092 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.8543 | 25.7109 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.4 | 101.8 | 1.1 | 1.0574 | 4.1 | 13 | 957 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0092 | 0.0092 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.1 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.4 | 101.8 | 1 | 1 | 4.1 | 12.9 | 957 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0092 | 0.009286 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 26.1 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.4 | 101.7 | 1.1 | 1.1 | 4.1 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2144 | 0.0092 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1001.856 | 25.9545 | 26.3728 | 100.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.4 | 101.7 | 1 | 1.0479 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0092 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.9 | 26.3484 | 100.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.3144 | 101.6 | 1 | 1.0431 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0092 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8388 | 26.3 | 100.3
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.3 | 101.7 | 1.1 | 1.0561 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2144 | 0.009286 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3102 | 100.3
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.2144 | 101.6 | 1 | 1 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0093 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.856 | 25.7 | 26.3894 | 100.2
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.2 | 101.6 | 1.1 | 1.0577 | 4.1 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2856 | 0.0093 | 0.0093 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7 | 26.3314 | 100.2
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.4 | 100.2 | 101.6 | 0.9 | 0.9392 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.2 | 0.0093 | 0.0093 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7956 | 26.3478 | 100.2
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.2 | 101.5 | 1.1 | 1.0506 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.2 | 0.0093 | 0.0093 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.846 | 26.4 | 100.2
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.2856 | 101.6 | 0.9 | 0.9378 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.2856 | 0.0093 | 0.009386 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3938 | 100.3
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.3 | 101.6 | 0.9 | 0.9279 | 4.1 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3146 | 100.3
2008-05-08 | 5:35 | 0 | 0 | 0 | 1.3 | 100.3 | 101.6 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.2 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.7 | 26.3645 | 100.3
2008-05-08 | 5:35 | 0 | 0.0008 | 8.159 | 1.3 | 100.3 | 101.6 | 1 | 1 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.7 | 26.4 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.3 | 100.2 | 101.5 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8557 | 25.7 | 26.4 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.3 | 100.2 | 101.5 | 0.9 | 0.94 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 26 | 5 | 195 | 59.3 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7 | 26.3979 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.3 | 100.2 | 101.5 | 1 | 0.9564 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2144 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7 | 26.3188 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.3 | 100.2 | 101.5 | 1 | 0.9759 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7 | 26.3604 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 1 | 0.9568 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2856 | 0.0093 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.5811 | 26.4 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.009386 | 0.0094 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.4942 | 26.3812 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.9487 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2144 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8901 | 26.302 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8228 | 26.3 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 1 | 1 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7436 | 26.3 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.943 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2144 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7354 | 26.3 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.942 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2856 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8285 | 26.3 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2 | 101.4 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9877 | 26.3 | 100.2
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2854 | 101.5 | 0.9 | 0.9372 | 4.1 | 12.9 | 957 | 962 | 946 | 21.4 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8538 | 26.3 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.2856 | 101.5 | 0.9 | 0.9277 | 4.1 | 12.9 | 957 | 962 | 946 | 21.6 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189.4 | 26 | 5 | 195 | 59.2856 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.5918 | 26.352 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.3 | 101.5 | 0.8 | 0.8 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2146 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5247 | 26.3688 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.3 | 101.5 | 0.9 | 0.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0095 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8214 | 26.3 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.3 | 101.4 | 0.8 | 0.8 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9797 | 26.3 | 100.3
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.9 | 0.8508 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0094 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7929 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.8 | 0.76 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0095 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.8 | 0.8 | 4.1 | 12.9 | 957 | 962 | 947 | 21.4 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0095 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7546 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.7 | 0.7416 | 4.1 | 12.9 | 957 | 962 | 947 | 21 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0095 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.8875 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.8 | 0.8374 | 4.1 | 12.9 | 957 | 962 | 947 | 21.6 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0095 | 0.0096 | 0.0059 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7286 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.1 | 100.4 | 101.5 | 0.8 | 0.8 | 4.1 | 12.9 | 957 | 962 | 947 | 21.4 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0095 | 0.0096 | 0.005986 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.6351 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 1.2 | 100.4 | 101.6 | 0.8 | 0.7807 | 4.1 | 12.9 | 957 | 962 | 947 | 21 | 0.37 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0095 | 0.0096 | 0.006 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7763 | 26.3 | 100.4
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 3 | 129.5576 | 132.5 | 0.9 | 0.9 | 4.46 | 12.9 | 957 | 962 | 947 | 21 | 0.38 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.009857 | 0.009686 | 0.007699 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9536 | 26.3 | 129.6
2008-05-08 | 5:35 | 0 | 0.001 | 10 | 3.5 | 148.8776 | 152.4 | 1 | 0.9701 | 5.27 | 12.9 | 958 | 962 | 947 | 21.6 | 0.39 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.011756 | 0.011613 | 0.008828 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7976 | 26.3 | 148.9
2008-05-08 | 5:35 | 0 | 0.0007 | 7.1917 | 3.4 | 149.9 | 153.3 | 1.3 | 1.2568 | 5.6 | 12.9 | 958 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.011986 | 0.0118 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7183 | 26.3 | 149.9
2008-05-08 | 5:35 | 0 | 0 | 0 | 3.2 | 149.9856 | 153.2 | 1.1 | 1.061 | 5.8 | 12.9 | 958 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.012086 | 0.011986 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3 | 150
2008-05-08 | 5:35 | 0 | 0 | 0 | 3.1 | 150 | 153.1 | 1.1 | 1.0539 | 5.88 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0122 | 0.012086 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8603 | 26.3 | 150
2008-05-08 | 5:35 | 0 | 0 | 0 | 3 | 150.1 | 153.1 | 0.9 | 0.9305 | 5.94 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.012385 | 0.012285 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0616 | 26.3192 | 150.1
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.8 | 150.1854 | 153 | 1 | 1 | 6 | 12.9 | 957 | 962 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.012485 | 0.012385 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9033 | 26.3983 | 150.2
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.8 | 150.1 | 152.9 | 0.9 | 0.9 | 6 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.012586 | 0.0125 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8224 | 26.3224 | 150.1
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.7 | 150.1 | 152.8 | 0.9 | 0.944 | 6.1 | 12.9 | 957 | 962 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0126 | 0.012586 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8557 | 25.6864 | 26.3 | 150.1
2008-05-08 | 5:35 | 0 | -0.0003 | -2.6829 | 2.6 | 150.2 | 152.8 | 1 | 1 | 6.1 | 12.9 | 957 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.012685 | 0.012685 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6716 | 26.3 | 150.2
2008-05-08 | 5:35 | 0 | -0.0019 | -18.9431 | 2.6 | 150.2 | 152.8 | 0.9 | 0.9415 | 6.1 | 12.9 | 957 | 962 | 947 | 21.4 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.012786 | 0.0127 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8293 | 26.3 | 150.2
2008-05-08 | 5:35 | 0 | -0.002 | -20 | 2.6 | 150.3 | 152.9 | 1 | 0.9572 | 6.1 | 12.9 | 957 | 962 | 947 | 21.6 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0128 | 0.0128 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1443 | 25.9887 | 26.3 | 150.3
2008-05-08 | 5:35 | 0 | -0.0014 | -13.8686 | 2.5 | 150.4 | 152.9 | 0.9 | 0.9 | 6.1 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0129 | 0.0129 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9264 | 26.3 | 150.4
2008-05-08 | 5:35 | 0 | -0.001 | -10 | 2.4 | 150.4 | 152.8 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.012986 | 0.012986 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7944 | 26.3 | 150.4
2008-05-08 | 5:35 | 0 | -0.0007 | -7.0833 | 2.4 | 150.4 | 152.8 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.013 | 0.013 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.332 | 150.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.4 | 150.4 | 152.8 | 1 | 0.9511 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.013086 | 0.0131 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7335 | 26.3888 | 150.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.4 | 150.5 | 152.9 | 1 | 0.9584 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0131 | 0.0131 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.971 | 26.3097 | 150.5
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.4 | 150.4 | 152.8 | 0.9 | 0.9423 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0131 | 0.0132 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7914 | 26.3 | 150.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.3 | 150.4 | 152.7 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0132 | 0.0132 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8462 | 26.3 | 150.4
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.3 | 150.4856 | 152.8 | 0.9 | 0.9372 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0132 | 0.0133 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 25.9722 | 26.3 | 150.5
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.3 | 150.5856 | 152.9 | 1 | 0.9518 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0132 | 0.0133 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.8932 | 26.3 | 150.6
2008-05-08 | 5:35 | 0 | 0 | 0 | 2.2 | 150.6 | 152.8 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0133 | 0.0133 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.8137 | 26.3 | 150.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.3 | 150.6 | 152.9 | 1 | 0.9777 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0133 | 0.0134 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8656 | 26.3 | 150.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.3 | 150.6856 | 153 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.013386 | 0.0134 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 25.8105 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.4 | 150.7 | 153.1 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0134 | 0.0134 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7956 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.7 | 150.6 | 153.3 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2146 | 0.013485 | 0.0134 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0876 | 26.3 | 150.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.8 | 150.6856 | 153.5 | 0.9 | 0.9488 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0135 | 0.0134 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7709 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.8 | 150.7 | 153.5 | 0.9 | 0.8567 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0135 | 0.013486 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8845 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 2.9 | 150.6145 | 153.5 | 0.9 | 0.9431 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2145 | 0.0135 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.3 | 150.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 150.6 | 153.8 | 0.9 | 0.9414 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.013586 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.3 | 150.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.6856 | 154 | 0.9 | 0.8878 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0136 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.7 | 154 | 1 | 0.9519 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0136 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7652 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.7 | 154 | 0.8 | 0.8435 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0136 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.585 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.7 | 154 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0136 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6833 | 26.3 | 150.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 150.8 | 154 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.013686 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9842 | 26.3 | 150.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.8856 | 154.2 | 1 | 0.9581 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0137 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 25.9049 | 26.3 | 150.9
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.9 | 154.2 | 0.9 | 0.8573 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0137 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.9 | 26.3 | 150.9
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 150.9856 | 154.3 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0137 | 0.0135 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8465 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 0.8 | 0.8197 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0137 | 0.013586 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151.1 | 154.4 | 0.9 | 0.9418 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0137 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7643 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151.1 | 154.4 | 0.9 | 0.939 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0137 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.8557 | 25.5266 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 0.9 | 0.9427 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2854 | 0.0137 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6406 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 1 | 0.9576 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.013786 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0138 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7856 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151 | 154.4 | 1 | 0.9572 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0138 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9926 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151 | 154.4 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2146 | 0.0138 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.913 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151 | 154.4 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0138 | 0.0136 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8337 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0138 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151 | 154.3 | 1 | 0.9513 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0138 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7753 | 26.3 | 151
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151.1 | 154.5 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0138 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7154 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151.1 | 154.5 | 0.9 | 0.9381 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0138 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0304 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 151.2 | 154.6 | 0.9 | 0.9474 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2144 | 0.0138 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 25.8517 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151.2 | 154.5 | 0.9 | 0.9413 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.2144 | 0.0139 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7415 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151.2 | 154.5 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.2144 | 0.0139 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.8 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 151.2 | 154.5 | 0.9 | 0.9443 | 6.2 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 26 | 5 | 195 | 59.2856 | 0.0139 | 0.0137 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 151.2 | 154.4 | 1.1 | 1.0581 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0139 | 0.013786 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.879 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 151.1144 | 154.3 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0139 | 0.0138 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7834 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 151.1 | 154.3 | 1 | 1 | 6.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2 | 0.0139 | 0.0138 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7751 | 26.3 | 151.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.1 | 151.2 | 154.3 | 0.9 | 0.9 | 6.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2857 | 0.0139 | 0.0138 | 0.0089 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8571 | 25.8664 | 26.3 | 151.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.8 | 177.8988 | 181.7 | 1 | 1.0479 | 6.44 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.014328 | 0.014142 | 0.010527 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1443 | 25.7081 | 26.3 | 177.9
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.9 | 197.1946 | 201.1 | 1.3 | 1.3418 | 7.25 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.015754 | 0.015625 | 0.011642 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8496 | 26.3 | 197.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.5 | 200.0856 | 203.6 | 1.3 | 1.3 | 7.63 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2 | 0.016257 | 0.016157 | 0.0118 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7917 | 26.3 | 200.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 200.8 | 204.2 | 1.1 | 1.1417 | 7.88 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2854 | 0.016385 | 0.016385 | 0.0118 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7336 | 26.3 | 200.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.1 | 201.6856 | 204.7 | 1 | 1.0413 | 8 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.016742 | 0.016828 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8 | 26.3 | 201.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 201.7 | 204.9 | 1.1 | 1.1 | 8.08 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.016886 | 0.016986 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8 | 26.3 | 201.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.5 | 201.8 | 205.3 | 1 | 1 | 8.1 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0171 | 0.0171 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 26.0135 | 26.3 | 201.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 201.9 | 205.3 | 1.2 | 1.1592 | 8.19 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0172 | 0.0172 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.9991 | 26.3 | 201.9
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 202.1 | 205.5 | 1 | 1.0293 | 8.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0173 | 0.0173 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9 | 26.3 | 202.1
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 202.1856 | 205.6 | 1.1 | 1.1 | 8.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0173 | 0.0173 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9088 | 26.3 | 202.2
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.6 | 202.2856 | 205.9 | 1.1 | 1.1 | 8.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.017486 | 0.0174 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.988 | 26.3 | 202.3
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.6 | 202.3856 | 206 | 1.1 | 1.097 | 8.2 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0175 | 0.0175 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.866 | 26.3 | 202.4
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.5 | 202.4856 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0176 | 0.017586 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9387 | 26.3 | 202.5
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.5 | 202.4144 | 205.9 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0176 | 0.0176 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9976 | 26.3 | 202.4
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 202.4 | 205.8 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.2146 | 0.0176 | 0.0176 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7048 | 26.3 | 202.4
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 202.4856 | 205.9 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.017686 | 0.017686 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7839 | 26.3 | 202.5
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.4 | 202.5 | 205.9 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0177 | 0.0177 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.3 | 202.5
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 202.6 | 205.9 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2854 | 0.017785 | 0.017785 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9276 | 26.3 | 202.6
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 202.6856 | 206 | 1.2 | 1.1574 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0178 | 0.0178 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 26.0349 | 26.3 | 202.7
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.3 | 202.8 | 206.1 | 1 | 1.0384 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0178 | 0.0178 | 0.0119 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.7991 | 26.3 | 202.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 202.8 | 206 | 1.1 | 1.142 | 8.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0178 | 0.017886 | 0.011986 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.7199 | 26.3 | 202.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 202.8 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0178 | 0.0179 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3 | 202.8
2008-05-08 | 5:36 | 0 | 0 | 0 | 3.2 | 202.8 | 206 | 1.2 | 1.2 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.017886 | 0.0179 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.815 | 26.3 | 202.8
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.2 | 202.8 | 206 | 1.2 | 1.1603 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0179 | 0.0179 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.3 | 202.8
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.2 | 202.9 | 206.1 | 1.1 | 1.1416 | 8.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.2857 | 0.0179 | 0.018 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.3 | 202.9
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.2 | 202.9 | 206.1 | 1.1 | 1.1433 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.2144 | 0.018 | 0.018 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7735 | 26.3 | 202.9
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.2 | 202.9854 | 206.2 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.018 | 0.0181 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.81 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.2 | 203 | 206.2 | 1.1 | 1.148 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.018 | 0.0181 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8654 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.1 | 203.1 | 206.2 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.018 | 0.0181 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.8 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.1 | 203.1 | 206.2 | 1.1 | 1.1 | 8.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0181 | 0.018186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.8 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 3 | 203.1 | 206.1 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.018185 | 0.018285 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.9444 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.9 | 203.1 | 206 | 1.1 | 1.1424 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0182 | 0.0183 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001.856 | 25.8974 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.9 | 203.1 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.3 | 0.0182 | 0.018386 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038.6752 | 1002 | 25.8914 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.9 | 203.1 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2146 | 0.018285 | 0.0184 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0607 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203 | 205.8 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0183 | 0.0184 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.856 | 25.744 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203 | 205.8 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0183 | 0.018486 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7677 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203 | 205.7 | 1.1 | 1.1444 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.018386 | 0.0185 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203 | 205.8 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0184 | 0.0185 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7733 | 26.3 | 203
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203.0856 | 205.9 | 1.1 | 1.1487 | 8.3 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0184 | 0.0185 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7059 | 26.3 | 203.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203.1856 | 206 | 1 | 1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0184 | 0.0185 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.7851 | 26.3 | 203.2
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203.3 | 206.1 | 1.1 | 1.1 | 8.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0184 | 0.018586 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9929 | 26.2357 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203.3856 | 206.2 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0184 | 0.0186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.144 | 26.0565 | 26.2435 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.8 | 203.4 | 206.2 | 1.1 | 1.0885 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0184 | 0.0186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001 | 25.9773 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203.3144 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.018486 | 0.0186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.8981 | 26.3 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203.3144 | 206 | 1.1 | 1.0511 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0185 | 0.0186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8188 | 26.3 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203.3 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0185 | 0.0186 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8604 | 26.3 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203.4 | 206.1 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0185 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1038.2898 | 1002 | 25.9 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.7 | 203.4 | 206.1 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0185 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8812 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.4 | 206 | 1 | 1.0421 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0185 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.802 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.4 | 206 | 1.1 | 1.1 | 8.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2854 | 0.0185 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8 | 26.2228 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.4144 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.018586 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8564 | 26.2564 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.4 | 206 | 1.1 | 1.14 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8292 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.4 | 206 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0187 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7143 | 26.3 | 203.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.3144 | 205.9 | 1.1 | 1.1431 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.018786 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002.856 | 25.794 | 26.3 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.6 | 203.3 | 205.9 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.8 | 26.3 | 203.3
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.4856 | 206 | 1.1 | 1.1489 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7475 | 26.3 | 203.5
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.5856 | 206.1 | 1.2 | 1.1574 | 8.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2144 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7952 | 26.3 | 203.6
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.6 | 206.1 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9891 | 26.2891 | 203.6
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.7 | 206.2 | 1.2 | 1.2 | 8.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.3 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9098 | 26.2098 | 203.7
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.7 | 206.2 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7613 | 26.2694 | 203.7
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.5 | 203.7 | 206.2 | 1.2 | 1.2 | 8.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0186 | 0.0188 | 0.012 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7486 | 26.3 | 203.7
2008-05-08 | 5:37 | 0 | 0 | 0 | 2.9 | 211.7464 | 214.7 | 1.4 | 1.3875 | 8.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0187 | 0.0188 | 0.012514 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8276 | 26.3 | 211.7
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.9 | 236.0907 | 241 | 1.5 | 1.4795 | 8.8 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.019405 | 0.019246 | 0.013951 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9 | 26.3 | 236.1
2008-05-08 | 5:37 | 0 | 0 | 0 | 5.4 | 244.8 | 250.2 | 1.4 | 1.4 | 9.45 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0199 | 0.0196 | 0.0144 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9 | 26.3 | 244.8
2008-05-08 | 5:37 | 0 | 0 | 0 | 5.1 | 253.5084 | 258.6 | 1.4 | 1.4423 | 9.83 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.020813 | 0.020685 | 0.014957 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9655 | 26.3 | 253.5
2008-05-08 | 5:37 | 0 | 0 | 0 | 5.1 | 254.2 | 259.3 | 1.3 | 1.3 | 10 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.021086 | 0.020986 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9106 | 26.2553 | 254.2
2008-05-08 | 5:37 | 0 | 0 | 0 | 5 | 254.3856 | 259.4 | 1.3 | 1.3 | 10.13 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0212 | 0.021086 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8717 | 26.2239 | 254.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.9 | 254.4856 | 259.4 | 1.3 | 1.3 | 10.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.6 | 195 | 59.2856 | 0.021386 | 0.021286 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.1 | 26.2969 | 254.5
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.4 | 254.4146 | 258.8 | 1.3 | 1.3 | 10.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.4 | 195 | 59.2 | 0.021785 | 0.021771 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.1 | 26.2177 | 254.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.4 | 254.4 | 258.8 | 1.4 | 1.3568 | 9.64 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0218 | 0.0218 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8539 | 26.2615 | 254.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.3 | 254.4 | 258.7 | 1.3 | 1.3447 | 9.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0219 | 0.0219 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7815 | 26.3 | 254.4
2008-05-08 | 5:37 | 0 | 0 | 0 | 4.2 | 254.4856 | 258.7 | 1.4 | 1.4 | 9.48 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.022 | 0.022 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8802 | 26.2802 | 254.5
2008-05-08 | 5:37 | 0 | 0 | 0 | 4 | 254.5856 | 258.6 | 1.4 | 1.3568 | 9.63 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0221 | 0.022186 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8016 | 26.2017 | 254.6
2008-05-08 | 5:37 | 0 | 0 | 0 | 4 | 254.7 | 258.7 | 1.4 | 1.4 | 9.79 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.022186 | 0.0222 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0348 | 26.2783 | 254.7
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.9 | 254.8856 | 258.8 | 1.3 | 1.3317 | 9.83 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0222 | 0.0223 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9847 | 26.3 | 254.9
2008-05-08 | 5:37 | 0 | 0.0003 | 2.5684 | 3.9 | 254.9 | 258.8 | 1.3 | 1.3415 | 9.98 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.022286 | 0.0223 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9 | 26.3 | 254.9
2008-05-08 | 5:37 | 0 | 0.001 | 10 | 3.8 | 255 | 258.8 | 1.4 | 1.3596 | 10 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2857 | 0.0223 | 0.0224 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.8839 | 26.3 | 255
2008-05-08 | 5:37 | 0 | 0.001 | 10 | 3.8 | 255.1 | 258.9 | 1.3 | 1.3468 | 10.1 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0224 | 0.0225 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.8048 | 26.3 | 255.1
2008-05-08 | 5:37 | 0 | 0.001 | 10 | 3.7 | 255.0144 | 258.7 | 1.4 | 1.3779 | 10.1 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.022486 | 0.0226 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.949 | 26.3 | 255
2008-05-08 | 5:37 | 0 | 0.001 | 10 | 3.6 | 255 | 258.6 | 1.3 | 1.3417 | 10.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0225 | 0.022686 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26 | 26.2463 | 255
2008-05-08 | 5:37 | 0 | 0.001 | 10 | 3.5 | 254.9144 | 258.4 | 1.4 | 1.4 | 10.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.022586 | 0.0227 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9342 | 26.2329 | 254.9
2008-05-08 | 5:37 | 0 | 0.0007 | 7.3773 | 3.5 | 255 | 258.5 | 1.3 | 1.3 | 10.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0226 | 0.0228 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.788 | 26.3 | 255
2008-05-08 | 5:37 | 0 | 0 | 0 | 3.4 | 254.9 | 258.3 | 1.5 | 1.4529 | 10.3 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.022686 | 0.0228 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.7094 | 26.3 | 254.9
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 254.9854 | 258.4 | 1.4 | 1.4 | 10.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2146 | 0.0227 | 0.0229 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9111 | 26.2296 | 255
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 254.9854 | 258.4 | 1.4 | 1.4429 | 10.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 27 | 5 | 195 | 59.2854 | 0.0227 | 0.0229 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9004 | 26.2498 | 255
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255 | 258.3 | 1.4 | 1.4 | 10.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.022786 | 0.022986 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.829 | 26.3 | 255
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.1 | 258.3 | 1.4 | 1.4 | 10.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0228 | 0.023 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9082 | 26.2918 | 255.1
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.1 | 258.3 | 1.4 | 1.4 | 10.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0228 | 0.0231 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9874 | 26.2126 | 255.1
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.1 | 255.1144 | 258.2 | 1.4 | 1.4 | 10.3 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0229 | 0.0231 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.8667 | 26.2666 | 255.1
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.1856 | 258.4 | 1.4 | 1.4 | 10.38 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0229 | 0.0231 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.8459 | 26.3 | 255.2
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.2856 | 258.6 | 1.3 | 1.2579 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0229 | 0.0231 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002.144 | 25.9502 | 26.3 | 255.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.3856 | 258.7 | 1.3 | 1.3 | 10.4 | 13 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0229 | 0.0231 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0914 | 26.3 | 255.4
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.5 | 258.8 | 1.3 | 1.2585 | 10.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.022986 | 0.023186 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001.144 | 25.9328 | 26.3 | 255.5
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.5 | 258.8 | 1.3 | 1.3 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0232 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9 | 26.3 | 255.5
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.6 | 259 | 1.3 | 1.2786 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.023 | 0.0232 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9 | 26.258 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.6 | 259 | 1.3 | 1.2805 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.023 | 0.0232 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.8576 | 26.2212 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.7 | 259.1 | 1.3 | 1.2991 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0232 | 0.015 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.7 | 26.3 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.7146 | 259.1 | 1.2 | 1.2435 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 27 | 5 | 195 | 59.2 | 0.023 | 0.0232 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.7 | 26.3 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.7 | 259.1 | 1.3 | 1.3 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2143 | 0.0231 | 0.023286 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038.7316 | 1002 | 25.7 | 26.3 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.7 | 259.1 | 1.3 | 1.3 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0231 | 0.0233 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.4 | 255.7 | 259.1 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0231 | 0.0233 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7518 | 26.2827 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.6 | 258.9 | 1.3 | 1.2601 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0231 | 0.0233 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9895 | 26.2035 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.5 | 258.8 | 1.4 | 1.3597 | 10.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.023186 | 0.023386 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 26.1503 | 26.2 | 255.5
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.6 | 258.9 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0232 | 0.0234 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.9268 | 26.2546 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.3 | 255.6 | 258.9 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0232 | 0.0234 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7342 | 26.2658 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.6856 | 258.9 | 1.4 | 1.358 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0232 | 0.0234 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.7866 | 26.2134 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.7859 | 259 | 1.3 | 1.3418 | 10.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0232 | 0.0234 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7074 | 26.2926 | 255.8
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.2 | 255.8 | 259 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0232 | 0.0235 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7718 | 26.2282 | 255.8
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.1 | 255.7 | 258.8 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0233 | 0.0235 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.251 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.1 | 255.6857 | 258.8 | 1.5 | 1.4562 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0233 | 0.0235 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1431 | 25.7698 | 26.2698 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.9 | 255.6144 | 258.5 | 1.4 | 1.3503 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0233 | 0.0235 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.856 | 25.7095 | 26.2 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.9 | 255.6 | 258.5 | 1.5 | 1.4592 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 27 | 5 | 195 | 59.3 | 0.0233 | 0.0236 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7887 | 26.2 | 255.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.9 | 255.7 | 258.6 | 1.3 | 1.3286 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0233 | 0.0236 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 26.0024 | 26.2 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7 | 258.5 | 1.5 | 1.459 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2854 | 0.0233 | 0.0236 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1457 | 26.0065 | 26.2468 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7861 | 258.6 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2139 | 0.0233 | 0.0236 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8736 | 26.2736 | 255.8
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7145 | 258.5 | 1.5 | 1.4523 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0233 | 0.0236 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7944 | 26.2056 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7 | 258.5 | 1.5 | 1.4558 | 10.4 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0233 | 0.0237 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7152 | 26.2848 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7 | 258.5 | 1.4 | 1.4416 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0234 | 0.0237 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7 | 26.3 | 255.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 255.7856 | 258.6 | 1.5 | 1.5 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0234 | 0.0237 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7432 | 26.2568 | 255.8
2008-05-08 | 5:38 | 0 | 0 | 0 | 2.8 | 257.3408 | 260.1 | 1.5 | 1.5 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0234 | 0.0237 | 0.0151 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8224 | 26.2224 | 257.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 5 | 278.8944 | 283.9 | 1.6 | 1.6435 | 10.59 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.023828 | 0.023786 | 0.016385 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9 | 26.2983 | 278.9
2008-05-08 | 5:38 | 0 | 0 | 0 | 5.7 | 314.0248 | 319.7 | 2 | 2.0444 | 11.89 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.026413 | 0.026298 | 0.01847 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9 | 26.219 | 314
2008-05-08 | 5:38 | 0 | 0 | 0 | 4.5 | 322.7464 | 327.3 | 2.1 | 2.1162 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0276 | 0.0278 | 0.019029 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.856 | 25.7806 | 26.2597 | 322.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 4.3 | 316.6 | 320.9 | 1.9 | 1.9 | 12.42 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0277 | 0.0279 | 0.0186 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.7778 | 26.2611 | 316.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 4.2 | 311.3 | 315.5 | 1.9 | 1.8597 | 12.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0277 | 0.0279 | 0.0183 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.9372 | 26.2186 | 311.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 4.1 | 309.6141 | 313.7 | 1.8 | 1.7556 | 12.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.027786 | 0.028 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 26.0956 | 26.2978 | 309.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 4 | 309.6144 | 313.6 | 1.7 | 1.7455 | 12.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0278 | 0.028 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8689 | 26.3 | 309.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 4 | 309.6856 | 313.7 | 1.7 | 1.7 | 12.38 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0278 | 0.028 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.2437 | 309.7
2008-05-08 | 5:38 | 0 | 0 | 0 | 4 | 308.9152 | 312.9 | 1.8 | 1.7698 | 12.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0278 | 0.028 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7646 | 26.2354 | 308.9
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.9 | 308.8 | 312.7 | 1.7 | 1.7372 | 12.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0278 | 0.0281 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.7584 | 26.2854 | 308.8
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.9 | 308.6144 | 312.5 | 1.7 | 1.7 | 12.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0278 | 0.0281 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0759 | 26.206 | 308.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.9 | 308.6 | 312.5 | 1.8 | 1.7573 | 12.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0279 | 0.0281 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9537 | 26.2732 | 308.6
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.5146 | 312.3 | 1.7 | 1.6526 | 12.42 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2146 | 0.0279 | 0.028185 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8481 | 26.3 | 308.5
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3144 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0279 | 0.0282 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.862 | 26.3 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.4 | 27 | 5 | 195 | 59.3 | 0.0279 | 0.0282 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9677 | 26.3 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3 | 312.1 | 1.8 | 1.7603 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.6 | 27 | 5 | 195 | 59.3 | 0.0279 | 0.0282 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.73 | 26.3 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.028 | 0.0282 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9076 | 26.2308 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.2856 | 312.1 | 1.7 | 1.6962 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.028 | 0.0282 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26 | 26.2484 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3 | 312.1 | 1.8 | 1.7593 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.028 | 0.028286 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9448 | 26.3 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.8 | 308.3 | 312.1 | 1.7 | 1.7427 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.028 | 0.0283 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.2933 | 308.3
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.7 | 308.2 | 311.9 | 1.9 | 1.9 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.028 | 0.0283 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.2139 | 308.2
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.7 | 308.2 | 311.9 | 1.8 | 1.8 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.4 | 27 | 5.4 | 195 | 59.2 | 0.028 | 0.0283 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8654 | 26.2 | 308.2
2008-05-08 | 5:38 | 0 | 0 | 0 | 3.6 | 308.1143 | 311.7 | 1.9 | 1.9 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.6 | 27 | 5 | 195 | 59.2 | 0.028 | 0.0283 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.8571 | 25.9892 | 26.2 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308 | 311.6 | 1.9 | 1.9 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.028 | 0.0283 | 0.0182 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 26.0287 | 26.2238 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.0856 | 311.7 | 2 | 1.9721 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.028 | 0.0283 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.8 | 26.297 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.5 | 308.1856 | 311.7 | 1.9 | 1.9 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2 | 0.028 | 0.0283 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.8 | 26.2178 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.5 | 308.2 | 311.7 | 1.9 | 1.9404 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.028 | 0.028386 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.2 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.5 | 308.2 | 311.7 | 1.9 | 1.9 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2 | 0.028 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8 | 26.2 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.5 | 308.3854 | 311.9 | 1.9 | 1.9 | 12.5 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.028 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.8397 | 26.2198 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.5 | 308.3144 | 311.8 | 1.9 | 1.8561 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2144 | 0.028086 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.9978 | 26.2989 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.3 | 311.9 | 1.7 | 1.7425 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 26.0783 | 26.3 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.3854 | 312 | 1.8 | 1.8 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 26.0424 | 26.2424 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.4 | 312 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.9265 | 26.2 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.3 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.816 | 26.216 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.3144 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8952 | 26.2952 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.4 | 312 | 1.7 | 1.7403 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0281 | 0.0284 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9 | 26.2256 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.5 | 312.1 | 1.6 | 1.6 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0281 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9 | 26.2536 | 308.5
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.4146 | 312.1 | 1.7 | 1.6581 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2146 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9329 | 26.2671 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.4 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.9639 | 26.212 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.4 | 312.1 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.7266 | 26.2911 | 308.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.3 | 312 | 1.7 | 1.6586 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8411 | 26.2295 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.3 | 312 | 1.6 | 1.6208 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8502 | 26.2498 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.2144 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.771 | 26.3 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.1144 | 311.8 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7247 | 26.2918 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.1854 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9622 | 26.2126 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.2 | 311.9 | 1.6 | 1.5508 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8667 | 26.2 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.3 | 312 | 1.6 | 1.6424 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8 | 26.2 | 308.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.7 | 308.2 | 311.9 | 1.6 | 1.6 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8754 | 26.2 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.1 | 311.7 | 1.7 | 1.6582 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 26.087 | 26.2043 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.1 | 311.7 | 1.7 | 1.6582 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0282 | 0.0285 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8498 | 26.2834 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.0144 | 311.6 | 1.6 | 1.6 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0282 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9254 | 26.2373 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 307.9 | 311.5 | 1.7 | 1.6561 | 12.5 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0282 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.958 | 26.2 | 307.9
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 307.9 | 311.5 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9424 | 26.2212 | 307.9
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 307.9856 | 311.6 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002.144 | 26.0984 | 26.2996 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.1712 | 311.8 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.2144 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7815 | 26.2204 | 308.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.1144 | 311.7 | 1.6 | 1.6415 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.2 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7588 | 26.2588 | 308.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308.0144 | 311.6 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.3 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.914 | 26.3 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308 | 311.6 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0481 | 26.3 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 308 | 311.6 | 1.6 | 1.6 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.2143 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8102 | 26.3 | 308
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 307.9144 | 311.5 | 1.7 | 1.6596 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8 | 26.2244 | 307.9
2008-05-08 | 5:39 | 0 | 0 | 0 | 3.6 | 307.8 | 311.4 | 1.7 | 1.7396 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 27 | 5 | 195 | 59.2854 | 0.0283 | 0.0286 | 0.0181 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.9647 | 26.2 | 307.8
2008-05-08 | 5:39 | 0 | 0 | 0 | 4 | 321.4632 | 325.5 | 1.8 | 1.8 | 12.66 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.028386 | 0.028686 | 0.018942 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 26.0658 | 26.2342 | 321.5
2008-05-08 | 5:39 | 0 | 0 | 0 | 5.5 | 343.8511 | 349.3 | 1.9 | 1.8574 | 13.24 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.029641 | 0.029727 | 0.02027 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002.1443 | 26 | 26.2866 | 343.9
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.9 | 350.6 | 355.5 | 1.8 | 1.8 | 13.78 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2146 | 0.030356 | 0.030556 | 0.0206 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26 | 26.2074 | 350.6
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.5 | 354.0712 | 358.6 | 1.9 | 1.9 | 14.02 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.030786 | 0.031086 | 0.0208 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8564 | 26.2 | 354.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.4 | 354.9557 | 359.4 | 1.8 | 1.8 | 14.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.030886 | 0.031186 | 0.020886 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8 | 26.2509 | 355
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.4 | 356.0712 | 360.5 | 1.9 | 1.9 | 14.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.031 | 0.0313 | 0.0209 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1001.144 | 25.8 | 26.2699 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.4 | 356.1 | 360.5 | 1.9 | 1.8509 | 14.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.3 | 0.0311 | 0.0314 | 0.0209 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.8095 | 26.2095 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.3 | 356.1 | 360.4 | 1.8 | 1.8416 | 14.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.0312 | 0.031586 | 0.0209 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.8887 | 26.2887 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.3 | 356.1 | 360.4 | 1.9 | 1.9 | 14.47 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2144 | 0.0313 | 0.0316 | 0.0209 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001.8557 | 25.8322 | 26.3 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.3 | 356.0144 | 360.3 | 1.9 | 1.9 | 14.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2856 | 0.031386 | 0.0317 | 0.0209 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9412 | 26.3 | 356
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.5 | 355.9144 | 360.4 | 1.9 | 1.9 | 14.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2144 | 0.031486 | 0.031786 | 0.020986 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002.856 | 26.0736 | 26.2736 | 355.9
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.5 | 355.9856 | 360.5 | 1.8 | 1.8 | 14.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0315 | 0.0318 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9945 | 26.2055 | 356
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.4 | 356.1 | 360.5 | 2 | 1.9588 | 14.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2 | 0.0316 | 0.0319 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9152 | 26.2848 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.4 | 356.1 | 360.5 | 1.8 | 1.8429 | 14.6 | 12.9 | 957 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.031686 | 0.032 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9 | 26.3 | 356.1
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.3 | 356.2 | 360.5 | 2 | 1.9958 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0317 | 0.032 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.986 | 26.257 | 356.2
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.2 | 356.2854 | 360.5 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2146 | 0.031785 | 0.0321 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0779 | 26.2221 | 356.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.4856 | 360.6 | 1.9 | 1.94 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0318 | 0.0322 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 26.0017 | 26.3 | 356.5
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.5 | 360.6 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0319 | 0.032286 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 26.0809 | 26.3 | 356.5
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.4 | 360.5 | 1.9 | 1.9411 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.2146 | 0.0319 | 0.0323 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.92 | 26.3 | 356.4
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.3 | 360.4 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.2144 | 0.031986 | 0.0323 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9176 | 26.2608 | 356.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.3 | 360.4 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.032 | 0.0324 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 26.0444 | 26.2 | 356.3
2008-05-08 | 5:39 | 0 | 0 | 0 | 4.1 | 356.3 | 360.4 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.032 | 0.0324 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.8069 | 26.2 | 356.3
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.3 | 360.4 | 1.9 | 1.9 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0321 | 0.032486 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0308 | 26.2769 | 356.3
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.2856 | 360.4 | 2 | 2 | 14.6 | 12.9 | 957 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0321 | 0.0325 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.1 | 26.2439 | 356.3
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.2144 | 360.3 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0321 | 0.0325 | 0.021 | 0.0001 |  | n/a |  | 30.0759 | 53 | 1039 | 1002 | 25.9943 | 26.2352 | 356.2
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.2 | 360.3 | 2 | 1.9586 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0322 | 0.0326 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8143 | 26.3 | 356.2
2008-05-08 | 5:40 | 0 | 0 | 0 | 4 | 356.1 | 360.1 | 1.9 | 1.9 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.2 | 0.0322 | 0.0326 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8938 | 26.3 | 356.1
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.0144 | 360.1 | 1.9 | 1.9412 | 14.6 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.3 | 0.0322 | 0.0326 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9731 | 26.3 | 356
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356 | 360.1 | 2 | 1.9508 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2856 | 0.032286 | 0.0326 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9478 | 26.3 | 356
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.2 | 360.3 | 2 | 2.0431 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.2856 | 0.0323 | 0.0327 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.9 | 26.3 | 356.2
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.3 | 360.4 | 2 | 2.0289 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.3 | 0.032386 | 0.0327 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8893 | 26.3 | 356.3
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.5711 | 360.7 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.2144 | 0.0324 | 0.032786 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.8101 | 26.3 | 356.6
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.6 | 360.7 | 2.1 | 2.0597 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0324 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8 | 26.3 | 356.6
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.6856 | 360.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0324 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.8 | 26.3 | 356.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.7 | 360.8 | 2.1 | 2.0592 | 14.6 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0324 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8274 | 26.3 | 356.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.2 | 356.6144 | 360.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0324 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9 | 26.3 | 356.6
2008-05-08 | 5:40 | 0 | -0.0011 | -11.4384 | 4.2 | 356.5 | 360.7 | 2.1 | 2.0735 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0324 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.9 | 26.3 | 356.5
2008-05-08 | 5:40 | 0 | -0.0016 | -15.8541 | 4.2 | 356.6 | 360.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0325 | 0.0328 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.9 | 26.3 | 356.6
2008-05-08 | 5:40 | 0 | -0.0008 | -7.6039 | 4.2 | 356.7 | 360.9 | 2 | 2.0413 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.3 | 0.0325 | 0.0329 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9444 | 26.2556 | 356.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.2 | 356.7 | 360.9 | 2 | 2.0484 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0325 | 0.0329 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0236 | 26.2236 | 356.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.2 | 356.8 | 361 | 2 | 2.0431 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0325 | 0.0329 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 26.0971 | 26.2971 | 356.8
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.2 | 356.9 | 361.1 | 2.1 | 2.06 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.032586 | 0.0329 | 0.021 | 0.0001 |  | n/a |  | 30.0751 | 53 | 1039 | 1002 | 26.018 | 26.218 | 356.9
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.9144 | 361 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2856 | 0.0326 | 0.0329 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1001.856 | 25.755 | 26.2613 | 356.9
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.2 | 356.7288 | 360.9 | 2.2 | 2.1565 | 14.6 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26 | 5 | 195 | 59.2 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.6405 | 26.3 | 356.7
2008-05-08 | 5:40 | 0 | 0.0007 | 7.4759 | 4.2 | 356.7 | 360.9 | 2 | 2.049 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.2144 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.759 | 26.2803 | 356.7
2008-05-08 | 5:40 | 0 | 0.0029 | 29.4305 | 4.2 | 356.6 | 360.8 | 2.1 | 2.1 | 14.6 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.996 | 26.2013 | 356.6
2008-05-08 | 5:40 | 0 | 0.0021 | 20.8024 | 4.2 | 356.4 | 360.6 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9219 | 26.2781 | 356.4
2008-05-08 | 5:40 | 0 | 0.002 | 20 | 4.2 | 356.4 | 360.6 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 26.0147 | 26.3 | 356.4
2008-05-08 | 5:40 | 0 | 0.0014 | 14.1982 | 4.2 | 356.4856 | 360.7 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 26.0275 | 26.3 | 356.5
2008-05-08 | 5:40 | 0 | 0.001 | 10 | 4.2 | 356.5 | 360.7 | 2.1 | 2.0768 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0326 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.9301 | 26.3 | 356.5
2008-05-08 | 5:40 | 0 | 0.001 | 10 | 4.2 | 356.6854 | 360.9 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0327 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 26.0899 | 26.3 | 356.7
2008-05-08 | 5:40 | 0 | 0.001 | 10 | 4.2 | 356.7856 | 361 | 2 | 2.0442 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0327 | 0.033 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002.144 | 25.9516 | 26.3 | 356.8
2008-05-08 | 5:40 | 0 | 0.001 | 10 | 4.2 | 356.8 | 361 | 2.1 | 2.0602 | 14.6 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0327 | 0.033086 | 0.021 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9534 | 26.3 | 356.8
2008-05-08 | 5:40 | 0 | 0.0003 | 2.8047 | 4.1 | 356.8 | 360.9 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2854 | 0.0327 | 0.0331 | 0.021 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.9674 | 26.3 | 356.8
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.7 | 360.8 | 2.1 | 2.059 | 14.6 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0327 | 0.0331 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.9 | 26.3 | 356.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.8 | 360.9 | 2 | 2 | 14.6 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0327 | 0.0331 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.9 | 26.3 | 356.8
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.1 | 356.8 | 360.9 | 2.1 | 2.1 | 14.6 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0328 | 0.0332 | 0.021 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1039 | 1002 | 25.8298 | 26.3 | 356.8
2008-05-08 | 5:40 | 0 | 0 | 0 | 4 | 356.9 | 360.9 | 2 | 2 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0328 | 0.0332 | 0.021 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.899 | 26.3 | 356.9
2008-05-08 | 5:40 | 0 | 0 | 0 | 4 | 356.9856 | 361 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0328 | 0.0332 | 0.021 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 26 | 26.2713 | 357
2008-05-08 | 5:40 | 0 | 0 | 0 | 4.9 | 386.3576 | 391.3 | 2.2 | 2.2403 | 14.87 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.032971 | 0.033286 | 0.022713 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.9921 | 26.2079 | 386.4
2008-05-08 | 5:40 | 0 | -0.0012 | -11.6874 | 5.8 | 405.3784 | 411.2 | 2.6 | 2.6445 | 15.74 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.034585 | 0.034785 | 0.023857 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1001 | 25.9129 | 26.2871 | 405.4
2008-05-08 | 5:40 | 0 | -0.0016 | -15.8333 | 5.3 | 406.3144 | 411.6 | 2.5 | 2.4998 | 16.12 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.035086 | 0.035386 | 0.0239 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001.856 | 25.7674 | 26.3 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.2 | 406.2 | 411.4 | 2.3 | 2.3 | 16.27 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0352 | 0.035586 | 0.0239 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.8367 | 26.3 | 406.2
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.2 | 406.2856 | 411.5 | 2.4 | 2.369 | 16.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.035286 | 0.035686 | 0.0239 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9504 | 26.3 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.7 | 406.4 | 411.1 | 2.3 | 2.2565 | 16.47 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2 | 0.035686 | 0.036186 | 0.0239 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8121 | 26.3 | 406.4
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.5 | 406.3 | 410.8 | 2.3 | 2.2692 | 16.5 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.3 | 0.0358 | 0.036286 | 0.0239 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0492 | 26.3 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.5 | 406.2854 | 410.8 | 2.2 | 2.2 | 16.57 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.2854 | 0.035985 | 0.0364 | 0.0239 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.9129 | 26.3 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.4 | 406.3 | 410.7 | 2.2 | 2.2418 | 16.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.2856 | 0.036086 | 0.0365 | 0.0239 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.8 | 26.3 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.5 | 406.2 | 410.7 | 2.3 | 2.2565 | 16.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0361 | 0.0366 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.8418 | 26.3 | 406.2
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 4.7 | 406.1146 | 410.8 | 2.3 | 2.2802 | 16.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2854 | 0.036185 | 0.0366 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 26 | 26.2999 | 406.1
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.1 | 406.2712 | 411.4 | 2.3 | 2.2586 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.036286 | 0.036686 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 26 | 26.2207 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.5 | 406.3 | 411.8 | 2.2 | 2.2422 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0364 | 0.0367 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.9415 | 26.2 | 406.3
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.5 | 406.5 | 412 | 2.3 | 2.25 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.036486 | 0.0367 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9377 | 26.2377 | 406.5
2008-05-08 | 5:40 | 0 | -0.001 | -10 | 5.5 | 406.4856 | 412 | 2.2 | 2.2 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0365 | 0.0368 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 25.9491 | 26.3 | 406.5
2008-05-08 | 5:40 | 0 | -0.0008 | -7.8649 | 5.5 | 406.5 | 412 | 2.3 | 2.2565 | 16.7 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0365 | 0.0368 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7114 | 26.3 | 406.5
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.5 | 406.5856 | 412.1 | 2.2 | 2.2 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.036586 | 0.036886 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.9258 | 26.2247 | 406.6
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.5 | 406.4144 | 411.9 | 2.3 | 2.3 | 16.7 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.036686 | 0.0369 | 0.024 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1039 | 1003 | 25.9455 | 26.2545 | 406.4
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.5 | 406.4856 | 412 | 2.2 | 2.2 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0367 | 0.037 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 25.9338 | 26.3 | 406.5
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.5 | 406.5 | 412 | 2.2 | 2.2429 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0367 | 0.037 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.987 | 26.3 | 406.5
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.5 | 406.6865 | 412.2 | 2.2 | 2.1815 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0368 | 0.037086 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 25.9078 | 26.3 | 406.7
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.4 | 406.7856 | 412.2 | 2.2 | 2.1615 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0368 | 0.0371 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9714 | 26.3 | 406.8
2008-05-08 | 5:40 | 0 | 0 | 0 | 5.4 | 406.8 | 412.2 | 2.2 | 2.1626 | 16.7 | 12.9 | 956 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0368 | 0.0371 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0506 | 26.2494 | 406.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.4 | 406.8854 | 412.3 | 2.1 | 2.1 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.036885 | 0.037185 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9805 | 26.2299 | 406.9
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.3 | 407.0856 | 412.4 | 2.2 | 2.1542 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0369 | 0.0372 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.7182 | 26.3 | 407.1
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.2 | 407.1856 | 412.4 | 2.1 | 2.1 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0369 | 0.0373 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002.1443 | 25.8766 | 26.3 | 407.2
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.2 | 407.4711 | 412.7 | 2.2 | 2.1565 | 16.7 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.037 | 0.0373 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.9675 | 26.3 | 407.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.2 | 407.5 | 412.7 | 2 | 2.0436 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.037 | 0.0373 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.9067 | 26.2533 | 407.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.2 | 407.5 | 412.7 | 2.1 | 2.1 | 16.7 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.037 | 0.0374 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.8518 | 26.2259 | 407.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.1 | 407.7 | 412.8 | 2 | 2 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.037086 | 0.0374 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1001 | 26 | 26.3 | 407.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.1 | 407.7856 | 412.9 | 2 | 2.0408 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0371 | 0.037486 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 26 | 26.3 | 407.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 5 | 407.8856 | 412.9 | 2 | 2 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 27 | 5 | 195 | 59.2 | 0.0371 | 0.0375 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.873 | 26.3 | 407.9
2008-05-08 | 5:41 | 0 | 0 | 0 | 5 | 407.8 | 412.8 | 2 | 2 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0372 | 0.0375 | 0.024 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1039 | 1001 | 25.8427 | 26.2573 | 407.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.7144 | 412.6 | 2 | 2 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30.0751 | 53 | 1039 | 1001.856 | 25.9 | 26.222 | 407.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.6144 | 412.5 | 1.9 | 1.94 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.8976 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.5 | 412.4 | 2 | 2 | 16.7 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7391 | 26.3 | 407.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.6 | 412.5 | 2 | 2 | 16.7 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.7596 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.5856 | 412.5 | 2 | 1.9571 | 16.7 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.8 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.9 | 407.6 | 412.5 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0372 | 0.0376 | 0.024 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.8 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.8 | 407.6 | 412.4 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2854 | 0.037285 | 0.0377 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1003 | 25.8 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.8 | 407.6 | 412.4 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.3 | 0.0373 | 0.0377 | 0.024 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.7236 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 407.7711 | 412.5 | 1.9 | 1.9424 | 16.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.2856 | 0.0373 | 0.0377 | 0.024 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002.1443 | 25.867 | 26.3 | 407.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 407.8 | 412.5 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0373 | 0.0378 | 0.024 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9302 | 26.3 | 407.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 407.6144 | 412.3 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0373 | 0.0378 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8423 | 26.3 | 407.6
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 407.7 | 412.4 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0374 | 0.0378 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0798 | 26.3 | 407.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 407.9712 | 412.7 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 27 | 5 | 195 | 59.2856 | 0.0374 | 0.0378 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 26.0275 | 26.2275 | 408
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 408 | 412.7 | 2 | 1.9536 | 16.7 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0374 | 0.0378 | 0.0241 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 26.0513 | 26.2513 | 408
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.7 | 408 | 412.7 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0374 | 0.0378 | 0.0241 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 26.0392 | 26.3 | 408
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.6 | 408.1 | 412.7 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0374 | 0.0379 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9 | 26.3 | 408.1
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.6 | 408.4 | 413 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0374 | 0.0379 | 0.0241 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1039 | 1002 | 25.9 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.5 | 413 | 1.9 | 1.9369 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.037486 | 0.0379 | 0.0241 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.763 | 26.3 | 408.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.4144 | 412.9 | 1.9 | 1.9 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.6 | 195 | 59.3 | 0.0375 | 0.0379 | 0.0241 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.7477 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.4 | 412.9 | 1.9 | 1.9412 | 16.7 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.0375 | 0.0379 | 0.0241 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.7731 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.4856 | 413 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.0375 | 0.037986 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7124 | 26.3 | 408.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.4 | 412.9 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2144 | 0.0375 | 0.038 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8707 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.5 | 408.4144 | 412.9 | 2 | 1.9786 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.0375 | 0.038 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9 | 26.2355 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.4 | 408.3144 | 412.7 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.4 | 195 | 59.3 | 0.0375 | 0.038 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8131 | 26.2435 | 408.3
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.4 | 408.4 | 412.8 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0375 | 0.038 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7225 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.3 | 408.4 | 412.7 | 1.9 | 1.9 | 16.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2146 | 0.0375 | 0.038 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8064 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.3 | 408.4 | 412.7 | 1.9 | 1.9421 | 16.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2143 | 0.0375 | 0.038086 | 0.0241 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0439 | 26.3 | 408.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 4.4 | 412.8497 | 417.2 | 1.9 | 1.9 | 16.71 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 27 | 5 | 195 | 59.2856 | 0.037586 | 0.0381 | 0.024357 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.8557 | 26.1606 | 26.3 | 412.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 8.2 | 453.1565 | 461.4 | 2.2 | 2.2231 | 17.18 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.038994 | 0.038877 | 0.026726 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0407 | 26.3 | 453.2
2008-05-08 | 5:41 | 0 | 0 | 0 | 7.5 | 480.3912 | 487.9 | 2.6 | 2.6475 | 18.55 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.041542 | 0.041613 | 0.028328 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8 | 26.3 | 480.4
2008-05-08 | 5:41 | 0 | -0.0002 | -2.296 | 7 | 473.5368 | 480.6 | 2.2 | 2.2432 | 19 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2 | 0.0421 | 0.0423 | 0.027872 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8 | 26.3 | 473.5
2008-05-08 | 5:41 | 0 | -0.001 | -10 | 6.7 | 467.7064 | 474.4 | 2.2 | 2.1587 | 18.88 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2 | 0.0421 | 0.042386 | 0.027543 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.1094 | 26.3 | 467.7
2008-05-08 | 5:41 | 0 | -0.001 | -10 | 6.3 | 462.1 | 468.4 | 2.1 | 2.1 | 18.8 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2144 | 0.0421 | 0.0424 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0868 | 26.3 | 462.1
2008-05-08 | 5:41 | 0 | -0.001 | -10 | 6.3 | 462 | 468.3 | 2.2 | 2.2 | 18.8 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.0422 | 0.0425 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9644 | 26.3 | 462
2008-05-08 | 5:41 | 0 | -0.001 | -10 | 6.3 | 462.1 | 468.4 | 2.2 | 2.2 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0422 | 0.0425 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9293 | 26.3 | 462.1
2008-05-08 | 5:41 | 0 | -0.0006 | -6.3083 | 6.2 | 462 | 468.2 | 2.3 | 2.2528 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0422 | 0.0425 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0884 | 26.3 | 462
2008-05-08 | 5:41 | 0 | 0 | 0 | 6.2 | 462.0856 | 468.3 | 2.2 | 2.2 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0422 | 0.0425 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001.1443 | 25.9533 | 26.3 | 462.1
2008-05-08 | 5:41 | 0 | 0 | 0 | 6.2 | 461.9288 | 468.1 | 2.3 | 2.2722 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0422 | 0.0425 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001.856 | 25.8474 | 26.3 | 461.9
2008-05-08 | 5:41 | 0 | 0 | 0 | 6.1 | 461.8 | 467.9 | 2.3 | 2.3 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0422 | 0.042586 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8955 | 26.3 | 461.8
2008-05-08 | 5:41 | 0 | 0 | 0 | 6.1 | 461.7 | 467.8 | 2.3 | 2.341 | 18.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.0422 | 0.0426 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.067 | 26.3 | 461.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 6 | 461.7 | 467.7 | 2.3 | 2.3 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0423 | 0.0426 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8295 | 26.3 | 461.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 6 | 461.7 | 467.7 | 2.3 | 2.342 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2854 | 0.0423 | 0.0426 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8694 | 26.2306 | 461.7
2008-05-08 | 5:41 | 0 | 0 | 0 | 6 | 461.5 | 467.5 | 2.4 | 2.4 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0423 | 0.0426 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9486 | 26.2486 | 461.5
2008-05-08 | 5:41 | 0 | 0 | 0 | 6 | 461.4144 | 467.4 | 2.4 | 2.4 | 18.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2856 | 0.0423 | 0.042686 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9445 | 26.3 | 461.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.9 | 461.4 | 467.3 | 2.5 | 2.4572 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8202 | 26.3 | 461.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.9 | 461.4 | 467.3 | 2.4 | 2.3587 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.2143 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0583 | 26.3 | 461.4
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.9 | 461.3 | 467.2 | 2.4 | 2.4 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8385 | 26.3 | 461.3
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.9 | 461.2144 | 467.1 | 2.4 | 2.4 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.8784 | 26.3 | 461.2
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.7 | 461.2143 | 466.9 | 2.4 | 2.4406 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.1 | 26.3 | 461.2
2008-05-08 | 5:41 | 0 | 0 | 0 | 5.7 | 461.1144 | 466.8 | 2.5 | 2.5 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 26.091 | 26.3 | 461.1
2008-05-08 | 5:42 | 0 | 0 | 0 | 5.7 | 461 | 466.7 | 2.5 | 2.4875 | 18.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8534 | 26.3 | 461
2008-05-08 | 5:42 | 0 | -0.0007 | -7.1002 | 5.7 | 460.9 | 466.6 | 2.5 | 2.5 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2856 | 0.0423 | 0.0427 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9841 | 26.3 | 460.9
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.7 | 460.7288 | 466.4 | 2.5 | 2.5418 | 18.9 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0423 | 0.042786 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.1 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.6 | 460.7 | 466.3 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0423 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0606 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.6 | 460.7856 | 466.4 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0423 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9022 | 26.3 | 460.8
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.6 | 460.8 | 466.4 | 2.7 | 2.7 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0423 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8219 | 26.3 | 460.8
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.5 | 460.5289 | 466 | 2.6 | 2.6284 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7426 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.5 | 460.5 | 466 | 2.7 | 2.7 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2854 | 0.0423 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1003 | 25.7 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 460.7712 | 466.2 | 2.6 | 2.5575 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0423 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002.144 | 25.7 | 26.3 | 460.8
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 460.9 | 466.3 | 2.6 | 2.6434 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.7 | 26.3 | 460.9
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 460.9 | 466.3 | 2.5 | 2.5435 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 26.0705 | 26.3 | 460.9
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 460.9856 | 466.4 | 2.5 | 2.5432 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9867 | 26.3 | 461
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 461 | 466.4 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8325 | 26.3 | 461
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 461.1 | 466.5 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0428 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.8767 | 26.3 | 461.1
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.3 | 461.1 | 466.4 | 2.7 | 2.6574 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.0424 | 0.042886 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.7191 | 26.3 | 461.1
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5 | 461.1 | 466.1 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0423 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.91 | 26.3 | 461.1
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5 | 461 | 466 | 2.7 | 2.7 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2144 | 0.0423 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 26.0493 | 26.3 | 461
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5 | 461 | 466 | 2.7 | 2.7 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.2856 | 0.042386 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1001 | 26.0432 | 26.3 | 461
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5 | 460.7288 | 465.7 | 2.8 | 2.7524 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0424 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8926 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 4.9 | 460.6144 | 465.5 | 2.7 | 2.6591 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0424 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8131 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | -0.0005 | -4.5382 | 4.8 | 460.6 | 465.4 | 2.7 | 2.7292 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9322 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.8 | 460.6 | 465.4 | 2.7 | 2.6578 | 18.9 | 12.9 | 956 | 963 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0429 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9547 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.8 | 460.6 | 465.4 | 2.7 | 2.6714 | 18.9 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.042986 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8511 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.8 | 460.6 | 465.4 | 2.7 | 2.7 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2146 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.7 | 460.6712 | 465.4 | 2.6 | 2.6402 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.7 | 460.7 | 465.4 | 2.7 | 2.7136 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2144 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9477 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.7 | 460.5 | 465.2 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.1 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.7 | 460.3144 | 465 | 2.7 | 2.6574 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0387 | 26.3 | 460.3
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.6 | 460.4 | 465 | 2.6 | 2.6 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8012 | 26.3 | 460.4
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.6 | 460.4 | 465 | 2.5 | 2.5399 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7212 | 26.3 | 460.4
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.6 | 460.4856 | 465.1 | 2.5 | 2.5 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.043 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.9321 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.5 | 460.5 | 465 | 2.5 | 2.54 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.1 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.4 | 460.7 | 465.1 | 2.5 | 2.5402 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0836 | 26.3 | 460.7
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.4 | 460.6144 | 465 | 2.5 | 2.5445 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0044 | 26.3 | 460.6
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.4 | 460.7712 | 465.2 | 2.5 | 2.5466 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26 | 26.3 | 460.8
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.4 | 460.8 | 465.2 | 2.5 | 2.5 | 18.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26 | 26.3 | 460.8
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.3 | 460.5144 | 464.8 | 2.7 | 2.6784 | 18.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0332 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.3 | 460.5 | 464.8 | 2.7 | 2.6747 | 18.9 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.0272 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0755 | 26.3 | 460.5
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.4 | 466.7488 | 471.1 | 3.1 | 3.1412 | 18.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0424 | 0.0431 | 0.027542 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9181 | 26.3 | 466.7
2008-05-08 | 5:42 | 0 | 0 | 0 | 6.1 | 482.7615 | 488.8 | 3.1 | 3.0573 | 19.31 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.042585 | 0.0432 | 0.028427 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9 | 26.3 | 482.8
2008-05-08 | 5:42 | -0.009 | 0 | 0 | 6.2 | 500.9536 | 507.2 | 3 | 3.0355 | 19.92 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.043542 | 0.044042 | 0.029571 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9499 | 26.3 | 501
2008-05-08 | 5:42 | -0.01 | 0 | 0 | 5.5 | 504.5856 | 510 | 3.1 | 3.1 | 20.31 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.044257 | 0.044842 | 0.0297 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.9419 | 26.3 | 504.6
2008-05-08 | 5:42 | -0.01 | 0 | 0 | 4.8 | 506.1834 | 510.9 | 3 | 3.0437 | 20.56 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.044856 | 0.045642 | 0.0298 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8082 | 26.3 | 506.2
2008-05-08 | 5:42 | -0.01 | 0 | 0 | 4.6 | 506.4856 | 511.1 | 3.1 | 3.1405 | 20.7 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.045086 | 0.045886 | 0.029886 | 0.000186 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8874 | 26.3 | 506.5
2008-05-08 | 5:42 | -0.01 | 0 | 0 | 4.8 | 507.2002 | 512 | 3.1 | 3.1 | 20.77 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.045286 | 0.046 | 0.0299 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7667 | 26.3 | 507.2
2008-05-08 | 5:42 | -0.01 | 0 | 0 | 5 | 507.1856 | 512.2 | 3.2 | 3.1534 | 20.8 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.2856 | 0.0454 | 0.0461 | 0.0299 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.3 | 507.2
2008-05-08 | 5:42 | -0.001 | -0.0002 | -2.2367 | 5.1 | 507.8 | 512.9 | 3 | 2.9584 | 20.86 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 27 | 5.4 | 195 | 59.3 | 0.0454 | 0.0461 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7751 | 26.3 | 507.8
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.4 | 508.0712 | 513.4 | 3 | 2.9598 | 20.9 | 12.9 | 957 | 963 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.6 | 195 | 59.3 | 0.0455 | 0.0462 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9958 | 26.3 | 508.1
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.3 | 508.1 | 513.4 | 2.9 | 2.9 | 20.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.0456 | 0.0462 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9169 | 26.3 | 508.1
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.3 | 508.1856 | 513.5 | 2.9 | 2.9 | 20.9 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.4 | 195 | 59.2856 | 0.0457 | 0.0463 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9624 | 26.3 | 508.2
2008-05-08 | 5:42 | 0 | -0.001 | -10 | 5.2 | 508.5712 | 513.8 | 2.8 | 2.7581 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.6 | 195 | 59.3 | 0.045786 | 0.0464 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0417 | 26.3 | 508.6
2008-05-08 | 5:42 | 0 | -0.0006 | -6.2779 | 5.1 | 508.8712 | 514 | 2.7 | 2.7492 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.0458 | 0.0465 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0791 | 26.3 | 508.9
2008-05-08 | 5:42 | 0 | 0 | 0 | 5.1 | 509.1 | 514.2 | 2.7 | 2.6584 | 21 | 12.9 | 956 | 963 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.0459 | 0.046586 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.9998 | 26.3 | 509.1
2008-05-08 | 5:42 | 0 | 0 | 0 | 5.1 | 509.1856 | 514.3 | 2.6 | 2.6 | 21 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.0459 | 0.0466 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.7622 | 26.3 | 509.2
2008-05-08 | 5:42 | 0 | 0 | 0 | 5.1 | 509.1 | 514.2 | 2.6 | 2.6429 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.0459 | 0.0466 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.8754 | 26.3 | 509.1
2008-05-08 | 5:42 | 0 | 0 | 0 | 5 | 509.1 | 514.1 | 2.5 | 2.5436 | 21 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.3 | 0.045986 | 0.046686 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8869 | 26.3377 | 509.1
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.9 | 509.2 | 514.1 | 2.7 | 2.6509 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.046 | 0.0467 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7169 | 26.3831 | 509.2
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.9 | 508.8433 | 513.7 | 2.6 | 2.6 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.4 | 195 | 59.3 | 0.046 | 0.0468 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.8557 | 25.796 | 26.304 | 508.8
2008-05-08 | 5:42 | 0 | 0 | 0 | 4.8 | 508.7146 | 513.5 | 2.7 | 2.7 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0461 | 0.0468 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8747 | 26.3 | 508.7
2008-05-08 | 5:42 | 0 | 0 | -0.20017 | 4.8 | 508.7 | 513.5 | 2.6 | 2.6 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0461 | 0.0468 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8459 | 26.3 | 508.7
2008-05-08 | 5:43 | 0 | -0.0009 | -8.5404 | 4.7 | 508.7855 | 513.5 | 2.7 | 2.7421 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0461 | 0.0469 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9007 | 26.3 | 508.8
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 4.6 | 508.5144 | 513.1 | 2.8 | 2.8 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.046186 | 0.0469 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003.856 | 26.0746 | 26.3 | 508.5
2008-05-08 | 5:43 | 0 | -0.0005 | -4.7483 | 4.6 | 508.6 | 513.2 | 2.8 | 2.752 | 21 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0462 | 0.047 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 25.9169 | 26.3 | 508.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.5 | 508.6856 | 513.2 | 2.8 | 2.8 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0462 | 0.047 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 26.0421 | 26.3 | 508.7
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.5 | 508.9857 | 513.5 | 2.7 | 2.7433 | 21 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0462 | 0.047 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9993 | 26.3 | 509
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.5 | 509.1 | 513.6 | 2.7 | 2.7403 | 21 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0462 | 0.047 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9591 | 26.3295 | 509.1
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.0144 | 513.4 | 2.7 | 2.7 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.2 | 0.0463 | 0.0471 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0913 | 26.4 | 509
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.2 | 513.6 | 2.7 | 2.7499 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0463 | 0.0471 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0121 | 26.4 | 509.2
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.4 | 513.7 | 2.6 | 2.6 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0463 | 0.0472 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8664 | 26.3332 | 509.4
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.4 | 513.7 | 2.7 | 2.7 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0464 | 0.0472 | 0.03 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.754 | 26.3 | 509.4
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.3 | 513.7 | 2.6 | 2.6046 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0464 | 0.0472 | 0.030086 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7508 | 26.3 | 509.3
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.3 | 513.7 | 2.7 | 2.7 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0464 | 0.0472 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9046 | 26.3046 | 509.3
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.4712 | 513.9 | 2.7 | 2.7 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0464 | 0.047286 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9837 | 26.3837 | 509.5
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.5 | 513.9 | 2.7 | 2.7496 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0464 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26 | 26.3371 | 509.5
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.5709 | 514 | 2.8 | 2.8 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.046485 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8736 | 26.3 | 509.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.6 | 514 | 2.7 | 2.742 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7639 | 26.3 | 509.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.8 | 514.2 | 2.9 | 2.8592 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26 | 26.3 | 509.8
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.7859 | 514.2 | 2.8 | 2.8391 | 21 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2859 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26 | 26.3 | 509.8
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.8 | 514.2 | 3 | 3 | 21 | 13 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8823 | 26.3 | 509.8
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 510.0712 | 514.5 | 3 | 3 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.914 | 26.338 | 510.1
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.8 | 514.2 | 3.1 | 3.1421 | 21 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26.0142 | 26.3828 | 509.8
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.7144 | 514.1 | 3.2 | 3.2 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0473 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6186 | 26.3037 | 509.7
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.7 | 514.1 | 3.3 | 3.3408 | 21 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.047386 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6752 | 26.3 | 509.7
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.5 | 513.9 | 3.5 | 3.5 | 21 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0474 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 25.7545 | 26.3 | 509.5
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.4 | 509.6 | 514 | 3.5 | 3.5 | 21 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0465 | 0.0474 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 25.7322 | 26.3339 | 509.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.6 | 513.9 | 3.6 | 3.5602 | 21 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0474 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 25.6131 | 26.3869 | 509.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.6 | 513.9 | 3.4 | 3.4 | 21 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0474 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1004 | 25.6922 | 26.3078 | 509.6
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.8 | 514.1 | 3.5 | 3.5 | 21 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0474 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.3 | 509.8
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.9144 | 514.2 | 3.4 | 3.356 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7506 | 26.3 | 509.9
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.9 | 514.2 | 3.4 | 3.37 | 21 | 13 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8297 | 26.3297 | 509.9
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 510 | 514.3 | 3.3 | 3.3 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9089 | 26.3911 | 510
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 510 | 514.3 | 3.3 | 3.3 | 21 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.988 | 26.312 | 510
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 510 | 514.3 | 3.3 | 3.3 | 21 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7983 | 26.3672 | 510
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.9146 | 514.2 | 3.3 | 3.2558 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2146 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.3536 | 509.9
2008-05-08 | 5:43 | 0 | 0 | 0 | 4.3 | 509.9 | 514.2 | 3.3 | 3.2553 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0466 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.3 | 509.9
2008-05-08 | 5:43 | 0 | 0 | -0.13333 | 4.3 | 509.9 | 514.2 | 3.2 | 3.2447 | 21 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.046686 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7047 | 26.3047 | 509.9
2008-05-08 | 5:43 | 0 | -0.0008 | -8.4667 | 4.3 | 509.9 | 514.2 | 3.3 | 3.3 | 21 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0467 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7835 | 26.3835 | 509.9
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 4.3 | 509.9 | 514.2 | 3.3 | 3.3 | 21 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0467 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8 | 26.3371 | 509.9
2008-05-08 | 5:43 | 0 | -0.0005 | -4.7925 | 4.3 | 510.0856 | 514.4 | 3.3 | 3.3385 | 21 | 13 | 954 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0467 | 0.0475 | 0.0301 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8423 | 26.3 | 510.1
2008-05-08 | 5:43 | 0 | 0 | 0 | 3.4 | 488.2336 | 491.6 | 3.2 | 3.2421 | 20.79 | 12.9 | 954 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.046529 | 0.0475 | 0.028844 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9 | 26.3 | 488.2
2008-05-08 | 5:43 | 0 | 0 | 0 | 2.5 | 470.313 | 472.8 | 3.4 | 3.4 | 20.03 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.045987 | 0.047072 | 0.027772 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8989 | 26.3 | 470.3
2008-05-08 | 5:43 | 0 | 0 | 0 | 2.5 | 468.3 | 470.8 | 3.4 | 3.4 | 19.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0456 | 0.0467 | 0.0276 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7404 | 26.3 | 468.3
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.9 | 464.8592 | 466.7 | 3.4 | 3.4453 | 19.4 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.045329 | 0.046514 | 0.027414 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.3589 | 464.9
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.8 | 464.5 | 466.3 | 3.4 | 3.3788 | 19.29 | 12.9 | 954 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0452 | 0.0464 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8523 | 26.3619 | 464.5
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.9 | 464.4 | 466.3 | 3.4 | 3.3683 | 19.2 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0451 | 0.0463 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 26.0307 | 26.3173 | 464.4
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.9 | 464.2288 | 466.1 | 3.4 | 3.3568 | 19.1 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.045014 | 0.0462 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.7142 | 26.3964 | 464.2
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.3 | 463.9416 | 465.2 | 3.3 | 3.3 | 19.1 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044814 | 0.0461 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9251 | 26.325 | 463.9
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.2 | 464.1 | 465.3 | 3.3 | 3.337 | 19.1 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0447 | 0.046 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 26 | 26.3 | 464.1
2008-05-08 | 5:43 | 0 | 0 | 0 | 1.3 | 464 | 465.3 | 3.3 | 3.3 | 19 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044614 | 0.045914 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9321 | 26.3 | 464
2008-05-08 | 5:43 | 0 | -0.0007 | -6.7723 | 1.4 | 463.9144 | 465.3 | 3.4 | 3.3587 | 19 | 13 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0446 | 0.045814 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7869 | 26.3 | 463.9
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.4 | 463.9 | 465.3 | 3.3 | 3.2629 | 19 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0445 | 0.0458 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7078 | 26.3 | 463.9
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.5 | 463.8 | 465.3 | 3.3 | 3.3 | 19 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0445 | 0.0457 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8429 | 26.3 | 463.8
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.6 | 463.6144 | 465.2 | 3.2 | 3.2 | 19 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0444 | 0.0456 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7987 | 26.3507 | 463.6
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.7 | 463.5 | 465.2 | 3.2 | 3.249 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0443 | 0.045514 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4 | 463.5
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.7 | 463.3 | 465 | 3.3 | 3.3 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0443 | 0.0455 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.709 | 26.391 | 463.3
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.8 | 463.2 | 465 | 3.2 | 3.2428 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0442 | 0.0454 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7881 | 26.3119 | 463.2
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.9 | 463.1 | 465 | 3.3 | 3.3392 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044114 | 0.0453 | 0.0274 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5316 | 26.3671 | 463.1
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 1.9 | 463.0144 | 464.9 | 3.3 | 3.3 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0441 | 0.0453 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3537 | 26.4 | 463
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 2 | 463.2 | 465.2 | 3.5 | 3.4557 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0441 | 0.0452 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.1437 | 25.4025 | 26.3744 | 463.2
2008-05-08 | 5:43 | 0 | -0.001 | -10 | 2 | 463.2 | 465.2 | 3.4 | 3.4 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044 | 0.0452 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.3047 | 463.2
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2 | 463.3 | 465.3 | 3.4 | 3.4435 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044 | 0.0452 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.3833 | 463.3
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.1 | 463.4 | 465.5 | 3.4 | 3.4281 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.044 | 0.0451 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 463.4
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.2 | 463.1291 | 465.3 | 3.5 | 3.4586 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.043915 | 0.045015 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7424 | 26.3576 | 463.1
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.2 | 462.8144 | 465 | 3.4 | 3.443 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0439 | 0.045 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8 | 26.3 | 462.8
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.3 | 462.6144 | 464.9 | 3.5 | 3.5438 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0439 | 0.045 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7994 | 26.3006 | 462.6
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.3 | 462.6 | 464.9 | 3.6 | 3.6 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0439 | 0.0449 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7202 | 26.3798 | 462.6
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.3 | 462.6 | 464.9 | 3.6 | 3.5892 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0438 | 0.0449 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5823 | 26.4 | 462.6
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.4 | 462.6 | 465 | 3.7 | 3.7 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0438 | 0.0449 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.538 | 26.4 | 462.6
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.5 | 462.7712 | 465.3 | 3.6 | 3.6 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0438 | 0.044814 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6173 | 26.3827 | 462.8
2008-05-08 | 5:44 | 0 | -0.001 | -10 | 2.5 | 462.6856 | 465.2 | 3.7 | 3.7 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0438 | 0.0448 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6964 | 26.3036 | 462.7
2008-05-08 | 5:44 | 0 | -0.0008 | -7.9968 | 2.5 | 462.7 | 465.2 | 3.6 | 3.6403 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.043714 | 0.0448 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5491 | 26.3754 | 462.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 462.5 | 465.1 | 3.7 | 3.7 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0437 | 0.0448 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7187 | 26.4 | 462.5
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 462.4288 | 465 | 3.6 | 3.6 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0437 | 0.0447 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8321 | 26.4 | 462.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.7 | 462.3144 | 465 | 3.7 | 3.7185 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.043614 | 0.0447 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 462.3
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.7 | 462.4 | 465.1 | 3.7 | 3.7 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0436 | 0.0446 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 462.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 462.4 | 465.2 | 3.7 | 3.6683 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 26.4 | 5 | 195 | 59.3 | 0.0436 | 0.0446 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 462.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 462.4 | 465.3 | 3.7 | 3.7 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 462.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 462.4 | 465.3 | 3.7 | 3.7 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6405 | 26.4 | 462.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 462.2 | 465.1 | 3.8 | 3.8 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5268 | 26.4 | 462.2
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 462.0288 | 464.9 | 3.8 | 3.8 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.7644 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.9 | 464.8 | 4 | 3.9584 | 18.9 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.9 | 464.8 | 3.9 | 3.9 | 18.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7539 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.9 | 464.8 | 3.9 | 3.9438 | 18.9 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.9 | 464.8 | 3.9 | 3.9 | 18.9 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.0445 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6953 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.8 | 464.7 | 4 | 4.0144 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0435 | 0.044414 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6161 | 26.4 | 461.8
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.9 | 461.9712 | 464.9 | 4 | 3.9581 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0435 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.726 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.9144 | 464.7 | 3.9 | 3.9414 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.043414 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6312 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.9 | 464.7 | 3.9 | 3.9442 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4641 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.9 | 464.7 | 3.9 | 3.9 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7014 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.9 | 464.7 | 3.9 | 3.8582 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.9391 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.8 | 464.6 | 3.7 | 3.7498 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7647 | 26.4 | 461.8
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.8711 | 464.7 | 3.7 | 3.7 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.562 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 462 | 464.8 | 3.6 | 3.561 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5172 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 462 | 464.8 | 3.5 | 3.5 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5964 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 462 | 464.8 | 3.4 | 3.3794 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7505 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 462 | 464.8 | 3.3 | 3.3395 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.691 | 26.4 | 462
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.9 | 464.7 | 3.3 | 3.2587 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0434 | 0.0444 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6676 | 26.4 | 461.9
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.8 | 461.8 | 464.6 | 3.1 | 3.1419 | 18.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.043315 | 0.0443 | 0.0273 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7742 | 26.4 | 461.8
2008-05-08 | 5:44 | 0 | 0 | 0 | 1.8 | 419.3104 | 421.1 | 3.1 | 3.1377 | 18.54 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.042016 | 0.043102 | 0.024788 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.6158 | 26.4 | 419.3
2008-05-08 | 5:44 | 0 | 0 | 0 | 1.9 | 391.6647 | 393.5 | 2.9 | 2.9401 | 17.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.038944 | 0.039873 | 0.023143 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6713 | 26.4 | 391.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.2 | 398.0218 | 400.2 | 3.1 | 3.1 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.038715 | 0.0397 | 0.023527 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7504 | 26.4 | 398
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.3 | 405.2 | 407.5 | 3 | 2.953 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.0239 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7408 | 26.4 | 405.2
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.4 | 407.5408 | 409.9 | 3 | 3 | 16.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.024086 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5826 | 26.4 | 407.5
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.4 | 408.7 | 411.1 | 2.9 | 2.9 | 16.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.0241 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4242 | 26.4 | 408.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 408.6144 | 411.1 | 3 | 2.9602 | 16.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 27 | 5 | 195 | 59.2144 | 0.0387 | 0.0396 | 0.0241 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.734 | 26.4 | 408.6
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 408.5 | 411 | 3 | 2.9601 | 16.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.024186 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.854 | 26.4 | 408.5
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 408.5 | 411 | 2.9 | 2.9447 | 16.9 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8253 | 26.4 | 408.5
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 409.2 | 411.7 | 2.9 | 2.9 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.891 | 26.4 | 409.2
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 409.3856 | 411.9 | 2.8 | 2.8414 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0396 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7329 | 26.4 | 409.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 409.6 | 412.1 | 2.9 | 2.9 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3856 | 0.0387 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7628 | 26.4 | 409.6
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 409.6 | 412.1 | 2.8 | 2.8 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0387 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.8 | 26.4 | 409.6
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.5 | 409.6 | 412.1 | 2.8 | 2.8 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7366 | 26.4 | 409.6
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.7 | 412.3 | 2.7 | 2.6522 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5006 | 26.4 | 409.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.7 | 412.3 | 2.7 | 2.7 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6589 | 26.4 | 409.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.5289 | 412.1 | 2.6 | 2.6 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4 | 409.5
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.4144 | 412 | 2.7 | 2.6558 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6245 | 26.4 | 409.4
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.6709 | 412.3 | 2.6 | 2.5604 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2854 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5337 | 26.4 | 409.7
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.8 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6921 | 26.4 | 409.8
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.8 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 409.8
2008-05-08 | 5:44 | 0 | 0 | 0 | 2.6 | 409.8 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5916 | 26.4 | 409.8
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.6 | 409.8 | 412.4 | 2.6 | 2.6 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.567 | 26.4 | 409.8
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.6 | 409.8 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4126 | 409.8
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.5 | 409.7 | 412.2 | 2.6 | 2.561 | 16.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 6 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4912 | 409.7
2008-05-08 | 5:45 | 0 | -0.0004 | -3.5838 | 2.5 | 409.7 | 412.2 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6292 | 26.4292 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.7 | 412.2 | 2.5 | 2.5423 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6501 | 26.4501 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.6 | 412.1 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.6 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6416 | 26.5 | 409.6
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7 | 412.3 | 2.6 | 2.5598 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5.4 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.4916 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7 | 412.3 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.4124 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.6144 | 412.2 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5667 | 26.4 | 409.6
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.6854 | 412.3 | 2.6 | 2.5606 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6458 | 26.4458 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7 | 412.3 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.5 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.6 | 412.2 | 2.6 | 2.6 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.4 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4959 | 409.6
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.6 | 412.2 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.4167 | 409.6
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.6857 | 412.2 | 2.6 | 2.6 | 16.8 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6376 | 26.4 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.7 | 412.2 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.6 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5585 | 26.4 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.7 | 412.2 | 2.5 | 2.5418 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0386 | 0.0395 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5206 | 26.4206 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.8712 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.039414 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.5995 | 26.4995 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 410 | 412.5 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8369 | 26.5 | 410
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.9142 | 412.4 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.8418 | 26.4418 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 410.0712 | 412.6 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0386 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7627 | 26.4 | 410.1
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 410.1 | 412.6 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0386 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6835 | 26.4 | 410.1
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 410.0144 | 412.5 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6044 | 26.4 | 410
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.8144 | 412.3 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.6 | 26.4 | 409.8
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.9 | 412.4 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7614 | 26.4 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 409.8144 | 412.3 | 2.5 | 2.47 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.834 | 26.4 | 409.8
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7146 | 412.3 | 2.3 | 2.3493 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6879 | 26.4121 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7 | 412.3 | 2.3 | 2.3392 | 16.8 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6091 | 26.4909 | 409.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.7856 | 412.4 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6 | 26.4297 | 409.8
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.6 | 409.9 | 412.5 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2146 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5008 | 26.4 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.7 | 409.8 | 412.5 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4862 | 26.4 | 409.8
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.7 | 409.9 | 412.6 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0394 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.7 | 409.9856 | 412.7 | 2.5 | 2.4567 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.039314 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4 | 410
2008-05-08 | 5:45 | 0 | -0.0005 | -4.974 | 2.7 | 410.2 | 412.9 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0393 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5677 | 26.4 | 410.2
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.8 | 409.9144 | 412.7 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0393 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5906 | 26.4 | 409.9
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.8 | 409.9856 | 412.8 | 2.4 | 2.4 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0393 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6756 | 26.4244 | 410
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.8 | 410 | 412.8 | 2.5 | 2.5413 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0393 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5964 | 26.4964 | 410
2008-05-08 | 5:45 | 0 | -0.0008 | -8.4034 | 2.8 | 409.9432 | 412.7 | 2.5 | 2.5 | 16.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0385 | 0.0393 | 0.0242 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5175 | 26.4175 | 409.9
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 1.6 | 386.5088 | 388.1 | 2.5 | 2.5 | 16.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.038329 | 0.0393 | 0.022873 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6237 | 26.4 | 386.5
2008-05-08 | 5:45 | 0 | -0.0005 | -4.9567 | 1.2 | 365.4944 | 366.7 | 2.5 | 2.4591 | 15.75 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.037558 | 0.038558 | 0.021658 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6589 | 26.4411 | 365.5
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.3 | 361.9856 | 364.2 | 2.5 | 2.4603 | 15.3 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.036243 | 0.036958 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5798 | 26.5 | 362
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.4 | 361.6288 | 364 | 2.5 | 2.5 | 15.1 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.036014 | 0.036714 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5007 | 26.5 | 361.6
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.4 | 361.6 | 364 | 2.4 | 2.4 | 15 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0359 | 0.0367 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4215 | 26.5 | 361.6
2008-05-08 | 5:45 | 0 | -0.0008 | -8.4034 | 2.4 | 361.6856 | 364.1 | 2.5 | 2.4586 | 14.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0359 | 0.0366 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5152 | 26.5 | 361.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.4 | 361.7 | 364.1 | 2.4 | 2.4 | 14.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.035714 | 0.036414 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6 | 26.4633 | 361.7
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.5 | 361.6 | 364.1 | 2.4 | 2.4 | 14.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0356 | 0.036315 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5842 | 26.4 | 361.6
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2.2 | 361.5 | 363.7 | 2.3 | 2.3 | 14.8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.035514 | 0.036214 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5049 | 26.4 | 361.5
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2 | 361.3144 | 363.3 | 2.2 | 2.2423 | 14.7 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.035314 | 0.036114 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6484 | 26.4742 | 361.3
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2 | 361.2144 | 363.2 | 2.2 | 2.2495 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0353 | 0.0361 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4467 | 361.2
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2 | 361.1144 | 363.1 | 2.2 | 2.2408 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.035214 | 0.0361 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7324 | 26.4 | 361.1
2008-05-08 | 5:45 | 0 | -0.001 | -10 | 2 | 361.0144 | 363 | 2.2 | 2.2 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0352 | 0.036 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7885 | 26.4115 | 361
2008-05-08 | 5:45 | 0 | -0.0005 | -4.9185 | 2 | 360.9 | 362.9 | 2.1 | 2.1 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 27 | 5 | 195 | 59.2144 | 0.035114 | 0.035914 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7094 | 26.4906 | 360.9
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.1 | 360.9 | 363 | 2.2 | 2.1625 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2861 | 0.0351 | 0.0359 | 0.0214 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4906 | 26.5 | 360.9
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.1 | 360.9 | 363 | 2 | 2.0488 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.035 | 0.0358 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.449 | 26.5 | 360.9
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.1 | 360.9 | 363 | 2.2 | 2.2 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.035 | 0.0358 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.472 | 26.5 | 360.9
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.2 | 360.8144 | 363 | 2.1 | 2.1 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0349 | 0.0357 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4205 | 26.5 | 360.8
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.3 | 360.6144 | 362.9 | 2.2 | 2.2 | 14.7 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.6 | 195 | 59.3 | 0.0349 | 0.0356 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6592 | 26.5 | 360.6
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.3 | 360.4144 | 362.7 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 6 | 195 | 59.2856 | 0.034814 | 0.0356 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6344 | 26.4344 | 360.4
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.4 | 360.5 | 362.9 | 2.1 | 2.125 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5.4 | 195 | 59.3 | 0.0348 | 0.0356 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6 | 26.4 | 360.5
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.4 | 360.3856 | 362.8 | 2.1 | 2.1431 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0348 | 0.0355 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.5285 | 26.4238 | 360.4
2008-05-08 | 5:45 | 0 | 0 | 0 | 2.4 | 360.3 | 362.7 | 2.1 | 2.0857 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0348 | 0.0355 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.312 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.4 | 360.4 | 362.8 | 2.2 | 2.1572 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0347 | 0.0355 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6285 | 26.5 | 360.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.3144 | 362.8 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0347 | 0.0354 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.3 | 362.8 | 2.2 | 2.1604 | 14.6 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0347 | 0.0354 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.579 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.3 | 362.8 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0347 | 0.0354 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4776 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.2856 | 362.8 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.034614 | 0.0354 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7931 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.3 | 362.8 | 2.1 | 2.0573 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0346 | 0.035314 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6445 | 26.5 | 360.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 360.1288 | 362.6 | 2.2 | 2.1629 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0346 | 0.0353 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.543 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 360.1 | 362.7 | 2.1 | 2.06 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0346 | 0.0353 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 360 | 362.6 | 2 | 2.0427 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0346 | 0.0353 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5152 | 26.5 | 360
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 360 | 362.6 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0345 | 0.0352 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5944 | 26.5 | 360
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 360.1 | 362.7 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0345 | 0.0352 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.453 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360 | 362.7 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0345 | 0.0352 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4526 | 26.5 | 360
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.9144 | 362.6 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.034414 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5635 | 26.5 | 359.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360 | 362.7 | 2.1 | 2.0792 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6563 | 26.5 | 360
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360.1 | 362.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3396 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360.2 | 362.9 | 2.1 | 2.1 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3692 | 26.5 | 360.2
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360.1146 | 362.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2854 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.545 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 360.1 | 362.8 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6176 | 26.5 | 360.1
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.9289 | 362.6 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.0351 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4197 | 26.5 | 359.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8146 | 362.5 | 2 | 2.0419 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6572 | 26.5 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8 | 362.5 | 2.1 | 2.0519 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0344 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.5 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8 | 362.5 | 2 | 2.0425 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.7 | 26.4562 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8 | 362.5 | 2.1 | 2.0588 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6544 | 26.4228 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8 | 362.5 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4977 | 26.5 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8858 | 362.6 | 2.1 | 2.1 | 14.6 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4185 | 26.5 | 359.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8856 | 362.6 | 2 | 2 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5212 | 26.5 | 359.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.8144 | 362.5 | 2 | 2.037 | 14.6 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6397 | 26.5 | 359.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.7 | 362.4 | 2 | 2 | 14.6 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0343 | 0.035 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.5 | 359.7
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.6 | 362.3 | 2 | 2 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.3 | 0.0343 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.7 | 26.5 | 359.6
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.5 | 362.2 | 2 | 1.9779 | 14.6 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0343 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5458 | 26.4229 | 359.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.4144 | 362.1 | 1.9 | 1.9408 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0343 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.4563 | 359.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.4 | 362.1 | 2.1 | 2.0508 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0343 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5705 | 26.5 | 359.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.3146 | 362 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.034215 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003.8543 | 25.6716 | 26.5 | 359.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.4 | 362.1 | 2.1 | 2.0556 | 14.6 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5127 | 26.5 | 359.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.5 | 362.2 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4272 | 26.5 | 359.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.7 | 359.5 | 362.2 | 2 | 2 | 14.6 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2854 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4516 | 26.4484 | 359.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 359.5856 | 362.2 | 2 | 2 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4695 | 26.4305 | 359.6
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 359.6856 | 362.3 | 2 | 2.0416 | 14.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.4102 | 26.5 | 359.7
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.6 | 359.6144 | 362.2 | 2 | 2 | 14.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0342 | 0.0349 | 0.0213 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.4893 | 26.5 | 359.6
2008-05-08 | 5:46 | 0 | 0 | 0 | 2.5 | 353.0648 | 355.6 | 2 | 2 | 14.6 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0342 | 0.0349 | 0.020858 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4316 | 26.5 | 353.1
2008-05-08 | 5:46 | 0 | 0 | 0 | -0.4 | 328.3096 | 328.3 | 2 | 1.956 | 14.06 | 12.9 | 954 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.033543 | 0.0348 | 0.019472 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4 | 26.5 | 328.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 0.4 | 315.5736 | 315.9 | 1.9 | 1.9 | 13.4 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.032243 | 0.033258 | 0.018714 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3733 | 26.5 | 315.6
2008-05-08 | 5:46 | 0 | 0 | 0 | 0.7 | 313.5024 | 314.2 | 1.9 | 1.9474 | 13.1 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.031943 | 0.032858 | 0.018614 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3234 | 26.5 | 313.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 0.8 | 312.3 | 313.1 | 1.9 | 1.8584 | 12.9 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.031714 | 0.032614 | 0.0185 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.6401 | 26.5 | 312.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 1 | 311.4 | 312.4 | 2 | 1.9586 | 12.8 | 13 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0316 | 0.032414 | 0.018414 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5075 | 26.5 | 311.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 1 | 311.4 | 312.4 | 1.9 | 1.9 | 12.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0315 | 0.0324 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.4 | 26.5 | 311.4
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.1 | 311.3144 | 312.4 | 1.9 | 1.94 | 12.7 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0314 | 0.032214 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3335 | 26.5 | 311.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.2 | 311.2856 | 312.5 | 1.9 | 1.9129 | 12.7 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0313 | 0.0321 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.1063 | 26.5 | 311.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.2 | 311.3 | 312.5 | 1.9 | 1.9 | 12.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.031215 | 0.032015 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4228 | 26.5 | 311.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.2 | 311.2144 | 312.4 | 1.9 | 1.9 | 12.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.031114 | 0.031914 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4403 | 26.5 | 311.2
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.2 | 311.0144 | 312.2 | 1.9 | 1.8607 | 12.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.031014 | 0.031814 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4776 | 26.5 | 311
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.3 | 310.9146 | 312.2 | 1.9 | 1.9 | 12.6 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.030915 | 0.031715 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.8543 | 25.5639 | 26.4819 | 310.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.5 | 310.9 | 312.4 | 1.9 | 1.9 | 12.6 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.030814 | 0.031614 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.4054 | 26.4027 | 310.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.5 | 310.9 | 312.4 | 1.8 | 1.8496 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0308 | 0.0315 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4 | 26.4764 | 310.9
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.6 | 310.8144 | 312.4 | 1.8 | 1.7597 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.030714 | 0.031414 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4555 | 26.4445 | 310.8
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.7 | 310.7 | 312.4 | 1.9 | 1.9 | 12.5 | 13 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0307 | 0.0314 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.4347 | 310.7
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.7 | 310.5144 | 312.2 | 1.8 | 1.8 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0306 | 0.031314 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.5 | 310.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.7 | 310.5 | 312.2 | 1.9 | 1.9 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0306 | 0.0313 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.5 | 310.5
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.8 | 310.3144 | 312.1 | 1.8 | 1.8 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0305 | 0.0312 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4279 | 26.5 | 310.3
2008-05-08 | 5:46 | 0 | 0 | 0 | 1.8 | 310.2144 | 312 | 1.8 | 1.7869 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0305 | 0.0312 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4511 | 26.5 | 310.2
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 310.1144 | 311.9 | 1.8 | 1.7592 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0305 | 0.031114 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5 | 26.5 | 310.1
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 310.2144 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0304 | 0.0311 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.4906 | 26.5 | 310.2
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 310.2 | 312 | 1.8 | 1.7603 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0304 | 0.0311 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4113 | 26.5 | 310.2
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.9 | 310.0856 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.031 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4 | 26.5 | 310.1
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.9 | 309.8291 | 311.7 | 1.8 | 1.8 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.031 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.493 | 26.5 | 309.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.8856 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.0309 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5481 | 26.5 | 309.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.8144 | 311.8 | 1.8 | 1.8 | 12.5 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0303 | 0.0309 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4051 | 26.5 | 309.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.9 | 309.8 | 311.7 | 1.8 | 1.8 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.030215 | 0.0309 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4843 | 26.5 | 309.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.9 | 309.8856 | 311.8 | 1.8 | 1.8 | 12.5 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0302 | 0.0309 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5634 | 26.5 | 309.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.9 | 309.9856 | 311.9 | 1.7 | 1.7 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0302 | 0.0309 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.1443 | 25.5574 | 26.5 | 310
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.9144 | 311.9 | 1.7 | 1.7442 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0302 | 0.030814 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.4783 | 26.5 | 309.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.7141 | 311.7 | 1.8 | 1.8 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.030114 | 0.0308 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3992 | 26.5 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.7 | 311.7 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0301 | 0.0308 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.32 | 26.5 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.7144 | 311.7 | 1.9 | 1.8566 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0301 | 0.0308 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5356 | 26.4411 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.6144 | 311.6 | 1.8 | 1.7586 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0301 | 0.030714 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.5859 | 26.4 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2 | 309.7 | 311.7 | 1.8 | 1.8 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0301 | 0.0307 | 0.0184 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3826 | 26.4174 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.1 | 309.6 | 311.7 | 1.8 | 1.8 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0307 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3035 | 26.4965 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.1 | 309.5144 | 311.6 | 1.8 | 1.8426 | 12.5 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0307 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.2244 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.1 | 309.5 | 311.6 | 1.7 | 1.7407 | 12.5 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0307 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.419 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.1 | 309.5 | 311.6 | 1.8 | 1.84 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.03 | 0.0306 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4982 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.1 | 309.5 | 311.6 | 1.8 | 1.7549 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.03 | 0.0306 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.2869 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.2 | 309.4 | 311.6 | 1.8 | 1.7883 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.3 | 0.029914 | 0.0306 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.2078 | 26.5 | 309.4
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.2 | 309.4856 | 311.7 | 1.8 | 1.8 | 12.5 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.3 | 0.0299 | 0.030514 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.2 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.2 | 309.5 | 311.7 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.3 | 0.0299 | 0.0305 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4021 | 26.4495 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.3 | 309.5856 | 311.9 | 1.8 | 1.76 | 12.5 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.3 | 0.0299 | 0.0305 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.5407 | 26.4297 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.3 | 309.7 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0299 | 0.0305 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4088 | 26.5 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.3 | 309.7854 | 312.1 | 1.8 | 1.7516 | 12.5 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.029815 | 0.0304 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4879 | 26.5 | 309.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.4 | 309.7144 | 312.1 | 1.7 | 1.6557 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0298 | 0.0304 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4331 | 26.5 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.4 | 309.6856 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0298 | 0.0304 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4461 | 26.4539 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.4 | 309.7 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.030314 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4494 | 26.4253 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.4 | 309.6144 | 312 | 1.7 | 1.6855 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0298 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.5144 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0298 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.6 | 312.1 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.7 | 312.2 | 1.8 | 1.7606 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0297 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 309.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.6 | 312.1 | 1.6 | 1.6 | 12.5 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.321 | 26.5 | 309.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.5 | 312 | 1.8 | 1.7605 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.0303 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3998 | 26.4998 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.5 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.030214 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3207 | 26.4207 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 2.5 | 309.5 | 312 | 1.7 | 1.7 | 12.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0297 | 0.0302 | 0.0183 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4169 | 26.4584 | 309.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.4 | 264.7934 | 266.2 | 1.6 | 1.5572 | 12.18 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.028331 | 0.028916 | 0.015703 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5 | 26.5 | 264.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.1 | 236.616 | 237.7 | 1.5 | 1.5426 | 10.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.025575 | 0.026189 | 0.014044 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4666 | 26.5 | 236.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.4 | 244.7352 | 246.1 | 1.6 | 1.6 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0254 | 0.025914 | 0.014542 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3082 | 26.5 | 244.7
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.6 | 255.2218 | 256.8 | 1.5 | 1.5435 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0253 | 0.025814 | 0.015157 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001 | 25.3 | 26.4253 | 255.2
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.8146 | 257.5 | 1.6 | 1.5551 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0253 | 0.0258 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001.8543 | 25.354 | 26.454 | 255.8
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.9 | 257.6 | 1.5 | 1.5 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0253 | 0.0258 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4333 | 26.5 | 255.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.9 | 257.6 | 1.6 | 1.5566 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.025214 | 0.0258 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5 | 26.5 | 255.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.9 | 257.6 | 1.5 | 1.5 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0258 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.8557 | 25.5 | 26.5 | 255.9
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.9856 | 257.7 | 1.6 | 1.5565 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4294 | 26.5 | 256
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.7 | 255.9856 | 257.7 | 1.5 | 1.458 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4 | 26.5 | 256
2008-05-08 | 5:47 | 0 | 0.0008 | 8.0051 | 1.7 | 256 | 257.7 | 1.4 | 1.4493 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4 | 26.5 | 256
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.7 | 256.1856 | 257.9 | 1.5 | 1.5 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.4 | 26.5 | 256.2
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.7 | 256.2856 | 258 | 1.5 | 1.5427 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4 | 26.5 | 256.3
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.7 | 256.4 | 258.1 | 1.4 | 1.4398 | 10.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3336 | 26.5 | 256.4
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.7 | 256.4856 | 258.2 | 1.4 | 1.4226 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0252 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2545 | 26.5 | 256.5
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.7 | 256.5 | 258.2 | 1.4 | 1.4498 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.025114 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2247 | 26.4753 | 256.5
2008-05-08 | 5:47 | 0 | 0.001 | 10 | 1.8 | 256.6 | 258.4 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0251 | 0.0257 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4 | 256.6
2008-05-08 | 5:47 | 0 | 0.0004 | 3.7007 | 1.8 | 256.6 | 258.4 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.3 | 0.0251 | 0.025615 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4 | 256.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 256.5144 | 258.3 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.3 | 26.462 | 256.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 256.5 | 258.3 | 1.6 | 1.5621 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3 | 26.5 | 256.5
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 256.6 | 258.4 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3 | 26.5 | 256.6
2008-05-08 | 5:47 | 0 | 0 | 0 | 1.8 | 256.6 | 258.4 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3 | 26.5 | 256.6
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 256.6 | 258.4 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3786 | 26.4214 | 256.6
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 256.7712 | 258.6 | 1.5 | 1.4871 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.2843 | 26.4578 | 256.8
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 256.8856 | 258.7 | 1.5 | 1.5422 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.4 | 26.6 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.237 | 26.463 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 256.9 | 258.7 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.4 | 5 | 195 | 59.3 | 0.0251 | 0.0256 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3161 | 26.4161 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.9 | 258.8 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.025014 | 0.025514 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3953 | 26.4953 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.9 | 258.8 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.2144 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3256 | 26.5 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.9 | 258.8 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.2858 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.9 | 258.8 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3 | 26.5 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.8 | 258.7 | 1.5 | 1.4795 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.6 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3117 | 26.5 | 256.8
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 256.9 | 258.8 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3905 | 26.5 | 256.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.0856 | 259 | 1.5 | 1.4982 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.26 | 26.5 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1854 | 259.1 | 1.5 | 1.462 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2493 | 26.5 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1143 | 259 | 1.4 | 1.3811 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.025 | 0.0255 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3284 | 26.4716 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1856 | 259.1 | 1.4 | 1.442 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.4 | 5 | 195 | 59.2144 | 0.025 | 0.025414 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3776 | 26.4075 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.2854 | 259.2 | 1.4 | 1.4411 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.024915 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1001.1457 | 25.1397 | 26.4868 | 257.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.3 | 259.2 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.015286 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2977 | 26.4341 | 257.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 257.2 | 259 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.6 | 5 | 195 | 59.2144 | 0.0249 | 0.0254 | 0.0153 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.355 | 26.445 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 257.2856 | 259.1 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0153 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.5 | 257.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 257.3 | 259.1 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0153 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2967 | 26.5 | 257.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 257.2144 | 259 | 1.4 | 1.4418 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.2175 | 26.5 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 257.1143 | 258.9 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.8566 | 25.2617 | 26.5 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1144 | 259 | 1.4 | 1.4479 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.3 | 26.4592 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.2 | 259.1 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.4 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4199 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1 | 259 | 1.4 | 1.441 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0249 | 0.0254 | 0.0152 | 0.0002 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3 | 26.4987 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.1 | 259 | 1.4 | 1.4402 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26.6 | 5 | 195 | 59.3 | 0.0249 | 0.025314 | 0.0152 | 0.000114 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3782 | 26.5 | 257.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.2 | 259.1 | 1.4 | 1.4 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4574 | 26.5 | 257.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.3856 | 259.3 | 1.5 | 1.4842 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0249 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4274 | 26.5 | 257.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.4 | 259.3 | 1.4 | 1.4212 | 10.4 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.2146 | 0.0249 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4849 | 257.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.3144 | 259.2 | 1.5 | 1.5 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.2856 | 0.024814 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4051 | 257.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.3856 | 259.3 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 27 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3741 | 26.4 | 257.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.9 | 257.4 | 259.3 | 1.4 | 1.4375 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 188 | 26.4 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4 | 26.4 | 257.4
2008-05-08 | 5:48 | 0 | -0.0006 | -6.2439 | 2 | 257.4 | 259.4 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3353 | 26.4 | 257.4
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.2144 | 259.2 | 1.4 | 1.4436 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.8557 | 25.2115 | 26.4 | 257.2
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.3 | 259.3 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.2904 | 26.4 | 257.3
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.3857 | 259.4 | 1.3 | 1.3419 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.1431 | 25.3697 | 26.4 | 257.4
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.4856 | 259.5 | 1.4 | 1.4405 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3511 | 26.4489 | 257.5
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.5 | 259.5 | 1.4 | 1.4 | 10.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.0253 | 0.0152 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3281 | 26.4719 | 257.5
2008-05-08 | 5:48 | 0 | -0.001 | -10 | 2 | 257.4856 | 259.5 | 1.5 | 1.4575 | 10.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.6 | 26 | 5 | 195 | 59.3 | 0.0248 | 0.025214 | 0.0153 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4 | 26.4072 | 257.5
2008-05-08 | 5:48 | 0 | -0.0006 | -5.5259 | 1.9 | 248.3968 | 250.3 | 1.3 | 1.3195 | 10.36 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188.4 | 26 | 5 | 195 | 59.3 | 0.024714 | 0.0252 | 0.014686 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4 | 26.4865 | 248.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 0.4 | 220.724 | 221.1 | 1.3 | 1.3 | 9.72 | 12.9 | 955 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.024143 | 0.024758 | 0.013072 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2031 | 26.5 | 220.7
2008-05-08 | 5:48 | 0 | 0 | 0 | 1 | 210.8904 | 211.9 | 1.2 | 1.2 | 9.16 | 12.9 | 956 | 962 | 947 | 21.4 | 0.37 | 0.01 | 985 | -32.6 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.023186 | 0.023701 | 0.012529 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2343 | 26.4552 | 210.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.3 | 210.2144 | 211.6 | 1.2 | 1.2431 | 8.9 | 12.9 | 956 | 962 | 947 | 21 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 188 | 26 | 5 | 195 | 59.3 | 0.022758 | 0.023258 | 0.0125 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3761 | 26.4 | 210.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 210.1144 | 211.6 | 1.2 | 1.2281 | 8.7 | 12.9 | 956 | 962 | 947 | 21 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 188.6 | 26 | 5 | 195 | 59.3 | 0.0227 | 0.0231 | 0.0125 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3 | 26.4 | 210.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.8 | 209.4 | 211.2 | 1.1 | 1.1487 | 8.6 | 12.9 | 957 | 962 | 947 | 21.6 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.022343 | 0.022658 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.144 | 25.3 | 26.4 | 209.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.7 | 209.2144 | 210.9 | 1.2 | 1.2 | 8.5 | 12.9 | 956 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.022114 | 0.022429 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 25.3 | 26.4 | 209.2
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 209.1 | 210.6 | 1.1 | 1.1 | 8.5 | 12.9 | 956 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.022 | 0.0224 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3405 | 26.4 | 209.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 209 | 210.5 | 1.1 | 1.14 | 8.46 | 12.9 | 956 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0219 | 0.0223 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3804 | 26.4196 | 209
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 209 | 210.5 | 1.1 | 1.1 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0218 | 0.0222 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3015 | 26.4985 | 209
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.9 | 210.4 | 1.2 | 1.1543 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0218 | 0.0222 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2221 | 26.5 | 208.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.6 | 208.7291 | 210.3 | 1.1 | 1.1 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.021715 | 0.022115 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.1428 | 26.5 | 208.7
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.7 | 210.2 | 1.2 | 1.1567 | 8.4 | 12.9 | 957 | 962 | 947 | 21.4 | 0.37 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.021614 | 0.022014 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.1721 | 26.4639 | 208.7
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.7 | 210.2 | 1.1 | 1.1 | 8.4 | 12.9 | 956 | 962 | 947 | 21 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0216 | 0.022 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3148 | 26.4 | 208.7
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.4 | 208.6144 | 210 | 1.1 | 1.0843 | 8.4 | 12.9 | 956 | 962 | 947 | 21 | 0.35 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.021514 | 0.022 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3946 | 26.4 | 208.6
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.4856 | 210 | 1.1 | 1.1 | 8.4 | 12.9 | 956 | 962 | 947 | 21.6 | 0.38 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0215 | 0.0219 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3263 | 26.4 | 208.5
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.4 | 209.9 | 1.1 | 1.1498 | 8.4 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0215 | 0.0219 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2471 | 26.4 | 208.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.4 | 209.9 | 1.1 | 1.1 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0214 | 0.0218 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2 | 26.4 | 208.4
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.5 | 208.3 | 209.8 | 1 | 1.0421 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0214 | 0.0218 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2 | 26.4112 | 208.3
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.6 | 208.1 | 209.7 | 1.1 | 1.143 | 8.4 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.021314 | 0.021714 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2 | 26.4904 | 208.1
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.6 | 207.9139 | 209.5 | 1.1 | 1.1 | 8.36 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0213 | 0.0217 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4088 | 26.4304 | 207.9
2008-05-08 | 5:48 | 0 | 0 | 0 | 1.7 | 207.9 | 209.6 | 1.2 | 1.2 | 7.38 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0213 | 0.021614 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4025 | 26.4 | 207.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.7 | 207.8 | 209.5 | 1.1 | 1.1 | 7.5 | 12.9 | 956 | 963 | 947 | 22 | 0.38 | 0.01 | 985 | -32.6 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0212 | 0.0216 | 0.0124 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3 | 26.4 | 207.8
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.7 | 207.8 | 209.5 | 1.1 | 1.1422 | 7.6 | 12.9 | 956 | 962 | 947 | 22 | 0.36 | 0.01 | 985 | -33 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.3 | 0.0212 | 0.0216 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3 | 26.4 | 207.8
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.8 | 207.8 | 209.6 | 1.1 | 1.1 | 7.74 | 12.9 | 957 | 962 | 947 | 22 | 0.39 | 0.01 | 985 | -32.4 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.3 | 0.0212 | 0.0215 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3 | 26.4 | 207.8
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.8 | 207.9 | 209.7 | 1.1 | 1.1435 | 7.8 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 26 | 5 | 195 | 59.3 | 0.0211 | 0.0215 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3653 | 26.4 | 207.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.9 | 207.8144 | 209.7 | 1.1 | 1.1 | 7.9 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0211 | 0.0214 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3557 | 26.4 | 207.8
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.9 | 207.7146 | 209.6 | 1.1 | 1.1 | 7.98 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.021015 | 0.0214 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.854 | 25.3233 | 26.4 | 207.7
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.7 | 209.7 | 1.2 | 1.159 | 8 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.021 | 0.0213 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4029 | 26.4 | 207.7
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.7 | 209.7 | 1.1 | 1.1 | 8.08 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.021 | 0.0213 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.482 | 26.4 | 207.7
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.6 | 209.6 | 1.1 | 1.1 | 8.1 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.3 | 0.020914 | 0.021214 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3776 | 26.4 | 207.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.6 | 209.6 | 1.1 | 1.1 | 8.1 | 12.9 | 957 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0209 | 0.0212 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4211 | 26.4 | 207.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.5144 | 209.6 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0209 | 0.0212 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.561 | 26.4 | 207.5
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.6 | 209.7 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0209 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4033 | 26.4 | 207.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.6 | 209.7 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0208 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.4 | 26.4 | 207.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.5144 | 209.6 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0208 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 25.343 | 26.4 | 207.5
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.5 | 209.5 | 1.1 | 1.0873 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0208 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3 | 26.4 | 207.5
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.4 | 209.4 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0208 | 0.0211 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3154 | 26.4 | 207.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 2 | 207.4 | 209.4 | 1.1 | 1.1 | 8.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0208 | 0.021014 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3945 | 26.4 | 207.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.1 | 8.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2526 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.0571 | 8.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4116 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5039 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.0577 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3 | 26.4112 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3854 | 209.5 | 1.2 | 1.1523 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3 | 26.4903 | 207.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.021 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3 | 26.4305 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.1 | 1.1429 | 8.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.020914 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4461 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.4 | 209.5 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0207 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.5167 | 26.4 | 207.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.4 | 209.6 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0206 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.3134 | 26.4 | 207.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.1 | 207.3 | 209.4 | 1.2 | 1.1568 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0206 | 0.0209 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4723 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.3 | 209.5 | 1.1 | 1.1 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0206 | 0.020814 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5654 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.3 | 209.5 | 1.2 | 1.2 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0206 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.144 | 25.6 | 26.4 | 207.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.2144 | 209.4 | 1.1 | 1.1 | 8.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.020514 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1038.8338 | 1001.856 | 25.6 | 26.4 | 207.2
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.2 | 209.4 | 1.2 | 1.2 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0205 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.6029 | 26.4 | 207.2
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.2 | 209.4 | 1.1 | 1.1193 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0205 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.682 | 26.4 | 207.2
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207.1144 | 209.3 | 1.2 | 1.2 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0205 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4552 | 26.4 | 207.1
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.2 | 207 | 209.2 | 1.2 | 1.1598 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0205 | 0.0208 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.856 | 25.3404 | 26.4 | 207
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.3 | 206.9 | 209.2 | 1.1 | 1.1426 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0205 | 0.020715 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4195 | 26.4 | 206.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.3 | 206.9 | 209.2 | 1.2 | 1.2 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0205 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4983 | 26.4 | 206.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 2.3 | 206.9 | 209.2 | 1.1 | 1.1406 | 8.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0205 | 0.0207 | 0.0123 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.4221 | 26.4 | 206.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 0.9 | 177.6804 | 178.6 | 1.2 | 1.1568 | 7.95 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.020329 | 0.0207 | 0.011187 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1003 | 25.3429 | 26.4 | 177.7
2008-05-08 | 5:49 | 0 | 0 | 0 | 0.5 | 159.9176 | 160.4 | 1.1 | 1.1 | 7.23 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.3 | 0.018958 | 0.019358 | 0.009529 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.144 | 25.3 | 26.4 | 159.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 0.7 | 157.0597 | 157.8 | 1.2 | 1.2 | 6.82 | 13 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.018514 | 0.0189 | 0.009314 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002.1443 | 25.3455 | 26.4 | 157.1
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.1 | 156.7 | 157.8 | 1.2 | 1.2 | 6.67 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0181 | 0.018414 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5838 | 26.4 | 156.7
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.5 | 156.6 | 158.1 | 1.4 | 1.3587 | 6.6 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.017714 | 0.017914 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4524 | 26.4 | 156.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.3 | 156.6 | 157.9 | 1.4 | 1.4 | 6.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0176 | 0.0178 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.453 | 26.4 | 156.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.4 | 156.6 | 158 | 1.4 | 1.4464 | 6.42 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0175 | 0.0177 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4678 | 26.4 | 156.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.4 | 156.6 | 158 | 1.5 | 1.4637 | 6.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.017314 | 0.017614 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3772 | 26.4 | 156.6
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.3 | 156.5144 | 157.8 | 1.5 | 1.4854 | 6.4 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.017214 | 0.0175 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.2191 | 26.4 | 156.5
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.3 | 156.4 | 157.7 | 1.5 | 1.5351 | 6.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.017114 | 0.017414 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.3393 | 26.4 | 156.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.3 | 156.4 | 157.7 | 1.5 | 1.5436 | 6.3 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0171 | 0.0174 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.144 | 25.4489 | 26.4 | 156.4
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.3 | 156.3 | 157.6 | 1.7 | 1.7 | 6.3 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.4 | 5 | 195 | 59.3 | 0.017014 | 0.0173 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001 | 25.4444 | 26.4 | 156.3
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.4 | 156.1288 | 157.5 | 1.8 | 1.7504 | 6.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 26.6 | 5 | 195 | 59.3 | 0.017 | 0.017214 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 25.3 | 26.4 | 156.1
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.4 | 156.0146 | 157.4 | 1.9 | 1.9 | 6.3 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0169 | 0.0172 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3 | 26.4 | 156
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.5 | 156 | 157.5 | 1.9 | 1.9 | 6.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016814 | 0.0171 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038.171 | 1002 | 25.3 | 26.4 | 156
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.6 | 155.9 | 157.5 | 2 | 2.017 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0168 | 0.017 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3447 | 26.4 | 155.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.7 | 155.9 | 157.6 | 2 | 2 | 6.2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0167 | 0.0169 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4479 | 26.3761 | 155.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.7 | 155.9 | 157.6 | 2 | 1.9506 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0167 | 0.0169 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.5968 | 26.3032 | 155.9
2008-05-08 | 5:49 | 0 | 0 | 0 | 1.7 | 155.9 | 157.6 | 2 | 1.9571 | 6.2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0166 | 0.0168 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.5177 | 26.3823 | 155.9
2008-05-08 | 5:49 | 0 | -0.0004 | -4.3178 | 1.8 | 155.9 | 157.7 | 1.9 | 1.9 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0166 | 0.0168 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4386 | 26.3386 | 155.9
2008-05-08 | 5:49 | 0 | -0.001 | -10 | 1.7 | 155.9 | 157.6 | 1.8 | 1.8265 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0166 | 0.0168 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.4406 | 26.3406 | 155.9
2008-05-08 | 5:49 | 0 | -0.001 | -10 | 1.7 | 155.8144 | 157.5 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016514 | 0.0167 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5 | 26.3803 | 155.8
2008-05-08 | 5:49 | 0 | -0.001 | -10 | 1.7 | 155.7 | 157.4 | 1.7 | 1.6582 | 6.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0165 | 0.0167 | 0.0093 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1039 | 1002 | 25.5 | 26.3013 | 155.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.8 | 155.6 | 157.4 | 1.5 | 1.5479 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0165 | 0.0166 | 0.0093 | 0.0001 |  | n/a |  | 30.0751 | 53 | 1039 | 1002 | 25.4219 | 26.3781 | 155.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.8 | 155.6 | 157.4 | 1.6 | 1.6 | 6.2 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016414 | 0.0166 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1002 | 25.4573 | 26.4 | 155.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.7 | 155.6856 | 157.4 | 1.6 | 1.5569 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0164 | 0.0166 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.5362 | 26.4 | 155.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.4 | 155.7 | 157.1 | 1.6 | 1.5613 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.016314 | 0.0166 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5552 | 26.4 | 155.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.4 | 155.7 | 157.1 | 1.5 | 1.5 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0163 | 0.0166 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3153 | 26.4 | 155.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.4 | 155.6144 | 157 | 1.5 | 1.5 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0163 | 0.0165 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.3741 | 26.3259 | 155.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.4 | 155.6 | 157 | 1.4 | 1.3798 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0163 | 0.0165 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1039 | 1002 | 25.5066 | 26.3533 | 155.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.4 | 155.5 | 156.9 | 1.3 | 1.3406 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0162 | 0.0165 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5675 | 26.4 | 155.5
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.5 | 157 | 1.4 | 1.4 | 6.2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0162 | 0.016414 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4884 | 26.3884 | 155.5
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.4146 | 156.9 | 1.4 | 1.4 | 6.2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0162 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.8543 | 25.4092 | 26.3092 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.4 | 156.9 | 1.5 | 1.4562 | 6.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0162 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.33 | 26.37 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.4 | 156.9 | 1.4 | 1.4 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0162 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3 | 26.4 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.4 | 156.9 | 1.5 | 1.5 | 6.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.0164 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.3 | 26.3718 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.4 | 156.9 | 1.5 | 1.5 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.016314 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.3142 | 26.3071 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.6 | 155.3144 | 156.9 | 1.6 | 1.5555 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.4734 | 26.3867 | 155.3
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.6 | 155.4 | 157 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2854 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1038 | 1002.1457 | 25.5 | 26.334 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.6 | 155.4 | 157 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.5903 | 26.3 | 155.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.5 | 155.3 | 156.8 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.6514 | 26.3243 | 155.3
2008-05-08 | 5:50 | 0 | -0.0002 | -2.3423 | 1.5 | 155.4 | 156.9 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4965 | 26.3965 | 155.4
2008-05-08 | 5:50 | 0 | -0.0006 | -5.945 | 1.5 | 155.4 | 156.9 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0161 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4175 | 26.3175 | 155.4
2008-05-08 | 5:50 | 0 | -0.0006 | -5.6191 | 1.5 | 155.4 | 156.9 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016014 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.4 | 26.3 | 155.4
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.5 | 155.3 | 156.8 | 1.6 | 1.6 | 6.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016 | 0.0163 | 0.0093 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.441 | 26.3 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.4 | 155.3 | 156.7 | 1.6 | 1.6 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4597 | 26.3 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.4 | 155.3 | 156.7 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.3016 | 26.3 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.4 | 155.3 | 156.7 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.016 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.5357 | 26.3786 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.3 | 155.3 | 156.6 | 1.8 | 1.8181 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.4266 | 26.3422 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.3 | 155.3 | 156.6 | 1.9 | 1.9 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.4109 | 26.337 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155.3856 | 156.6 | 1.9 | 1.8501 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.5515 | 26.3838 | 155.4
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155.3146 | 156.5 | 1.9 | 1.9 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0162 | 0.0093 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.314 | 26.3047 | 155.3
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155.2144 | 156.4 | 1.8 | 1.8 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0162 | 0.0092 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1002 | 25.449 | 26.3 | 155.2
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155.1 | 156.3 | 1.9 | 1.9 | 6.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0159 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.5537 | 26.3 | 155.1
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155.1 | 156.3 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 27 | 5 | 195 | 59.3 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30.9241 | 53 | 1039 | 1002.856 | 25.5342 | 26.3 | 155.1
2008-05-08 | 5:50 | 0 | 0 | 0 | 1.2 | 155 | 156.2 | 1.8 | 1.7554 | 6.2 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.3759 | 26.312 | 155
2008-05-08 | 5:50 | 0 | 0.0003 | 2.5812 | 1.2 | 155 | 156.2 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.2187 | 26.3906 | 155
2008-05-08 | 5:50 | 0 | 0.001 | 10 | 1.2 | 155 | 156.2 | 1.7 | 1.7431 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1003 | 25.2703 | 26.4 | 155
2008-05-08 | 5:50 | 0 | 0.001 | 10 | 1.2 | 155.1856 | 156.4 | 1.7 | 1.7 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30.076 | 53 | 1039 | 1002 | 25.3496 | 26.3504 | 155.2
2008-05-08 | 5:50 | 0 | 0.001 | 10 | 1.3 | 155.2 | 156.5 | 1.8 | 1.7617 | 6.2 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0158 | 0.0161 | 0.0092 | 0.0001 |  | n/a |  | 30.924 | 53 | 1039 | 1001.1443 | 25.4576 | 26.3 | 155.2
2008-05-08 | 5:50 | 0 | 0.001 | 10 | 1.1 | 146.8279 | 148 | 1.7 | 1.6571 | 6.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0158 | 0.0161 | 0.008773 | 0.0001 |  | n/a |  | 30 | 53 | 1039 | 1001 | 25.6 | 26.3 | 146.8
2008-05-08 | 5:50 | 0 | 0.001 | 10 | -0.2 | 110.5288 | 110.5 | 1.5 | 1.5469 | 5.41 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.01453 | 0.014915 | 0.006615 | 0.0001 |  | n/a |  | 30 | 53 | 1038.689 | 1001 | 25.6 | 26.3 | 110.5
2008-05-08 | 5:50 | 0 | 0.001 | 10 | 0.4 | 105.2144 | 105.6 | 1.6 | 1.6 | 4.82 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.013358 | 0.013658 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1001.856 | 25.5337 | 26.3 | 105.2
2008-05-08 | 5:50 | 0 | 0.0009 | 8.7316 | 0.5 | 105.1144 | 105.6 | 1.6 | 1.6 | 4.6 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.4 | 5 | 195 | 59.3 | 0.0132 | 0.013414 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5911 | 26.3 | 105.1
2008-05-08 | 5:50 | 0 | 0.0001 | 0.54828 | 0.7 | 105 | 105.7 | 1.7 | 1.7 | 4.5 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.3 | 0.013015 | 0.013215 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6752 | 26.3 | 105
2008-05-08 | 5:50 | 0 | 0 | 0 | 0.7 | 105 | 105.7 | 1.6 | 1.6 | 4.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.3 | 0.013 | 0.0132 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.592 | 26.3 | 105
2008-05-08 | 5:50 | 0 | 0 | 0 | 0.5 | 104.9 | 105.4 | 1.7 | 1.7 | 4.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26 | 5 | 195 | 59.3 | 0.012814 | 0.013114 | 0.0063 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.4342 | 26.3 | 104.9
2008-05-08 | 5:50 | 0 | -0.0004 | -4.2308 | 0.5 | 104.9 | 105.4 | 1.7 | 1.7 | 4.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 26.6 | 5 | 195 | 59.3 | 0.012715 | 0.013015 | 0.0063 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4622 | 26.3 | 104.9
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 0.6 | 104.9 | 105.5 | 1.7 | 1.742 | 4.3 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.4 | 27 | 5 | 195 | 59.2856 | 0.0127 | 0.012914 | 0.0063 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.3753 | 26.3 | 104.9
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 0.7 | 104.9 | 105.6 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0126 | 0.0128 | 0.0063 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.2207 | 26.3 | 104.9
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 0.8 | 104.8 | 105.6 | 1.8 | 1.7575 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0125 | 0.0127 | 0.0063 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.2998 | 26.3 | 104.8
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 0.9 | 104.7 | 105.6 | 1.8 | 1.7602 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0124 | 0.0126 | 0.0063 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.6163 | 26.3 | 104.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 0.9 | 104.7 | 105.6 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2144 | 0.0124 | 0.012514 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5834 | 26.3 | 104.7
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1 | 104.6 | 105.6 | 1.8 | 1.759 | 4.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.2856 | 0.0123 | 0.0125 | 0.0063 | 0.0001 |  | n/a |  | 30.0759 | 53 | 1038 | 1002 | 25.5 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.1 | 104.6 | 105.7 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.012214 | 0.0124 | 0.0063 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1038 | 1002 | 25.4664 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.6 | 105.8 | 1.8 | 1.8 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26 | 5 | 195 | 59.3 | 0.0122 | 0.0123 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3081 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.6 | 105.8 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.6 | 5 | 195 | 59.3 | 0.012114 | 0.012214 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.375 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.6 | 105.8 | 1.8 | 1.7788 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0122 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3458 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.6 | 105.8 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0121 | 0.0122 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3 | 26.3 | 104.6
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.4144 | 105.6 | 1.7 | 1.7433 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 26.4 | 5 | 195 | 59.2856 | 0.012 | 0.012114 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3253 | 26.3 | 104.4
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.2 | 104.3146 | 105.5 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.4 | 26 | 5 | 195 | 59.2854 | 0.012 | 0.0121 | 0.0063 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.4825 | 26.3 | 104.3
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.1 | 104.3 | 105.4 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.012 | 0.0121 | 0.0063 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.3582 | 26.3 | 104.3
2008-05-08 | 5:50 | 0 | -0.001 | -10 | 1.1 | 104.3 | 105.4 | 1.8 | 1.7604 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.011914 | 0.0121 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1443 | 25.3 | 26.3 | 104.3
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.1 | 104.4 | 105.5 | 1.7 | 1.7 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2 | 0.0119 | 0.0121 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.359 | 26.3 | 104.4
2008-05-08 | 5:51 | 0 | -0.0004 | -4.0887 | 1.1 | 104.3144 | 105.4 | 1.8 | 1.7583 | 4.2 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.3 | 0.0119 | 0.012 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4913 | 26.3 | 104.3
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.1 | 104.3 | 105.4 | 1.7 | 1.7 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0118 | 0.012 | 0.0063 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.412 | 26.3 | 104.3
2008-05-08 | 5:51 | 0 | 0 | 0 | 1 | 104.1144 | 105.1 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0118 | 0.012 | 0.006214 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3329 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1 | 104.1 | 105.1 | 1.7 | 1.7 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0118 | 0.0119 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3926 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1 | 104.1 | 105.1 | 1.7 | 1.7432 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0118 | 0.0119 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.449 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1 | 104 | 105 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.4 | 5 | 195 | 59.3 | 0.0117 | 0.0119 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.3046 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1 | 104 | 105 | 1.7 | 1.7418 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0117 | 0.0119 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.3835 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.1 | 104.0856 | 105.2 | 1.8 | 1.7584 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0117 | 0.011814 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.4629 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.1 | 104.1 | 105.2 | 1.8 | 1.7594 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.2856 | 0.0117 | 0.0118 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.2 | 104.1 | 105.3 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26 | 5 | 195 | 59.3 | 0.0116 | 0.0118 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4785 | 26.3 | 104.1
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.2 | 104.0144 | 105.2 | 1.8 | 1.7509 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 26.6 | 5 | 195 | 59.3 | 0.0116 | 0.0117 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4007 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.2 | 104 | 105.2 | 1.9 | 1.856 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0116 | 0.0117 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.48 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.2 | 104 | 105.2 | 1.7 | 1.7214 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0116 | 0.0117 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5589 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.3 | 104 | 105.3 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0116 | 0.0117 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6 | 26.3 | 104
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.3 | 104 | 105.3 | 1.7 | 1.7 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.011514 | 0.011614 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.5475 | 26.3 | 104
2008-05-08 | 5:51 | 0 | -0.0001 | -0.81826 | 1.3 | 104 | 105.3 | 1.7 | 1.7405 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0115 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1457 | 25.3097 | 26.3 | 104
2008-05-08 | 5:51 | 0 | -0.0009 | -9.4315 | 1.3 | 104 | 105.3 | 1.7 | 1.6525 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0115 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5269 | 26.3 | 104
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.3 | 103.9 | 105.2 | 1.6 | 1.64 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0115 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4901 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.3 | 103.9 | 105.2 | 1.7 | 1.6572 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0115 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.3314 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.3 | 103.9 | 105.2 | 1.6 | 1.6 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0115 | 0.0116 | 0.0062 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.2271 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.3 | 103.9 | 105.2 | 1.7 | 1.6583 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0115 | 0.011514 | 0.0062 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3856 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.9 | 105.3 | 1.6 | 1.6 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5439 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.9 | 105.3 | 1.7 | 1.7 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1038 | 1003 | 25.6 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.9 | 105.3 | 1.7 | 1.7 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2854 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5699 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.9 | 105.3 | 1.8 | 1.8 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5192 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.9 | 105.3 | 1.8 | 1.7577 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.6775 | 26.3 | 103.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.8 | 105.2 | 1.8 | 1.8433 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.0114 | 0.0115 | 0.0062 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.3 | 103.8
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.8 | 105.2 | 1.9 | 1.9 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.011414 | 0.0062 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.6058 | 26.3 | 103.8
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.7144 | 105.1 | 1.8 | 1.8206 | 4.1 | 12.9 | 957 | 963 | 947 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0114 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.4736 | 26.3 | 103.7
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 103.7 | 105.1 | 1.9 | 1.9 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.011314 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002.856 | 25.4 | 26.3 | 103.7
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.6856 | 105.2 | 1.9 | 1.9 | 4.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0113 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.1443 | 25.4 | 26.3 | 103.7
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.6 | 105.1 | 2 | 1.9606 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0113 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.4 | 26.3 | 103.6
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.5 | 105 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0113 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3569 | 26.3 | 103.5
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.4 | 104.9 | 1.9 | 1.9 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0113 | 0.0114 | 0.0062 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.3 | 26.3 | 103.4
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.4 | 104.9 | 1.8 | 1.8 | 4.1 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0113 | 0.011314 | 0.0062 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002.1354 | 25.3063 | 26.3 | 103.4
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.5 | 103.4 | 104.9 | 1.9 | 1.8604 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0113 | 0.0113 | 0.0062 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6233 | 26.3 | 103.4
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.4 | 94.0696 | 95.5 | 1.8 | 1.7624 | 4.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0113 | 0.0113 | 0.005686 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4599 | 26.3 | 94.1
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 1.3 | 43.1304 | 44.4 | 1.6 | 1.6 | 3.26 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.008889 | 0.008732 | 0.002631 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3392 | 26.3 | 43.1
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.2 | 33.252 | 35.5 | 1.6 | 1.6 | 2.14 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007358 | 0.007072 | 0.002071 | 0.0001 |  | n/a |  | 30.9249 | 53 | 1038 | 1002 | 25.4 | 26.3 | 33.3
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.5 | 43.4512 | 45.9 | 1.6 | 1.64 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007214 | 0.0069 | 0.002657 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4 | 26.3 | 43.5
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.6 | 48.8 | 51.4 | 1.7 | 1.7 | 2.18 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0072 | 0.0069 | 0.003 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.856 | 25.6306 | 26.3 | 48.8
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.5 | 48.9 | 51.4 | 1.6 | 1.6396 | 2.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0072 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.5878 | 26.3 | 48.9
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.5 | 48.9856 | 51.5 | 1.8 | 1.7592 | 2.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2856 | 0.007114 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.5353 | 26.3 | 49
2008-05-08 | 5:51 | 0 | -0.001 | -10 | 2.4 | 49.0856 | 51.5 | 1.6 | 1.6108 | 2.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0071 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.144 | 25.5713 | 26.3 | 49.1
2008-05-08 | 5:51 | 0 | -0.0007 | -7.4673 | 2.1 | 49.2 | 51.3 | 1.7 | 1.7 | 2.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.007014 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.4127 | 26.3 | 49.2
2008-05-08 | 5:51 | 0 | 0 | 0 | 2 | 49.3 | 51.3 | 1.7 | 1.6605 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.007 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.4 | 26.3 | 49.3
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 49.3856 | 51.3 | 1.7 | 1.7145 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2 | 0.007 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.4521 | 26.3 | 49.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 49.4 | 51.3 | 1.7 | 1.6555 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.5 | 26.3 | 49.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.8 | 49.4 | 51.2 | 1.6 | 1.6453 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007 | 0.0068 | 0.003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4895 | 26.3 | 49.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.8 | 50.3 | 52.1 | 1.6 | 1.6478 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007 | 0.0068 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4102 | 26.3 | 50.3
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 50.4 | 52.3 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.007 | 0.006715 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4 | 26.3 | 50.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 50.4 | 52.3 | 1.7 | 1.6598 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.006914 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3519 | 26.3 | 50.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 50.4 | 52.3 | 1.6 | 1.6 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0069 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.144 | 25.3273 | 26.3 | 50.4
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.9 | 50.4856 | 52.4 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0069 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3935 | 26.3 | 50.5
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.8 | 50.5 | 52.3 | 1.7 | 1.6572 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.2144 | 0.0069 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3143 | 26.3 | 50.5
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.8 | 50.5 | 52.3 | 1.7 | 1.6587 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0069 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4298 | 26.3 | 50.5
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.8 | 50.5 | 52.3 | 1.7 | 1.6583 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0069 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5877 | 26.3 | 50.5
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.7 | 50.5 | 52.2 | 1.6 | 1.6494 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6542 | 26.3 | 50.5
2008-05-08 | 5:51 | 0 | 0 | 0 | 1.7 | 50.4144 | 52.1 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.856 | 25.5 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.7 | 50.4 | 52.1 | 1.6 | 1.6429 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.5 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.7 | 50.4856 | 52.2 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.439 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.6 | 50.4 | 52 | 1.6 | 1.6 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0067 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.6 | 50.5 | 52.1 | 1.7 | 1.6715 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 189.6 | 27 | 5 | 195 | 59.3 | 0.0068 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.3612 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.5 | 50.5 | 52 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2147 | 0.0068 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.2027 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.5 | 50.5 | 52 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.006714 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4335 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | -0.0008 | -7.7318 | 1.5 | 50.5 | 52 | 1.6 | 1.6483 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1002 | 25.5 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.5 | 50.4854 | 52 | 1.7 | 1.7424 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002.1457 | 25.5 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | -0.0006 | -5.7983 | 1.4 | 50.5 | 51.9 | 1.7 | 1.6588 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.53 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.4 | 51.8 | 1.7 | 1.6611 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.6893 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.4 | 51.8 | 1.7 | 1.6578 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1003 | 25.7 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.3 | 50.4 | 51.7 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.924 | 53 | 1038 | 1003 | 25.5938 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.3 | 50.4 | 51.7 | 1.6 | 1.6479 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.4677 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.3 | 50.4 | 51.7 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1003 | 25.4 | 26.3 | 50.4
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.3 | 50.4856 | 51.8 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002.144 | 25.4 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.3 | 50.5 | 51.8 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4698 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | -0.0001 | -0.96121 | 1.3 | 50.5 | 51.8 | 1.8 | 1.8 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2144 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.5983 | 26.3 | 50.5
2008-05-08 | 5:52 | 0 | -0.0009 | -9.3929 | 1.3 | 50.5856 | 51.9 | 1.7 | 1.7 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.6154 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.7 | 1.6871 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30 | 53 | 1038 | 1002 | 25.4071 | 26.2929 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.7 | 1.6588 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0067 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 30.076 | 53 | 1038 | 1002 | 25.4867 | 26.2133 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.006615 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.566 | 26.266 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.7 | 1.6626 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0066 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5512 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.6 | 1.649 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3964 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190 | 27 | 5 | 195 | 59.2856 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3175 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.7 | 1.7 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 190.6 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4858 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.3 | 50.6 | 51.9 | 1.6 | 1.5592 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.856 | 25.5175 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.4 | 50.6 | 52 | 1.6 | 1.6 | 2.1 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4204 | 26.2796 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.4 | 50.6 | 52 | 1.6 | 1.6 | 2.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4993 | 26.2007 | 50.6
2008-05-08 | 5:52 | 0 | -0.001 | -10 | 1.4 | 50.6 | 52 | 1.6 | 1.6398 | 1.86 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5.6 | 195 | 59.2144 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5 | 26.2 | 50.6
2008-05-08 | 5:52 | 0 | -0.0002 | -2.3954 | 1.4 | 50.6 | 52 | 1.7 | 1.7 | 1.8 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5.4 | 195 | 59.2856 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5 | 26.2581 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.6 | 52 | 1.7 | 1.6964 | 1.88 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5373 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.6856 | 52.1 | 1.8 | 1.7574 | 1.9 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5835 | 26.3 | 50.7
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.7 | 52.1 | 1.7 | 1.7248 | 1.9 | 12.9 | 957 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 25.5044 | 26.3 | 50.7
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.6 | 52 | 1.9 | 1.8574 | 1.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3508 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.6 | 52 | 1.8 | 1.8289 | 1.9 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3539 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.4 | 50.6 | 52 | 1.9 | 1.9 | 2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0066 | 0.0065 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4667 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.5 | 50.6 | 52.1 | 1.9 | 1.9 | 2 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.006414 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5874 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 1.5 | 50.6 | 52.1 | 1.9 | 1.9417 | 2 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0066 | 0.0064 | 0.0031 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.8557 | 25.5085 | 26.3 | 50.6
2008-05-08 | 5:52 | 0 | 0 | 0 | 0.1 | 38.2345 | 38.3 | 2 | 1.9529 | 1.93 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.006515 | 0.0064 | 0.002358 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3583 | 26.3 | 38.2
2008-05-08 | 5:52 | 0 | 0 | 0 | -1.1 | 13.08 | 13.1 | 1.8 | 1.8387 | 1.26 | 12.9 | 955 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.005758 | 0.006043 | 0.000872 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.144 | 25.4505 | 26.2498 | 13.1
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 2.962 | 3 | 1.8 | 1.758 | 0.74 | 12.9 | 955 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.004742 | 0.004756 | 0.000328 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6587 | 26.2 | 3
2008-05-08 | 5:52 | 0 | 0 | 0 | 0.3 | 2.3 | 2.6 | 1.7 | 1.7 | 0.58 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.004014 | 0.003914 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7658 | 26.2086 | 2.3
2008-05-08 | 5:52 | 0 | 0 | 0 | 0.2 | 2.3 | 2.5 | 1.8 | 1.7564 | 0.4 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.6 | 27 | 5 | 195 | 59.2856 | 0.003914 | 0.0039 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4488 | 26.2878 | 2.3
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.3 | 2.2 | 2.2 | 1.7 | 1.7 | 0.3 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0037 | 0.003714 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6002 | 26.3 | 2.2
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.2 | 2.1 | 2.1 | 1.7 | 1.7483 | 0.3 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0036 | 0.003614 | 0.0003 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.654 | 26.3 | 2.1
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.2 | 2 | 2 | 1.7 | 1.7 | 0.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0035 | 0.003514 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.549 | 26.3 | 2
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.2 | 1.8144 | 1.8 | 1.8 | 1.8 | 0.2 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0034 | 0.003414 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4047 | 26.3 | 1.8
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.8 | 1.8 | 1.7 | 1.741 | 0.2 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0034 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4837 | 26.3 | 1.8
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.7144 | 1.7 | 1.8 | 1.8417 | 0.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.003314 | 0.0034 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.374 | 26.3 | 1.7
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.6146 | 1.6 | 1.9 | 1.9 | 0.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.003215 | 0.0033 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3846 | 26.3 | 1.6
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.5144 | 1.5 | 1.8 | 1.8385 | 0.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.003114 | 0.0032 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5429 | 26.3 | 1.5
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.4 | 1.4 | 1.9 | 1.9 | 0.1 | 12.9 | 955 | 961 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0031 | 0.0031 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6993 | 26.3 | 1.4
2008-05-08 | 5:52 | 0 | 0 | 0 | -0.1 | 1.3 | 1.3 | 1.9 | 1.9 | 0.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.003 | 0.003014 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6201 | 26.3 | 1.3
2008-05-08 | 5:52 | 0 | 0 | 0 | 0 | 1.1144 | 1.1 | 2 | 1.9573 | 0.1 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.6 | 195 | 59.3 | 0.003 | 0.003 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5411 | 26.3 | 1.1
2008-05-08 | 5:52 | 0 | 0 | 0 | 0 | 1.1 | 1.1 | 1.8 | 1.7604 | 0.1 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.4 | 195 | 59.3 | 0.002915 | 0.002915 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5 | 26.3 | 1.1
2008-05-08 | 5:52 | 0 | 0 | 0 | 0.1 | 1 | 1.1 | 1.8 | 1.8 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0029 | 0.0029 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4648 | 26.3 | 1
2008-05-08 | 5:52 | 0 | 0.0003 | 2.504 | 0.1 | 0.91443 | 1 | 1.7 | 1.7 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.6 | 195 | 59.2144 | 0.0029 | 0.002814 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3065 | 26.3 | 0.9
2008-05-08 | 5:52 | 0 | 0.001 | 10 | 0.1 | 0.8144 | 0.9 | 1.7 | 1.7417 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5.4 | 195 | 59.2144 | 0.0028 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6037 | 26.3 | 0.8
2008-05-08 | 5:52 | 0 | 0.001 | 10 | 0.1 | 0.8 | 0.9 | 1.6 | 1.6434 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0028 | 0.0028 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.3 | 0.8
2008-05-08 | 5:53 | 0 | 0.0003 | 2.725 | 0.1 | 0.8 | 0.9 | 1.6 | 1.6413 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0028 | 0.002714 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.2656 | 0.8
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.8 | 0.9 | 1.6 | 1.6492 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0027 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.2136 | 0.8
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.7144 | 0.8 | 1.5 | 1.5428 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0027 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.2928 | 0.7
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.7 | 0.8 | 1.6 | 1.5574 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0027 | 0.0027 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.3 | 0.7
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.7 | 0.8 | 1.5 | 1.5 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.002614 | 0.002614 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1001.856 | 25.7 | 26.3 | 0.7
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.7 | 0.8 | 1.6 | 1.6 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0026 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.3 | 0.7
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.6144 | 0.7 | 1.6 | 1.5568 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0026 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6613 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | 1.6 | 1.6 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0026 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.856 | 25.3444 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | 1.6 | 1.6 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0026 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | 1.6 | 1.6413 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.002514 | 0.0026 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4892 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | -0.1 | 0.6 | 0.6 | 1.6 | 1.6 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0025 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6735 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | -0.1 | 0.6 | 0.6 | 1.6 | 1.6366 | 0 | 12.9 | 955 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0025 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5943 | 26.2943 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | -0.1 | 0.6 | 0.6 | 1.6 | 1.6489 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0025 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5151 | 26.2151 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | -0.1 | 0.6 | 0.6 | 1.7 | 1.7 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0025 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5642 | 26.2642 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | 1.7 | 1.7 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2146 | 0.002415 | 0.0025 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.5134 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.6 | 0.6 | 1.7 | 1.7 | 0 | 12.9 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0024 | 0.002414 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.4 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | -0.0001 | -0.56667 | 0 | 0.6 | 0.6 | 1.8 | 1.8 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0024 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4036 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | -0.0009 | -8.9 | -0.1 | 0.6 | 0.6 | 1.7 | 1.7 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0024 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.562 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | -0.001 | -10 | -0.1 | 0.6 | 0.6 | 1.8 | 1.8 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0024 | 0.0024 | 0.0002 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.5398 | 26.3 | 0.6
2008-05-08 | 5:53 | 0 | -0.001 | -10 | -0.1 | 0.5144 | 0.5 | 1.8 | 1.8 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0024 | 0.0024 | 0.000114 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.5789 | 26.3 | 0.5
2008-05-08 | 5:53 | 0 | -0.001 | -10 | 0 | 0.5 | 0.5 | 1.8 | 1.8372 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0024 | 0.0024 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.6814 | 26.3 | 0.5
2008-05-08 | 5:53 | 0 | -0.0007 | -7.4145 | 0 | 0.4 | 0.4 | 1.8 | 1.7834 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0023 | 0.0024 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.6022 | 26.3 | 0.4
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.4 | 0.4 | 1.7 | 1.7433 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.677 | 26.3 | 0.4
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.4 | 0.4 | 1.8 | 1.8 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.7 | 26.3 | 0.4
2008-05-08 | 5:53 | 0 | 0 | 0 | 0 | 0.4 | 0.4 | 1.7 | 1.7416 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2146 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6646 | 26.3 | 0.4
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.4 | 0.5 | 1.9 | 1.8573 | 0 | 13 | 956 | 962 | 945 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2856 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6146 | 26.3 | 0.4
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.3 | 0.4 | 1.8 | 1.7993 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.2144 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.6939 | 26.3 | 0.3
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.1 | 0.3 | 0.4 | 1.9 | 1.8591 | 0 | 13 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0023 | 0.0023 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002.8543 | 25.4808 | 26.3 | 0.3
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | 1.8 | 1.8 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 192 | 27 | 5 | 195 | 59.3 | 0.0023 | 0.002214 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4519 | 26.3 | 0.2
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | 1.8 | 1.8473 | 0 | 13 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191.4 | 27 | 5 | 195 | 59.2856 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.469 | 26.3 | 0.2
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.11457 | 0.3 | 1.9 | 1.9 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4107 | 26.3 | 0.1
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.1 | 0.3 | 1.9 | 1.9424 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2856 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4899 | 26.3 | 0.1
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.1 | 0.3 | 1.9 | 1.94 | 0 | 12.9 | 956 | 961 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.3 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.5 | 26.2308 | 0.1
2008-05-08 | 5:53 | 0 | 0 | 0 | 0.2 | 0.1 | 0.3 | 1.9 | 1.943 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0022 | 0.0022 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.4032 | 26.2484 | 0.1
2008-05-08 | 5:53 | 0 | -0.0001 | -0.85278 | 0.3 | 0.1 | 0.4 | 2 | 2 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2144 | 0.0022 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1003 | 25.3276 | 26.3 | 0.1
2008-05-08 | 5:53 | 0 | -0.0009 | -8.8978 | 0.3 | 0.1 | 0.4 | 2 | 2 | 0 | 12.9 | 956 | 962 | 946 | 22 | 0.4 | 0.01 | 985 | -32 | 0 | 30 | 191 | 27 | 5 | 195 | 59.2 | 0.0022 | 0.0021 | 0.0001 | 0.0001 |  | n/a |  | 31 | 53 | 1038 | 1002 | 25.3932 | 26.3 | 0.1
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
No Processing Errors Detected
Test Date | 2008-05-08
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | D06-SDS01
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 50568
Version | 6.907
CO Span(%) | 0.5
CO2 Span(%) | 6
C6H14 Span(ppm) | 325
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 208
Version | 1.611
NO Span(ppm) | 301
NO2 Span(ppm) | 505
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Diesel_FID
Serial | 123
Range(ppmC)1 | 100.00 Bottle(ppmC) = 0000000
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 605.0
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description:
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | D06-SDS01
AMBII RPM Multiplier | 2
Torque (ecm or calc) | calc
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 5:32
Test End Time | 5:53
Test Duration (s) | 1261
Average Ambient Temperature (deg C) | -32.037
Average Ambient Pressure (mbar) | 985
Average Relative Humidity (%) | 21.99
Average Absolute Humidity (grains/lb dry air) | 0.4
Average Kh Factor | 0.886
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Percent Time in NTE Zone | 0
Total Number of NTE Windows | 0
Total Number of non-consecutive NTE Windows | 0
| Avg. NTE Emissions | %: Maximum NTE Emissions; %: Std. Dev.
CO2 (g/kW-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/08/2008 05:32:05.181 - None Found
Warnings:
0X0000 - 05/08/2008 05:32:05.181 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 1261
Brake Specific Dropout Limit Count | 1261
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 |  |  |  | -1
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Test Information:
SEMTECH_DATA_FILE | 2008-05-08 | 5:31 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]

## NO2 Linearity
Test Date:
Test Time:
5/8/2008
5:32:05
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## NDUV
Test Date:
Test Time:
5/8/2008
5:32:05
SEMTECH Serial Number:
D06-SDS01
NDUV Serial Number:
208

## O2 Linearity
Test Date:
Test Time:
5/8/2008
5:32:05
SEMTECH Serial Number:
D06-SDS01

## O2
Test Date:
Test Time:
5/8/2008
5:32:05
SEMTECH Serial Number:
D06-SDS01