## Test stand info
Enter the SEMTECH serial numbers, gas dividers and bottle data below for the test stand
Customer: | EPA
Gas divider: | Serial number | Last calibration date | Calibration due date
Model:
Horiba SGD-710C | GDU-000 | 2007-09-20 | 2008-09-20
Horiba SGD-A10 | GDU-003 | 2007-07-24 | 2008-08-24
Horiba SGD-710C | GDU-002 | 39457 | 39823
0 | 0 | 0 | 0
SEMTECH test software: | 10.09 SP1
SEMTECH Serial Number: | G05-SDS02
AMBII Serial Number: | 26374
NDUV Serial Number: | 210
FID Serial Number: | 66
Heated Filter Serial number: | 0
|  | Cylinder supplier | Cylinder # | Cylinder accuracy | Expiration date
Span bottle | Actual bottle value |  |  |  |  | THC
300 ppm Prop | 319 | Scott gas | ALM023745 | ±1% | 2011-04-02 | 957
1200 ppm CO | 1201
12% CO2 | 12
1500 ppm NO | 1495
Audit bottle | [5 empty] | THC
33 ppm Prop | 31.04 | Scott gas | ALM035016 | ±1% | 2010-01-29 | 93.12
200 ppm CO | 200.6
6% CO2 | 6
300 ppm NO | 298.1
BAR97 High | [5 empty] | THC
3200 ppm Prop | 3219 |  |  |  |  | 9657
80000 ppm CO | 80200
12% CO2 | 12.02
3000 ppm NO | 3024
BAR97 Low | [5 empty] | THC
200 ppm Prop | 201.6 |  |  |  |  | 604.8
5000 ppm CO | 4994
6% CO2 | 6
300 ppm NO | 300
THC Bottles
Propane value | [5 empty] | THC
13333.3 ppm | 13050 |  |  |  |  | 39150
NO2 bottles
80 - 150 ppm | 117
500 ppm | 505 | Scott gas | ALM011971 | ±1% | 2010-04-08
Formulas:
Last calibration date = '[1]Test stand info'!$C8
Calibration due date = '[1]Test stand info'!$D8
Expiration date = '[1]Test stand info'!$F23

## View These Results
SEMTECH Serial Number: | G05-SDS02 |  | CO2 - AMBII Serial Number: |  |  | 26374 |  |  | CO - AMBII Serial Number: |  |  | 26374 |  | HC - AMBII Serial Number: |  |  | 26374 |  | NO - NDUV Serial Number: |  |  | 210 |  | THC - FID Serial Number: |  |  | 66
|  |  | Bottle Value % | Horiba SGD-710C | Absolute Error Limit % | Relative Error Limit % |  |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppmC3 |  | Absolute Error Limit ppmC6 | Relative Error Limit % |  | Bottle Value ppm |  | Absolute Error Limit ppm | Relative Error Limit % |  | Bottle Value ppm | Range ppmC | Absolute Error Limit ppm | Relative Error Limit %
Test Date: | 2008-05-07 |  | 6 | X | 0.1 | 3 |  |  | 200.6 |  | 50 | 3 |  | 31.04 |  | 4 | 3 |  | 298.1 |  | 15 | 3 |  | 93.12 | 100 | 5 | 2
Test Time: | 13:58
Test Software: | 10.09 SP1
Gas Divider Step |  |  | CO2 Accuracy Test Results | [5 empty] | CO Accuracy Test Results |  |  |  |  | HC Accuracy Test Results |  |  |  |  | NO Accuracy Test Results |  |  |  |  | THC Accuracy Test Results
Start | End | Step | Divider Value % | Corrected Divider % | Measured % | Absolute Error % | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppm | Measured ppmC3 | Absolute Error ppmC3 | Relative Error % of reading | Status (Pass if blank) | Divider Values ppm | Measured ppm | Absolute Error ppm | Relative Error % of reading | Status (Pass if blank) | Divider Value ppmC | Measured ppmC | Absolute Error ppmC | Relative Error % of reading | Status (Pass if blank)
30 | 45 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | -14 | -14 |  |  | 0 | 985 | 985 |  |  | 0 | 0.2 | 0.2 |  |  | 0 | -0.89 | -0.89
90 | 105 | 1 | 0.6 | 0.596 | 0.64 | 0.044 | 7.38 |  | 20.06 | 20 | 0 | 0 |  | 3.104 | 985 | 981.9 | 31633.38 | Fail | 29.81 | 30.1 | 0.3 | 1.01 |  | 9.312 | 8.29 | -1.02 | -10.95
150 | 165 | 2 | 1.2 | 1.193 | 1.24 | 0.047 | 3.94 |  | 40.12 | 41 | 1 | 2.49 |  | 6.208 | 985.3 | 979.1 | 15771.59 | Fail | 59.62 | 62.1 | 2.5 | 4.19 |  | 18.624 | 17.69 | -0.93 | -4.99
210 | 225 | 3 | 1.8 | 1.791 | 1.837 | 0.046 | 2.57 |  | 60.18 | 63 | 3 | 4.99 |  | 9.312 | 985.5 | 976.2 | 10483.25 | Fail | 89.43 | 91.5 | 2.1 | 2.35 |  | 27.936 | 26.99 | -0.95 | -3.4
270 | 285 | 4 | 2.4 | 2.39 | 2.444 | 0.054 | 2.26 |  | 80.24 | 88 | 8 | 9.97 |  | 12.416 | 985.4 | 973 | 7836.66 | Fail | 119.24 | 121 | 1.8 | 1.51 |  | 37.248 | 35.96 | -1.29 | -3.46
330 | 345 | 5 | 3 | 2.989 | 3.042 | 0.053 | 1.77 |  | 100.3 | 100 | 0 | 0 |  | 15.52 | 985.3 | 969.8 | 6248.71 | Fail | 149.05 | 152 | 2.9 | 1.95 |  | 46.56 | 45.08 | -1.48 | -3.18
390 | 405 | 6 | 3.6 | 3.59 | 3.63 | 0.04 | 1.11 |  | 120.36 | 122 | 2 | 1.66 |  | 18.624 | 985.2 | 966.6 | 5190.08 | Fail | 178.86 | 181 | 2.1 | 1.17 |  | 55.872 | 54.5 | -1.37 | -2.45
450 | 465 | 7 | 4.2 | 4.191 | 4.216 | 0.025 | 0.6 |  | 140.42 | 140 | 0 | 0 |  | 21.728 | 985.2 | 963.5 | 4434.37 | Fail | 208.67 | 209.9 | 1.2 | 0.58 |  | 65.184 | 63.81 | -1.37 | -2.1
510 | 525 | 8 | 4.8 | 4.793 | 4.793 | 0 | 0 |  | 160.48 | 154 | -6 | -3.74 |  | 24.832 | 985.3 | 960.5 | 3867.99 | Fail | 238.48 | 239.9 | 1.4 | 0.59 |  | 74.496 | 73.2 | -1.3 | -1.75
570 | 585 | 9 | 5.4 | 5.396 | 5.389 | -0.007 | -0.13 |  | 180.54 | 170 | -11 | -6.09 |  | 27.936 | 985.6 | 957.7 | 3428.19 | Fail | 268.29 | 271 | 2.7 | 1.01 |  | 83.808 | 82.69 | -1.12 | -1.34
630 | 645 | 10 | 6 | 6 | 5.991 | -0.009 | -0.15 |  | 200.6 | 190 | -11 | -5.48 |  | 31.04 | 985.5 | 954.5 | 3075.06 | Fail | 298.1 | 300.1 | 2 | 0.67 |  | 93.12 | 92 | -1.12 | -1.2
690 | 705 | 9 | 5.4 | 5.396 | 5.385 | -0.011 | -0.2 |  | 180.54 | 179 | -2 | -1.11 |  | 27.936 | 985.4 | 957.5 | 3427.48 | Fail | 268.29 | 271.6 | 3.3 | 1.23 |  | 83.808 | 83.31 | -0.5 | -0.6
750 | 765 | 8 | 4.8 | 4.793 | 4.793 | 0 | 0 |  | 160.48 | 164 | 4 | 2.49 |  | 24.832 | 985.5 | 960.7 | 3868.8 | Fail | 238.48 | 240.2 | 1.7 | 0.71 |  | 74.496 | 74.4 | -0.1 | -0.13
810 | 825 | 7 | 4.2 | 4.191 | 4.22 | 0.029 | 0.69 |  | 140.42 | 140 | 0 | 0 |  | 21.728 | 985.5 | 963.8 | 4435.75 | Fail | 208.67 | 211.6 | 2.9 | 1.39 |  | 65.184 | 65.33 | 0.15 | 0.23
870 | 885 | 6 | 3.6 | 3.59 | 3.631 | 0.041 | 1.14 |  | 120.36 | 112 | -8 | -6.65 |  | 18.624 | 985.4 | 966.8 | 5191.15 | Fail | 178.86 | 181.9 | 3 | 1.68 |  | 55.872 | 56.83 | 0.96 | 1.72
930 | 945 | 5 | 3 | 2.989 | 3.03 | 0.041 | 1.37 |  | 100.3 | 100 | 0 | 0 |  | 15.52 | 985.7 | 970.2 | 6251.29 | Fail | 149.05 | 150.9 | 1.8 | 1.21 |  | 46.56 | 46.2 | -0.36 | -0.77
990 | 1005 | 4 | 2.4 | 2.39 | 2.431 | 0.041 | 1.72 |  | 80.24 | 80 | 0 | 0 |  | 12.416 | 985.3 | 972.9 | 7835.86 | Fail | 119.24 | 122.4 | 3.2 | 2.68 |  | 37.248 | 36.9 | -0.35 | -0.94
1050 | 1065 | 3 | 1.8 | 1.791 | 1.839 | 0.048 | 2.68 |  | 60.18 | 64 | 4 | 6.65 |  | 9.312 | 985.3 | 976 | 10481.1 | Fail | 89.43 | 91.9 | 2.5 | 2.8 |  | 27.936 | 27.38 | -0.56 | -2
1110 | 1125 | 2 | 1.2 | 1.193 | 1.232 | 0.039 | 3.27 |  | 40.12 | 35 | -5 | -12.46 |  | 6.208 | 985.8 | 979.6 | 15779.64 | Fail | 59.62 | 62.5 | 2.9 | 4.86 |  | 18.624 | 17.7 | -0.92 | -4.94
1170 | 1185 | 1 | 0.6 | 0.596 | 0.641 | 0.045 | 7.55 |  | 20.06 | 13 | -7 | -34.9 |  | 3.104 | 985.1 | 982 | 31636.6 | Fail | 29.81 | 31 | 1.2 | 4.03 |  | 9.312 | 8.26 | -1.05 | -11.28
1230 | 1245 | 0 | 0 | 0 | -0.005 | -0.005 |  |  | 0 | -7 | -7 |  |  | 0 | 985.3 | 985.3 |  |  | 0 | 1 | 1 |  |  | 0 | -0.96 | -0.96
Comments |  |  | CO2 Linearity Test Results | [5 empty] | CO Linearity Test Results |  |  |  |  | HC Linearity Test Results |  |  |  |  | NO Linearity Test Results |  |  |  |  | THC Linearity Test Results
|  |  |  | Actual | Criteria | Status (Pass if blank) |  |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank) |  |  | Actual | Criteria | Status (Pass if blank)
|  |  | Intercept | 0.043 | 0.06 |  |  |  | Intercept | -2 | 2 |  |  | Intercept | 985 | 0 | Fail |  | Intercept | 1.4 | 3 |  |  | Intercept | -0.9 | 0.93
|  |  | Slope | 0.99 | 0.99 - 1.01 |  |  |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 0.01 | 0.99 - 1.01 | Fail |  | Slope | 1 | 0.99 - 1.01 |  |  | Slope | 1 | 0.99 - 1.01
|  |  | SEE | 0.0217 | 0.06 |  |  |  | SEE | 5.7539 | 2.006 | Fail |  | SEE | 0.1932 | 0.3104 |  |  | SEE | 0.8357 | 2.981 |  |  | SEE | 0.6035 | 0.9312
|  |  | r2 | 0.9999 | 0.998 |  |  |  | r2 | 0.992 | 0.998 | Fail |  | r2 | 0.1648 | 0.998 | Fail |  | r2 | 0.9999 | 0.998 |  |  | r2 | 0.9996 | 0.998
For Reference Only
For Reference Only

## Copy Data Here
DATE | TIME | CO2 | CO | CO | NO | NO2 Raw | NOx | THC | THC | O2 | Power Supply Voltage | Sample Pump Pressure | Drain Pump 1 Pressure | Drain Pump 2 Pressure | Relative Humidity | Absolute Humidity | Volume Humidity | Local Ambient Pressure | Local Ambient Temperature | Auxiliary Temperature | CJC Temperature | Heated Filter Temperature | External Line Temperature | Chiller Temperature | THC Oven Temperature | NDUV Block1 Temperature | NO Raw Modulation | Ref1 Raw Modulation | NO2 Raw Modulation | Ref2 Raw Modulation |  | Not Available | NMHC | NDUV Ambient Temperature | AMBII Ambient Temperature | AMBII Pressure | NDUV Pressure | Manifold Relative Humidity | Manifold Temperature | NO2
sDATE | sTIME | iAMBII_CO2 | iAMBII_CO | iAMBII_COPPM | iNDUV_NO | iNDUV_NO2_RAW | iNOxz | iFID_THC | iHCz | iAMBII_O2 | iSCB_PSV | iSCB_SPP | iSCB_DP1P | iSCB_DP2P | iSCB_RH | iHum_Abs | iHum_Vol | iSCB_LAP | iSCB_LAT | iSCB_ET | iSCB_CJCT | iSCB_FT | iSCB_ELT | iSCB_CT | iFID_OT | iNDUV_BT1 | iNDUV_NO_RAWMOD | iNDUV_REF1_RAWMOD | iNDUV_NO2_RAWMOD | iNDUV_REF2_RAWMOD | iFID2_CH4 | iFID2_OT | iNMHCz | iNDUV_ATEMP | iAMBII_ATEMP | iAMBII_Pressure | iNDUV_Pressure | iHUM_HUMID | iHUM_TEMP | iNDUV_NO2
mm/dd/yyyy | hh:mm:ss.xxx | % | % | ppm | ppm | ppm | ppm | ppmC | ppmC | % | Vdc | mbar | mbar | mbar | % | grains/lb dry air | % | mbar | deg C | deg C | deg C | deg C | deg C | deg C | deg C | deg C | [5 empty] | n/a | ppmC | deg C | deg C | mbar | mbar | % | deg C | ppm
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.8 | 1.2 | -0.5 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0001 | 0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.7 | 26.6 | 0.8
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.7 | 1.1 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0001 | 5.7e-05 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.7 | 26.6 | 0.7
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.7 | 1.1 | -0.5 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.3569 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.7 | 26.6 | 0.7
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.6 | 1 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.637 | 26.6 | 0.6
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.5 | 0.9 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6417 | 26.6 | 0.5
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.4 | 0.5 | 0.9 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 984.568 | 29.7412 | 26.6 | 0.5
2008-05-07 | 13:58 | 0.01 | 0 | 0 | 0.3 | 0.5 | 0.8 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 984 | 29.8983 | 26.6 | 0.5
2008-05-07 | 13:58 | 0.005 | 0 | 0 | 0.3 | 0.5 | 0.8 | -0.5 | 0 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6667 | 26.6 | 0.5
2008-05-07 | 13:58 | 0 | 0 | 0 | 0.3 | 0.5 | 0.8 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.769 | 26.6 | 0.5
2008-05-07 | 13:58 | 0 | 0 | 0 | 0.3 | 0.45703 | 0.8 | -0.5 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8651 | 26.6 | 0.5
2008-05-07 | 13:58 | 0 | 0 | 0 | 0.3 | 0.4 | 0.7 | -0.5 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.7865 | 26.6 | 0.4
2008-05-07 | 13:58 | 0 | -0.0004 | -4.2411 | 0.3 | 0.4 | 0.7 | -0.5 | 0 | 0 | 12.7 | 910 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.7079 | 26.6 | 0.4
2008-05-07 | 13:58 | 0 | -0.003 | -30 | 0.3 | 0.4 | 0.7 | -0.5 | 0 | 0 | 12.7 | 910 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3565 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8414 | 26.6 | 0.4
2008-05-07 | 13:58 | 0 | -0.003 | -30 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.7 | 910 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0001 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8507 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.003 | -30 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.6 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 5.7e-05 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6279 | 0.3
2008-05-07 | 13:58 | 0 | -0.0025 | -25.1712 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.7 | 909 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8065 | 26.6935 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8851 | 26.6149 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8365 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.7 | 0 | 0 | 12.9 | 908 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 984.432 | 29.8 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.7 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.6 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3432 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.3 | 0.5 | -0.7 | 0 | 0 | 12.9 | 908 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8566 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.25689 | 0.5 | -0.7 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8647 | 26.6 | 0.3
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.7 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.7 | 0 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.5691 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.7 | 0 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.432 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.1 | 0.2 | 0.3 | -0.7 | 0 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.1 | 0.2 | 0.3 | -0.8 | 0 | 0 | 12.9 | 909 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.1 | 0.2 | 0.3 | -0.8 | 0 | 0 | 12.8 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.8 | 0 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.8 | 0 | 0 | 12.8 | 910 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.928 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.8 | 0 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | 0 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9148 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.8 | 0 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8211 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -0.8 | 0 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8998 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.0019 | -18.6689 | 0.2 | 0.2 | 0.4 | -0.9 | 0 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9783 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10.1365 | 0.2 | 0.2 | 0.4 | -0.9 | 0 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9431 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.2 | 0.2 | 0.4 | -0.9 | 0 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8645 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.2 | 0.3 | -0.9 | 0 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8141 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.2 | 0.3 | -0.9 | 0 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8926 | 26.6 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.1568 | 0.3 | -0.9 | 0 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.829 | 26.671 | 0.2
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6503 | 0.1
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.1
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.6 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.1
2008-05-07 | 13:58 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -4.3e-05 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1.1 | 0 | 0 | 12.7 | 911 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.8643 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.6 | 911 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9428 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.1 | 0.1 | 0.2 | -1 | 0 | 0 | 12.7 | 911 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985.568 | 29.9572 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | -1 | 0 | 0 | 12.6 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.8 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | -1.1 | 0 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3432 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.9563 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.2 | 0.1 | 0.3 | -1.1 | 0 | 0 | 12.6 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0569 | 26.6 | 0.1
2008-05-07 | 13:59 | 0 | -0.001 | -10 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0286 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.0007 | -7.0726 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3568 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9144 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 911 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9929 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 910 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0709 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0504 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | 0 | 0 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.6 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.0009 | -9.1681 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.7 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9929 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.9143 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.002 | -20 | 0.2 | 0.2 | 0.4 | -1.1 | 0 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.9643 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.002 | -20 | 0.1 | 0.2 | 0.3 | -1.1 | 0 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 986 | 29.9571 | 26.6 | 0.2
2008-05-07 | 13:59 | 0 | -0.0012 | -11.7936 | 0.1 | 0.2 | 0.3 | -1.1 | 0 | 0 | 12.8 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 986 | 29.9215 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.126 | -0.001 | -10 | 0.1 | 0.2 | 0.3 | -1.1 | 0 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0001 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.432 | -0.0005 | -5.4406 | 0.7 | 0.2 | 0.9 | -1.1 | 0 | 0 | 12.8 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 3e-05 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.621 | 0.0003 | 2.7273 | 13.6 | 0.2 | 13.8 | -1.1 | 0 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.002234 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.0572 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.63 | 0.001 | 10 | 23.5 | 0.2 | 23.7 | 2.2 | 2.2292 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.003902 | -0.0001 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.63 | 0.001 | 10 | 27.5 | 0.2 | 27.7 | 6.3 | 6.3141 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0046 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0856 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.63 | 0.001 | 10 | 28.5 | 0.2 | 28.7 | 8.4 | 8.4373 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.004743 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 984 | 30.0071 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.633 | 0.001 | 10 | 29 | 0.2 | 29.2 | 9 | 9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0048 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 984 | 30.0713 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.001 | 10 | 29.1 | 0.2 | 29.3 | 8.7 | 8.716 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0048 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 984.432 | 30.0501 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.001 | 10 | 29.1 | 0.2 | 29.3 | 8.4 | 8.4101 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0048 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.001 | 10 | 29.2 | 0.2 | 29.4 | 8.4 | 8.4035 | 0 | 12.7 | 912 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.004843 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9928 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.001 | 10 | 29.3 | 0.2 | 29.5 | 8.5 | 8.5 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9142 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.0018 | 17.7543 | 29.3 | 0.2 | 29.5 | 8.4 | 8.4105 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0288 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.4 | 0.2 | 29.6 | 8.3 | 8.2769 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.057 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.5 | 0.2432 | 29.7 | 8.2 | 8.2 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3568 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.5 | 0.3 | 29.8 | 8.2 | 8.2 | 0 | 12.7 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0001 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.6 | 0.3 | 29.9 | 8.3 | 8.3 | 0 | 12.6 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0784 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.7 | 0.3 | 30 | 8.3 | 8.3 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9287 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.7 | 0.3 | 30 | 8.4 | 8.4 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8716 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.7 | 0.3 | 30 | 8.4 | 8.4 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.7 | 0.3 | 30 | 8.4 | 8.4 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.8 | 0.3 | 30.1 | 8.4 | 8.4 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.004943 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.3 | 30.2 | 8.4 | 8.4 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0498 | 26.6 | 0.3
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.4 | 8.4 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.4 | 8.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0928 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.4 | 8.4455 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0142 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.5 | 8.5 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.5 | 8.4904 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 30 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.5 | 8.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 30.0215 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 29.9 | 0.4 | 30.3 | 8.4 | 8.4079 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 30.0999 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30 | 0.4 | 30.4 | 8.4 | 8.4057 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0219 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30 | 0.4 | 30.4 | 8.4 | 8.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.057 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30 | 0.4 | 30.4 | 8.4 | 8.4 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30 | 0.4432 | 30.4 | 8.3 | 8.3104 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.000143 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.4
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30 | 0.5 | 30.5 | 8.3 | 8.3 | 0 | 12.6 | 913 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0285 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0501 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.208 | 0 | 12.6 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0426 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.2 | 0 | 12.6 | 913 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3568 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9073 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.9859 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.2 | 0 | 12.7 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.2 | 0 | 12.7 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.1 | 0.5 | 30.6 | 8.2 | 8.2 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.065 | 26.6 | 0.5
2008-05-07 | 13:59 | 0.64 | 0.002 | 20 | 30.2 | 0.5 | 30.7 | 8.2 | 8.2 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2997 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.002 | 20 | 30.2 | 0.5 | 30.7 | 8.2 | 8.2 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1428 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.002 | 20 | 30.3 | 0.5 | 30.8 | 8.2 | 8.2 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.002 | 20 | 30.3 | 0.5 | 30.8 | 8.2 | 8.2 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.002 | 20 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.002 | 20 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.0026 | 26.2284 | 30.4 | 0.5 | 30.9 | 8.1 | 8.1066 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.05 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9425 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8221 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.4 | 0.5 | 30.9 | 8.2 | 8.2 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0576 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.4 | 0.5 | 30.9 | 8.1 | 8.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1645 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.1569 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.0035 | 34.5185 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1217 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.0037 | 37.3605 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1998 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022.6742 | 986 | 30.1216 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.64 | 0.003 | 30 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 0.825 | 0.003 | 30 | 30.5 | 0.5 | 31 | 8.1 | 8.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.197 | 0.006 | 59.8507 | 34.5 | 0.5 | 35 | 8.2 | 8.2 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.00569 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.58 | 0.0071 | 71.4333 | 54.6 | 0.5 | 55.1 | 8.2 | 8.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.009121 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 984.432 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.312 | 0.0055 | 54.5815 | 73.4 | 0.5 | 73.9 | 13.8 | 13.8444 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.012343 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.267 | 0.004 | 40 | 71.5 | 0.5 | 72 | 22.3 | 22.3018 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.011984 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1499 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.26 | 0.004 | 40 | 67.1 | 0.5 | 67.6 | 20.1 | 20.1326 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.011414 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1712 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.255 | 0.004 | 40 | 66.1 | 0.5 | 66.6 | 18.5 | 18.5237 | 0 | 12.6 | 912 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.011114 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1074 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.25 | 0.004 | 40 | 64.8 | 0.5 | 65.3 | 17.9 | 17.914 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010857 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1859 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.25 | 0.004 | 40 | 63.8 | 0.5 | 64.3 | 17.7 | 17.6747 | 0 | 12.6 | 912 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.010757 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.071 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.25 | 0.004 | 40 | 63.6 | 0.5 | 64.1 | 17.6 | 17.6 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.242 | 0.004 | 40 | 63.5 | 0.5 | 64 | 17.6 | 17.6 | 0 | 12.6 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0217 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.5 | 0.5 | 64 | 17.6 | 17.6 | 0 | 12.7 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.5 | 0.5 | 64 | 17.6 | 17.6 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190.3 | 6 | 195 | 59.4 | 0.0107 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.4 | 0.5 | 63.9 | 17.6 | 17.6 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 30.1575 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.3 | 0.4568 | 63.7 | 17.6 | 17.6 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190.7 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2 | 26.6 | 0.5
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.2 | 0.4 | 63.6 | 17.6 | 17.6 | 0 | 12.6 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1854 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63.1 | 0.4 | 63.5 | 17.6 | 17.6 | 0 | 12.6 | 912 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 30.1068 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 63 | 0.4 | 63.4 | 17.5 | 17.5049 | 0 | 12.6 | 912 | 933 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 30.1718 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.0048 | 47.651 | 63 | 0.4 | 63.4 | 17.5 | 17.5 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1496 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.8 | 0.4 | 63.2 | 17.5 | 17.5 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.010557 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.4309 | 30.1 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.6 | 0.4 | 63 | 17.6 | 17.5927 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0105 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1149 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.4 | 0.4 | 62.8 | 17.6 | 17.6 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0105 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.432 | 30.2718 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.3 | 0.4 | 62.7 | 17.6 | 17.5932 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0105 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.568 | 30.1708 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.3 | 0.4 | 62.7 | 17.6 | 17.6 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0105 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1433 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.3 | 0.4 | 62.7 | 17.6 | 17.6 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0105 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1782 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.005 | 50 | 62.2 | 0.4 | 62.6 | 17.6 | 17.6 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.568 | 30.1008 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.0045 | 45.4056 | 62.2 | 0.4 | 62.6 | 17.6 | 17.6 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2573 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.2 | 0.4 | 62.6 | 17.7 | 17.6946 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1852 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.3 | 0.4 | 62.7 | 17.6 | 17.6 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0639 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.3 | 0.4 | 62.7 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.2 | 0.4 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.2 | 0.4 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1438 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.2 | 0.4 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1495 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62.1 | 0.4 | 62.5 | 17.7 | 17.7 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0709 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.0041 | 40.8765 | 62.1 | 0.4 | 62.5 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022.2699 | 985 | 30.0076 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.0049 | 49.2237 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0859 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.0042 | 42.3589 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0354 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.7 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.13 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.2781 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1995 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1216 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1362 | 26.6 | 0.4
2008-05-07 | 14:00 | 1.24 | 0.004 | 40 | 62 | 0.4 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2 | 26.6 | 0.4
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.3568 | 62.4 | 17.7 | 17.7 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2 | 26.6 | 0.4
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.3 | 62.3 | 17.7 | 17.7 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.128 | 26.6 | 0.3
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.3 | 62.2 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1505 | 26.6 | 0.3
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.3 | 62.2 | 17.7 | 17.7 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1419 | 26.6 | 0.3
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.3 | 62.2 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0229 | 26.6 | 0.3
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.2 | 62.1 | 17.7 | 17.7 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.3569 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2585 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.2 | 62.1 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1704 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.2 | 62.1 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0104 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 61.9 | 0.2 | 62.1 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.010357 | -0.0002 | 4.3e-05 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.122 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1995 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.3431 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1211 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022.2316 | 985 | 30.2152 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62 | 0.2 | 62.2 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62.1 | 0.2 | 62.3 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2703 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 17.7 | 17.7079 | 0 | 12.8 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.010357 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1132 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.24 | 0.004 | 40 | 62.2 | 0.2 | 62.4 | 17.8 | 17.7934 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2439 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.508 | 0.0063 | 62.8874 | 62.2 | 0.2 | 62.4 | 17.8 | 17.8 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0103 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1483 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.87 | 0.007 | 70 | 73 | 0.2 | 73.2 | 17.8 | 17.8 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.011394 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 14:01 | 1.869 | 0.007 | 70 | 87 | 0.1 | 87.1 | 18.5 | 18.5415 | 0 | 12.6 | 914 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.014559 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0232 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.861 | 0.007 | 70 | 90.5 | 0.1 | 90.6 | 26.2 | 26.2484 | 0 | 12.6 | 914 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.015143 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2578 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.852 | 0.007 | 70 | 91 | 0.1 | 91.1 | 27.4 | 27.3866 | 0 | 12.6 | 913 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0152 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 30.1706 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.1 | 0.1 | 91.2 | 27.5 | 27.5059 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1869 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.1 | 0.1 | 91.2 | 27.3 | 27.3088 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.256 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.1 | 0.1 | 91.2 | 27.2 | 27.2 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0152 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1005 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.2 | 27.2 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.015243 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1787 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.2 | 27.2 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.2 | 27.2 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2363 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.85 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.2 | 27.2 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2851 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.846 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.2 | 27.2 | 0 | 12.7 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0003 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2065 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.2 | 0.1 | 91.3 | 27.1 | 27.1 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.128 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.1 | 27.1 | 0 | 12.6 | 912 | 933 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2518 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.3 | 0.1 | 91.4 | 27.1 | 27.1 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3353 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.0064 | 63.99 | 91.5 | 0.1 | 91.6 | 27.1 | 27.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2131 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.0064 | 64.3515 | 91.5 | 0.1 | 91.6 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.5 | 0.1 | 91.6 | 27.1 | 27.1 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1005 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.5 | 0.1 | 91.6 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1791 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3731 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.5 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0.1 | 91.5 | 27.1 | 27.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 30.4851 | 26.6 | 0.1
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0 | 91.4 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31.7932 | 59 | 1022 | 985 | 30.4065 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0 | 91.4 | 27.1 | 27.0939 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0 | 91.4 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2988 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.2583 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.007 | 70 | 91.4 | 0 | 91.4 | 27.1 | 27.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4077 | 26.6 | 0
2008-05-07 | 14:01 | 1.838 | 0.007 | 70 | 91.5 | 0 | 91.5 | 27 | 27.0048 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.4862 | 26.6 | 0
2008-05-07 | 14:01 | 1.831 | 0.0066 | 65.7595 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3703 | 26.6 | 0
2008-05-07 | 14:01 | 1.839 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.7 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.344 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0153 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4987 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.015343 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3423 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2424 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0.0001 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2363 | 26.6 | 0
2008-05-07 | 14:01 | 1.84 | 0.006 | 60 | 91.5 | 0 | 91.5 | 27 | 27 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3 | 26.6 | 0
2008-05-07 | 14:01 | 1.832 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5689 | 30.3 | 26.6 | 0
2008-05-07 | 14:01 | 1.83 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.372 | 26.6 | 0
2008-05-07 | 14:01 | 1.83 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3494 | 26.6 | 0
2008-05-07 | 14:01 | 1.83 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31.2062 | 59 | 1023 | 985 | 30.3 | 26.6 | 0
2008-05-07 | 14:02 | 1.831 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3077 | 26.6 | 0
2008-05-07 | 14:02 | 1.839 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3859 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3566 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.322 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.4 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.4 | 26.6 | 0
2008-05-07 | 14:02 | 1.84 | 0.006 | 60 | 91.6 | 0 | 91.6 | 26.9 | 26.9 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4575 | 26.6 | 0
2008-05-07 | 14:02 | 1.832 | 0.006 | 60 | 91.6 | 0 | 91.6 | 27 | 26.9674 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4637 | 26.6 | 0
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | 0 | 91.7 | 26.9 | 26.9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4 | 26.6 | 0
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | 0 | 91.7 | 27 | 26.9954 | 0 | 12.8 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4 | 26.6 | 0
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 26.9 | 26.9013 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.472 | 26.6 | -0.1
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 27 | 26.9903 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 986 | 30.3987 | 26.6 | -0.1
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 27 | 26.9903 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 30.2708 | 26.6 | -0.1
2008-05-07 | 14:02 | 1.83 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 26.9 | 26.9 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2232 | 26.6 | -0.1
2008-05-07 | 14:02 | 1.854 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 26.9 | 26.9 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4578 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.091 | 0.006 | 60 | 91.7 | -0.1 | 91.7 | 26.9 | 26.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0154 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4353 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.402 | 0.0078 | 78.0602 | 103.7 | -0.1 | 103.7 | 26.9 | 26.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.017389 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3565 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 116.8 | -0.1 | 116.8 | 29.3 | 29.3124 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.01963 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3442 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 119.5 | -0.1 | 119.5 | 35.3 | 35.2657 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.020043 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4993 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.2 | -0.1 | 120.2 | 36.5 | 36.4629 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.020143 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4209 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.4 | -0.1 | 120.4 | 36.3 | 36.3079 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0202 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4578 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.4 | -0.1 | 120.4 | 36.2 | 36.1911 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0202 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4636 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.5 | -0.1 | 120.5 | 36.1 | 36.1 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0202 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3851 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.6 | -0.1432 | 120.6 | 36.1 | 36.1 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.020243 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3064 | 26.6 | -0.1
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.7 | -0.2 | 120.7 | 36.2 | 36.1783 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4442 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.7 | -0.2 | 120.7 | 36.2 | 36.1893 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190.3 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3987 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.7 | -0.2 | 120.7 | 36.1 | 36.1451 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 190 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3585 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.7 | -0.2 | 120.7 | 36.1 | 36.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 190 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.5157 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.0084 | 84.1922 | 120.8 | -0.2 | 120.8 | 36.1 | 36.1 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022.2401 | 985 | 30.6724 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 120.8 | -0.2 | 120.8 | 36.1 | 36.1 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.4309 | 30.6351 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 120.8 | -0.2 | 120.8 | 36.1 | 36.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.5689 | 30.5128 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 120.8 | -0.2432 | 120.8 | 36.1 | 36.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4 | 26.6 | -0.2
2008-05-07 | 14:02 | 2.446 | 0.008 | 80 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.4021 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.446 | 0.0095 | 95.1682 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 190.7 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.6367 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.0088 | 88 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5848 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5364 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.585 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.0081 | 80.6255 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.5064 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.0089 | 88.9658 | 120.9 | -0.3 | 120.9 | 36 | 36 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985.432 | 30.4278 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.8 | -0.3 | 120.8 | 36 | 36 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985.568 | 30.4 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.9 | -0.3 | 120.9 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 30.3707 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.9 | -0.3 | 120.9 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3158 | 26.6 | -0.3
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.9 | -0.4 | 120.9 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.4719 | 26.6 | -0.4
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 120.9 | -0.4 | 120.9 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0 |  | n/a |  | 31 | 59 | 1022.2113 | 985 | 30.5 | 26.6 | -0.4
2008-05-07 | 14:02 | 2.446 | 0.009 | 90 | 121 | -0.5 | 121 | 36 | 36 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 4.3e-05 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.4128 | 26.6 | -0.5
2008-05-07 | 14:02 | 2.446 | 0.009 | 90 | 121 | -0.5 | 121 | 36 | 36 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3222 | 26.6 | -0.5
2008-05-07 | 14:02 | 2.44 | 0.009 | 90 | 121 | -0.5 | 121 | 36 | 36 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4007 | 26.6 | -0.5
2008-05-07 | 14:02 | 2.441 | 0.009 | 90 | 121 | -0.5 | 121 | 36 | 36 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4794 | 26.6 | -0.5
2008-05-07 | 14:02 | 2.449 | 0.009 | 90 | 120.9 | -0.6 | 120.9 | 36 | 36 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.5 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.442 | 0.009 | 90 | 121 | -0.6 | 121 | 36 | 36 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3432 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4269 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.446 | 0.009 | 90 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.45 | 0.009 | 90 | 121 | -0.6 | 121 | 35.9 | 35.9076 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.45 | 0.009 | 90 | 121 | -0.6 | 121 | 36 | 35.9947 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.4443 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.449 | 0.009 | 90 | 121 | -0.6 | 121 | 36 | 36 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3985 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.441 | 0.009 | 90 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.44 | 0.0089 | 89.3838 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.44 | 0.0081 | 81.0575 | 121 | -0.6 | 121 | 35.9 | 35.9081 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.44 | 0.008 | 80 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0203 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:02 | 2.442 | 0.008 | 80 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.020343 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.3437 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.008 | 80 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.008 | 80 | 121 | -0.6 | 121 | 35.9 | 35.9 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985.568 | 30.3992 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.0081 | 80.5843 | 121.1 | -0.6 | 121.1 | 35.9 | 35.9 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3208 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.0089 | 88.9316 | 121.1 | -0.6 | 121.1 | 35.9 | 35.9 | 0 | 12.8 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4311 | 30.3 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.009 | 90 | 121.1 | -0.6 | 121.1 | 35.9 | 35.9 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.2269 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.009 | 90 | 121.1 | -0.6 | 121.1 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.4331 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.009 | 90 | 121.2 | -0.6 | 121.2 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.45 | 0.009 | 90 | 121.3 | -0.6 | 121.3 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.432 | 30.3888 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.442 | 0.009 | 90 | 121.3 | -0.5 | 121.3 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3476 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.44 | 0.009 | 90 | 121.4 | -0.5 | 121.4 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2294 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.44 | 0.009 | 90 | 121.4 | -0.5 | 121.4 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.308 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.44 | 0.009 | 90 | 121.4 | -0.6 | 121.4 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3862 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.44 | 0.0086 | 85.8217 | 121.4 | -0.6 | 121.4 | 35.9 | 35.9 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.335 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.44 | 0.008 | 80 | 121.5 | -0.6 | 121.5 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.3437 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.44 | 0.008 | 80 | 121.5 | -0.6 | 121.5 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.44 | 0.008 | 80 | 121.4 | -0.6 | 121.4 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.3432 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.44 | 0.008 | 80 | 121.5 | -0.6 | 121.5 | 35.9 | 35.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.4309 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.525 | 0.0086 | 85.9385 | 121.5 | -0.6 | 121.5 | 35.9 | 35.9 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0204 | -0.0001 | 0 | 4.3e-05 |  | n/a |  | 31.793 | 59 | 1022 | 985.432 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 2.81 | 0.0094 | 94.2407 | 123.5 | -0.5 | 123.5 | 35.9 | 35.9 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.020486 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022.1857 | 986 | 30.473 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.969 | 0.0098 | 97.5342 | 135.9 | -0.5 | 135.9 | 36.1 | 36.086 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.022532 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.5394 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.99 | 0.009 | 90 | 143.7 | -0.5432 | 143.7 | 39.4 | 39.4165 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.024089 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.2251 | 26.6 | -0.5
2008-05-07 | 14:03 | 2.996 | 0.009 | 90 | 149.1 | -0.6 | 149.1 | 43.3 | 43.293 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.024943 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2723 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.008 | 0.009 | 90 | 149.9 | -0.6 | 149.9 | 44.8 | 44.7692 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0251 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.4311 | 30.3508 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.02 | 0.009 | 90 | 151.1 | -0.6 | 151.1 | 45.1 | 45.1245 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025286 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.02 | 0.009 | 90 | 151.8 | -0.6 | 151.8 | 45.2 | 45.2 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.4329 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.02 | 0.009 | 90 | 151.8 | -0.6 | 151.8 | 45.2 | 45.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.027 | 0.009 | 90 | 151.9 | -0.6 | 151.9 | 45.2 | 45.2 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2703 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.009 | 90 | 152 | -0.6 | 152 | 45.2 | 45.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2874 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.0097 | 97.2594 | 152 | -0.6 | 152 | 45.2 | 45.2 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.3778 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.6 | 152 | 45.2 | 45.2 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3009 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.6 | 152 | 45.2 | 45.2 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.3568 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3794 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.6 | 152 | 45.2 | 45.2 | 0 | 12.8 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.4 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.6 | 152 | 45.2 | 45.1913 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3635 | 26.6 | -0.6
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.7 | 152 | 45.2 | 45.2 | 0 | 12.8 | 910 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 986 | 30.3 | 26.6 | -0.7
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152 | -0.7 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 986 | 30.3 | 26.6 | -0.7
2008-05-07 | 14:03 | 3.028 | 0.01 | 100 | 152 | -0.7 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 986 | 30.4445 | 26.6 | -0.7
2008-05-07 | 14:03 | 3.021 | 0.01 | 100 | 152 | -0.7 | 152 | 45.1 | 45.1 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.4492 | 26.6 | -0.7
2008-05-07 | 14:03 | 3.029 | 0.01 | 100 | 152.1 | -0.8 | 152.1 | 45.1 | 45.1 | 0 | 12.9 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.4 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152.1 | -0.8 | 152.1 | 45.1 | 45.1 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 30.3921 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152.1 | -0.8 | 152.1 | 45.1 | 45.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.3136 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.03 | 0.01 | 100 | 152.1 | -0.8 | 152.1 | 45 | 45.0058 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.3 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.032 | 0.01 | 100 | 152.1 | -0.8 | 152.1 | 45.1 | 45.1 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3874 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.8 | 152 | 45.1 | 45.0913 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.4555 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.84311 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3 | 26.6 | -0.8
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.3578 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.044 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3269 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.05 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 30.2302 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.049 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.8 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 30.3873 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.041 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 30.2555 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.2508 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.3294 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.7 | 912 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.3761 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.5691 | 30.1415 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 44.9 | 44.9076 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.4311 | 30.165 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45 | 45 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.0923 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2222 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022.3333 | 985 | 30.2977 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.0633 | 26.6 | -0.9
2008-05-07 | 14:03 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45.1 | 45.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1731 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45 | 45.0041 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 152 | -0.9 | 152 | 45 | 45 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.3 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.033 | 0.01 | 100 | 152 | -0.9 | 152 | 45 | 45.0036 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985.432 | 30.3 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 152 | -0.9 | 152 | 45 | 45 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1555 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 151.9 | -0.9 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 985.432 | 30.1508 | 26.6 | -0.9
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.1412 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1022 | 986 | 30.0237 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1022 | 986 | 30.2578 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.03 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1704 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.036 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45 | 45.0062 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.4343 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1445 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.8 | 151.9 | 45.1 | 45.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2984 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.75646 | 151.9 | 45 | 45 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1413 | 26.6 | -0.8
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.7 | 151.9 | 45 | 45 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0421 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.7 | 151.9 | 45 | 45 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.073 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.7 | 151.9 | 45 | 45 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2151 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.04 | 0.01 | 100 | 151.9 | -0.7 | 151.9 | 45 | 45 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022.155 | 985 | 30.2936 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.24 | 0.0105 | 105.416 | 151.9 | -0.7 | 151.9 | 45 | 45 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0254 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1556 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.889 | 0.0122 | 121.6021 | 152.3 | -0.7 | 152.3 | 45 | 45 | 0 | 12.7 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025486 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2015 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.881 | 0.0135 | 135.2941 | 179.3 | -0.7 | 179.3 | 45 | 45 | 0 | 12.8 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.029834 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2414 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.706 | 0.0121 | 120.5489 | 191.2 | -0.7 | 191.2 | 52.6 | 52.6281 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0317 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.64 | 0.0129 | 129.1252 | 189.5 | -0.7 | 189.5 | 59.7 | 59.6795 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.03144 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.64 | 0.013 | 130 | 185.5 | -0.7 | 185.5 | 56.9 | 56.9162 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0308 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.638 | 0.0125 | 124.5878 | 183.1 | -0.7 | 183.1 | 55.2 | 55.177 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030457 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 987 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 182.8 | -0.7 | 182.8 | 54.6 | 54.607 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0304 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986.432 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 182.3 | -0.7 | 182.3 | 54.5 | 54.5 | 0 | 12.9 | 909 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030357 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 987 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.8 | -0.7 | 181.8 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 987 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.5 | 54.5 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 31.7932 | 59 | 1023 | 986 | 30.1727 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 31.2062 | 59 | 1023 | 986 | 30.2851 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2064 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1279 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4 | 0 | 12.8 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.0493 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.7 | -0.7 | 181.7 | 54.4 | 54.4149 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0292 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.6 | -0.7 | 181.6 | 54.5 | 54.5 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.0922 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.6 | -0.7 | 181.6 | 54.5 | 54.5 | 0 | 12.7 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0139 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.6 | -0.7 | 181.6 | 54.5 | 54.5 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1296 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.6 | -0.7 | 181.6 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6568 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.178 | 26.6 | -0.7
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.2 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1011 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.5 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2577 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195.9 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1273 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195.8 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0726 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 181.5 | -0.6 | 181.5 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.5 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1852 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.0121 | 120.5242 | 181.4 | -0.6 | 181.4 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195.1 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1066 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.0129 | 128.7142 | 181.4 | -0.6 | 181.4 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.013 | 130 | 181.4 | -0.6 | 181.4 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1505 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.013 | 130 | 181.4 | -0.6 | 181.4 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0303 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1709 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.0126 | 126.0288 | 181.1 | -0.6 | 181.1 | 54.5 | 54.5 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.0353 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.0434 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.7 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1005 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.9 | -0.6 | 180.9 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1786 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.7 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 5.7e-05 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1426 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.7 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:04 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.2438 | 26.6 | -0.6
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.8 | -0.6 | 180.8 | 54.5 | 54.5 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1991 | 26.6 | -0.6
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.7 | -0.64311 | 180.7 | 54.5 | 54.5 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.4311 | 30.1581 | 26.6 | -0.6
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.7 | -0.7 | 180.7 | 54.5 | 54.5 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2848 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.6 | -0.7 | 180.6 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1282 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.6 | -0.7 | 180.6 | 54.5 | 54.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.63 | 0.012 | 120 | 180.6 | -0.7 | 180.6 | 54.5 | 54.5 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.631 | 0.012 | 120 | 180.6 | -0.7 | 180.6 | 54.5 | 54.5 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1218 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.639 | 0.012 | 120 | 180.6 | -0.7 | 180.6 | 54.5 | 54.5 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1996 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.0119 | 119.4622 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1217 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.0111 | 111.0588 | 180.5 | -0.7 | 180.5 | 54.4 | 54.4 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1723 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2705 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1134 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.64 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4148 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 3.719 | 0.011 | 110 | 180.6 | -0.7 | 180.6 | 54.4 | 54.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0302 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 4.063 | 0.011 | 110 | 183.3 | -0.7 | 183.3 | 54.5 | 54.4947 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.030646 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1647 | 26.6 | -0.7
2008-05-07 | 14:05 | 4.188 | 0.011 | 110 | 194.7 | -0.7 | 194.7 | 55.5 | 55.4513 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.032445 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1568 | 26.6 | -0.7
2008-05-07 | 14:05 | 4.21 | 0.0111 | 111.0339 | 204.7 | -0.7 | 204.7 | 59.7 | 59.7224 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.033987 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:05 | 4.21 | 0.0127 | 127.189 | 206.2 | -0.6 | 206.2 | 62.6 | 62.6474 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.034243 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1004 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.214 | 0.0137 | 136.8891 | 206.7 | -0.6 | 206.7 | 63.6 | 63.582 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0343 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1788 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 207.5 | -0.6 | 207.5 | 63.7 | 63.7 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.034443 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1426 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 208.2 | -0.6 | 208.2 | 63.8 | 63.7936 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.034543 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 208.5 | -0.6 | 208.5 | 63.8 | 63.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0346 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0854 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 208.6 | -0.6 | 208.6 | 63.8 | 63.8 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0346 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0068 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 208.6 | -0.6 | 208.6 | 63.8 | 63.8 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0346 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1431 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.0133 | 133.161 | 208.7 | -0.6 | 208.7 | 63.8 | 63.8 | 0 | 12.7 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0346 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0997 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.013 | 130 | 208.8 | -0.6 | 208.8 | 63.8 | 63.8 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0346 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0866 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.013 | 130 | 209.2 | -0.6 | 209.2 | 63.8 | 63.8 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.285 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.0132 | 131.9171 | 209.2 | -0.6 | 209.2 | 63.8 | 63.8 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1282 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.3 | -0.6 | 209.3 | 63.8 | 63.8 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.3 | -0.6 | 209.3 | 63.9 | 63.8919 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.3 | -0.6 | 209.3 | 63.9 | 63.8917 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.3 | -0.6 | 209.3 | 63.9 | 63.8933 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.4 | -0.6 | 209.4 | 63.8 | 63.8077 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.0142 | 142.2065 | 209.4 | -0.6 | 209.4 | 63.9 | 63.9 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.0149 | 149.4855 | 209.4 | -0.6 | 209.4 | 63.9 | 63.9 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.0141 | 141.4469 | 209.4 | -0.6 | 209.4 | 63.9 | 63.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.3 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.5 | -0.6 | 209.5 | 63.8 | 63.8075 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.5 | -0.6 | 209.5 | 63.8 | 63.821 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 190.7 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0287 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.5 | -0.5568 | 209.5 | 63.9 | 63.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.05 | 26.6 | -0.6
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.5 | -0.5 | 209.5 | 63.9 | 63.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.0347 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1577 | 26.6 | -0.5
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.7 | -0.5 | 209.7 | 63.8 | 63.8038 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2852 | 26.6 | -0.5
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.7 | -0.5 | 209.7 | 63.8 | 63.8464 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1283 | 26.6 | -0.5
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.8 | -0.4 | 209.8 | 63.8 | 63.8067 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.4
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.8 | -0.3 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -0.3
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.8 | -0.3 | 209.8 | 63.9 | 63.8909 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0784 | 26.6 | -0.3
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.8 | -0.3 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 913 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0003 | 26.6 | -0.3
2008-05-07 | 14:05 | 4.22 | 0.014 | 140 | 209.8 | -0.2568 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 913 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1566 | 26.6 | -0.3
2008-05-07 | 14:05 | 4.215 | 0.014 | 140 | 209.8 | -0.2 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 913 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1429 | 26.6 | -0.2
2008-05-07 | 14:05 | 4.214 | 0.014 | 140 | 209.8 | -0.2 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9922 | 26.6 | -0.2
2008-05-07 | 14:05 | 4.218 | 0.014 | 140 | 209.8 | -0.2 | 209.8 | 63.8 | 63.8 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8581 | 26.6 | -0.2
2008-05-07 | 14:05 | 4.21 | 0.014 | 140 | 209.9 | -0.2 | 209.9 | 63.8 | 63.8 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1723 | 26.6 | -0.2
2008-05-07 | 14:05 | 4.21 | 0.014 | 140 | 210 | -0.2 | 210 | 63.8 | 63.8 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1289 | 26.6 | -0.2
2008-05-07 | 14:05 | 4.217 | 0.014 | 140 | 210 | -0.1 | 210 | 63.8 | 63.8 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:05 | 4.215 | 0.014 | 140 | 210 | -0.1 | 210 | 63.8 | 63.8 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3568 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:05 | 4.21 | 0.014 | 140 | 210.1 | -0.1 | 210.1 | 63.8 | 63.8 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:05 | 4.21 | 0.014 | 140 | 210.1 | -0.1 | 210.1 | 63.8 | 63.8 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.21 | 0.014 | 140 | 210.1 | -0.1 | 210.1 | 63.8 | 63.8 | 0 | 12.6 | 912 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.21 | 0.014 | 140 | 210.2 | -0.1 | 210.2 | 63.8 | 63.8 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.057 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.21 | 0.014 | 140 | 210.2 | -0.1 | 210.2 | 63.8 | 63.8 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0648 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.21 | 0.014 | 140 | 210.1 | -0.1 | 210.1 | 63.8 | 63.8 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2997 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.213 | 0.014 | 140 | 210.2 | -0.1 | 210.2 | 63.8 | 63.8 | 0 | 12.7 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1437 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.22 | 0.0135 | 134.6624 | 210.2 | -0.1 | 210.2 | 63.8 | 63.8 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4431 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.22 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0348 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.22 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.213 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0286 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.2 | -0.2 | 210.2 | 63.8 | 63.8 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0287 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.3 | -0.1432 | 210.3 | 63.8 | 63.8 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.3 | -0.2 | 210.3 | 63.8 | 63.8 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.4 | -0.1568 | 210.4 | 63.8 | 63.8 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0357 | 26.6 | -0.2
2008-05-07 | 14:06 | 4.21 | 0.013 | 130 | 210.5 | -0.1 | 210.5 | 63.8 | 63.8 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0429 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.437 | 0.0137 | 136.8276 | 210.5 | -0.1 | 210.5 | 63.8 | 63.8 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0349 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0786 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.858 | 0.0161 | 160.5819 | 215.8 | -0.1 | 215.8 | 63.8 | 63.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.035116 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.849 | 0.0177 | 176.5217 | 232.1 | -0.1 | 232.1 | 63.8 | 63.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.038148 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.841 | 0.017 | 170 | 238.8 | -0.1 | 238.8 | 71.1 | 71.0583 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0392 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0571 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.84 | 0.017 | 170 | 239.1 | -0.1 | 239.1 | 73.8 | 73.843 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0392 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.835 | 0.017 | 170 | 239.1 | -0.056886 | 239.1 | 73.4 | 73.4188 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0392 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4311 | 30.0857 | 26.6 | -0.1
2008-05-07 | 14:06 | 4.83 | 0.017 | 170 | 239.2 | 0 | 239.2 | 73.3 | 73.3 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0071 | 26.6 | 0
2008-05-07 | 14:06 | 4.826 | 0.017 | 170 | 239.2 | 0.1 | 239.3 | 73.3 | 73.3 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0712 | 26.6 | 0.1
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.2 | 0.14311 | 239.4 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 30.1 | 26.6 | 0.1
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.3 | 0.2 | 239.5 | 73.2 | 73.2 | 0 | 12.7 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4309 | 30.0716 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.3 | 0.2 | 239.5 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.3 | 0.2 | 239.5 | 73.2 | 73.2 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.4 | 0.2 | 239.6 | 73.2 | 73.2 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.4 | 0.2 | 239.6 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5663 | 30.0427 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.5 | 0.2 | 239.7 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.81 | 0.017 | 170 | 239.5 | 0.2 | 239.7 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.803 | 0.017 | 169.6188 | 239.5 | 0.2 | 239.7 | 73.2 | 73.2 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0217 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.0162 | 161.5085 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9431 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9356 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.3 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.6 | 0.2 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4331 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.7 | 0.2 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.7 | 0.2 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.7 | 0.2 | 239.9 | 73.2 | 73.2071 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 190 | 6 | 195 | 59.4 | 0.0393 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.8 | 0.016 | 160 | 239.8 | 0.2 | 240 | 73.3 | 73.2941 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.039343 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4309 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.791 | 0.016 | 160 | 239.8 | 0.2 | 240 | 73.3 | 73.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.79 | 0.016 | 160 | 239.8 | 0.2 | 240 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0212 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.79 | 0.016 | 160 | 239.8 | 0.2 | 240 | 73.2 | 73.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 190 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0997 | 26.6 | 0.2
2008-05-07 | 14:06 | 4.79 | 0.016 | 160 | 239.8 | 0.1 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0219 | 26.6 | 0.1
2008-05-07 | 14:06 | 4.79 | 0.016 | 160 | 239.8 | 0.1 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0.1
2008-05-07 | 14:06 | 4.79 | 0.016 | 160 | 239.8 | 0 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.798 | 0.016 | 160 | 239.8 | 0 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.8 | 0.016 | 159.5432 | 239.8 | 0 | 239.8 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.8 | 0.0151 | 151.2375 | 239.9 | 0 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.8 | 0.015 | 150 | 239.9 | 0 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.8 | 0.015 | 150 | 239.9 | 0 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0001 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.799 | 0.015 | 150 | 239.9 | 0 | 239.9 | 73.2 | 73.2 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9865 | 26.6 | 0
2008-05-07 | 14:06 | 4.783 | 0.015 | 150 | 239.9 | 0 | 239.9 | 73.2 | 73.2 | 0 | 12.7 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4431 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8292 | 26.6 | 0
2008-05-07 | 14:06 | 4.787 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9279 | 26.6 | 0
2008-05-07 | 14:06 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:06 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240.1 | 0 | 240.1 | 73.2 | 73.2 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.1 | 73.1061 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9433 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4346 | 29.9353 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4432 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9861 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9075 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9711 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 4.79 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.0068 | 26.6 | 0
2008-05-07 | 14:07 | 4.8 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0853 | 26.6 | 0
2008-05-07 | 14:07 | 4.81 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | 0
2008-05-07 | 14:07 | 4.808 | 0.015 | 150 | 240 | 0 | 240 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5691 | 30.0162 | 26.6 | 0
2008-05-07 | 14:07 | 4.8 | 0.015 | 150 | 240.1 | 0 | 240.1 | 73.2 | 73.2 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.039443 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8789 | 26.6 | 0
2008-05-07 | 14:07 | 4.8 | 0.015 | 150 | 240.2 | 0 | 240.2 | 73.2 | 73.2 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8004 | 26.6 | 0
2008-05-07 | 14:07 | 4.8 | 0.015 | 150 | 240.3 | 0 | 240.3 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9563 | 26.6 | 0
2008-05-07 | 14:07 | 4.8 | 0.015 | 150 | 240.3 | 0 | 240.3 | 73.2 | 73.2 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.084 | 0.015 | 150 | 240.3 | 0 | 240.3 | 73.2 | 73.2 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9294 | 26.6 | 0
2008-05-07 | 14:07 | 5.582 | 0.018 | 179.9501 | 244.1 | 0 | 244.1 | 73.2 | 73.2 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.040062 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.8139 | 26.6 | 0
2008-05-07 | 14:07 | 5.587 | 0.02 | 199.907 | 269.9 | 0 | 269.9 | 73.4 | 73.3914 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.043134 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.8925 | 26.6 | 0
2008-05-07 | 14:07 | 5.473 | 0.018 | 180 | 282.1 | 0 | 282.1 | 83.1 | 83.1252 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0457 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | 0
2008-05-07 | 14:07 | 5.443 | 0.018 | 180 | 279.1 | 0 | 279.1 | 88 | 88.0247 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.045198 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9495 | 26.6 | 0
2008-05-07 | 14:07 | 5.43 | 0.018 | 180 | 275.4 | 0 | 275.4 | 84.7 | 84.6878 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.044657 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9436 | 26.6 | 0
2008-05-07 | 14:07 | 5.42 | 0.018 | 180 | 274.3 | 0 | 274.3 | 83.1 | 83.1343 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.044457 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8137 | 26.6 | 0
2008-05-07 | 14:07 | 5.418 | 0.018 | 180 | 273.8 | 0 | 273.8 | 82.8 | 82.8086 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0444 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.9708 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.018 | 180 | 273.1 | 0 | 273.1 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.3568 | 0.0443 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.872 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.018 | 180 | 273 | 0 | 273 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8426 | 26.6 | 0
2008-05-07 | 14:07 | 5.417 | 0.018 | 180 | 272.6 | 0 | 272.6 | 82.7 | 82.7 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.044257 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9211 | 26.6 | 0
2008-05-07 | 14:07 | 5.42 | 0.018 | 180 | 272.1 | 0 | 272.1 | 82.7 | 82.7 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9997 | 26.6 | 0
2008-05-07 | 14:07 | 5.417 | 0.0173 | 173.3389 | 272 | 0 | 272 | 82.7 | 82.7 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 29.8439 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.017 | 170 | 272 | 0 | 272 | 82.6 | 82.6063 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.017 | 170 | 271.9 | 0 | 271.9 | 82.7 | 82.7 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.871 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.017 | 170 | 271.7 | 0 | 271.7 | 82.6 | 82.6 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.41 | 0.017 | 170 | 271.7 | 0 | 271.7 | 82.7 | 82.6909 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.405 | 0.017 | 170 | 271.7 | 0 | 271.7 | 82.6 | 82.6068 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.4 | 0.0177 | 176.6472 | 271.6 | 0 | 271.6 | 82.7 | 82.6939 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.4 | 0.018 | 180 | 271.6 | 0 | 271.6 | 82.6 | 82.6072 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | 0
2008-05-07 | 14:07 | 5.4 | 0.018 | 180 | 271.6 | 0 | 271.6 | 82.6 | 82.6 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9932 | 26.6 | 0
2008-05-07 | 14:07 | 5.4 | 0.018 | 180 | 271.5 | -0.1 | 271.5 | 82.6 | 82.6042 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9145 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.373 | 0.018 | 180 | 271.5 | -0.1 | 271.5 | 82.6 | 82.6 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.37 | 0.018 | 180 | 271.5 | -0.1 | 271.5 | 82.7 | 82.6989 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9426 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.38 | 0.018 | 180 | 271.5 | -0.1 | 271.5 | 82.6 | 82.6 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.38 | 0.018 | 180 | 271.5 | -0.1 | 271.5 | 82.7 | 82.6915 | 0 | 12.6 | 912 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.38 | 0.018 | 180 | 271.4 | -0.1 | 271.4 | 82.7 | 82.7 | 0 | 12.7 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.38 | 0.0178 | 178.1712 | 271.4 | -0.1 | 271.4 | 82.7 | 82.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.38 | 0.017 | 170 | 271.3 | -0.1 | 271.3 | 82.7 | 82.7 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.385 | 0.017 | 170 | 271.2 | -0.1 | 271.2 | 82.7 | 82.6933 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.9719 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 271.2 | -0.1 | 271.2 | 82.7 | 82.7 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8148 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 271.1 | -0.1 | 271.1 | 82.7 | 82.6941 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9423 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 271.1 | -0.1 | 271.1 | 82.7 | 82.7 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9005 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 271 | -0.1 | 271 | 82.7 | 82.7 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8284 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.9 | -0.1 | 270.9 | 82.7 | 82.7 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.9 | -0.1 | 270.9 | 82.7 | 82.7 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.9 | -0.1 | 270.9 | 82.7 | 82.7 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.9 | -0.1 | 270.9 | 82.7 | 82.7 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.8 | -0.1 | 270.8 | 82.7 | 82.7 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0441 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9214 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.8 | -0.1 | 270.8 | 82.6 | 82.605 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9999 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.8 | -0.1 | 270.8 | 82.7 | 82.6917 | 0 | 12.9 | 909 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9218 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.8 | -0.1 | 270.8 | 82.7 | 82.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8431 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.017 | 170 | 270.8 | -0.1 | 270.8 | 82.7 | 82.7 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -0.1
2008-05-07 | 14:07 | 5.39 | 0.0175 | 175.0426 | 270.8 | -0.1 | 270.8 | 82.7 | 82.7 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8284 | 26.6 | -0.1
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.8 | -0.1 | 270.8 | 82.7 | 82.7 | 0 | 12.6 | 912 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 29.9855 | 26.6 | -0.1
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.8 | -0.1432 | 270.8 | 82.7 | 82.7 | 0 | 12.6 | 913 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.858 | 26.6 | -0.1
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.8 | -0.2 | 270.8 | 82.7 | 82.7 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8994 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.8 | -0.2 | 270.8 | 82.7 | 82.7 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 29.9714 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.2 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.2 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.2 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9643 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.2 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9142 | 26.6 | -0.2
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.3 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.843 | 26.6 | -0.3
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.3 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 29.9999 | 26.6 | -0.3
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.3432 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9213 | 26.6 | -0.3
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.4 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.0146 | 26.6 | -0.4
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.4 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.0641 | 26.6 | -0.4
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.4432 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.3 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9711 | 26.6 | -0.4
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.5 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8138 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.5 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9429 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.39 | 0.018 | 180 | 270.7 | -0.5 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5689 | 29.8998 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.428 | 0.0185 | 185.0642 | 270.7 | -0.5 | 270.7 | 82.7 | 82.7 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.044 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.605 | 0.0187 | 186.6953 | 271.3 | -0.5 | 271.3 | 82.7 | 82.6936 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.044086 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.4309 | 29.8074 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.853 | 0.0178 | 178.1997 | 279.7 | -0.5 | 279.7 | 82.9 | 82.8799 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.045317 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8859 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.913 | 0.017 | 170.0164 | 290.2 | -0.5 | 290.2 | 85.1 | 85.1024 | 0 | 12.7 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.046816 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9645 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.95 | 0.017 | 170 | 297.8 | -0.5 | 297.8 | 89.3 | 89.2549 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.047943 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9137 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.98 | 0.017 | 170 | 299 | -0.5 | 299 | 91.4 | 91.3513 | 0 | 12.7 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.982 | 0.017 | 170 | 299 | -0.5 | 299 | 91.9 | 91.9017 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 5.99 | 0.0173 | 173.2733 | 299.1 | -0.5 | 299.1 | 92 | 92 | 0 | 12.7 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 6.007 | 0.018 | 180 | 299.2 | -0.5 | 299.2 | 92 | 92 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9147 | 26.6 | -0.5
2008-05-07 | 14:08 | 6.01 | 0.018 | 180 | 299.3 | -0.5 | 299.3 | 92 | 92 | 0 | 12.6 | 913 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9277 | 26.6 | -0.5
2008-05-07 | 14:08 | 6.005 | 0.018 | 180 | 299.3 | -0.5 | 299.3 | 92 | 92 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8147 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.018 | 180 | 299.3 | -0.5 | 299.3 | 92.1 | 92.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 29.8933 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.018 | 180 | 299.3 | -0.5 | 299.3 | 92.1 | 92.1 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0481 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9719 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.018 | 180 | 299.3 | -0.5 | 299.3 | 92 | 92.004 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8991 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.018 | 180 | 299.4 | -0.5 | 299.4 | 92.1 | 92.1 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.018 | 180.1463 | 299.4 | -0.5 | 299.4 | 92.1 | 92.1 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.0188 | 188.2764 | 299.5 | -0.5 | 299.5 | 92.1 | 92.1 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8 | 26.6 | -0.5
2008-05-07 | 14:08 | 6 | 0.019 | 190 | 299.6 | -0.6 | 299.6 | 92 | 92 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.998 | 0.019 | 190 | 299.6 | -0.6 | 299.6 | 92 | 92 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.8869 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 299.6 | -0.6 | 299.6 | 92 | 92 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9559 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.982 | 0.019 | 190 | 299.7 | -0.6 | 299.7 | 92 | 92 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5691 | 29.8 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.98 | 0.019 | 190 | 299.8 | -0.6 | 299.8 | 92 | 92 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.8 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.98 | 0.019 | 190 | 299.9 | -0.6 | 299.9 | 92 | 92 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.9154 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.983 | 0.019 | 190 | 299.9 | -0.6 | 299.9 | 92 | 92.0089 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.9636 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.6 | 300 | 92 | 92 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.915 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.6432 | 300 | 92 | 92 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.9936 | 26.6 | -0.6
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985.432 | 29.9278 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.8492 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985 | 29.8588 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0482 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.048243 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 29.935 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 29.9437 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.1 | -0.7 | 300.1 | 92 | 92 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1024 | 986 | 29.9778 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300 | -0.7 | 300 | 92 | 92 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023.4896 | 986 | 29.9 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.1 | -0.7 | 300.1 | 92 | 92 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.9 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.1 | -0.7 | 300.1 | 92 | 92 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5691 | 29.9577 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.2 | -0.7 | 300.2 | 92 | 92 | 0 | 12.7 | 912 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.9634 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.998 | 0.019 | 190 | 300.2 | -0.7 | 300.2 | 92 | 92 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 29.9152 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.994 | 0.019 | 190 | 300.3 | -0.7 | 300.3 | 92 | 92 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9937 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0723 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9981 | 26.6 | -0.7
2008-05-07 | 14:08 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9295 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 190.3 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0162 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 91.9938 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1733 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0695 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.995 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.7 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9848 | 26.6 | -0.7
2008-05-07 | 14:09 | 6 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9063 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.998 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0447 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92 | 92 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 190 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.998 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.3 | -0.7 | 300.3 | 92.1 | 92.0922 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9296 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92.1 | 92.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0081 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.3 | -0.7 | 300.3 | 92.1 | 92.1 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0864 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.99 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92.1 | 92.1 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.979 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92.1 | 92.1 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.924 | 0.019 | 190 | 300.4 | -0.7 | 300.4 | 92.1 | 92.1032 | 0 | 12.6 | 911 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4433 | 0.0483 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0551 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.525 | 0.0192 | 191.7902 | 296.1 | -0.7 | 296.1 | 92.2 | 92.2 | 0 | 12.6 | 910 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.048084 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.901 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.44 | 0.02 | 200 | 283.6 | -0.7432 | 283.6 | 91.8 | 91.7571 | 0 | 12.6 | 910 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.046141 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9791 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.416 | 0.02 | 200 | 278 | -0.8 | 278 | 87.2 | 87.1694 | 0 | 12.7 | 910 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.045084 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1159 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.02 | 200 | 274.7 | -0.8 | 274.7 | 84.3 | 84.3324 | 0 | 12.6 | 911 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.044657 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1632 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.02 | 200 | 273.9 | -0.8 | 273.9 | 83.4 | 83.4117 | 0 | 12.7 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0445 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0846 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.0194 | 193.634 | 273.2 | -0.8 | 273.2 | 83.1 | 83.1055 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0444 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0063 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 272.9 | -0.8 | 272.9 | 83 | 83 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0444 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1449 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.018 | 179.8617 | 272.7 | -0.8 | 272.7 | 83 | 83 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.044357 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1489 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.0172 | 171.725 | 272.6 | -0.8 | 272.6 | 82.9 | 82.9 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 272.4 | -0.8 | 272.4 | 82.9 | 82.9 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 272.4 | -0.8 | 272.4 | 82.9 | 82.9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 272.3 | -0.8 | 272.3 | 82.9 | 82.9 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 272.3 | -0.8 | 272.3 | 82.9 | 82.9 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.384 | 0.017 | 170 | 272.2 | -0.8 | 272.2 | 82.9 | 82.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0443 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.017 | 170 | 272 | -0.8 | 272 | 82.9 | 82.9 | 0 | 12.8 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.044257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.017 | 170 | 271.9 | -0.8 | 271.9 | 82.9 | 82.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.017 | 170 | 271.9 | -0.8 | 271.9 | 83 | 82.9909 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0417 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.017 | 170 | 271.9 | -0.8 | 271.9 | 83 | 83 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0368 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.388 | 0.017 | 170 | 271.8 | -0.8 | 271.8 | 83.1 | 83.0934 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 271.8 | -0.8 | 271.8 | 83.1 | 83.1 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.017 | 170 | 271.9 | -0.75689 | 271.9 | 83.2 | 83.2 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.0176 | 176.4432 | 271.8 | -0.7 | 271.8 | 83.2 | 83.205 | 0 | 12.7 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0489 | 26.6 | -0.7
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.8 | -0.8 | 271.8 | 83.3 | 83.3 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.8 | -0.8 | 271.8 | 83.4 | 83.3919 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0165 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.7 | -0.8 | 271.7 | 83.4 | 83.4 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1733 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.383 | 0.018 | 180 | 271.7 | -0.8 | 271.7 | 83.4 | 83.4 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1346 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.018 | 180 | 271.7 | -0.8 | 271.7 | 83.4 | 83.4 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.056 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.018 | 180 | 271.7 | -0.8 | 271.7 | 83.4 | 83.409 | 0 | 12.6 | 911 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.0175 | 175.3078 | 271.7 | -0.8 | 271.7 | 83.5 | 83.4953 | 0 | 12.7 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.001 | 26.6 | -0.8
2008-05-07 | 14:09 | 5.38 | 0.0173 | 173.0167 | 271.7 | -0.9 | 271.7 | 83.4 | 83.4 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0797 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.38 | 0.018 | 180 | 271.7 | -0.9 | 271.7 | 83.4 | 83.4 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.38 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.4 | 83.4 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.385 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.4 | 83.4 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.3 | 83.3056 | 0 | 12.6 | 912 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.5 | -0.9 | 271.5 | 83.3 | 83.3 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0275 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.5 | -0.9 | 271.5 | 83.3 | 83.2774 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.382 | 0.018 | 180 | 271.5 | -0.9 | 271.5 | 83.2 | 83.2 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 986 | 30.0298 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.386 | 0.018 | 180 | 271.5 | -0.9 | 271.5 | 83.2 | 83.2 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 986 | 30.1166 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985.4309 | 30.2728 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1695 | 26.6 | -0.9
2008-05-07 | 14:09 | 5.39 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.056 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.018 | 180 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.0451 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.0178 | 178.4251 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1989 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83 | 83.0037 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022.9887 | 985 | 30.1203 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.387 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83.1 | 83.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 30.1369 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.381 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83.2 | 83.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83.2 | 83.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.6 | -0.9 | 271.6 | 83.3 | 83.2935 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.055 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.5 | -0.9 | 271.5 | 83.3 | 83.3 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0512 | 26.6 | -0.9
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.4 | -1 | 271.4 | 83.4 | 83.4 | 0 | 12.8 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 5.39 | 0.017 | 170 | 271.4 | -1 | 271.4 | 83.5 | 83.5 | 0 | 12.8 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 5.389 | 0.017 | 170 | 271.4 | -1 | 271.4 | 83.6 | 83.5924 | 0 | 12.9 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 5.38 | 0.017 | 170 | 271.3 | -1 | 271.3 | 83.7 | 83.6905 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.1654 | 26.6 | -1
2008-05-07 | 14:10 | 5.38 | 0.017 | 170 | 271.2 | -1 | 271.2 | 83.7 | 83.7035 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985.5691 | 30.156 | 26.6 | -1
2008-05-07 | 14:10 | 5.38 | 0.017 | 170 | 271.3 | -1 | 271.3 | 83.8 | 83.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 985.4311 | 30.1226 | 26.6 | -1
2008-05-07 | 14:10 | 5.385 | 0.017 | 170 | 271.2 | -1 | 271.2 | 83.8 | 83.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.2 | 26.6 | -1
2008-05-07 | 14:10 | 5.1 | 0.0164 | 163.5621 | 271.1 | -1.1 | 271.1 | 83.9 | 83.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0442 | 0.0003 | -4.3e-05 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.2 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.447 | 0.0169 | 169.3223 | 267.7 | -1.1 | 267.7 | 83.9 | 83.9 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.043683 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1416 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.688 | 0.0177 | 177.0549 | 245.6 | -1.1 | 245.6 | 82.7 | 82.6663 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.040409 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.764 | 0.0169 | 168.7086 | 232.2 | -1.1 | 232.2 | 72.3 | 72.2606 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0384 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.78 | 0.016 | 160.2124 | 236 | -1.1 | 236 | 71.1 | 71.1022 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.038986 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.78 | 0.016 | 160 | 238.6 | -1.1 | 238.6 | 73.6 | 73.6331 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.0274 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.78 | 0.016 | 160 | 239 | -1.1 | 239 | 74.4 | 74.3853 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0394 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.0512 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.78 | 0.016 | 160 | 239.4 | -1.2 | 239.4 | 74.6 | 74.5883 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.78 | 0.016 | 160 | 239.7 | -1.2 | 239.7 | 74.6 | 74.6 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.788 | 0.016 | 160 | 239.8 | -1.2 | 239.8 | 74.6 | 74.6 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0395 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.79 | 0.016 | 160 | 239.9 | -1.2 | 239.9 | 74.6 | 74.6 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023.1204 | 986 | 30.1654 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.799 | 0.016 | 160 | 240 | -1.2 | 240 | 74.6 | 74.6 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1559 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.1 | -1.2 | 240.1 | 74.6 | 74.6 | 0 | 12.8 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0773 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.2 | -1.2 | 240.2 | 74.6 | 74.6 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0011 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.2 | -1.2 | 240.2 | 74.6 | 74.6 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.0798 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.3 | -1.2 | 240.3 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.3 | -1.2 | 240.3 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.4 | -1.2 | 240.4 | 74.6 | 74.5949 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.4 | -1.2 | 240.4 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -0.0001 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.811 | 0.016 | 160 | 240.4 | -1.2 | 240.4 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | -5.7e-05 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.82 | 0.016 | 160 | 240.4 | -1.1568 | 240.4 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -1.2
2008-05-07 | 14:10 | 4.812 | 0.016 | 160 | 240.4 | -1.1 | 240.4 | 74.5 | 74.5 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.4354 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.3 | -1.1 | 240.3 | 74.5 | 74.4764 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.2 | -1.1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:10 | 4.81 | 0.016 | 160 | 240.1 | -1 | 240.1 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 4.809 | 0.016 | 160 | 240.1 | -1 | 240.1 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 4.8 | 0.016 | 160 | 240.1 | -1 | 240.1 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.3 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1227 | 26.6 | -1
2008-05-07 | 14:10 | 4.8 | 0.016 | 160 | 240 | -1 | 240 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1987 | 26.6 | -1
2008-05-07 | 14:10 | 4.8 | 0.016 | 160 | 240 | -1 | 240 | 74.4 | 74.4 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 190.7 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1201 | 26.6 | -1
2008-05-07 | 14:10 | 4.795 | 0.016 | 160 | 240.1 | -1 | 240.1 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 4.793 | 0.016 | 160 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5689 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 4.8 | 0.016 | 160 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1156 | 26.6 | -1
2008-05-07 | 14:10 | 4.8 | 0.016 | 160 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5646 | 30.1942 | 26.6 | -1
2008-05-07 | 14:10 | 4.792 | 0.016 | 160 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1273 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.016 | 160 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.016 | 160 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1299 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.016 | 160 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1915 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.0165 | 164.623 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1131 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.2312 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.2115 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5691 | 30.0771 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0014 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.4 | 74.4 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.08 | 26.6 | -1
2008-05-07 | 14:10 | 4.79 | 0.017 | 170 | 240.2 | -1 | 240.2 | 74.4 | 74.4 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5689 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1728 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.017 | 170 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1485 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.0165 | 165.4917 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.3 | -1 | 240.3 | 74.5 | 74.5 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.95691 | 240.2 | 74.5 | 74.5 | 0 | 12.8 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.9 | 240.2 | 74.5 | 74.5 | 0 | 12.9 | 910 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1655 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.3 | -0.9 | 240.3 | 74.5 | 74.5 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1557 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.9 | 240.2 | 74.6 | 74.5894 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.9 | 240.2 | 74.6 | 74.6 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0986 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.9 | 240.2 | 74.6 | 74.6 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.02 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.2 | -0.9 | 240.2 | 74.6 | 74.6 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.3568 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1173 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.79 | 0.016 | 160 | 240.3 | -0.9 | 240.3 | 74.6 | 74.5915 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1628 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.524 | 0.0152 | 151.7985 | 240.3 | -0.9 | 240.3 | 74.6 | 74.5891 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0396 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.19 | 0.0144 | 143.5738 | 231.4 | -0.9 | 231.4 | 74.5 | 74.5039 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.038298 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.19 | 0.014 | 140 | 219.8 | -0.9 | 219.8 | 73.7 | 73.6681 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0365 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.19 | 0.014 | 140 | 215 | -0.9 | 215 | 66.4 | 66.4244 | 0 | 12.9 | 909 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4568 | 0.035814 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1024 | 986 | 30.1515 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.198 | 0.014 | 140 | 213.2 | -0.9 | 213.2 | 65.3 | 65.3 | 0 | 12.9 | 908 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.035557 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023.6217 | 986 | 30.2 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.2 | 0.014 | 140 | 212.7 | -0.9 | 212.7 | 65.3 | 65.3 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.035457 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5691 | 30.1914 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.2 | 0.014 | 140 | 212.6 | -0.9 | 212.6 | 65.3 | 65.3 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1128 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.203 | 0.014 | 140 | 212.6 | -0.9 | 212.6 | 65.3 | 65.2779 | 0 | 12.9 | 909 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.5 | -0.8 | 212.5 | 65.2 | 65.2 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.4 | -0.8 | 212.4 | 65.2 | 65.2 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.123 | 26.6 | -0.8
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.4 | -0.8 | 212.4 | 65.2 | 65.2 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.8
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.3 | -0.84329 | 212.3 | 65.2 | 65.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2 | 26.6 | -0.8
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.2 | -0.9 | 212.2 | 65.2 | 65.2 | 0 | 12.6 | 913 | 936 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1412 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.2 | -0.9 | 212.2 | 65.2 | 65.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1373 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.2 | -0.9 | 212.2 | 65.2 | 65.2 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1683 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.21 | 0.014 | 140 | 212.2 | -0.9 | 212.2 | 65.2 | 65.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0115 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.226 | 0.014 | 140 | 212.1 | -0.9 | 212.1 | 65.2 | 65.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.073 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.23 | 0.014 | 140 | 212.1 | -0.9 | 212.1 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0484 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.225 | 0.014 | 140 | 212.1 | -0.9 | 212.1 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0354 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.0302 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 212 | -0.9 | 212 | 65.3 | 65.3 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 190.3 | 6 | 195 | 59.4 | 0.0354 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 212 | -0.9 | 212 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0354 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.9 | -0.9 | 211.9 | 65.3 | 65.3 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.236 | 0.014 | 140 | 211.8 | -0.9 | 211.8 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.24 | 0.014 | 140 | 211.8 | -0.9 | 211.8 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1463 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.235 | 0.014 | 140 | 211.8 | -0.9 | 211.8 | 65.3 | 65.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.3 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2969 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.227 | 0.014 | 140 | 211.7 | -0.9 | 211.7 | 65.3 | 65.3 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1395 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.7 | -0.9 | 211.7 | 65.3 | 65.3 | 0 | 12.6 | 913 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.7 | -0.9 | 211.7 | 65.3 | 65.3 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1749 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.2681 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.2 | 65.2239 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.112 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.2 | 65.2 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985.432 | 30.1518 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.7 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1696 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.2 | 65.2 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 30.1179 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2752 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1677 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.4309 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.6 | -0.9 | 211.6 | 65.3 | 65.3 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5697 | 30.1036 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.4 | 65.4 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.5646 | 30.261 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.4 | 65.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 31.2068 | 59 | 1023 | 986 | 30.1819 | 26.5409 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.4 | 65.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 30.1 | 26.5376 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.4 | 65.4 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1161 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.4 | 65.4 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1945 | 26.6 | -0.9
2008-05-07 | 14:11 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.5 | 65.5 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1267 | 26.6 | -0.9
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.5 | 65.5 | 0 | 12.9 | 911 | 933 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.5 | 65.5 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.5 | -0.9 | 211.5 | 65.5 | 65.5 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0909 | 26.6 | -0.9
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.5 | -0.9432 | 211.5 | 65.5 | 65.5 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.0123 | 26.6 | -0.9
2008-05-07 | 14:12 | 4.217 | 0.014 | 140 | 211.5 | -1 | 211.5 | 65.5 | 65.5 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1988 | 26.6 | -1
2008-05-07 | 14:12 | 4.21 | 0.014 | 140 | 211.5 | -1 | 211.5 | 65.6 | 65.6 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2102 | 26.6 | -1
2008-05-07 | 14:12 | 4.21 | 0.014 | 140 | 211.5 | -1 | 211.5 | 65.6 | 65.6 | 0 | 12.8 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0353 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1234 | 26.6 | -1
2008-05-07 | 14:12 | 4.21 | 0.014 | 140 | 211.5 | -1 | 211.5 | 65.6 | 65.6 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2 | 26.5984 | -1
2008-05-07 | 14:12 | 4.204 | 0.014 | 140 | 211.3 | -1 | 211.3 | 65.6 | 65.6 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2 | 26.5195 | -1
2008-05-07 | 14:12 | 4.209 | 0.014 | 140 | 211.3 | -1 | 211.3 | 65.6 | 65.6 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1408 | 26.5592 | -1
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.2 | -1 | 211.2 | 65.7 | 65.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986.4346 | 30.0623 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211.1 | -1 | 211.1 | 65.7 | 65.7 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0162 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 211 | -1 | 211 | 65.7 | 65.7 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.000343 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0942 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 210.9 | -1 | 210.9 | 65.8 | 65.7935 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.014 | 140 | 210.9 | -1 | 210.9 | 65.8 | 65.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.0132 | 132.2446 | 210.9 | -1 | 210.9 | 65.8 | 65.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0693 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.013 | 130 | 210.9 | -1 | 210.9 | 65.9 | 65.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0093 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.013 | 130 | 210.9 | -1 | 210.9 | 65.9 | 65.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0878 | 26.6 | -1
2008-05-07 | 14:12 | 4.22 | 0.013 | 130 | 210.9 | -1 | 210.9 | 66 | 65.9864 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0352 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.84 | 0.0127 | 127 | 209.4 | -1 | 209.4 | 66 | 66 | 0 | 12.7 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.034984 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.145 | 26.6 | -1
2008-05-07 | 14:12 | 3.674 | 0.0102 | 102 | 198.2 | -1 | 198.2 | 65.9 | 65.9091 | 0 | 12.6 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.033209 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1764 | 26.6 | -1
2008-05-07 | 14:12 | 3.64 | 0.0108 | 107.8165 | 186.4 | -1 | 186.4 | 62.1 | 62.0738 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.031384 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4329 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.64 | 0.011 | 110 | 184 | -1 | 184 | 58.4 | 58.4308 | 0 | 12.6 | 911 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.031 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.64 | 0.0115 | 114.6414 | 183.7 | -1 | 183.7 | 57.1 | 57.1403 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0309 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 183.4 | -1 | 183.4 | 56.8 | 56.8117 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0309 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 183 | -1 | 183 | 56.7 | 56.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.030857 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5835 | -1
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 182.8 | -1 | 182.8 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.0308 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.5052 | -1
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 182.5 | -1 | 182.5 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0308 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5691 | 30.3206 | 26.5735 | -1
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 182.2 | -1 | 182.2 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3478 | 26.5478 | -1
2008-05-07 | 14:12 | 3.638 | 0.012 | 120 | 182.2 | -1 | 182.2 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2386 | 26.5307 | -1
2008-05-07 | 14:12 | 3.621 | 0.012 | 120 | 182.2 | -1 | 182.2 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.62 | 0.012 | 120 | 182.2 | -0.9568 | 182.2 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1
2008-05-07 | 14:12 | 3.62 | 0.012 | 120 | 182.2 | -0.9 | 182.2 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1665 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.62 | 0.012 | 120 | 182.1 | -0.9 | 182.1 | 56.7 | 56.7 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.155 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.62 | 0.012 | 120 | 182.1 | -0.9 | 182.1 | 56.7 | 56.7 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0765 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.62 | 0.012 | 120 | 182.1 | -0.9 | 182.1 | 56.8 | 56.7911 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0036 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.62 | 0.0117 | 117.1672 | 182 | -0.9 | 182 | 56.8 | 56.8 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1615 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.62 | 0.011 | 110 | 182 | -0.9 | 182 | 56.8 | 56.8 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2594 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.626 | 0.011 | 110 | 181.9 | -0.9 | 181.9 | 56.8 | 56.8 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1862 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0164 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0945 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.0113 | 112.6373 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0478 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.012 | 120 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0308 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.636 | 0.012 | 120 | 181.9 | -0.9 | 181.9 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1094 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.64 | 0.012 | 120 | 181.9 | -0.9 | 181.9 | 56.9 | 56.9 | 0 | 12.6 | 912 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.188 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.637 | 0.012 | 120 | 181.9 | -0.9 | 181.9 | 56.9 | 56.9 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5691 | 30.1334 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.0116 | 115.5314 | 181.9 | -0.9 | 181.9 | 56.8 | 56.8 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.9 | 181.9 | 56.8 | 56.8 | 0 | 12.7 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1471 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.9 | 181.9 | 56.8 | 56.8 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2968 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0307 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1385 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.8 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0262 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.8 | 181.9 | 56.8 | 56.8 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0524 | 26.6 | -0.8
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.9 | 181.9 | 56.8 | 56.8 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.9 | -0.9 | 181.9 | 56.9 | 56.8968 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.5906 | -0.9
2008-05-07 | 14:12 | 3.63 | 0.011 | 110 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.5119 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.011 | 110 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1667 | 26.5667 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.0111 | 110.932 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.2452 | 26.5548 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.0119 | 119.3283 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31.2068 | 59 | 1023 | 985 | 30.2525 | 26.5237 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.8 | -0.9 | 181.8 | 56.9 | 56.9 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31.7938 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.8 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.1 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 30.219 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 30.1856 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0166 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -0.9 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0948 | 26.6 | -0.9
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -1 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -1.1 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.1
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -1.2 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.069 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -1.2 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0283 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.63 | 0.012 | 120 | 181.7 | -1.2 | 181.7 | 56.9 | 56.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2644 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.636 | 0.012 | 120 | 181.8 | -1.2 | 181.8 | 57 | 56.9959 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1666 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.64 | 0.012 | 120 | 181.9 | -1.2 | 181.9 | 57 | 57 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1906 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.381 | 0.0114 | 114.0741 | 181.9 | -1.2 | 181.9 | 57 | 57 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0306 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3238 | 26.6 | -1.2
2008-05-07 | 14:13 | 2.655 | 0.0102 | 101.5833 | 175.1 | -1.3 | 175.1 | 57.1 | 57.1 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4568 | 0.029579 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.3945 | 26.6 | -1.3
2008-05-07 | 14:13 | 2.897 | 0.009 | 90 | 159.6 | -1.3 | 159.6 | 55.9 | 55.9006 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.027078 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1572 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.004 | 0.0089 | 88.9599 | 145.6 | -1.3 | 145.6 | 45.2 | 45.1588 | 0 | 12.6 | 911 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0248 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2192 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.0081 | 80.7699 | 146.5 | -1.3 | 146.5 | 44.1 | 44.1175 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.024986 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2618 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.0103 | 102.9174 | 148.1 | -1.3 | 148.1 | 46.4 | 46.44 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.025243 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1664 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.011 | 110 | 149 | -1.3 | 149 | 47.2 | 47.1834 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0254 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0093 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.0106 | 105.7202 | 149.7 | -1.3 | 149.7 | 47.3 | 47.285 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1478 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 149.7 | -1.3 | 149.7 | 47.2 | 47.2 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2525 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 149.8 | -1.3 | 149.8 | 47.1 | 47.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2377 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 149.9 | -1.3 | 149.9 | 47 | 47.0056 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1097 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150 | -1.2 | 150 | 46.9 | 46.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0255 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1883 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.4 | -1.2 | 150.4 | 46.8 | 46.8 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.4 | -1.2 | 150.4 | 46.7 | 46.7069 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1092 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.5 | -1.2 | 150.5 | 46.6 | 46.6025 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0718 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.5 | -1.2 | 150.5 | 46.6 | 46.6 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2953 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.5 | -1.2 | 150.5 | 46.5 | 46.5 | 0 | 12.8 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1378 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.034 | 0.01 | 100 | 150.5 | -1.2 | 150.5 | 46.5 | 46.5 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2194 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.04 | 0.01 | 100 | 150.6 | -1.2 | 150.6 | 46.5 | 46.5 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3766 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.04 | 0.01 | 100 | 150.5 | -1.2 | 150.5 | 46.5 | 46.5 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.4662 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.04 | 0.01 | 100 | 150.6 | -1.2 | 150.6 | 46.4 | 46.4063 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.309 | 26.6 | -1.2
2008-05-07 | 14:13 | 3.04 | 0.01 | 100 | 150.6 | -1.2 | 150.6 | 46.4 | 46.4 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.4482 | 26.5259 | -1.2
2008-05-07 | 14:13 | 3.04 | 0.01 | 100 | 150.7 | -1.2435 | 150.7 | 46.4 | 46.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5654 | 30.3421 | 26.5526 | -1.2
2008-05-07 | 14:13 | 3.036 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.4 | 46.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1689 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.4 | 46.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.4 | 46.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.3 | 46.3 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.3007 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.3 | 46.3 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.309 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.3 | 46.3 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1761 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.3 | 46.2742 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.2 | 46.2 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.2 | 46.2 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2196 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.3 | 150.8 | 46.2 | 46.2 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2233 | 26.6 | -1.3
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.8 | -1.4 | 150.8 | 46.2 | 46.2 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.9 | -1.4 | 150.9 | 46.2 | 46.2 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 150.9 | -1.4 | 150.9 | 46.2 | 46.2 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1526 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1907 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1116 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1241 | 26.6 | -1.4
2008-05-07 | 14:13 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1973 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.3 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1188 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.7 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.2197 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1848 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0339 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.191 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.031 | 0.01 | 100 | 151 | -1.4 | 151 | 46.1 | 46.1 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1259 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.039 | 0.01 | 100 | 151.1 | -1.4 | 151.1 | 46.1 | 46.1 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.04 | 0.01 | 100 | 151.1 | -1.4 | 151.1 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1624 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.04 | 0.01 | 100 | 151 | -1.4 | 151 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2806 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.04 | 0.01 | 100 | 151.1 | -1.4431 | 151.1 | 46.2 | 46.2 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1233 | 26.6 | -1.4
2008-05-07 | 14:14 | 3.04 | 0.01 | 100 | 151.1 | -1.5 | 151.1 | 46.2 | 46.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.039 | 0.01 | 100 | 151.1 | -1.5 | 151.1 | 46.3 | 46.3 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1911 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.031 | 0.01 | 100 | 151.1 | -1.5 | 151.1 | 46.3 | 46.3 | 0 | 12.6 | 913 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0256 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2519 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151.1 | -1.5 | 151.1 | 46.3 | 46.3 | 0 | 12.7 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.025643 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1047 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151.1 | -1.5 | 151.1 | 46.3 | 46.3 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.2623 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.03 | 0.01 | 100 | 151.2 | -1.5 | 151.2 | 46.4 | 46.4 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1803 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.032 | 0.01 | 100 | 151.2 | -1.5 | 151.2 | 46.4 | 46.4 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1768 | 26.6 | -1.5
2008-05-07 | 14:14 | 3.019 | 0.01 | 100 | 151.2 | -1.5 | 151.2 | 46.4 | 46.4 | 0 | 12.7 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.266 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.795 | 0.0098 | 97.5874 | 151.1 | -1.5 | 151.1 | 46.4 | 46.4 | 0 | 12.6 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0257 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1088 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.565 | 0.0091 | 90.7407 | 146.4 | -1.5 | 146.4 | 46.4 | 46.4 | 0 | 12.6 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.024925 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1742 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.482 | 0.0099 | 99.1582 | 135.6 | -1.5 | 135.6 | 45.4 | 45.3859 | 0 | 12.6 | 911 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.023095 | 0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1473 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.466 | 0.01 | 100 | 127.5 | -1.5 | 127.5 | 40.2 | 40.2387 | 0 | 12.7 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.02177 | 0.000257 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1311 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.46 | 0.01 | 100 | 125.6 | -1.5 | 125.6 | 38.1 | 38.0517 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0215 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1905 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.46 | 0.0096 | 95.9333 | 125.1 | -1.5 | 125.1 | 37.4 | 37.4194 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0214 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1117 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.46 | 0.009 | 90 | 125.1 | -1.5 | 125.1 | 37.2 | 37.2042 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0214 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.452 | 0.0089 | 88.9254 | 125 | -1.5 | 125 | 37.1 | 37.1 | 0 | 12.8 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0214 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1911 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.444 | 0.0081 | 80.9062 | 124.9 | -1.5 | 124.9 | 37.1 | 37.1 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0214 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.252 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 124.7 | -1.5 | 124.7 | 37.1 | 37.1 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4432 | 0.0213 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 124.2 | -1.5 | 124.2 | 37 | 37 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.021257 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 124 | -1.5 | 124 | 37 | 37 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 124 | -1.5 | 124 | 37 | 37 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.9 | -1.5 | 123.9 | 36.9 | 36.9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.8 | -1.5 | 123.8 | 36.9 | 36.9 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.8 | -1.5 | 123.8 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.8 | -1.5 | 123.8 | 36.9 | 36.9 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.0076 | 75.9011 | 123.7 | -1.5 | 123.7 | 37 | 36.9944 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.0072 | 72.4463 | 123.7 | -1.5 | 123.7 | 36.9 | 36.9103 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0212 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.6 | -1.5 | 123.6 | 36.9 | 36.9 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.021157 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.5 | -1.5 | 123.5 | 36.9 | 36.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.44 | 0.008 | 80 | 123.4 | -1.5 | 123.4 | 36.9 | 36.9 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0211 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.438 | 0.008 | 80 | 123.2 | -1.5 | 123.2 | 36.9 | 36.9 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021057 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1482 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.8 | -1.4569 | 122.8 | 36.9 | 36.9 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.4309 | 30.2947 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.6 | -1.4 | 122.6 | 36.9 | 36.9 | 0 | 12.7 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 30.1375 | 26.6 | -1.4
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.6 | -1.4 | 122.6 | 36.9 | 36.9 | 0 | 12.7 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.4
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.5 | -1.4 | 122.5 | 36.9 | 36.9 | 0 | 12.6 | 912 | 936 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0616 | 26.6 | -1.4
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.4 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | -1.4
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.4432 | 122.4 | 36.9 | 36.9058 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.021 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | -1.4
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.5 | -1.5 | 122.5 | 36.9 | 36.9 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0742 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.5 | -1.5 | 122.5 | 36.9 | 36.9 | 0 | 12.6 | 912 | 934 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0903 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0116 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.3431 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.0455 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.438 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 190.7 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.434 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.432 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0402 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.43 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.6 | -1.5
2008-05-07 | 14:14 | 2.431 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.439 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 914 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.0741 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.5 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 30.1 | 26.6 | -1.5
2008-05-07 | 14:15 | 2.44 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.009 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.452 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.9 | 36.9 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9478 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.46 | 0.008 | 80 | 122.5 | -1.6 | 122.5 | 36.9 | 36.9 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1021 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.46 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.9 | 36.8912 | 0 | 12.7 | 911 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.181 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.46 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1403 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.46 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.46 | 0.008 | 80 | 122.4 | -1.6 | 122.4 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4431 | 0.0209 | 0.0002 | 0 | 0 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0832 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.425 | 0.008 | 80 | 122.3 | -1.6 | 122.3 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 4.3e-05 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0046 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.4 | 0.008 | 80 | 122.3 | -1.6 | 122.3 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0739 | 26.6 | -1.6
2008-05-07 | 14:15 | 2.272 | 0.0068 | 68.02 | 122.3 | -1.6 | 122.3 | 36.8 | 36.8 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0209 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0475 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.859 | 0.0064 | 64.2114 | 115 | -1.6 | 115 | 36.8 | 36.8 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.019687 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.851 | 0.007 | 70 | 104.2 | -1.6 | 104.2 | 36.1 | 36.0553 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.017895 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0094 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.85 | 0.007 | 70 | 97 | -1.6 | 97 | 29.6 | 29.628 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.016741 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0882 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.85 | 0.007 | 70 | 94.5 | -1.6 | 94.5 | 27.9 | 27.8524 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.016257 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.1 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.85 | 0.007 | 70 | 94.1 | -1.6 | 94.1 | 27.5 | 27.4963 | 0 | 12.9 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0162 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0093 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.85 | 0.007 | 70 | 94.1 | -1.6 | 94.1 | 27.4 | 27.4 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.3568 | 0.0162 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9475 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.849 | 0.007 | 70 | 94 | -1.6 | 94 | 27.4 | 27.4 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0162 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0965 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.841 | 0.007 | 70 | 93.9 | -1.6 | 93.9 | 27.4 | 27.4 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0162 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9382 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.848 | 0.0068 | 68.1169 | 93.9 | -1.6 | 93.9 | 27.3 | 27.3134 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0162 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0192 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.832 | 0.0042 | 42.3216 | 93.8 | -1.6 | 93.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.016157 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0236 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.824 | 0.0048 | 47.6151 | 93.6 | -1.7 | 93.6 | 27.3 | 27.3 | 0 | 12.7 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0161 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.832 | 0.0044 | 44.2524 | 93.6 | -1.7 | 93.6 | 27.3 | 27.3 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0161 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.84 | 0.0044 | 43.9814 | 93.5 | -1.7 | 93.5 | 27.3 | 27.3 | 0 | 12.6 | 914 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0161 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.974 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.84 | 0.005 | 50 | 93.5 | -1.7 | 93.5 | 27.3 | 27.3 | 0 | 12.6 | 914 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0161 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.84 | 0.005 | 50 | 92.7 | -1.7 | 92.7 | 27.3 | 27.3 | 0 | 12.6 | 914 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.015956 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.5646 | 29.9689 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.84 | 0.005 | 50 | 92.4 | -1.7 | 92.4 | 27.3 | 27.3 | 0 | 12.6 | 914 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9 | 26.6 | -1.7
2008-05-07 | 14:15 | 1.84 | 0.005 | 50 | 92.4 | -1.6 | 92.4 | 27.3 | 27.3 | 0 | 12.6 | 914 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.0056 | 55.9733 | 92.4 | -1.6 | 92.4 | 27.3 | 27.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0337 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.3 | -1.6 | 92.3 | 27.3 | 27.3 | 0 | 12.6 | 914 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.3 | -1.6 | 92.3 | 27.3 | 27.3 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.1 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.3 | -1.6 | 92.3 | 27.3 | 27.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0976 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.3 | -1.6 | 92.3 | 27.4 | 27.3928 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.019 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.2 | -1.6 | 92.2 | 27.3 | 27.3 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0596 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.2 | -1.6 | 92.2 | 27.4 | 27.3901 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0236 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.2 | -1.6 | 92.2 | 27.4 | 27.4 | 0 | 12.6 | 913 | 936 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9168 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.1 | -1.6 | 92.1 | 27.4 | 27.4 | 0 | 12.7 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9954 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.1 | -1.6 | 92.1 | 27.4 | 27.4 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0159 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92.1 | -1.6 | 92.1 | 27.4 | 27.3952 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.015857 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 92 | -1.6 | 92 | 27.4 | 27.4 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0309 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.9 | -1.6 | 91.9 | 27.4 | 27.4 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30.0719 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.9 | -1.6 | 91.9 | 27.4 | 27.4 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8355 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.9 | -1.6 | 91.9 | 27.4 | 27.4 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8668 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.9 | -1.6 | 91.9 | 27.4 | 27.4 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9238 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.006 | 60 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.0064 | 64.1269 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.4343 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.007 | 70 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.7 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.007 | 70 | 91.8 | -1.6 | 91.8 | 27.4 | 27.4 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.007 | 70 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.84 | 0.007 | 70 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.838 | 0.007 | 70 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:15 | 1.831 | 0.0066 | 66.0544 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.839 | 0.006 | 60 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.8 | -1.6433 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.8 | -1.7 | 91.8 | 27.3 | 27.3 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.8 | -1.6 | 91.8 | 27.3 | 27.3 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4568 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.8 | -1.7 | 91.8 | 27.3 | 27.3 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 30 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.9 | -1.7 | 91.9 | 27.3 | 27.2926 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.9965 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 91.9 | -1.7 | 91.9 | 27.2 | 27.2 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.8387 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.84 | 0.006 | 60 | 92 | -1.7 | 92 | 27.2 | 27.2 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8594 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.834 | 0.006 | 60 | 92 | -1.7 | 92 | 27.2 | 27.2 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.976 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92 | -1.7 | 92 | 27.2 | 27.2 | 0 | 12.9 | 910 | 932 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30.0672 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92 | -1.7 | 92 | 27.1 | 27.1 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9111 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92 | -1.7 | 92 | 27.1 | 27.1 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9735 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92 | -1.7 | 92 | 27.1 | 27.1 | 0 | 12.8 | 911 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 30 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92 | -1.7 | 92 | 27.1 | 27.1 | 0 | 12.8 | 911 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9693 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92.1 | -1.7432 | 92.1 | 27.1 | 27.1 | 0 | 12.8 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8908 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.83 | 0.006 | 60 | 92.1 | -1.8 | 92.1 | 27.1 | 27.1 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8121 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.682 | 0.006 | 60 | 92.2 | -1.8 | 92.2 | 27.1 | 27.1 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0158 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.8 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.201 | 0.0052 | 51.8205 | 90.1 | -1.8 | 90.1 | 27.1 | 27.1 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.015454 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.432 | 29.8901 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.218 | 0.004 | 40 | 75.8 | -1.8 | 75.8 | 27.1 | 27.1 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.013026 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 30 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.227 | 0.004 | 40 | 66.5 | -1.8 | 66.5 | 21.2 | 21.2186 | 0 | 12.6 | 912 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.011427 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.9984 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 64 | -1.8 | 64 | 18 | 18.0153 | 0 | 12.6 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0.011 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9194 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.242 | 0.004 | 40 | 63.8 | -1.8 | 63.8 | 17.9 | 17.9 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.011 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.26 | 0.004 | 40 | 63.5 | -1.8 | 63.5 | 17.8 | 17.8 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.259 | 0.004 | 40 | 63.3 | -1.8 | 63.3 | 17.8 | 17.8 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8835 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.242 | 0.004 | 40 | 63.3 | -1.8 | 63.3 | 17.8 | 17.8 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.568 | 29.805 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63.2 | -1.8 | 63.2 | 17.8 | 17.8 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63.2 | -1.8 | 63.2 | 17.8 | 17.8 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63.1 | -1.8 | 63.1 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8 | 26.6 | -1.8
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63.1 | -1.7 | 63.1 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190.3 | 6 | 195 | 59.4 | 0.0109 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8189 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63 | -1.6568 | 63 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 190 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.976 | 26.6 | -1.7
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63 | -1.6 | 63 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 190 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 30 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 63 | -1.6 | 63 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.7 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.955 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.902 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.24 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9807 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.239 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0108 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8814 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.231 | 0.0042 | 42.3415 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8379 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.005 | 50 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9164 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.005 | 50 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.9948 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.005 | 50 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.7794 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.0044 | 44.3997 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8043 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.7 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.8693 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.8093 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.7 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.8879 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9664 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.9101 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.0034 | 34.4598 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8235 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.003 | 30 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8983 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.003 | 30 | 62.6 | -1.6 | 62.6 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8193 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.003 | 30 | 62.5 | -1.6 | 62.5 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.23 | 0.003 | 30 | 62.6 | -1.5 | 62.6 | 17.7 | 17.7 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.003 | 30 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0107 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8164 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.0036 | 35.6375 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8945 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.5689 | 29.8266 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7384 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.23 | 0.004 | 40 | 62.5 | -1.5 | 62.5 | 17.7 | 17.7 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.6283 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.234 | 0.0033 | 32.8125 | 62.4 | -1.5 | 62.4 | 17.7 | 17.7 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.8639 | 26.6 | -1.5
2008-05-07 | 14:16 | 1.24 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.9665 | 26.6 | -1.6
2008-05-07 | 14:16 | 1.24 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.8 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8647 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.24 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.7 | 911 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.7 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.233 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.7016 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.23 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.7807 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.234 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.24 | 0.003 | 30 | 62.4 | -1.6 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.6
2008-05-07 | 14:17 | 1.24 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.0031 | 30.5447 | 62.4 | -1.7 | 62.4 | 17.7 | 17.7 | 0 | 12.7 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.0039 | 39.0553 | 62.4 | -1.7 | 62.4 | 17.8 | 17.7952 | 0 | 12.6 | 913 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.004 | 40 | 62.4 | -1.7 | 62.4 | 17.7 | 17.7 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | 0 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.004 | 40 | 62.4 | -1.7 | 62.4 | 17.8 | 17.789 | 0 | 12.6 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -4.3e-05 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.0036 | 36.0237 | 62.4 | -1.7 | 62.4 | 17.7 | 17.7067 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.24 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.8 | 17.8 | 0 | 12.6 | 912 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.226 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.8 | 17.8 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.22 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.8 | 17.8 | 0 | 12.6 | 913 | 934 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.22 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.8 | 17.8 | 0 | 12.6 | 913 | 933 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 1.015 | 0.003 | 30 | 62.4 | -1.7 | 62.4 | 17.8 | 17.7951 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0106 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.355 | 0.0014 | 14.1766 | 56.6 | -1.7 | 56.6 | 17.7 | 17.7098 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.010384 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.437 | -0.001 | -10 | 48.9 | -1.7 | 48.9 | 17.7 | 17.7 | 0 | 12.7 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3 | 0.0083 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.57 | -0.001 | -10 | 26.2 | -1.7 | 26.2 | 9.5 | 9.5298 | 0 | 12.6 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.004327 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.626 | -0.001 | -10 | 24 | -1.7 | 24 | 4.4 | 4.4128 | 0 | 12.7 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.003943 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | -0.0003 | -3.0219 | 27.9 | -1.7 | 27.9 | 6.5 | 6.5042 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.004643 | -0.0001 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0 | 0 | 29.4 | -1.7 | 29.4 | 8 | 8.0199 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.004843 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0 | 0 | 29.7 | -1.7 | 29.7 | 8.4 | 8.3925 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0 | 0 | 29.8 | -1.7 | 29.8 | 8.4 | 8.4 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0 | 0 | 29.8 | -1.7 | 29.8 | 8.4 | 8.4 | 0 | 12.9 | 910 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0049 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0 | 0 | 30 | -1.7 | 30 | 8.4 | 8.4 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.004944 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.7
2008-05-07 | 14:17 | 0.63 | 0.0007 | 6.9558 | 30.2 | -1.6 | 30.2 | 8.4 | 8.4 | 0 | 12.9 | 910 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.8 | 26.6 | -1.6
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.2 | -1.5569 | 30.2 | 8.3 | 8.3102 | 0 | 12.8 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.7519 | 26.6 | -1.6
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.3 | -1.4565 | 30.3 | 8.3 | 8.3 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.6052 | 26.6 | -1.5
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.3 | -1.4 | 30.3 | 8.3 | 8.3 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.7623 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.3 | -1.4 | 30.3 | 8.3 | 8.3 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.6805 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.3 | -1.4 | 30.3 | 8.3 | 8.3 | 0 | 12.8 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6383 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.4 | -1.4 | 30.4 | 8.3 | 8.3 | 0 | 12.9 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.7169 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.5 | -1.4 | 30.5 | 8.3 | 8.3 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7954 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.6 | -1.4 | 30.6 | 8.3 | 8.3 | 0 | 12.7 | 912 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.6 | -1.4 | 30.6 | 8.3 | 8.3 | 0 | 12.7 | 912 | 935 | 925 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.6 | -1.3 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.63 | 0.001 | 10 | 30.6 | -1.3432 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7903 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.634 | 0.001 | 10 | 30.6 | -1.4 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 913 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7116 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.6 | -1.4 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 912 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.767 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.6 | -1.4 | 30.6 | 8.3 | 8.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.4311 | 29.8 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.3 | 0 | 12.7 | 912 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.3 | 0 | 12.6 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.5703 | 29.7953 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.3 | 0 | 12.7 | 912 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.005043 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6377 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.3 | 0 | 12.6 | 913 | 936 | 921 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6768 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.7 | -1.4 | 30.7 | 8.3 | 8.2749 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.766 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.8 | -1.4 | 30.8 | 8.2 | 8.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6088 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.8 | -1.3568 | 30.8 | 8.2 | 8.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7483 | 26.6 | -1.4
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.8 | -1.3 | 30.8 | 8.2 | 8.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 29.8 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.644 | 0.001 | 10 | 30.9 | -1.3 | 30.9 | 8.2 | 8.2 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 29.7376 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.648 | 0.001 | 10 | 30.9 | -1.3 | 30.9 | 8.2 | 8.2 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 29.6095 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.001 | 10 | 30.9 | -1.3 | 30.9 | 8.2 | 8.2 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.6884 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.0012 | 12.2772 | 31 | -1.3 | 31 | 8.3 | 8.2919 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.7 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.002 | 20 | 31.1 | -1.3 | 31.1 | 8.2 | 8.2 | 0 | 12.6 | 913 | 936 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0002 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 985 | 29.6544 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.002 | 20 | 31.3 | -1.3 | 31.3 | 8.3 | 8.2961 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.000243 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1023 | 985 | 29.6481 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.002 | 20 | 31.3 | -1.3 | 31.3 | 8.3 | 8.3 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.7959 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.002 | 20 | 31.4 | -1.3 | 31.4 | 8.3 | 8.3 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.432 | 29.6374 | 26.6 | -1.3
2008-05-07 | 14:17 | 0.64 | 0.002 | 20 | 31.4 | -1.3 | 31.4 | 8.3 | 8.3 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.72 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.0016 | 16.0963 | 31.4 | -1.3 | 31.4 | 8.3 | 8.3 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7229 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.0001 | 1.0756 | 31.4 | -1.3 | 31.4 | 8.3 | 8.3 | 0 | 12.8 | 912 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.5829 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.0018 | 17.8824 | 31.5 | -1.4 | 31.5 | 8.3 | 8.3 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.5042 | 26.6 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.4 | 31.5 | 8.3 | 8.3 | 0 | 12.8 | 912 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6486 | 26.6 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.4 | 31.5 | 8.3 | 8.3 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.7 | 26.6 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.4 | 31.5 | 8.3 | 8.3 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6686 | 26.6314 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.4 | 31.5 | 8.3 | 8.3 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6901 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.4 | 31.5 | 8.4 | 8.3938 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6114 | -1.4
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.3913 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7344 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7083 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6243 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.3 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 909 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.2569 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 985.4309 | 29.6827 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.64 | 0.002 | 20 | 31.5 | -1.2 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 5.7e-05 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6041 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.618 | 0.0016 | 16.2151 | 31.5 | -1.2 | 31.5 | 8.4 | 8.4 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985.568 | 29.6745 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.28 | 0.0012 | 11.8333 | 31.5 | -1.2 | 31.5 | 8.3 | 8.3083 | 0 | 12.8 | 911 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.0051 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.6469 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.096 | 0.002 | 20 | 25.7 | -1.2 | 25.7 | 8.4 | 8.4 | 0 | 12.7 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | 0.004535 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.024 | 0.002 | 20 | 8.4 | -1.2 | 8.4 | 6.1 | 6.1156 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | 0.001145 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0.0013 | 12.7013 | 4.7 | -1.2 | 4.7 | 1.3 | 1.2758 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.000557 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0.001 | 10 | 3.1 | -1.2 | 3.1 | -0.3 | 0 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.000257 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6675 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0.001 | 10 | 2.5 | -1.2 | 2.5 | -0.7 | 0 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.654 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0.0008 | 7.649 | 2.4 | -1.2 | 2.4 | -0.8 | 0 | 0 | 12.6 | 914 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.5754 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2.4 | -1.2 | 2.4 | -0.8 | 0 | 0 | 12.6 | 914 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.5032 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2.4 | -1.2 | 2.4 | -0.9 | 0 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.5818 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2.3 | -1.2 | 2.3 | -0.9 | 0 | 0 | 12.6 | 914 | 936 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | 0.0001 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2.2 | -1.2 | 2.2 | -0.8 | 0 | 0 | 12.6 | 913 | 935 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | 5.7e-05 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2.2 | -1.2 | 2.2 | -0.8 | 0 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190.3 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2 | -1.2 | 2 | -0.8 | 0 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2 | -1.2 | 2 | -0.8 | 0 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 190 | 6 | 195 | 59.3569 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.01 | 0 | 0 | 2 | -1.2 | 2 | -0.8 | 0 | 0 | 12.6 | 914 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 190 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31.207 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.006 | 0 | 0 | 1.9 | -1.2 | 1.9 | -0.8 | 0 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31.793 | 59 | 1022 | 985 | 29.5682 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.002 | -0.0006 | -5.5245 | 1.8 | -1.2 | 1.8 | -0.8 | 0 | 0 | 12.6 | 913 | 934 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.7 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 31.2068 | 59 | 1022 | 985 | 29.5102 | 26.6 | -1.2
2008-05-07 | 14:18 | 0.009 | -0.0006 | -5.9333 | 1.8 | -1.3432 | 1.8 | -0.8 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 190.3 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.589 | 26.6 | -1.3
2008-05-07 | 14:18 | 0.001 | 0 | 0 | 1.7 | -1.4 | 1.7 | -0.8 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 190.7 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.432 | 29.6677 | 26.6 | -1.4
2008-05-07 | 14:18 | 0 | 0 | 0 | 1.7 | -1.4 | 1.7 | -0.8 | 0 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | 0 | -0.0003 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 985.432 | 29.7 | 26.6 | -1.4
2008-05-07 | 14:18 | 0 | 0 | 0 | 1.8 | -1.5 | 1.8 | -0.8 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0004 | 0 | 0 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.6751 | 26.6 | -1.5
2008-05-07 | 14:18 | 0 | 0 | 0 | 1.8 | -1.5 | 1.8 | -0.8 | 0 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0004 | 0 | 4.3e-05 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.5965 | 26.6 | -1.5
2008-05-07 | 14:18 | 0 | -0.0005 | -5.4678 | 1.8 | -1.5 | 1.8 | -0.8 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0004 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.5179 | 26.6 | -1.5
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 1.9 | -1.5 | 1.9 | -0.8 | 0 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0004 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.5607 | 26.6 | -1.5
2008-05-07 | 14:18 | 0 | -0.0008 | -8.0482 | 2.1 | -1.6 | 2.1 | -0.8 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0004 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 986 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.007 | 0 | 0 | 2.1 | -1.6 | 2.1 | -0.9 | 0 | 0 | 12.7 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.000443 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985.568 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | 0 | 0 | 2.2 | -1.6 | 2.2 | -0.9 | 0 | 0 | 12.6 | 912 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | 0 | 0 | 2.3 | -1.6 | 2.3 | -0.9 | 0 | 0 | 12.7 | 913 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | -0.0001 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | 0 | 0 | 1.8 | -1.6 | 1.8 | -0.9 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0002 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | 0 | 0 | 1.3 | -1.6 | 1.3 | -0.9 | 0 | 0 | 12.6 | 912 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022 | 985 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | -0.0002 | -2.0156 | 1.3 | -1.6 | 1.3 | -0.9 | 0 | 0 | 12.6 | 913 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.3 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1022.5859 | 985 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | -0.001 | -10 | 1.2 | -1.6 | 1.2 | -0.9 | 0 | 0 | 12.6 | 913 | 935 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985.568 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.01 | -0.001 | -10 | 1.2 | -1.6 | 1.2 | -0.9 | 0 | 0 | 12.6 | 913 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | -0.006 | -0.001 | -10 | 1.2 | -1.6 | 1.2 | -1 | 0 | 0 | 12.7 | 912 | 935 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.8 | -1.6 | 0.8 | -1 | 0 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.3 | 191 | 6 | 195 | 59.4 | -0.0003 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.5496 | 26.6252 | -1.6
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.7 | -1.6 | 0.7 | -1 | 0 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0004 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.4154 | 26.6961 | -1.6
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.7 | -1.7 | 0.7 | -1 | 0 | 0 | 12.7 | 912 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0004 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7299 | 26.6175 | -1.7
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.4 | -1.7 | 0.4 | -1 | 0 | 0 | 12.8 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0004 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6778 | 26.6 | -1.7
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.3 | -1.7 | 0.3 | -1 | 0 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189 | 191 | 6 | 195 | 59.4 | -0.0004 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1 | 0 | 0 | 12.8 | 911 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 189.7 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1 | 0 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:18 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1.1 | 0 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1.1 | 0 | 0 | 12.9 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1.1 | 0 | 0 | 12.9 | 911 | 934 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 190.7 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31.793 | 59 | 1023 | 986 | 29.6654 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.2 | -1.7 | 0.2 | -1.1 | 0 | 0 | 12.8 | 911 | 934 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.7774 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.7 | 0.1 | -1.1 | 0 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6201 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.7 | 0.1 | -1.1 | 0 | 0 | 12.9 | 911 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.7 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.7 | 0.1 | -1.1 | 0 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.7 | 0.1 | -1.1 | 0 | 0 | 12.9 | 911 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.7 | 0.1 | -1.1 | 0 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.7
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0.1 | -1.6 | 0.1 | -1.1 | 0 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.7 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0 | -1.5568 | 0 | -1.2 | 0 | 0 | 12.9 | 910 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 31.207 | 59 | 1023 | 986 | 29.6 | 26.6 | -1.6
2008-05-07 | 14:19 | 0 | -0.001 | -10 | 0 | -1.5 | 0 | -1.1 | 0 | 0 | 12.9 | 910 | 933 | 922 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 193 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 986 | 29.52 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.1 | 0 | 0 | 12.8 | 911 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192.3 | 191 | 6 | 195 | 59.3435 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.4369 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.7 | 912 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.5948 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.7 | 912 | 932 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.6 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 192 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7088 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.6 | 913 | 933 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191.3 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.767 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.6 | 913 | 932 | 924 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.4 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7118 | 26.6 | -1.5
2008-05-07 | 14:19 | 0 | -0.001 | -10 | -0.1 | -1.5 | 0 | -1.2 | 0 | 0 | 12.6 | 913 | 933 | 923 | 0 | 0 | 0 | 976 | -40 | 0 | 28 | 191 | 191 | 6 | 195 | 59.3568 | -0.0005 | -0.0005 | 0 | 0.0001 |  | n/a |  | 32 | 59 | 1023 | 985 | 29.7904 | 26.6 | -1.5
Summary Information:
Post Processor DLL Version | 6.81
Status: | MD5 digest is valid
Torque from ecm but iENG_TORQUE not available
Test Date | 2008-05-07
System Information:
Name | SEMTECH-DS GAS ANALYZER
Model | SEMTECH-DS
Serial | G05-SDS02
Version | 2.018 22
-----------------------------------------------------------------
Name | AUTOMOTIVE MICROBENCH II
Model | AMBII
Serial | 26374
Version | 6.907
CO Span(%) | 0.02
CO2 Span(%) | 6
C6H14 Span(ppm) | 2004
-----------------------------------------------------------------
Name | NDUV NO/NO2 ANALYZER
Model | NDUV-NO/NO2
Serial | 210
Version | 1.611
NO Span(ppm) | 298
NO2 Span(ppm) | 504
-----------------------------------------------------------------
Name | THC FID
Model | SEMTECH_DS_Dual
Serial | 66
Range(ppmC)1 | 100.00 Bottle(ppmC) = 93.12
Range(ppmC)2 | 1000.0 Bottle(ppmC) = 0000000
Range(ppmC)3 | 10000 Bottle(ppmC) = 0000000
Range(ppmC)4 | 40000 Bottle(ppmC) = 0000000
-----------------------------------------------------------------
Vehicle Description: | CAT TH83
License Plate
Engine Displacement | 0
Rated Horsepower
Rated RPM | 0
Fuel Specific Gravity | 0.85
H/C Molar Ratio of Fuel | 1.8
SEMTECH Serial Number | G05-SDS02
AMBII RPM Multiplier | 4
Torque (ecm or calc) | ecm
Use Frictional Torque | NO
Mass Calc Method | EXH_FLOW
NDIR Delay (s) | 4
NDUV Delay (s) | 4
THC FID Delay (s) | 3
Methane FID Delay (s) | 5
SEMTECH EFM Delay (s) | 2
Vehicle Interface Delay (s) | 0
Engine Speed Delay (s) | none
Environmental Delay (s) | 0
Aux Temp Delay (s) | none
EAI1 Delay (s) | 0
EAI2 Delay (s) | 0
EAI3 Delay (s) | 0
Methane FID PF-CH4 value | 0
Methane FID PF-C2H6 value | 0
Vehicle Interface Type | Not Enabled -
Flow Meter Type | Not Enabled
NOx Kh Calculation | 86.1370-2007 NTE
Curb Idle Load (%) | 0
Test Start Time | 13:58
Test End Time | 14:19
Test Duration (s) | 1263
Average Ambient Temperature (deg C) | -40
Average Ambient Pressure (mbar) | 976
Average Relative Humidity (%) | 0
Average Absolute Humidity (grains/lb dry air) | 0
Average Kh Factor | 0.885
Overrides:
iSCB_RH
iSCB_LAT
iVEH_SPEED_USED | iVEH_SPEED
iENG_SPEED_USED | iENG_SPEED
iVEH_SPEED | iSCB_EAI1_XF
Overall Test Results:
Total Distance Traveled (mi) | 0
Total Fuel Consumed (gal) | 0
Overall Fuel Economy (mpg) | 0
Total Work (bhp-hr) | 0
Overall Emissions (Distance Specific):
CO2 (g/mi) | 0
CO (g/mi) | 0
NOx (g/mi) | 0
kNOx (g/mi) (corrected NOx) | 0
THC (g/mi) | 0
CH4 (g/mi) | 0
NMHC (g/mi) | 0
C6H14 (g/mi) | 0
Overall Emissions (Brake Specific):
CO2 (g/bhp-hr) | 0
CO (g/bhp-hr) | 0
NOx (g/bhp-hr) | 0
kNOx (g/bhp-hr) (corrected NOx) | 0
THC (g/bhp-hr) | 0
CH4 (g/bhp-hr) | 0
NMHC (g/bhp-hr) | 0
C6H14 (g/bhp-hr) | 0
NOx + NMHC (g/bhp-hr) | 0
NTE Test Results:
Total Number of NTE Windows | 0
| Avg. NTE Emissions | Maximum NTE Emissions | Std. Dev.
CO2 (g/bhp-hr)
CO (g/bhp-hr)
NOx (g/bhp-hr)
kNOx (g/bhp-hr) (corrected NOx)
THC (g/bhp-hr)
CH4 (g/bhp-hr)
NMHC (g/bhp-hr)
C6H14 (g/bhp-hr)
NTE Window # | Start Date | Start Time | End Date | End Time | Duration (s) | Avg CO2 (%) | Avg CO (%) | Avg NOx (ppm) | Avg kNOx (ppm) | Avg THC (ppmC) | Avg CH4 (ppmC) | Avg NMHC (ppmC) | Avg C6H14 (ppm) | Avg Exhaust Flow (kg/hr) | Avg Exhaust Temperature (deg C) | Total Mass CO2 (g) | Total Mass CO (g) | Total Mass NOx (g) | Total Mass kNOx (g) | Total Mass THC (g) | Total Mass CH4 (g) | Total Mass NMHC (g) | Total Mass C6H14 (g) | Total Work (bhp-hr) | CO2 (g/bhp-hr) | CO (g/bhp-hr) | NOx (g/bhp-hr) | kNOx (g/bhp-hr) | THC (g/bhp-hr) | CH4 (g/bhp-hr) | NMHC (g/bhp-hr) | C6H14 (g/bhp-hr) | PM (g/bhp-hr) | AVL MSS (g/bhp-hr) | Flow Weighted Fuel-Specific CO2 (g) | Flow Weighted Fuel-Specific CO (g) | Flow Weighted Fuel-Specific NOx (g) | Flow Weighted Fuel-Specific kNOx (g) | Flow Weighted Fuel-Specific THC (g) | Flow Weighted Fuel-Specific CH4 (g) | Flow Weighted Fuel-Specific NMHC (g) | Flow Weighted Fuel-Specific C6H14 (g)
Detection Limits:
CO Limit (%) | 0
CO2 Limit (%) | 0
NO Limit (ppm) | 0
NO2 Limit (ppm) | 0
HC Limit (ppmC) | 0
Methane Limit (ppmC) | 0
Hexane Limit (ppm) | 0
AVL MSS Concentraiton Limit (mg/m3) | 0
AVL MSS Dilution Ratio Limit | 0
Faults:
0X0000 - 05/07/2008 13:53:19.800 - None Found
0X0000 - 05/07/2008 13:57:45.440 - None Found
Warnings:
0X0000 - 05/07/2008 13:53:19.801 - None Found
0X0000 - 05/07/2008 13:57:45.441 - None Found
Post Processor Limits:
Engine Speed Limit (rpm/s) | 1000
Vehicle Speed Limit (mph/s) | 21
Fuel Rate Limit (gal/s) | 0.05
Reference Torque Limit (lb-ft) | 10000
Fuel Specific Dropout Limit(% C) | 0.5
Brake Specific Dropout Limit (bhp-h) | 0.005
FID Range Change Ignore | 4
Post Processor Limit Events:
Engine Speed Limit Count | 0
Vehicle Speed Limit Count | 0
GPS Speed Limit Count | 0
Fuel Rate Limit Count | 0
Reference Torque Limit Count | 0
Fuel Specific Dropout Limit Count | 126
Brake Specific Dropout Limit Count | 1263
FID Range Change Ignore Count | 0
External Input Configuration:
| ID | Description | Units | Polynomial Order | x^0 | x^1 | x^2 | x^3 | x^4 | x^5 | x^6 | x^7 | x^8 | x^9
EAI1 | iVEH_SPEED | Vehicle Speed (Dyno) | MPH | 1 | x^0: 0; x^1: 6
EAI2 |  |  |  | -1
EAI3 |  |  |  | -1
Audit/Span/Zero Information:
Zero
InfoVer | 1
Date | 2008-05-07
Time | 13:53
Purge Delay | 30
Gas Path | Zero
Gas | Previous | Current | Difference
CO(ppm) | 13.125 | 0 | 13.125
CO2(%) | 0 | 0 | 0
NO(ppm) | -0.34375 | 0.12667 | -0.47042
NO2(ppm) | -0.33125 | 0.39333 | -0.72458
THC(ppmC) | 0.2 | -0.3 | 0.5
Test Information:
SEMTECH_DATA_FILE | 2008-05-07 | 13:52 | [RELEASE_VER=2.018 BUILD=22 BDATE=01/29/2008 IP=10.10.1.55]
No Information Available

## THC Lin
Test Date:
Test Time:
5/7/2008
13:57:45
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## CO2 Lin
Test Date:
Test Time:
5/7/2008
13:57:45
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## CO Lin
Test Date:
Test Time:
5/7/2008
13:57:45
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## NO Lin
Test Date:
Test Time:
5/7/2008
13:57:45
NDUV Serial No:
210
SEMTECH Serial Number:
G05-SDS02

## FID
Bottle Concentrations:
HC: ppm propane
THC: ppmC
Test Date:
Test Time:
5/7/2008
13:57:45
FID Serial No:
66
SEMTECH Serial Number:
G05-SDS02

## AMBII and NDUV
Bottle Concentrations:
CO: ppm
CO2: %
NO: ppm
Test Date:
Test Time:
5/7/2008
13:57:45
NDUV Serial No:
210
AMBII Serial No:
26374
SEMTECH Serial Number:
G05-SDS02

## CO2 Correction
| Divide point on divider | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90
% CO2 gas
6 | Rm = | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277 | 0.99277
| C.F.= | 0.099349 | 0.19884 | 0.29848 | 0.39826 | 0.49819 | 0.59826 | 0.69847 | 0.79884 | 0.89935
| gas = | 0.59609 | 1.193 | 1.7909 | 2.3896 | 2.9891 | 3.5895 | 4.1908 | 4.793 | 5.3961
Ro for CO2= | 0.8795